https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9589376&pc=QCB | 2023826-2020-00100 | 2019-11-19 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+3.0/081 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND FOREIGN RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.0/+3.0/081 | OD |  | 12/18/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9608209&pc=QCB | 2023826-2020-00170 | 2019-10-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2019 A 13.2MM, VTICMO13.2, -11.5/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) AND REMOVED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. A REPLACEMENT LENS WAS LATER IMPLANTED AND IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/1.5/089 | OD |  | 12/25/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9635037&pc=QCB | 2023826-2020-00198 | 2019-12-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/085 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+1.5/085 | OD |  | 12/30/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9664306&pc=QCB | 2023826-2020-00243 | 2019-12-23 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 12.6MM TMICL12.6, -10.00/1.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. A BACK-UP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6, -10.00/1.5/080 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LENS FLIPPED AFTER INSERTION. LENS WAS REMOVED INTRAOPERATIVELY. LENS TORE BEFORE/DURING 2ND LOADING. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE BACK-UP LENS WAS IMPLANTED ON (B)(6) 2019. THE LENS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. THIS RESOLVED THE PROBLEM. REPORTEDLY, "LENS FLIPPED AFTER INSERTION, REMOVED LENS, WHEN RE-LOADING LENS IT RIPPED." (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED DRY, INSIDE A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CORRECTED DATA: D11 - INJECTOR MODEL-MSI-PF; LOT#-UNK SHOULD BE REMOVED AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/1.5/080 | OS |  | 01/16/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9717754&pc=QCB | 2023826-2020-00255 | 2020-01-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -10.00/1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS BEING IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS STOPPED AT THE INCISION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. WHEN REMOVING THE LENS THE SURGEON "WAS NOT CAREFUL AND FOUND THAT THE HAPITC WAS TORN AND COULD NOT BE IMPLANTED." A SMALLER LENGTH BACK-UP LENS WAS IMPLANTED DURING THE INITIAL SURGERY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/1.5/094 | OS |  | 01/30/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9741504&pc=QCB | 2023826-2020-00327 | 2020-01-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -12.00/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE "LENS DID NOT UNFOLD PROPERLY AND ROTATED UPON INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. WHILE REMOVING THE LENS TO RELOAD, THE LENS TORE. BACK-UP LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE LENS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY (ON ((B)(6) 2020). Manufacturer Narrative: ADDITIONAL DATA: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6- RESULT CODE 114 IS ADDED BASED ON THE PRODUCT EVALUATION RESULTS IN PREVIOUS MDR. (B)(4). Manufacturer Narrative: CORRECTED DATA: H10-CLAIM NUMBER IN PREVIOUS MDR CORRECTED TO (B)(4). CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D4: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/2.0/089 | OD |  | 02/05/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9753284&pc=QCB | 2023826-2020-00367 | 2020-01-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), IT TORE. IT WAS REMOVED INTRAOPERATIVELY AND THERE WAS NO PATIENT INJURY. THIS OCCURRED ON (B)(6) 2020. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. ON (B)(6) 2020 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: CORRECTED DATA: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC TORN. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/+1.0/093 | OD |  | 01/30/2020 | 2020 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9759156&pc=QCB | 2023826-2020-00386 | 2020-01-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.0/+1.0/85 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, THERE WAS NO PATIENT INJURY. ON (B)(6) 2020 A REPLACEMENT LENS WAS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS USER ERROR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.0/+1.0/85 | NA |  | 01/31/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781530&pc=QCB | 2023826-2020-00340 | 2019-12-28 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.00/2.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED INTRAOPERATIVELY. A LENS WITH A SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) WAS LATER IMPLANTED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. REPORTEDLY, "LENS GOT STUCK WITH PLUNGER AND DAMAGED WHILE INJECTING INTO THE EYE, WHICH WAS REMOVED ON THE SAME DAY. SECOND LENS IMPLANTATION WAS SOOTH AND PATIENT IS EXTREMELY HAPPY." Manufacturer Narrative: ADDITIONAL INFORMATION H3- DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND OPTIC TORN/HAPTIC TORN AND BENT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/2.5/094 | OD |  | 02/11/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781698&pc=QCB | 2023826-2020-00393 | 2019-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -14.0/+2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), IT TORE. IT WAS REMOVED INTRAOPERATIVELY AND THERE WAS NO PATIENT INJURY. THIS OCCURRED ON (B)(6) 2019. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON (B)(6) 2019 AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE AND DEBRIS IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FOREIGN RESIDUE AND DEBRIS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.0/+2.5/091 | OS |  | 02/10/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9837837&pc=QCB | 2023826-2020-00528 | 2019-08-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+2.5/091 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.5/+2.5/091 | OS |  | 02/18/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9913475&pc=QCB | 2023826-2020-00727 | 2019-04-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/+3.5/081 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENT'S RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR THE BACKUP LENS, WHICH WAS A LONGER LENS, WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.50/+3.5/081 | OD |  | 03/08/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9938153&pc=QCB | 2023826-2020-00794 | 2020-03-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -5.50/1.5/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. PER REPORTER ON (B)(6) 2020, ALTERNATE LENS IS TO BE IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY, "LENS PRESENTED STRANGE FOLD." Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN WITH DEBRIS ON THE LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.50/1.5/102 | OD |  | 03/12/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9944276&pc=QCB | 2023826-2020-00812 | 2019-12-30 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/072 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENT'S LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER LENS (SHORTER LENS), WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE SURGEON ACCIDENTALLY TORE THE LENS WHILE ADJUSTING THE POSITION AFTER IMPLANTING IT INTO THE EYE AND REPLACED IT WITH A SPARE ONE ON THE SPOT. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.0/+1.5/072 | OS |  | 03/17/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9954795&pc=QCB | 2023826-2020-00887 | 2019-12-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.5/0.5/027 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE ON (B)(6) 2019. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY. ON (B)(6) 2019 A REPLACEMENT LENS WAS IMPLANTED. CAUSE OF EVENT UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/0.5/027 | NA |  | 03/17/2020 | 2020 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10025649&pc=QCB | 2023826-2020-01008 | 2019-03-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0/+3.5/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), IT TORE. IT WAS REMOVED INTRAOPERATIVELY AND THERE WAS NO PATIENT INJURY. THIS OCCURRED ON (B)(6) 2019. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+3.5/082 | OS |  | 04/08/2020 | 2020 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10025948&pc=QCB | 2023826-2020-01018 | 2020-03-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.00/+1.5/109 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENT'S RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS, ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.00/+1.5/109 | OD |  | 04/15/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10032660&pc=QCB | 2023826-2020-01023 | 2020-03-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-05-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -11.00/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH A SHORTER LENGTH LENS INTRAOPERATIVELY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN AND DEVICE. REPORTEDLY, 'WHEN LENS WAS HALF WAY INJECTED INTO AC, IT WAS NOTED THERE WAS A TEAR ON THE LENS, SO THE LENS WAS WITHDRAWN WITH NO CONTACT WITH ANTERIOR LENS CAPSULE OR CORNEAL ENDOTHELIUM." Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. DEVICE RETURNED WITH CLEAR SURGICAL RESIDUE IN A POUCH. VISUAL INSPECTION OF THE DEVICE FOUND NO VISIBLE DAMAGE AND FOREGIN RESIDUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.00/1.0/085 | OS |  | 04/16/2020 | 2020 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10075021&pc=QCB | 2023826-2020-01089 | 2020-04-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 7.00/+1.0/115 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO THE SURGEON NOTED THE LENS WAS TORN AFTER INSERTION INTO THE PATIENT'S RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL AND IS HAPPY. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | - 7.00/+1.0/115 | OD |  | 04/29/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10077243&pc=QCB | 2023826-2020-01128 | 2020-04-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.0/1.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED WITHIN THE SAME INITIAL SURGERY. THERE WAS NO PATIENT INJURY AND THE CAUSE OF THE EVENT WAS UNKNOWN. A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.0/1.5/084 | OS |  | 04/25/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10093696&pc=QCB | 2023826-2020-01081 | 2020-05-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2MM, -10.00/3.0/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED AND EXCHANGED DURING THE INITIAL SURGERY WITH SAME LENGTH LENS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN WITH DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/3.0/069 | OS |  | 05/07/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10094032&pc=QCB | 2023826-2020-01125 | 2018-01-30 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.00/+0.5/093 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENT'S RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS, ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.00/+0.5/093 | OD |  | 05/11/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10105251&pc=QCB | 2023826-2020-01080 | 2018-12-20 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 13.2MM VTICMO13.2, -18.00/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED INTRAOPERATIVELY, AND LATER REPLACED ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND OPTIC TORN/HAPTIC TORN. CLAIM# : (B)(4). Manufacturer Narrative: H10 IN INITIAL MDR PATIENT CODE 3191- SECONDARY SURGICAL INTERVENTION, LENS EXPLANT IS NOT APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.00/2.5/091 | OD |  | 05/11/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10116218&pc=QCB | 2023826-2020-01157 | 2020-03-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY. REPORTEDLY, THE LENS WAS STUCK IN THE CARTRIDGE/INJECTION SYSTEM. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AT A LATER DATE AND THE PROBLEM WAS RESOLVED. NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.0/+1.5/089 | OD |  | 05/11/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10125556&pc=QCB | 2023826-2020-01215 | 2020-05-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, TMICL12.1, -9.00/1.5/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN ( UNFOLDED AND FLIPPED) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED WITH THE BACKUP LENS WITHIN THE SAME SURGERY WITH NO PATIENT INJURY. PROBLEM RESOLVED, NO ISSUES, VISION DOING WELL. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.00/1.5/107 | OD |  | 05/08/2020 | 2020 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10158197&pc=QCB | 2023826-2020-01239 | 2020-04-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/131 (SPHERE/CYLINDER/AXIS), BUT THE PATIENTS RIGHT EYE (OD) MOVED. THE LENS WAS PARTIALLY IN THE EYE AND WAS REMOVED WITH NO PATIENT INJURY. THE LENS WAS EXCHANGED DURING THE SAME SURGERY FOR A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED DRY IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.0/131 | OD |  | 05/22/2020 | 2020 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10162510&pc=QCB | 2023826-2020-01293 | 2020-05-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT ON 18 MAY 2020 A 13.7MM, VTICMO13.7, -15.0/2.0/85 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) AND REMOVED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE, NO PATIENT INJURY. ON (B)(6) 27 A REPLACEMENT LENS WAS IMPLANTED AND IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED, PATIENT IS HAPPY. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.0/2.0/85 | OS |  | 05/18/2020 | 2020 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163020&pc=QCB | 2023826-2020-01251 | 2020-04-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED UPSIDE DOWN A 12.6MM VTICM5_12.6, -1.50/1.5/138 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS WAS REMOVED AND DURING REMOVAL OF THE LENS (INTRAOPERATIVELY) THE LENS TORE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. A NEW LENS WAS IMPLANTED ON (B)(6) 2020. REPORTEDLY, "PATIENT IS HAPPY." Manufacturer Narrative: ADDITIONAL INFORMATION. H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -1.50/1.5/138 | OS |  | 05/28/2020 | 2020 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10171665&pc=QCB | 2023826-2020-01269 | 2020-05-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Improper or Incorrect Procedure or Method; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2020. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY. NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. THE PATIENT DECLINED A REPLACEMENT LENS. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+2.0/089 | OD |  | 05/26/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10189490&pc=QCB | 2023826-2020-01290 | 2020-01-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+1.0/116 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENT'S RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY AND WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN, BENT AND DEFORMED WITH RESIDUE ON THE LENS. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.00/+1.0/116 | OD |  | 06/01/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10271255&pc=QCB | 2023826-2020-01478 | 2020-06-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 13.2MM VTICM5_13.2, -12.50/1.0/114 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. DURING THE INITIAL SURGERY ON (B)(6) 2020, A REPLACEMENT LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CORRECTED DATA: B5 - "DURING THE INITIAL SURGERY ON (B)(6) 2020, A REPLACEMENT LENS WAS IMPLANTED." SHOULD BE CORRECTED TO "ON (B)(6) 2020 A REPLACEMENT LENS WAS IMPLANTED". IT WAS NOT INDICATED IF THE REPLACEMENT OCCURRED INTRAOPERATIVELY AS REMOVAL AS THE INITIAL LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.50/1.0/114 | OD |  | 06/26/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10285741&pc=QCB | 2023826-2020-01484 | 2020-06-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Improper or Incorrect Procedure or Method; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/122 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY ON (B)(6) 2020. REPORTEDLY, THERE WAS NO PATIENT INJURY AND THE PATIENT DOES NOT WANT A REPLACEMENT LENS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN AND BENT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+1.0/122 | OS |  | 06/21/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10317973&pc=QCB | 2023826-2020-01532 | 2020-06-23 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-23 | QCB | IMPLANTABLE COLLAMER LENS | Improper or Incorrect Procedure or Method; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.5/+4.5/124 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), IT TORE. IT WAS REMOVED INTRAOPERATIVELY AND THERE WAS NO PATIENT INJURY. THIS OCCURRED ON (B)(6) 2020. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+4.5/124 | OS | Physician | 06/24/2020 | 2020 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10332960&pc=QCB | 2023826-2020-01573 | 2020-06-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -10.50/1.5/127 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.50/1.5/127 | OD |  | 07/07/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10334285&pc=QCB | 2023826-2020-01586 | 2020-01-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.0/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME INITIAL SURGERY AND THE PROBLEM WAS RESOLVED. THERE WAS NO PATIENT INJURY AND THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.0/1.5/088 | OD |  | 06/29/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401660&pc=QCB | 2023826-2020-01696 | 2020-06-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.00/1.0/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN INTO THE PATIENT'S LEFT EYE (OS) AND REMOVED ON (B)(6) 2020.THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. A SIMILAR LENS WAS IMPLANTED ON (B)(6) 2020, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CORRECTED DATA H6- INVESTIGATION FINDING IN PREVIOUS MDR IS CORRECTED FROM 114 TO 213. CLAIM#(B)(4). Manufacturer Narrative: CORRECTED DATA: B5 IN INITIAL MDR REMOVAL DATE (B)(6) 2020 IS CORRECTED TO (B)(6) 2020. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRA-OPERATIVELY (DURING THE INITIAL SURGERY) ON (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO [OR#] SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.00/1.0/107 | OS |  | 07/22/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403846&pc=QCB | 2023826-2020-01746 | 2020-06-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.5/+2.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS NOT IMPLANTED BUT THERE WAS PATIENT CONTACT WITH THE LENS. REPORTEDLY, THERE WAS NO PATIENT INJURY AND A REPLACEMENT LENS WAS SUCCESSFULLY IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.5/+2.5/087 | OS |  | 07/21/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403885&pc=QCB | 2023826-2020-01725 | 2020-07-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Improper or Incorrect Procedure or Method; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT, WHILE ATTEMPTING TO IMPLANT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/003 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE, IT TORE. THIS OCCURRED ON (B)(6) 2020. LENS WAS NOT IMPLANTED BUT THERE WAS PATIENT CONTACT; NO PATIENT INURY. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.5/003 | NA |  | 07/21/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10406728&pc=QCB | 2023826-2020-01754 | 2020-07-16 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, TMICL12.1, -16.00/3.50/093, IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME SURGERY. IT WAS REPORTED THAT THE PROBLEM RESOLVED, PATIENT IS HAPPY. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.00/3.50/093 | NA | Physician | 07/16/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10418290&pc=QCB | 2023826-2020-01782 | 2020-05-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2020 A 13.2MM, VTICMO13.2, -13.5/1.5/75 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) AND REMOVED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. ON (B)(6) 2020 A REPLACEMENT LENS WAS IMPLANTED AND IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/1.5/75 | OS |  | 07/18/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429483&pc=QCB | 2023826-2020-01799 | 2020-07-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2020, A 13.2MM, VTICM5_13.2, -11.5/1.0/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) . THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.5/1.0/078 | OS |  | 07/22/2020 | 2020 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10434291&pc=QCB | 2023826-2020-01806 | 2020-07-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2020 A 13.2MM, VTICM5_13.2, -9.0/0.5/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.0/0.5/102 | OD |  | 07/23/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10436172&pc=QCB | 2023826-2020-01773 | 2020-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.50/4.5/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED AND EXCHANGED DURING THE INITIAL SURGERY ON (B)(6) 2020 WITH A SAME SIZE, SIMILAR (SPHERE/CYLINDER/AXIS) LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.50/4.5/099 | OD |  | 07/30/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10495984&pc=QCB | 2023826-2020-02016 | 2020-08-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.5/2.0/81 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020 AND REMOVED WITHIN THE SAME SURGERY. THE REPORTER STATED " LENS FLIPPED UPSIDE DOWN IN THE EYE WHILE THE LENS UNFOLDED. HAPTIC BROKE WHILE TRYING TO REMOVE THE ICL." NO PATIENT INJURY AND USER ERROR WAS REPORTED. ON (B)(6) 2020 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CORRECTED DATA: D11 - CARTRIDGE MODEL-SFC-45; LOT#-UNK SHOULD BE REMOVED FROM INITIAL MEDWATCH REPORT AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | -12.5/2.0/81 | OS |  | 08/06/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9427123&pc=QCB | 2023826-2019-02267 | 2019-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.50/1.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. ON (B)(6) 2019, THE LENS WAS REMOVED AND A REPLACEMENT LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS, AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH RESIDUE ON THE LENS. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) WAS REPORTED IN ADDITIONAL MANUFACTURER NARRATIVE BUT SHOULD HAVE ALSO BEEN SELECTED AS A PATIENT CODE IN H6. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.50/1.0/086 | OD |  | 11/18/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9963615&pc=QCB | 2023826-2020-00891 | 2020-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 11.6MM, VTICH11.6, -10.00/2.5/177 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT BUT NO INJURY. IT WAS REPORTED THAT THE SURGEON REMOVED THE LENS WITHIN THE SAME SURGERY. A REPLACEMENT LENS WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. PATIENT IS HAPPY. CAUSE OF EVENT " LENS GOT STUCK IN THE PLUNGER." Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 11.6MM, VTICH11.6, -10.00/2.5/117 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION INTO THE EYE. THE LENS WAS LEFT IN THE EYE AND REMOVED AND REPLACED ON (B)(6) 2020 WITHIN THE SAME SURGERY AND THE PROBLEM RESOLVED, PATIENT IS HAPPY. THERE WAS NO PATIENT INJURY. CAUSE OD EVENT "LENS GOT STUCK IN THE PLUNGER". H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXCHANGE. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: "10.00/2.5/117 (SPHERE / CYLINDER / AXIS)" SHOULD BE CORRECTED TO "+10.00/2.5/117 (SPHERE / CYLINDER / AXIS)" IN SUPPLEMENTAL MDR #1. H1: "MALFUNCTION" SHOULD BE CORRECTED TO "SERIOUS INJURY" IN SUPPLEMENTAL MDR #2. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/2.5/177 | NA |  | 03/19/2020 | 2020 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10332961&pc=QCB | 2023826-2020-01574 | 2020-06-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.00/2.0/092 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.00/2.0/092 | OS |  | 07/07/2020 | 2020 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652169&pc=QCB | 2023826-2019-00967 | 2019-04-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient; No Known Impact Or Consequence To Patient | P030016 | Event Description: THE REPORTER INDICATED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 9.50/+1.5/090 (SPHERE/CYLINDER/AXIS), WAS TORN/BROKEN DUE TO A LOADING ERROR AND THERE WAS PATIENT CONTACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR, A NEW TECHNICIAN WAS LOADING. ANOTHER SAME MODEL,LENGTH,DIOPTER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED WRAPPED IN GAUZE IN A LENS CASE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | - 9.50/+1.5/090 | NA |  | 05/10/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9131632&pc=QCB | 2023826-2019-01793 | 2019-07-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; No Apparent Adverse Event | No Consequences Or Impact To Patient; No Known Impact Or Consequence To Patient | P030016 | Event Description: RECEIVED A DAMAGED 13.2MM , TMICL13.2, -11.0/2.0/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WITH NO REPORTED PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA DEVICE CODE 2923: LENS ROTATION. (B)(4). Manufacturer Narrative: CORRECTED DATA: EVENT: LENS REPOSITIONING WAS ALSO ATTEMPTED. DEVICE CODE: 2923 LENS ROTATION IN PREVIOUS MDR IS NOT APPLICABLE. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -11.00/2.0/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME SURGERY DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. NO PATIENT INJURY WAS REPORTED. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS USER ERROR. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: LENS WAS RETURNED DRY ,IN CARTRIDGE CASE. VISUAL INSPECTION FOUND RESIDUE ON LENS SURFACE AND A TORN HAPTIC. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/2.0/068 | OS |  | 08/28/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9668837&pc=QCB | 2023826-2020-00242 | 2019-12-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Insufficient Information | No Information | N/A | Event Description: THE SURGEON REPORTED A 13.2MM VTICM5_13.2, -9.5/+1.5/102 (SPHERE/CYLINDER/AXIS) AND THE PATIENT'S RIGHT EYE (OD) ON A COMPLAINT QUESTIONNAIRE. IT IS UNKNOWN AS TO WHETHER THE LENS WAS IMPLANTED AND/OR REMOVED. A SHORTER REPLACEMENT LENS IS ALSO REPORTED. IT IS UNKNOWN AS TO WHETHER THIS LENS WAS IMPLANTED AND THE REASON FOR THE COMPLAINT IS ALSO NOT REPORTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 13.2MM VTICM5_13.2, -9.5/+1.5/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019 INTRAOPERATIVELY. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATES THAT THE "PATIENT IS HAPPY!" THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1- "ADVERSE EVENT" AHOULD BE CORRECTED TO "PRODUCT PROBLEM" IN THE PREVIOUS MDR H1- "SERIOUS INJURY" SHOULD BE CORRECTED TO "MALFUNCTION" IN THE PREVIOUS MDR. ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH RESIDUE AND MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/+1.5/102 | OD |  | 01/14/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7765074&pc=QCB | 2023826-2018-01176 | 2018-06-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT 12.6MM, VTICM5_12.6, -10.00/+1.0/80 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE EYE IS PRESENTING VERY LOW VAULT. THE VAULT IS REALLY LOW AND THE SPACE BETWEEN THE IRIS AND THE ICL IS ALSO LOW. LENS REMAINS IMPLANTED. IT WAS REPORTED THAT THE EYE WILL BE RE-EVALUATED WITHIN THE NEXT WEEKS. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: REPORTER STATED " THE VAULTS ARE VERY LOW. ACUITIES GOOD AND STABLE. LENS IS OKAY. DOCTOR IS CONCERNED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CORRECTED: A1 - PATIENT IDENTIFIER: (B)(6). D2- COMMON DEVICE NAME: FROM PHAKIC INTRAOCULAR LENS, PRODUCT CODE MTA TO PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE QCB. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. "WORD" ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.00/+1.0/80 | OD |  | 07/09/2018 | 2018 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8254867&pc=QCB | 2023826-2019-00093 | 2018-10-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+3.5/180 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON NOTES EXCESSIVE VAULTING. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: DIOPTER CORRECTED TO " -16.50/+3.5/108 (SPHERE/CYLINDER/AXIS) " COMMON DEVICE NAME: CORRECTED TO: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.; (B)(4). | -16.5/+3.5/180 | OD |  | 12/20/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8256516&pc=QCB | 2023826-2019-00038 | 2018-09-10 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT ON 10/SEP/2018 A VTICMO 12.1, -8.50/2.0/064 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE EYE. LENS WAS REMOVED INTRAOPERATIVELY. A REPLACEMENT LENS WAS LATER IMPLANTED ON (B)(6) 2018. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: CORRECTED DATA: CORRECTED TO: "02-JAN-2019" IN THE INITIAL MDR. CORRECTED TO: "02-JAN-2019" IN THE INITIAL MDR. CORRECTED TO: "PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB" (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.50/2.0/064 | NA | Physician | 01/02/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372538&pc=QCB | 2023826-2019-00329 | 2018-11-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON ATTEMPTED TO IMPLANT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/+3.0/064 (SPHERE/CYLINDAR/AXIS) INTO THE PATIENT'S LEFT EYE (OS), HE DISCOVERED IT WAS TORN/BROKEN. REPORTEDLY, THE LENS TORE DURING LOADING AND THERE WAS PATIENT CONTACT WITH THE LENS HOWEVER, IT WAS NOT FULLY IMPLANTED. THIS OCCURRED ON (B)(6) 2018. INTRAOPERATIVELY AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PREVIOUSLY STATED THE LENS TORE DURING LOADING. UPDATED INFORMATION RECEIVED: THE LENS TORE DURING INJECTION. IN ADDITION, THE SURGEON BELIEVES THE CAUSE OF THE EVENT IS DUE TO A "DEFECT IN THE FOAM TIP." NO IDENTIFYING INFORMATION IS AVAILABLE FOR THE FOAM TIP AS IT WAS DISCARDED. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN ALONG WITH A PIECE OF THE LENS MISSING. COMMON DEVICE NAME SHOULD BE CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.0/+3.0/064 | OS |  | 02/01/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8422982&pc=QCB | 2023826-2019-00408 | 2018-08-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+3.0/101 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO LOW VAULTING. THERE WAS NO REPORT OF ANY INJURY. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGCIAL RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+3.0/101 | NA |  | 02/12/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8422985&pc=QCB | 2023826-2019-00409 | 2018-12-06 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/084 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THERE WAS NO REPORT OF ANY INJURY. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: DEVICE EVALUATION: "HAPTIC TORN AND DEFORMED" SHOULD BE CORRECTED TO "HAPTIC BENT AND DEFORMED" IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DEFORMED WITH FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+2.5/084 | NA |  | 02/12/2019 | 2019 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8453004&pc=QCB | 2023826-2019-00490 | 2018-09-18 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-26 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT, WHILE LOADING A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/072 (SPHERE/CYLINDAR/AXIS) INTO THE CARTRIDGE, IT TORE/BROKE. THIS WAS MEANT TO BE IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THERE WAS NO PATIENT CONTACT WITH THE LENS AND THIS OCCURRED ON (B)(6) 2018. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. ON (B)(6) 2018 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.0/072 | OS | Physician | 09/20/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8475388&pc=QCB | 2023826-2019-00525 | 2019-03-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+4.0/091 (SPHERE/CYLINDER/AXIS), FROM THE LENS VIAL, HE REALIZED IT WAS TORN. THIS OCCURRED ON (B)(6) 2019 PRIOR TO THE LOADING PROCESS. THE LENS WAS SCHEDULED TO BE IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+4.0/091 | OD |  | 03/08/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8557287&pc=QCB | 2023826-2019-00705 | 2019-03-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: IT WAS REPORTED THAT A 13.2MM, VTICM5_13.2, -8.50/1.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT LEFT EYE (OS) AS A REPLACEMENT LENS DUE TO LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: RECEIVED VIA EMAIL ON (B)(6) 2019 "INDEED THIS WAS A MISUNDERSTANDING. PATIENT HAPPY AFTER EXCHANGE. VISION 10/10 -0.25(-0.5X155)." "EXCHANGE DID NOT RESOLVE THE PROBLEM" IN ORIGINAL MDR IS NOT APPLICABLE. (B)(4). Manufacturer Narrative: DATE OF EVENT IN ORIGINAL MDR IS NOT APPLICABLE. IN PREVIOUS MDR CORRECTED TO ADDITIONAL INFORMATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.50/1.5/081 | OS |  | 03/28/2019 | 2019 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8563074&pc=QCB | 2023826-2019-00709 | 2019-03-18 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-29 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS LOADING A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -11.0/+4.0/092 (SPHERE/CYLINDER/AXIS) INTO THE INJECTOR, THE HAPTIC BROKE. "NO INJURY TO PATIENT" IS NOTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+4.0/092 | NA | Physician | 04/05/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8564639&pc=QCB | 2023826-2019-00686 | 2019-02-25 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/2.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM.. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/2.0/081 | OS |  | 04/17/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8564646&pc=QCB | 2023826-2019-00689 | 2018-10-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -7.00/1.0/077 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: CORRECTION - SERIOUS INJURY SHOULD BE CORRECTED TO MALFUNCTION IN ORIGINAL MEDWATCH REPORT. DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -7.00/1.0/077 | OS |  | 04/17/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8564647&pc=QCB | 2023826-2019-00690 | 2018-10-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -7.50/3.0/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED FOR A LONGER LENGTH LENS DURING INITIAL SURGERY DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PRODUCT EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.50/3.0/110 | OS |  | 04/17/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8564649&pc=QCB | 2023826-2019-00688 | 2018-10-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/1.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED FOR A LONGER LENGTH LENS DURING INITIAL SURGERY DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PRODUCT EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/1.5/085 | OS |  | 04/17/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8570979&pc=QCB | 2023826-2019-00595 | 2018-12-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A VTICMO13.2, -18.00/3.0/093 (SPHERE/CYLINDER/AXIS), INTRAOCULAR LENS WAS TORN BEFORE LOADING, AFTER OPENING VIAL. BACKUP LENS WAS USED, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18.00/3.0/093 | NA |  | 04/02/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8573314&pc=QCB | 2023826-2019-00721 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -14.5/4.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE AND THEN NOTICED THERE WAS NO ASTIGMATISM MARKER LINE ON LENS. IT WAS REPORTER THAT THE PATIENT IS IN GOOD STATUS AS OF NOW AND THE DOCTOR SUGGESTED NOT TO EXPLANT/EXCHANGE THE LENS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.5/4.0/085 | NA |  | 04/03/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8609189&pc=QCB | 2023826-2019-00827 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-05-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -9.5/+1.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. THE SURGEON STATES THIS IS AN UNEXPECTED CASE OF HIGH VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION OR UPDATES HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/+1.0/095 | OS | Physician | 04/19/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8633176&pc=QCB | 2023826-2019-00881 | 2019-04-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT, WHILE LOADING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+2.5/084 (SPHERE/CYLINDAR/AXIS) INTO THE CARTRIDGE, THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2019. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC WAS TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.0/+2.5/084 | NA |  | 04/26/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8660761&pc=QCB | 2023826-2019-00969 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-05-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Difficult to Fold, Unfold or Collapse | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.00/+2.0/086 (SPHERE/CYLINDER/AXIS), WAS NOT EASY TO FOLD. THE SURGEON TRIED TWO TIMES BUT LENS WAS RATHER ROUGH. THE BACKUP LENS WAS USED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.00/+2.0/086 | NA | Other Health Care Professional | 05/06/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8663478&pc=QCB | 2023826-2019-00988 | 2019-02-19 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.5/1.0/79 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO OBSERVATION OF EXCESSIVE VAULT. ON (B)(6) 2019 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/1.0/79 | OD |  | 05/06/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8692894&pc=QCB | 2023826-2019-01068 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-06-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+3.5/102 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON STATES THERE IS A "LITTLE BIT OF VAULTING" HOWEVER IT SEEMS TO BE RESOLVING ON ITS OWN. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -18.0/+3.5/102 | OD | Other Health Care Professional | 05/17/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8699720&pc=QCB | 2023826-2019-01051 | 2019-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Unintended Movement | No Known Impact Or Consequence To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT THE HAPTICS OF A TMICL12.6, -11.50/1.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS STUCK TOGETHER AND UNFOLDED UPSIDE DOWN UPON INSERTION INTO PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS TORE DURING REMOVAL, BACKUP LENS WAS IMPLANTED DURING INITIAL SURGERY AND RESOLVED PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# : (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/1.5/085 | OS |  | 05/23/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8716003&pc=QCB | 2023826-2019-01087 | 2019-05-28 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -18.00/2.0/130 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED DURING INITIAL DUE TO EXCESSIVE VAULT. PER REPORTER, A NEW LENS HAS BEEN IMPLANTED AND PER SURGEONS COMMENT "PERFECT". CAUSE OF THE EVENT IS REPORTED AS "SIZING PROBLEM". Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -18.00/2.0/130 | OD | Physician | 06/03/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8781312&pc=QCB | 2023826-2019-01200 | 2019-03-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/1.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT ID: 1351770167 SHOULD BE CORRECTED TO (B)(6). DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/1.5/087 | OD |  | 07/02/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8806866&pc=QCB | 2023826-2019-01202 | 2018-10-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 A VTICMO12.6, -5.50/0.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, ON (B)(6) 2018. A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2018. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.50/0.5/087 | NA |  | 07/02/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8806867&pc=QCB | 2023826-2019-01214 | 2019-03-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/1.5/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. PATIENT ID DOES NOT FIT INTO ESUBMITTER FIELD. CORRECTED TO (B)(6). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/1.5/103 | OD |  | 07/02/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8806870&pc=QCB | 2023826-2019-01203 | 2019-06-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.50/1.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. AFTER INSERTION, THE LENS "FLIPPED IN THE EYE." INTRAOPERATIVELY, THE LENS WAS REMOVED AND SUCCESSFULLY EXCHANGED WITH AN ALTERNATE LENS. REPORTEDLY, NO PATIENT INJURY OCCURRED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.50/1.0/093 | OD |  | 07/03/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8848655&pc=QCB | 2023826-2019-01322 | 2019-04-01 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -4.50/+1.5/147 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINS IMPLANTED. THE REPORTER INDICATED THE SURGEON HAS DECIDED TO MONITOR THE PATIENT CLOSELY AND WILL NOT EXCHANGE THE LENS RIGHT NOW. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -4.50/+1.5/147 | OD | Physician | 07/03/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8871976&pc=QCB | 2023826-2019-01396 | 2019-04-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.5/091 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS DURING THE SAME SURGERY BUT THE PROBLEM WAS NOT RESOLVED. SEE MFR. REPORT # 2023826-2019-01397 FOR THE EXCHANGED LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE,VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF SURGICAL RESIDUE,DEBRIS ON THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+2.5/091 | NA |  | 07/09/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8894308&pc=QCB | 2023826-2019-01443 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-08-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/047 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS HE WANTS TO EXCHANGE THE LENS FOR A SHORTER ONE. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: AGE OR DATE OF BIRTH, SEX, WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.0/+1.5/047 | OS | Physician | 06/27/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8894312&pc=QCB | 2023826-2019-01442 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-08-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/115 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS HE WANTS TO EXCHANGE THE LENS FOR A SHORTER ONE. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: AGE, SEX, WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT:- UNK. IMPLANTED: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.0/+1.5/115 | OD | Physician | 06/27/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8910116&pc=QCB | 2023826-2019-01466 | 2019-07-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -10.00/1.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE BEFORE/DURING LOADING. PER THE REPORTER ON (B)(6) 2019, "THE LENS WAS NOT IMPLANTED AND HAD ONLY BEEN TOUCHED USING A WET PLUNGER PRIOR TO NOTICING THE TEAR. A SMALL TEAR WAS DIFFICULT TO SEE IN OR" NO PATIENT INJURY. BACK-UP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A CARTRIDGE CASE, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: CORRECTED DATA: MSIPF, SFC45, AND FTP INFORMATION SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/1.0/105 | NA |  | 07/29/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8912040&pc=QCB | 2023826-2019-01520 | 2019-03-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO_12.6, -13.00/3.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO OBSERVATION OF EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/3.5/088 | OS |  | 07/22/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8930916&pc=QCB | 2023826-2019-01499 | 2019-06-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.00/1.0/057 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTER INDICATED ON (B)(6) 2019, "THE SURGEON PLANS TO REMOVE AND EXCHANGE LENS IN EARLY (B)(6)." WITH A LONGER LENGTH LENS. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: "AFTER A FOLLOW UP WITH THE PATIENT, THE SURGEON REPORTED THAT THE PATIENT HAD GOOD VAULT THEREFORE THE LENS WILL NOT BE EXCHANGED." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.00/1.0/057 | OD |  | 08/05/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8932862&pc=QCB | 2023826-2019-01546 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/072 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS: UNK. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+1.0/072 | OD |  | 07/29/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8957668&pc=QCB | 2023826-2019-01532 | 2019-08-05 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -14.00/2.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. PER THE REPORTER, "SURGEON STILL MONITORING EYE." PER REPORTER ON (B)(6) 2019, "AFTER REVIEWING, THE PUPIL IS REACTIVE, ANGLE WAS MORE THAN 20, IOP WAS 15. ALL FINDINGS SHOW (B)(6) RESULTS. SURGEON HAS DECIDED NOT TO REMOVE THE LENS. LENS REMAINS IMPLANTED. "HENCE, CONCERNS RESOLVED." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.00/2.5/095 | OS | Physician | 08/07/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966520&pc=QCB | 2023826-2019-01629 | 2019-08-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Unintended Movement | No Known Impact Or Consequence To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. THE SURGEON REPORTS THE LENS FLIPPED UPSIDE DOWN IN THE ANTERIOR CHAMBER OF THE EYE IMMEDIATELY AFTER INJECTION. INTRAOPERATIVELY, THE SURGEON REMOVED THE LENS AND SUCCESSFULLY IMPLANTED A BACK-UP LENS WITHOUT INCIDENT. THE EYE IS REPORTEDLY BLURRY AFTER SURGERY BUT THIS WAS EXPECTED TO CLEAR. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. (B)(4) (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -9.0/+1.5/089 | OS |  | 08/09/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966525&pc=QCB | 2023826-2019-01632 | 2019-06-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+2.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.5/+2.0/093 | OS |  | 08/08/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9031793&pc=QCB | 2023826-2019-01663 | 2019-06-29 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.00/2.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. PER THE REPORTER, "THE VAULT DECREASED AFTER ONE WEEK" BUT SURGEON PREFERS "TO EXCHANGE LENS FOR SHORTER LENGTH LENS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/2.5/094 | OD | Physician | 08/19/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9031824&pc=QCB | 2023826-2019-01673 | 2019-07-26 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2019 A 12.6 MM, VTICMO12.6, -18.00/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING LOADING. NO PATIENT CONTACT. REPLACEMENT LENS IMPLANTED ON 14/AUG/2019, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -18.00/2.0/094 | NA | Physician | 08/23/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9050221&pc=QCB | 2023826-2019-01684 | 2019-07-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.00/+2.50/094 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.00/+2.50/094 | NA |  | 08/21/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9072242&pc=QCB | 2023826-2019-01658 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-09-16 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | Event Description: IT WAS REPORTED THAT A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -12.00/+3.5/076 (SPHERE/CYLINDER/AXIS) IS SCHEDULED TO BE EXCHANGED WITH A SHORTER LENGTH LENS. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: AGE, SEX, WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/+3.5/076 | NA | Physician | 08/26/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9071745&pc=QCB | 2023826-2019-01659 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | Event Description: IT WAS REPORTED THAT THE SURGEON PLANS ON EXCHANGING A 13.2 MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -12.50/+3.5/093 (SPHERE/CYLINDER/AXIS) FOR A SHORTER LENGTH LENS. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: UNK. AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.50/+3.5/093 | NA |  | 08/26/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9115517&pc=QCB | 2023826-2019-01207 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-09-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2 MM, VTICMO13.2, -16.50/6.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IMPLANTED INTO PATIENT'S RIGHT EYE WAS "NOT ACHIEVING THE DESIRED RESULT." REPORTEDLY THIS WAS DUE TO PATIENTS "BIG ASTIGMATISM." THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: AGE/DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANT DATE: (B)(6) 2019. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -16.50/6.0/093 | OD | Non-Healthcare Professional | 09/02/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9115533&pc=QCB | 2023826-2019-01778 | 2019-09-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT ON 05/SEP/2019 A VTICM5_12.1, -02.25/3.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, ON (B)(6) 2019. A REPLACEMENT LENS WAS IMPLANTED ON 18/SEP/2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS INJECTOR OR LENS. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT ON (B)(6) 2019 THE LENS TORE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED INTRA-OPERATIVELY (WITHIN THE SAME SURGERY) ON (B)(6) 2019. PRE-EXISTING CONDITIONS: CATARACT; RESIDUAL CORNEAL ASTIGMATISM IMPLANTED DATE: (B)(6) 2019 IS CORRECTED TO (B)(6) 2019. EXPLANT DATE: (B)(6) 2019 IS CORRECTED TO (B)(6) 2019. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: INJECTOR MODEL; INDIGO; LOT # UNK SFC-45 OR SFC-25; LOT # UNK. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION OF THE LENS FOUND THE OPTIC TORN, HAPTIC TORN, RESIDUE ON THE LENS AND A MISSING PIECE OF LENS IS STUCK IN CARTRIDGE. THE CARTRIDGE WAS RETURNED IN A BIOHAZARD BAG WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION OF THE CARTRIDGE FOUND NO VISIBLE DAMAGE TO DEVICE, PLUNGER OVERRIDDEN, RESIDUE ON THE LENS AND PORTION OF THE LENS STUCK IN CARTRIDGE. (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. WEIGHT: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -02.25/3.0/074 | NA |  | 09/04/2019 | 2019 | 77
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9118464&pc=QCB | 2023826-2019-01782 | 2019-08-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6 MM, VTICM5_12.6, -12.5/2.00/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS REMOVED WITHIN THE SAME SURGERY DUE TO LOW VAULT WITH ROTATION. ON (B)(6) 2019 A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED BUT IT DID NOT RESOLVE THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: DEVICE CODE 2923 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.5/2.00/093 | OD |  | 08/30/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9181367&pc=QCB | 2023826-2019-01945 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-10-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -10.50/1.0/129 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE WHILE LOADING. REPORTER STATED " UNSURE IF PATIENT CONTACT". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. EXPLANTED: NA. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.50/1.0/129 | NA | Physician | 09/16/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9186723&pc=QCB | 2023826-2019-01879 | 2019-05-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/1.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6)2019. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT . IT WAS REPORTER THAT THE PROBLEM RESOLVED AFTER THE LENS REMOVAL, PATIENTS CONDITION IS GOOD. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN LIQUID INA LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/1.0/093 | OS |  | 09/10/2019 | 2019 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9188655&pc=QCB | 2023826-2019-01947 | 2019-07-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED, THE REPORTER NOTING, "DUE TO THE STABLE CONDITION, THE SURGEON DECIDED TO MAKE CLOSE FOLLOW-UP." THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS CLAIM WAS REPORTED AS A DUPLICATE; REFERENCE MDR# 2023826-2019-01777 FOR THE ORIGINAL COMPLAINT EVENT. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+1.5/091 | OS |  | 09/19/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9252491&pc=QCB | 2023826-2019-02038 | 2019-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICM5_ 12.6, -9.50/1.50/093 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A LONGER LENGTH LENS WITHIN THE SAME SURGERY DUE TO LOW VAULT. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM, ALL IS FINE. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.50/093 | OS |  | 10/10/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9284947&pc=QCB | 2023826-2019-02054 | 2019-06-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -8.5/2.0/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/2.0/071 | OS |  | 10/21/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9293583&pc=QCB | 2023826-2019-02101 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-11-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | Event Description: A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -1.50/+3.0/180 (SPHERE/CYLINDER/AXIS) WAS RETURNED TO THE MANUFACTURER DAMAGED. ADDITIONAL INFORMATION WAS REQUESTED PERTAINING TO THE LENS DAMAGE BUT NO MORE INFORMATION WAS PROVIDED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. EXPLANTED: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND A PIECE OF ONE HAPTIC TORN OFF. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -1.50/+3.0/180 | NA | Physician | 10/17/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9292262&pc=QCB | 2023826-2019-01977 | 2019-03-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A VTICM5_12.6, -13.00/2.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE BEFORE/DURING LOADING ON (B)(6) 2019. DAMAGE WAS NOTED WHILE PULLING THE LENS INTO CARTRIDGE TIP. REPLACEMENT LENS WAS SUCCESSFULLY IMPLANTED. "THE PATIENT CONDITION NOW IS STABLE, NO PROBLEMS." CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN LIQUID IN LENS VIAL IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL OBSERVATION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.00/2.0/081 | NA |  | 10/16/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317883&pc=QCB | 2023826-2019-02133 | 2019-04-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2019. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+2.0/093 | OS |  | 10/18/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9319128&pc=QCB | 2023826-2019-02165 | 2019-10-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Difficult to Fold, Unfold or Collapse; Adverse Event Without Identified Device or Use Problem | No Known Impact Or Consequence To Patient | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 12.00/+1.5/087 (SPHERE/CYLINDER/AXIS), DUE TO THE LENS UNFOLDED INCORRECTLY IN THE EYE. THE LENS WAS REPLACED WITH ANOTHER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: B5 - PER UPDATED INFORMATION, IT WAS REPORTED THAT THE LENS TORE DURING LOADING. THERE WAS NO PATIENT CONTACT. THE CAUSE WAS REPORTED AS USER ERROR. (B)(4). Manufacturer Narrative: SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | - 12.00/+1.5/087 | NA |  | 10/17/2019 | 2019 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9336767&pc=QCB | 2023826-2019-02175 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.5/+3.5/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2019. THE SURGEON REPORTS EXCESSIVE VAULT, BUT THE PATIENT IS ASYMPTOMATIC. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: H6 - RESULT CODE: 114 SHOULD BE CORRECTED BACK TO 213, AS 114 IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - RESULT CODE: 213 SHOULD BE CORRECTED TO 114. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+3.5/089 | OD |  | 10/22/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9351287&pc=QCB | 2023826-2019-02183 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-11-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | Event Description: THE REPORTER REPORTED A LOADING ISSUE WITH A 13.2MM, TMICL13.2, -7.00/2.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | -7.00/2.0/094 | NA | Other Health Care Professional | 10/25/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9361667&pc=QCB | 2023826-2019-02154 | 2019-07-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.00/2.5/093 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B3- "(B)(6) 2019" SHOULD BE CORRECTED TO "(B)(6) 2019" IN THE INITIAL MDR. B4- "26-JUL-2019" SHOULD BE CORRECTED TO "04-NOV-2019" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/2.5/093 | OD |  | 11/04/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9412441&pc=QCB | 2023826-2019-02284 | 2019-07-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+2.0/085 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2019. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -8.5/+2.0/085 | OS |  | 11/11/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9412449&pc=QCB | 2023826-2019-02280 | 2019-08-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -18.0/+2.0/086 | OS |  | 11/11/2019 | 2019 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9441182&pc=QCB | 2023826-2019-02352 | 2019-11-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT AFTER THE SURGEON REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/092 FROM THE VIAL, IT WAS DISCOVERED THAT THE LENS WAS TORN/BROKEN. THIS OCCURRED ON (B)(6) 2019. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON A LATER DATE AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K) -THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.0/+2.0/092 | NA |  | 11/14/2019 | 2019 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9454477&pc=QCB | 2023826-2019-02367 | 2019-08-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -4.5/+4.0/094 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -4.5/+4.0/094 | NA |  | 11/11/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9460094&pc=QCB | 2023826-2019-02363 | 2019-11-22 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -13.50/1.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2019, SURGEON DECIDED TO REPLACE LENS WITH SMALLER ONE." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - 'PER THE REPORTER ON (B)(6) 2019, SURGEON DECIDED TO REPLACE LENS WITH SMALLER ONE.' CORRECTED TO 'PER THE REPORTER ON (B)(6) 2019, SURGEON DECIDED TO REPLACE LENS WITH SMALLER ONE.' CLAIM# (B)(4). Manufacturer Narrative: B5 - PER UPDATED INFORMATION, THE VAULT HAS RETURNED TO NORMAL AND THE LENS WILL NOT BE EXCHANGED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.50/1.0/079 | OS |  | 11/22/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9483559&pc=QCB | 2023826-2019-02394 | 2019-11-06 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.50/1.0/100 , (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2019. EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/1.0/100 | OS | Physician | 12/02/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9516551&pc=QCB | 2023826-2019-02449 | 2019-11-26 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6, -11.50/1.0/067 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. REPORTEDLY, "EXCESSIVE VAULT AS PUPIL COULD NOT BE BROUGHT DOWN. LENS EXPLANTED REPLACED WITH 12.1MM LENS DURING SAME SURGERY. NO POST PROCEDURE COMPLAINTS." Manufacturer Narrative: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). | -11.50/1.0/067 | OD | Nurse | 12/04/2019 | 2019 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9516547&pc=QCB | 2023826-2019-02492 | 2019-11-11 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5 12.6, -10.00/2.5/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE BEFORE/DURING LOADING ON (B)(6) 2019. DAMAGE WAS NOTED WHILE REMOVING FROM VIAL. REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. REPORTEDLY, LENS "CHIP WHILE DOING SURGERY." PER THE REPORTER ON (B)(6) 2019, "IT IS A MATTER OF CHANCE THAT SUCH AN ISSUE HAPPENED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE WITH FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/2.5/071 | NA |  | 12/03/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9533649&pc=QCB | 2023826-2019-02516 | 2019-11-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -15.50/4.0/121 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. IT WAS REPORTED THAT EXCESSIVE VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS REPORTED THE PATIENT ALSO EXPERIENCED LENS ROTATION. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - DEVICE CODE: 2923 SHOULD HAVE BEEN INCLUDED IN INITAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.50/4.0/121 | OD |  | 12/02/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9563657&pc=QCB | 2023826-2020-00059 | 2019-10-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT 13.2MM, TMICL13.2, -10.00/1.0/107 (SPHERE/CYLINDER AXIS) IMPLANTABLE COLLAMER LENS TORE DURING LOADING INTO INJECTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5- "THE REPORTER INDICATED THAT A 13.2MM TMICL 13.2, -10.00/1.0/107 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE DURING LOADING INTO THE INJECTOR FOR THE LEFT EYE (OS) ON (B)(6) 2019. THERE WAS NO PATIENT CONTACT. A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DUE TO LOADING ERROR." SHOULD BE IN THE INITIAL MDR. H3- DEVICE EVALUATION: "LIWUID" SHOULD BE CORRECTED TO "LIQUID" IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIWUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#(B)(4). Manufacturer Narrative: ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/1.0/107 | OS |  | 12/10/2019 | 2019 | 71
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9582438&pc=QCB | 2023826-2020-00073 | 2019-08-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/1.5/075 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#:(B)(4). Manufacturer Narrative: CORRECTED DATA: G4 - (B)(6) 2019 SHOULD BE CORRECTED TO (B)(6).2019. B4 - (B)(6).2019 SHOULD BE CORRECTED TO (B)(6).2019. B5 - THE LENS WAS REMOVAL AND REPLACEMENT OCCURRED DURING THE SAME SURGERY. H6 - DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -11.5/1.5/075 | OS |  | 12/18/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9583119&pc=QCB | 2023826-2020-00074 | 2019-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.0/1.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC DEFORMED WITH FOREIGN RESIDUE ON THE LENS. CLAIM# : (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- "06-JUN-2019" SHOULD BE CORRECTED TO "03-JUN-2019" IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B4- (B)(6) 2019" SHOULD BE CORRECTED TO "(B)(6) 2019" IN THE INITIAL MDR. G4- (B)(6)2019" SHOULD BE CORRECTED TO "(B)(6) 2019" IN THE INITIAL MDR. CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -10.0/1.0/086 | OS |  | 12/18/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9587682&pc=QCB | 2023826-2020-00122 | 2019-11-20 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.0/1.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -7.0/1.0/089 | OD |  | 12/20/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9608210&pc=QCB | 2023826-2020-00163 | 2019-12-23 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -08.50/1.0/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2019, "THE DOCTOR WANTS TO EXCHANGE LENSES." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -08.50/1.0/080 | OS | Physician | 01/02/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9608211&pc=QCB | 2023826-2020-00164 | 2019-12-20 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -09.00/1.5/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2019, "THE DOCTOR WANTS TO EXCHANGE LENSES." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -09.00/1.5/101 | OD | Physician | 01/02/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9614042&pc=QCB | 2023826-2020-00133 | 2019-12-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -0.50/+3.0/165 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. A SHORTER LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -0.50/+3.0/165 | NA |  | 12/24/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9618777&pc=QCB | 2023826-2020-00134 | 2019-12-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+4.0/073 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.00/+4.0/073 | NA |  | 12/24/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9656158&pc=QCB | 2023826-2020-00211 | 2019-10-23 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, - 7.50/+1.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS EYE. THE REPORTER INDICATED THE LENS WILL BE EXPLANTED DUE TO TOO MUCH VAULTING. THE LENS WILL BE EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | - 7.50/+1.0/092 | NA | Other Health Care Professional | 01/13/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9696241&pc=QCB | 2023826-2020-00224 | 2020-01-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient |  | Event Description: THE REPORTER INDICATED THAT A TORIC ICL 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -14.5/2.0/094 (SPHERE/CYLINDER/AXIS) TORE IN THE INJECTOR DURING SURGERY ON THE PATIENT'S RIGHT EYE (OD). THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY ON (B)(6) 2020. ON A LATER DATE AN ALTERNATE LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN AS WELL AS RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.5/2.0/094 | OD |  | 01/16/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9705873&pc=QCB | 2023826-2020-00226 | 2018-08-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/1.5/087 (SPHERE/CYLINDER/AXIS) TORE/BROKE BEFORE/DURING LOADING AFTER OPENING THE VIAL FOR THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON (B)(6) 2018. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.0/1.5/087 | OS |  | 01/14/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9717757&pc=QCB | 2023826-2020-00293 | 2020-01-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -16.00/0.5/031 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A LONGER LENGTH LENS DURING INITIAL SURGERY DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY "PATIENT IS HAPPY." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON THE LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: H3 - "FOREIGN MATERIAL ON THE LENS SURFACE" SHOULD BE REMOVED AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.00/0.5/031 | OS |  | 01/31/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9721912&pc=QCB | 2023826-2020-00328 | 2019-12-09 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-18 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | Event Description: A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -9.0/+2.0/089 (SPHERE/CYLINDER/AXIS) WAS RETURNED TO STAAR SURGICAL. THE WORDS "UPSIDE DOWN" AND "DID NOT USE-OD" WERE WRITTEN ON THE BOX. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: ON (B)(6) 2019 THE TMICL LENS WAS IMPLANTED UPSIDE DOWN. IT WAS REMOVED AND REPLACED SUCCESSFULLY WITH A BACK-UP LENS. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED ALL WITHIN THE SAME SURGERY. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | -9.0/+2.0/089 | NA | Other Health Care Professional | 01/28/2020 | 2020 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9782978&pc=QCB | 2023826-2020-00410 | 2020-01-31 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -1.0/+1.0/005 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -1.0/+1.0/005 | OS | Physician | 02/07/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781525&pc=QCB | 2023826-2020-00339 | 2020-02-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A VTICM5_12.6, -09.50/0.5/022 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE BEFORE/DURING LOADING. DAMAGE WAS NOTED IN CARTRIDGE. "SURGERY RESCHEDULED." CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -09.50/0.5/022 | NA |  | 02/11/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781714&pc=QCB | 2023826-2020-00420 | 2019-10-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/078 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2019. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# : (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510K - THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -8.5/+1.0/078 | OD |  | 02/17/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9789273&pc=QCB | 2023826-2020-00412 | 2020-01-28 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020 AS A REPLACEMENT LENS - SEE MFR. REPORT 2023826-2020-00411 FOR INITIAL LENS. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.0/+2.0/091 | OS | Physician | 02/09/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9789280&pc=QCB | 2023826-2020-00413 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-03-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.0/+2.0/096 | OD | Physician | 02/09/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824533&pc=QCB | 2023826-2020-00358 | 2019-08-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9..50/3.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H10: ADDITIONAL NARRATIVES/DATA: "H6: PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE)" SHOULD BE REMOVED FROM INITIAL MEDWATCH REPORT AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9..50/3.0/095 | OS |  | 02/18/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824542&pc=QCB | 2023826-2020-00360 | 2019-10-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/1.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.50/1.5/080 | OS |  | 02/18/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824550&pc=QCB | 2023826-2020-00362 | 2019-11-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/2.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "THE DENTAL ARCH HEIGHT WAS SIGNIFICANTLY HIGHER AFTER IMPLANTATION, AND THE ANTERIOR CHAMBER WAS SHALLOW. TO BE ON THE SAFE SIDE, REMOVE THE 13.2 SIZE LENS AND IMPLANTED THE 12.6 SIZE LENS." Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/2.5/084 | OS |  | 02/18/2020 | 2020 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824551&pc=QCB | 2023826-2020-00364 | 2019-07-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -11.50/2.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REMOVED FROM THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019 DUE TO EXCESSIVE VAULT. A SHORTER LENGTH LENS WAS IMPLANTED ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/2.0/085 | OS |  | 02/18/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824555&pc=QCB | 2023826-2020-00363 | 2019-09-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -11.00/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REMOVED FROM THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019 DUE TO EXCESSIVE VAULT. A SHORTER LENGTH LENS WAS IMPLANTED ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/1.5/093 | OD |  | 02/18/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824568&pc=QCB | 2023826-2020-00486 | 2019-11-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/2.0/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.00/2.0/080 | OD |  | 02/20/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9836557&pc=QCB | 2023826-2020-00541 | 2019-04-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.5/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS LATER REMOVED AND REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1-IN INITIAL MDR SHOULD BE CORRECTED TO PRODUCT PROBLEM. B2-IN INITIAL MDR REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT /DAMAGE IS NOT APPLICABLE. B3- DATE OF EVENT IS UNK. B5-THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.5/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS IMPLANTED/REMOVED AND REPLACED INTRAOPERATIVELY WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. H1-IN INITIAL MDR AND SUPPLEMENTAL #1; SERIOUS INJURY SHOULD BE CORRECTED TO MALFUNCTION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/1.5/089 | OD |  | 02/18/2020 | 2020 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9837917&pc=QCB | 2023826-2020-00529 | 2019-10-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/+1.0/102 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/+1.0/102 | NA |  | 02/20/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9845342&pc=QCB | 2023826-2020-00575 | 2019-11-27 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER STATED THAT A 13.2MM, VTICMO13.2, -16.5/4.0/114 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. IT WAS REPORTED THAT MEDICAL OR SURGICAL INTERVENTION WAS NOT NECESSARY. " VAULT IS 2+ ICL THICKNESS. VA HAS IMPROVED ENOUGH TI WARRANT NO FURTHER TX AT THIS TIME." CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.5/4.0/114 | OD | Other Health Care Professional | 01/03/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9841763&pc=QCB | 2023826-2020-00551 | 2019-07-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.0/1.0/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS LATER REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: B5 - THE LENS WAS IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/1.0/099 | OS |  | 02/20/2020 | 2020 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842793&pc=QCB | 2023826-2020-00511 | 2019-11-18 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/059 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+1.0/059 | OS |  | 02/20/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9914135&pc=QCB | 2023826-2020-00776 | 2020-02-28 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT 12.6MM, VTICM5_12.6, -7.50/1.00/87 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. LOW VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. CAUSE OF EVENT IS UNKNOWN. THE REPORTER STATED " THE VAULTS ARE VERY LOW. THE ACUITIES ARE STABLE AND GOOD AND THE LENS ARE OK." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.50/1.00/87 | OS | Physician | 03/04/2020 | 2020 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9913329&pc=QCB | 2023826-2020-00738 | 2020-03-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -8.00/+2.0/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE LENS WAS EXPLANTED ON (B)(6) 2020 AND A LONGER LENGTH NON-STAAR IPCL LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. REPORTEDLY, "LASER TOUCH UP IS PLANNED AND PATIENT IS HAPPY." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.00/+2.0/083 | OS |  | 03/12/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9916653&pc=QCB | 2023826-2020-00732 | 2019-11-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+3.5/085 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+3.5/085 | NA |  | 03/11/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9916672&pc=QCB | 2023826-2020-00733 | 2019-08-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+3.0/080 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+3.0/080 | NA |  | 03/11/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9928353&pc=QCB | 2023826-2020-00711 | 2020-01-31 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.0/3.5/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18.0/3.5/103 | OD |  | 03/11/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9928355&pc=QCB | 2023826-2020-00712 | 2019-11-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.50/3.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.50/3.5/091 | OD |  | 03/11/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9940832&pc=QCB | 2023826-2020-00844 | 2020-02-14 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -10.5/1.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WILL BE REPLACED WITH A SHORTER LENGTH LENS. THE REPORTER STATED THAT THE LENS IS TOO BIG FOR THE PATIENTS EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/1.5/091 | NA | Physician | 03/11/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9940866&pc=QCB | 2023826-2020-00845 | 2020-02-14 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -10.5/1.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WILL BE REPLACED WITH A SHORTER LENGTH LENS. THE REPORTER STATED THAT THE LENS IS TOO BIG FOR THE PATIENTS EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/1.5/090 | NA | Physician | 03/11/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9938080&pc=QCB | 2023826-2020-00710 | 2019-07-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -17.00/2.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17.00/2.5/085 | OS |  | 03/11/2020 | 2020 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9938130&pc=QCB | 2023826-2020-00709 | 2019-11-19 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/1.5/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/1.5/068 | OS |  | 03/11/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9940468&pc=QCB | 2023826-2020-00834 | 2019-09-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.5/4.5/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENTS LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT WAS OBSERVED AND THE LENS WAS REPLACED FOR A SHORTER LENGTH LENS WITHIN THE SAME SURGERY. IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/4.5/079 | OS |  | 03/11/2020 | 2020 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9955260&pc=QCB | 2023826-2020-00888 | 2019-12-31 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-14.5/2.5/007 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN. CORRECTED DATA: B7-CATARACT IN THE OPERATIVE EYE. H6-DEVICE CODE 1494 OFF-LABEL; PRE-EXISTING CATARACT IN THE OPERATIVE EYE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.5/2.5/007 | OS |  | 03/17/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9955312&pc=QCB | 2023826-2020-00889 | 2019-12-31 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-13.5/2.0/008 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT/DEFORMED AND RESIDUE ON LENS. CLAIM# :(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.5/2.0/008 | OD |  | 03/17/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957590&pc=QCB | 2023826-2020-00778 | 2019-01-29 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -9.50/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM.. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H1 - CORRECTED TO MALFUNCTION IN INITIAL MDR. H6 - PATIENT CODE 3191 IS NOT APPLICABLE IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.0/088 | OS |  | 03/17/2020 | 2020 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9977665&pc=QCB | 2023826-2020-00915 | 2020-03-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.50/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS IMPLANTED IN THE PATIENT WAS NOT THE PLANNED LENS(DIFFERENT LENS). LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: "THE SURGEON "MADE THE MISTAKE OF IMPLANTING THE LENS OF ANOTHER PATIENT." " SHOULD BE ADDED TO INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.50/1.5/090 | OD |  | 03/30/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9986805&pc=QCB | 2023826-2020-00942 | 2020-01-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-12.5/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. THE PROBLEM RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.5/1.5/091 | OD |  | 03/24/2020 | 2020 | 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9991678&pc=QCB | 2023826-2020-00974 | 2019-08-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.00/+6.0/096 (SPHERE/CYLINDER/AXIS), FROM THE PATIENTS LEFT EYE (OS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.00/+6.0/096 | OS |  | 04/07/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9991680&pc=QCB | 2023826-2020-00975 | 2019-08-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.00/+6.0/099 (SPHERE/CYLINDER/AXIS), FROM THE PATIENTS RIGHT EYE (OD), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.00/+6.0/099 | OD |  | 04/07/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10048954&pc=QCB | 2023826-2020-01041 | 2020-04-16 | Malfunction | STAAR SURGICAL COMPANY | 2020-05-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.0/+3.0/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS "OTHER" WITH NO FURTHER INFORMATION. Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.0/+3.0/073 | OD | Physician | 04/21/2020 | 2020 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10116220&pc=QCB | 2023826-2020-01162 | 2020-03-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. WITHIN THE SAME SURGERY, THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -8.0/+1.5/094 | OS |  | 05/11/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10116222&pc=QCB | 2023826-2020-01163 | 2020-03-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/078 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. WITHIN THE SAME SURGERY, THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND THAT THE HAPTIC WAS BENT AND DEFORMED. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -8.5/+1.5/078 | OD |  | 05/12/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10116227&pc=QCB | 2023826-2020-01161 | 2019-09-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+3.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. WITHIN THE SAME SURGERY THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE VAULT VALUE WENT FROM 230UM PRIOR TO EXCHANGE, TO 440UM AFTER EXCHANGE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: PROBLEM WAS RESOLVED WITH EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.0/+3.0/082 | OS |  | 05/11/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10135946&pc=QCB | 2023826-2020-01170 | 2020-03-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON WAS INSERTING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.00/+1.5/007 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), AND NOTED A SMALL DEFORMITY AT THE TOP EDGE OF THE LENS. THE LENS WAS PARTIALLY INSERTED INTO THE EYE AND WAS REMOVED WITHIN THE SAME SURGERY. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.00/+1.5/007 | OD |  | 05/15/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138549&pc=QCB | 2023826-2020-01180 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 13.2MM, VTICMO13.2, -11.00/2.5/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THIS RESOLVED THE PROBLEM. REPORTEDLY, "INITIAL LENS WAS FOUND BROKEN AFTER IMPLANTATION AND REMOVED DURING INITIAL SURGERY. REPLACEMENT LENS NOT YET IMPLANTED DUE TO PATIENT'S ISSUES." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN AND HAPTIC BENT WITH DEBRIS ON THE LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.00/2.5/107 | OS |  | 05/18/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138558&pc=QCB | 2023826-2020-01138 | 2020-01-30 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -17.50/2.5/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REMOVED FROM THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020 DUE TO EXCESSIVE VAULT. A SHORTER LENGTH LENS WAS IMPLANTED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17.50/2.5/099 | OD |  | 05/18/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10189314&pc=QCB | 2023826-2020-01275 | 2020-05-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | Event Description: THE REPORTER INDICATED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 15.50/+1.0/087 (SPHERE/CYLINDER/AXIS), WAS TORN DURING LOADING. THE SURGEON IMPLANTED ANOTHER SAME MODEL/LENGTH LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#:(B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | - 15.50/+1.0/087 | NA |  | 05/29/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10197347&pc=QCB | 2023826-2020-01330 | 2019-08-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.0/+2.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. WITHIN THE SAME SURGERY, THE LENS WAS EXPLANTED DUE TO LOW VAULT. ON THE SAME DATE A LONGER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -14.0/+2.5/093 | OD |  | 06/01/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10197350&pc=QCB | 2023826-2020-01325 | 2020-05-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.5/+3.0/113 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B3- "(B)(6) 2020" SHOULD BE CORRECTED TO "(B)(6) 2020" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+3.0/113 | OS |  | 05/28/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10197357&pc=QCB | 2023826-2020-01328 | 2019-09-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.0/+3.0/076 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. WITHIN THE SAME SURGERY, THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.0/+3.0/076 | OD |  | 06/01/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10222172&pc=QCB | 2023826-2020-01411 | 2020-03-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -12.0/+2.0/090 | OS |  | 06/05/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10222181&pc=QCB | 2023826-2020-01413 | 2020-02-24 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0/+1.5/081 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.0/+1.5/081 | OS |  | 06/05/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10224705&pc=QCB | 2023826-2020-01391 | 2020-06-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.50/1.5/167 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - CORRECTED TO: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.50/1.5/167 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2020. REPORTEDLY, THE "PATIENT IS OK" AFTER LENS REPLACEMENT. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.50/1.5/167 | OD |  | 06/11/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10242498&pc=QCB | 2023826-2020-01421 | 2020-06-03 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON WAS ATTEMPTING TO INSERT A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -14.00/+1.0/040 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) AND NOTED THE LENS WAS TORN BEFORE IMPLANTING. THE BACKUP LENS WAS USED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.00/+1.0/040 | OS | Other Health Care Professional | 06/09/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10271257&pc=QCB | 2023826-2020-01479 | 2020-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Material Opacification | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -4.00/1.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS OPACITY WAS OBSERVED ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS REMOVED AND CATARACT SURGERY WAS PERFORMED. REPORTEDLY, "PATIENT IS HAPPY." CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: B4: DATE OF THIS REPORT IN INITIAL MDR 26-JUN-2020 IS CORRECTED TO 25-JUN-2020. G4 IN INITIAL MDR 26-JUN-2020 IS CORRECTED TO 25-JUN-2020. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). SECONDARY SURGICAL INTERVENTION, LENS EXPLANT, CATARACT SURGERY. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -4.00/1.0/084 | OS |  | 06/25/2020 | 2020 | 54
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308836&pc=QCB | 2023826-2020-01529 | 2020-07-01 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.50/1.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. AN "OIL-LIKE ATTACHMENTS WERE OBSERVED ON THE SURFACE OF THE LENS." LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2020, "PATIENT'S ACUITY WAS 1.0 ON (B)(6) 2020. LITTLE ASTIGMATISM, SURGEON BELIEVES CAUSED BY INCISION. EVERYTHING ELSE OK, AC WAS QUIET. SURGEON WILL CONTINUE TO OBSERVE CLOSELY. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.50/1.0/083 | OS | Physician | 07/01/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308790&pc=QCB | 2023826-2020-01519 | 2020-05-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.50/1.5/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY ON (B)(6) 2020 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION; H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.50/1.5/078 | OD |  | 06/29/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308807&pc=QCB | 2023826-2020-01520 | 2020-06-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/3.5/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY ON (B)(6) 2020 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK, ETHNICITY - UNK, RACE - UNK, PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -12.50/3.5/107 | OD |  | 06/29/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308810&pc=QCB | 2023826-2020-01518 | 2020-03-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -6.50/0.5/130 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.50/0.5/130 | OD |  | 06/29/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308829&pc=QCB | 2023826-2020-01523 | 2019-10-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 12.1MM VTICMO12.1, -10.00/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. ON (B)(6) 2019, A REPLACEMENT LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/1.0/087 | OS |  | 06/29/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335345&pc=QCB | 2023826-2020-01531 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-07-28 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS LOADING A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/088 (SPHERE/CYLINDER/AXIS) IT TORE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.0/+1.0/088 | NA | Physician | 06/30/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10332986&pc=QCB | 2023826-2020-01583 | 2020-06-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -13.00/2.0/088 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS "REMOVED FROM VIAL AND FOUND THAT THERE WERE CRACKS IN THE LENS. THE LENS WAS NOT IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.00/2.0/088 | NA |  | 07/07/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335336&pc=QCB | 2023826-2020-01545 | 2020-03-06 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+2.5/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE, RESIDUE, AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE WITH RESIDUE AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+2.5/084 | OS |  | 06/29/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10356719&pc=QCB | 2023826-2020-01606 | 2020-07-08 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.50/+1.5/090 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO A BLEMISH/SCRATCH ON THE OPTIC, NOTED ONCE THE LENS WAS INSERTED. THE BACKUP LENS WAS IMPLANTED AND THE EXCHANGE WAS UNEVENTFUL WITH NO SEQUELAE. THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS WAS LOST. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.50/+1.5/090 | NA | Other Health Care Professional | 07/10/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10361982&pc=QCB | 2023826-2020-01623 | 2020-02-03 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.50/0.5/146 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: UNK. UNK. UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.50/0.5/146 | OD | Physician | 07/13/2020 | 2020 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10361986&pc=QCB | 2023826-2020-01622 | 2020-02-03 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.50/0.5/180 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.50/0.5/180 | OS | Physician | 07/13/2020 | 2020 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10361989&pc=QCB | 2023826-2020-01626 | 2020-06-24 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -3.00/2.5/062 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. REPORTEDLY, "THE LENS IS IN THE RIGHT AXIS (18") BUT THE ASTIGMATISM IS NOT CRRECTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -3.00/2.5/062 | OS | Physician | 07/16/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10398653&pc=QCB | 2023826-2020-01747 | 2020-01-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.0/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020 AND REMOVED WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. ON (B)(6) 2020 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6-DEVICE EVALUATION: LENS RETURNED IN A MIRCO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.0/1.5/091 | OD |  | 07/15/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10405558&pc=QCB | 2023826-2020-01753 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | Event Description: A DAMAGED 12.6MM, TMICL12.6, -14.0/1.5/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS RETUNED TO US. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, INSIDE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. CLAIM # (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.0/1.5/072 | NA |  | 07/14/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10406783&pc=QCB | 2023826-2020-01751 | 2020-07-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -10.5/5.5/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: "PATIENT'S LAST VISIT WAS ON (B)(6) 2020. BCVA 20/20; VAULT 900. PATIENT IS HAPPY. LENS REMAINS IMPLANTED." (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.5/5.5/102 | OD |  | 07/17/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429647&pc=QCB | 2023826-2020-01803 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-08-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Known Impact Or Consequence To Patient | NA | Event Description: IT WAS REPORTED THAT A 13.2MM, VTICMO13.2, -9.5/2/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT LEFT EYE (OS) AS A REPLACEMENT LENS DUE TO LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM (VAULT IS STILL NOT IDEAL). ON (B)(6) 2020 THE REPORTER STATED " THE PATIENT CONDITION IS GOOD NOW EXCEPT THE VAULT IS STILL LOW: ABOUT 200 MICROMETER. THE CAUSE OF THE UNIDEAL VAULT AFTER THE CHANGED ICL SURGERY MIGHT BE THE STRUCTURE OF THE PATIENT'S EYE IS ABNORMAL. THE DOCTOR HASN'T DONE THE INTERVENTION FOR THE LOW VAULT AND DECIDES TO DO THE FOLLOW-UP CLOSELY. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/2/111 | OS | Physician | 07/21/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429503&pc=QCB | 2023826-2020-01793 | 2019-06-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.5/1.0/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. ON (B)(6) 2019 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/1.0/108 | OD |  | 07/21/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10440595&pc=QCB | 2023826-2020-01808 | 2020-07-16 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/102 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS DURING THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. | -10.00/+1.0/102 | NA | Physician | 07/29/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476241&pc=QCB | 2023826-2020-01932 | 2020-06-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -7.50/+1.5/077(SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY ON (B)(6) 2020, DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3. DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -7.50/+1.5/077 | NA |  | 08/04/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10480709&pc=QCB | 2023826-2020-01998 | 2020-06-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -12.0/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/1.5/090 | OD |  | 08/04/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10491990&pc=QCB | 2023826-2020-01960 | 2019-08-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/175 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.5/175 | OD |  | 08/05/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10495978&pc=QCB | 2023826-2020-02014 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-09-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.50/1.00/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020 AND EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. THE REPORTER STATED " PLEASE NOTE THAT WE ARE STILL MONITORING THE PATIENT. NO FURTHER INTERVENTION DONE." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | -8.50/1.00/084 | OD | Physician | 08/06/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10495983&pc=QCB | 2023826-2020-02015 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-09-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.50/1.5/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020 AND EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. THE REPORTER STATED " PLEASE NOTE THAT WE ARE STILL MONITORING THE PATIENT. NO FURTHER INTERVENTION DONE." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | -8.50/1.5/103 | OS | Physician | 08/06/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8912816&pc=QCB | 2023826-2019-01509 | 2019-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -14.0/+2.5/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON NOTES EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF ICA. REPORTEDLY THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF -14.00/2.5/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON NOTES EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF ICA. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. D6B- EXPLANT DATE: (B)(6) 2019. H6- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H10. IN INITAL MDR CLAIM # (B)(4) IS NOT APPLICABLE, INITIAL MDR IS FOR CLAIM# (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.0/+2.5/084 | OD |  | 07/26/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10271249&pc=QCB | 2023826-2020-01473 | 2020-06-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.00/3.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WITH ROTATION WAS OBSERVED ON (B)(6) 2020. LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2020, LENS TO BE EXCHANGED WITH A LONGER LENGTH LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION. B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE . B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.00/3.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTER STATES LOW VAULT; AND THE LENS WAS ROTATED 40 DEGREES 1 AFTER 1 WEEK. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. D6B- EXPLANT DATE: (B)(6) 2020. H6- CLINICAL CODE 4582: VAULTING. CORRECTED DATA: H6- RESULT FINDINGS IN INITIAL MDR 114 SHOULD BE CORRECTED TO 213. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 STATEMENT IN INITIAL MDR SHOULD INCLUDE THAT THE PATIENT ALSO EXPERIENCED LOW VISION WITH LOW VAULT. H6-CLINICAL CODE 4581- LOW VISION. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.00/3.0/092 | OD |  | 06/22/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652115&pc=QCB | 2023826-2019-00928 | 2018-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/107 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THIS DID NOT RESOLVE THE PROBLEM OF LOW VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. REFERENCE MFR REPORT# 2023826-2019-00927 FOR INITIAL IMPLANTED LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON 22/FEB/2019 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -8.5/+1.5/107 | OD |  | 05/06/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9608594&pc=QCB | 2023826-2020-00174 | 2019-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -16.5/5.0/138 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. EXCESSIVE VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 -16.50/5.0/138 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLAMER LENS INTO THE PATIENT RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. D6B- EXPLANT DATE (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT WITH DEBRIS ON LENS. (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. D6B- EXPLANT DATE (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.5/5.0/138 | OD |  | 12/27/2019 | 2019 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9608630&pc=QCB | 2023826-2020-00175 | 2019-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -14.5/5.0/51 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE BS-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 -14.50/5.0/051 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLAMER LENS INTO THE PATIENT LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. D6B- EXPLANT DATE (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT WITH RESIDUE ON LENS AND LENS IS CURVED. (B)(4). Manufacturer Narrative: CORRECTED DATA: D6B- EXPLANT DATE (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.5/5.0/51 | OS |  | 12/27/2019 | 2019 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9782975&pc=QCB | 2023826-2020-00409 | 2020-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -1.0/+1.0/012 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5 - IT WAS REPORTED THAT ON (B)(6) 2020 THE LENS WAS EXPLANTED AND ON THE SAME DATE, THE LENS WAS REPLACED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -1.0/+1.0/012 | OD |  | 02/07/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10361997&pc=QCB | 2023826-2020-01631 | 2020-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/1.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY ON (B)(6) 2020 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CORRECTED DATA: B5 - LENS BEING REPLACED DURING INITIAL SURGERY SHOULD BE CORRECTED TO LENS WAS EXCHANGED LATER ON THE SAME DATE ON (B)(6) 2020 IN A SECOND SURGERY. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/1.0/109 | OD |  | 07/18/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10392937&pc=QCB | 2023826-2020-01670 | 2020-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Known Impact Or Consequence To Patient; No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.50/+0.5/076 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS). THE REPORTER INDICATED THE LENS TORE DURING IMPLANTATION AND A LARGE PART OF THE LENS REMAINED IN THE EYE. THE PLAN IS TO EXPLANT THE LENS AT A LATER DATE. THE REPORTER INDICATED PART OF THE LENS WAS STUCK IN THE PLUNGER AND THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE LENS WAS IMPLANTED ON (B)(6) 2020 AND EXPLANTED ON (B)(6) 2020 DUE TO LENS TORE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.50/+0.5/076 | OS |  | 07/22/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10416196&pc=QCB | 2023826-2020-01721 | 2020-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -14.0/+3.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS HAD LOW VAULTING AND REMAINED IMPLANTED. THE PLAN IS TO EXPLANT/EXCHANGE THE LENS AT A LATER DATE. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2020 AND WAS EXCHANGED FOR A LONGER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.0/+3.0/088 | OS |  | 07/25/2020 | 2020 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8433521&pc=QCB | 2023826-2019-00443 | 2019-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.50/1.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. REPORTEDLY, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADVERSE EVENT OR PRODUCT PROBLEM UPDATE - FROM PRODUCT PROBLEM TO ADVERSE EVENT UPDATE TO INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) NOTED AS "REMAINS IMPLANTED" - UPDATE - THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. PREVIOUSLY REPORTED DISTRIBUTOR. CORRECTION, DISTRIBUTOR SHOULD NOT BE REPORTED. PREVIOUSLY NOT REPORTED - CORRECTION, THIS PRODUCT IS NOT MARKETED IN THE US CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -9.50/1.0/082 | OD |  | 02/19/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8475385&pc=QCB | 2023826-2019-00524 | 2019-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/020 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON NOTES EXCESSIVE VAULTING. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED INFORMATION: ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS. THE PROBLEM WAS RESOLVED. PATIENT CODE UPDATED TO 3191, NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/+1.0/020 | OD |  | 03/08/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8479491&pc=QCB | 2023826-2019-00534 | 2019-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -11.5/3/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. EXCESSIVE VAULT HAS BEEN OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THIS EVENT IS UNKNOWN. IT WAS REPORTED THAT A LENS EXCHANGE IS PLANNED IN ABOUT 2 MONTHS. THE PATIENT IS CURRENTLY FEELING WELL (NO PUPIL OR HIGH IOP ISSUES), CURRENT BCVA IS 20/20. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICTED THAT A 13.7MM, VTICM5_13.7, -11.5/3/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. IT WAS REPORTED THAT THE PATIENT IS HAPPY. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION,EXCHANGE. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/3/095 | OD |  | 03/06/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8591938&pc=QCB | 2023826-2019-00779 | 2019-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Known Impact Or Consequence To Patient; No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -9.5/2.5/067 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. AFTER INSERTION, THE LENS "FLIPPED IN THE EYE." THE LENS WAS SUCCESSFULLY EXCHANGED WITH AN ALTERNATE LENS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EXPLANT DATE: (B)(6) 20019. PATIENT CODE 2692: NO KNOWN IMPACT OR CONSEQUENCE TO PATIENT. PATIENT CODE 3191 IN ORIGINAL MDR IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICTED THAT THE SURGEON IMPLANTED A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -9.5/2.5/067 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. AFTER INSERTION THE LENS "FLIPPED IN THE EYE". "THE LENS WAS TAKEN OUT AND REPLACED WITH A DIFFERENT ICL." (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). | -9.5/2.5/067 | OS |  | 04/09/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8592238&pc=QCB | 2023826-2019-00638 | 2019-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -13.00/1.5/090 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS REMAINS IMPLANTED. REPORTER INDICATED LENS EXCHANGE PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THE LENS WAS REMOVED AND EXCHANGED ON (B)(6)B 2019, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. "PATIENT AND DOCTOR ARE HAPPY." ADDED EXPLANT DATE OF (B)(6) 2019. PATIENT CODE - 2692 NO KNOWN IMPACT OR CONSEQUENCE TO PATIENT TO 3191 NO CODE AVAILABLE - SECONDARY SURGERY, LENS EXCHANGE. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.00/1.5/090 | OS |  | 04/09/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8631512&pc=QCB | 2023826-2019-00848 | 2019-03-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXPLANT SHOULD BE CORRECTED TO PATIENT CODE: 2692 - NO KNOWN IMPACT OR CONSEQUENCE TO PATIENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/1.5/091 | OS |  | 05/06/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8631532&pc=QCB | 2023826-2019-00856 | 2019-03-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/2.0/089 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/2.0/089 | OS |  | 05/06/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8633183&pc=QCB | 2023826-2019-00887 |  | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.0/+1.0/146 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS LENS DECENTRATION INFERIORLY AND TEMPORALLY AND LENS ROTATION. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: "REPORTER INDICATED THAT THE LENS WAS EXPLANTED" SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL #1 MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: EXPLANT DATE: UNK. PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: DATE RECEIVED BY MANUFACTURER: 13-SEP-2019 SHOULD BE CORRECTED TO 10-SEP-2019 IN SUPPLEMENTAL MEDWATCH REPORT #1. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.0/+1.0/146 | OD |  | 04/30/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8648029&pc=QCB | 2023826-2019-00966 | 2019-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13. 2 IMPLANTABLE COLLAMER LENS, - 11.00/+1.0/112 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON IS PLANNING TO EXPLANT AND EXCHANGE THE LENS FOR A SMALLER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED ON 11-SEP-2019. REPORTEDLY, THE LENS WAS TOO LARGE. ON (B)(6) 2019 THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENGTH LENS. THE PROBLEM WAS RESOLVED. PATIENT CODE 2692 NO LONGER APPLICABLE DUE TO ADDITIONAL INFO. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXPLANT. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.00/+1.0/112 | OD |  | 05/08/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8736582&pc=QCB | 2023826-2019-01124 | 2019-06-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -8.0/+3.0/163 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON NOTES EXCESSIVE VAULT. REPORTEDLY THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S., CLAIM# (B)(4). | -8.0/+3.0/163 | OD |  | 06/04/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8772293&pc=QCB | 2023826-2019-01198 | 2019-06-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -5.00/1.5/080 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019.THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT CODE: 3191 SHOULD BE CORRECTED TO 2692 IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.00/1.5/080 | OS |  | 07/02/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8819695&pc=QCB | 2023826-2019-01302 | 2019-02-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/101 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2019. THE REPORTER INDICATED THE LENS HAD A LOW VAULT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH DEBRIS ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC BENT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.5/+1.0/101 | OS |  | 06/27/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8819703&pc=QCB | 2023826-2019-01305 | 2019-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/055 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS IMPLANTED ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED, THE VAULT WAS STILL LOW. THE CAUSE OF THE EVENT WAS UNKNOWN. SEE MFR REPORT # 2023826-2019-01546 FOR THE EXCHANGED LENS. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+1.0/055 | OD |  | 06/28/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8883523&pc=QCB | 2023826-2019-01357 | 2019-03-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/2.5/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A LONGER LENGTH LENS DURING INITIAL SURGERY DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/2.5/073 | OS |  | 07/22/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8894305&pc=QCB | 2023826-2019-01431 | 2019-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM TMICL13.7 LENS, -12.5/+2.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. THE SURGEON STATES THE LENS IS TOO LARGE AND THE VAULT IS TOO HIGH. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION OR UPDATES HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED INFORMATION: THE LENS WAS RETURNED TO THE MANUFACTURER AND EVALUATED, THEREFORE LENS HAS BEEN EXPLANTED. NO ADDITIONAL INFORMATION AVAILABLE. IS THIS A SINGLE-USE DEVICE THAT WAS REPROCESSED AND REUSED ON A PATIENT?: UNK. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL, INSIDE OF A STERILIZATION PACK. VISUAL INSPECTION FOUND THAT THE THERE WAS A TEAR IN THE HAPTIC. PATIENT CODE UPDATED: 3191 NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.5/+2.5/089 | OD |  | 07/23/2019 | 2019 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8894309&pc=QCB | 2023826-2019-01432 | 2019-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM TMICL13.7 LENS, -6.5/+1.5/096 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. THE SURGEON STATES THE LENS IS TOO LARGE AND THE VAULT IS TOO HIGH. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED INFORMATION: THE LENS HAS BEEN RETURNED AND EVALUATED, THEREFORE THE LENS HAS BEEN EXPLANTED. NO ADDITIONAL INFORMATION IS KNOWN. EXPLANT DATE: UNK. DEVICE EVALUATION: THE LENS WAS RETURNED IN A LENS VIAL, IN A STERILIZATION POUCH. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. PATIENT CODE UPDATED: 3191 NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.5/+1.5/096 | OS |  | 07/23/2019 | 2019 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8910141&pc=QCB | 2023826-2019-01447 | 2019-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.50/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LOW VAULT WITH ROTATION WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTER INDICATED ON (B)(6) 2019 THAT A LENS REMOVAL AND EXCHANGE ARE PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.50/2.5/088 | OS |  | 07/30/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8965327&pc=QCB | 2023826-2019-01614 |  | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Material Opacification; No Apparent Adverse Event | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -4.00/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). WHITE ACCRETION ON THE LENS AND LENS OPACITY WERE OBSERVED. THE LENS REMAINS IMPLANTED. PER REPORTER, "THE PATIENT HAD NO SUBJECTIVE SYMTOMS AND THE LENS OPACITY WAS MILD AND IMPROVED WITH FLUOROMETHOLONE EYE DROPS IN 1 MONTH. CURRENTLY, THE LENS OPACITY REMAINS SLIGHTLY." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADVERSE EVENT OR PRODUCT PROBLEM: SERIOUS. OUTCOMES ATTRIBUTED TO ADVERSE EVENT: REQUIRED INTERVENTION TO PREVENT PERMANENT. IMPAIREMENT/DAMAGE (DEVICE). PATIENT CODE 3191: EYE TREATED WITH FLUOROMETHOLONE EYE DROPS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE 3189 IS CORRECTED TO 1426 FOR OPACITY ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). | -4.00/1.5/088 | OS |  | 08/08/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8965332&pc=QCB | 2023826-2019-01613 |  | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Material Opacification; No Apparent Adverse Event | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -4.50/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). WHITE ACCRETION ON THE LENS AND LENS OPACITY WERE OBSERVED. THE LENS REMAINS IMPLANTED. PER THE REPORTER, "THE PATIENT HAD NO SUBJECTIVE SYMPTOMS AND THE LENS OPACITY WAS MILD AND IMPROVED WITH FLUOROMETHOLONE EYE DROPS IN 1 MONTH. CURRENTLY, THE LENS OPACITY REMAINS SLIGHTLY." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE: 2692 SHOULD BE CORRECTED TO 3191: NO CODE AVAILABLE (MEDICAL INTERVENTION, TOPICAL STEROIDS). DEVICE CODE: 3189 SHOULD BE CORRECTED TO 1426. PATIENT IDENTIFIER: (B)(6) SHOULD BE CORRECTED TO (B)(6). PRODUCT PROBLEM SHOULD BE CORRECTED TO ADVERSE EVENT. REQUIRED INTERVENTION SHOULD BE MARKED. AT A FOLLOW UP VISIT IT WAS REPORTED THAT THE WHITE ACCRETIONS OF THE PATIENTS GOT SMALLER WITHOUT SPECIFIC MEDICAL TREATMENT. THE DOCTOR THINKS THE RISK OF RECURRENCE WILL BE LOW. REPORTED LENS OPACITY MEANS THAT ICL LENS OPACITY DUE TO WHITE ACCRETIONS AND NOT THE PATIENT'S CATARACT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). | -4.50/1.0/085 | OD |  | 08/08/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9112033&pc=QCB | 2023826-2019-01777 | 2019-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -18.0/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT WAS OBSERVED BUT DUE TO STABLE CONDITION THE SURGEON DECIDED TO LEAVE THE LENS IN THE EYE AND MONITOR THE PATIENT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: "(B)(4)" SHOULD BE CORRECTED TO "(B)(4)" IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CLAIM#:(B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.0/1.5/091 | OS |  | 08/25/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9115525&pc=QCB | 2023826-2019-01730 | 2019-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -15.50/4.5/152 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -15.50/4.5/152 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; LENS EXCHANGE CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.50/4.5/152 | OS |  | 09/04/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9131806&pc=QCB | 2023826-2019-01759 | 2019-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50/+5.5/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING. THE PLAN IS TO EXPLANT THE LENS BUT TO DATE THE LENS REMAINED IMPLANTED. Manufacturer Narrative: REC'D BY MFR: CORRECTED TO 26-OCT-2019 IN SUPPLEMENTAL MEDWATCH REPORT #1. DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED TO SERIOUS INJURY IN SUPPLEMENTAL MEDWATCH REPORT #2. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2019. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.50/+5.5/086 | OS |  | 09/02/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9131823&pc=QCB | 2023826-2019-01760 | 2019-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.00/+5.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING. THE PLAN IS TO EXPLANT THE LENS BUT TO DATE THE LENS REMAINED IMPLANTED. Manufacturer Narrative: TYPE OF REPORTABLE EVENT: CORRECTED TO SERIOUS INJURY IN SUPPLEMENTAL MEDWATCH REPORT #2. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6)2019. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED (B)(4). Manufacturer Narrative: PMA\510K: CORRECTED TO (B)(6) 2019 IN SUPPLEMENTAL MEDWATCH REPORT #1 DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.00/+5.5/092 | OD |  | 09/02/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9139622&pc=QCB | 2023826-2019-01818 | 2019-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -07.50/3.5/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WAS OBSERVED, THE LENS REMAINS IMPLANTED. AN ADDITION OF A NEW PI WAS PERFORMED. REPORTEDLY THE PATIENT IS "ASYMPTOMATIC." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A CASE/ VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AS WELL AS THE HAPTIC BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED INFORMATION ON 19/SEP/2019 THE LENS WAS REMOVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, LENS EXPLANT. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT PROBLEM CORRECTED TO ADVERSE EVENT IN INITIAL MDR. REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. EXPLANT DATE: (B)(6) 2019 SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MEDWATCH REPORT #1. PATIENT CODE 2692 CORRECTED TO PATIENT CODE 3191 IN INITIAL MDR. PATIENT CODE 3191 CORRECTED TO 'SECONDARY SURGICAL INTERVENTION: SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, SECONDARY YAG' IN INITIAL MDR. 'LENS EXCHANGE' PREVIOUSLY STATED IS NOT APPLICABLE. PER NEW INFORMATION RECEIVED, THE LENS WAS EXCHANGED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS REPORTEDLY HAPPY AND SATISFIED WITH THE FINAL RESULT. PATIENT CODE 3191 - SECONDARY SURGICAL INTERVENTION: LENS EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). WORK ORDER SEARCH: NO [OR#] SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -07.50/3.5/101 | OD |  | 09/03/2019 | 2019 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9147370&pc=QCB | 2023826-2019-01800 | 2017-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.00/3.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED, BUT DID NOT RESOLVE THE PROBLEM. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EVENT DESCRIPTION: "BUT DID NOT RESOLVE THE PROBLEM" SHOULD BE REMOVED FROM INITIAL MEDWATCH REPORT, AS IT IS NOT APPLICABLE. PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS REPOSITIONING). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - LENS WAS IMPLANTED ON (B)(6)2017. LENS WAS REPOSITIONED ON (B)(6)2017 AND THE PROBLEM WAS RESOLVED. POSTOPERATIVE REPORT STATED "PATIENT IS HAPPY". CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.00/3.0/089 | OD |  | 09/11/2019 | 2019 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9147383&pc=QCB | 2023826-2019-01801 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -6.00/2.0/057 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED . LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: "PRODUCT PROBLEM" SHOULD BE CORRECTED TO "ADVERSE EVENT" IN THE INITIAL MDR. "REQUIRED INTERVENTION" SHOULD BE MARKED IN THE INITIAL MDR. "LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED." SHOULD BE REMOVED FROM THE INITIAL MDR. PATIENT CODE 2692 IS NOT APPLICABLE AND SHOULD BE CORRECTED TO PATIENT CODE 3191- SECONDARY SURGICAL INTERVENTION; REPOSITIONING IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.00/2.0/057 | OS |  | 09/11/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9147400&pc=QCB | 2023826-2019-01802 | 2018-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.00/2.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED, BUT DID NOT RESOLVE THE PROBLEM. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EVENT: LENS REPOSITIONING RESOLVED THE PROBLEM. PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS REPOSITIONING). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - LENS WAS IMPLANTED ON (B)(6) 2018. THE LENS WAS REPOSITIONED ON (B)(6) 2018. POSTOPERATIVE REPORT STATED "PATIENT IS HAPPY". CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANTED: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.00/2.0/087 | OD |  | 09/11/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9147401&pc=QCB | 2023826-2019-01803 | 2017-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -7.00/2.0/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED, BUT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: "PRODUCT PROBLEM" SHOULD BE CORRECTED TO "ADVERSE EVENT" IN INITIAL MDR. "REQUIRED INTERVENTION" SHOULD BE CHECKED IN THE INITIAL MDR. "BUT DID NOT RESOLVE THE PROBLEM" SHOULD BE REMOVED FROM THE INITIAL MDR. PATIENT CODE: 2692 IS NOT APPLICABLE AND SHOULD BE CORRECTED TO "3191: SECONDARY SURGICAL INTERVENTION; REPOSITIONING" IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B3- (B)(6) 2017. B5- " LENS WAS REPOSITIONED ON (B)(6) 2017 AND PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY.". D6- (B)(6) 2017. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION: LENS REPOSITIONING. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.00/2.0/078 | OS |  | 09/11/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9226358&pc=QCB | 2023826-2019-02019 | 2019-09-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICM5_13.2, -7.50/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2019 AND LOW VAULT WAS OBSERVED. PATIENTS CURRENT BCVA WAS REPORTED TO BE 20/20 AND PATIENT IS DOING WELL. LENS REMAINS IMPLANTED. THE SURGEON PLANS TO EXCHANGE THE LENS WITH A LONGER LENGTH LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: "LENS REMAINS IMPLANTED. THE SURGEON PLANS TO EXCHANGE THE LENS WITH A LONGER LENS." SHOULD BE REMOVED AND CORRECTED TO "THE LENS WAS EXCHANGED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM." IN THE INITIAL MDR. TYPE OF REPORTABLE EVENT: "MALFUNCTION" SHOULD BE CORRECTED TO "SERIOUS". (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: FOLLOW-UP TYPE: "CORRECTION" SHOULD BE CORRECTED TO "ADDITIONAL INFORMATION" IN SUPPLEMENTAL MEDWATCH REPORT #1. "CORRECTED DATA" SHOULD BE CORRECTED TO "ADDITIONAL MANUFACTURER NARRATIVE" IN SUPPLEMENTAL MEDWATCH REPORT #1. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.50/2.0/090 | OS |  | 09/25/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9227249&pc=QCB | 2023826-2019-02022 | 2019-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICM5_12.6, -10.5/3/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) AND LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED, LENS EXCHANGE IS PLANNED FOR NOVEMBER. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PER UPDATED INFORMATION, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. PATIENT CODE 3191 (SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE). (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. IMPLANTED:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/3/098 | OS |  | 09/27/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9246813&pc=QCB | 2023826-2019-02029 | 2019-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+0.5/163 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT. THE LENS CURRENTLY REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED INFORMATION: ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM IS RESOLVED. PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). (B)(4). Manufacturer Narrative: PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; REPOSITIONING IS NOT APPLICABLE IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/+0.5/163 | OD |  | 10/03/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9271546&pc=QCB | 2023826-2019-02040 | 2019-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.00/5.0/091 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY THE SURGEON PLANS TO EXCHANGE THE LENS WITH A LONGER LENGTH LENS. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.00/5.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE REPORTER STATED "ALL IS FINE- PATIENT IS HAPPY!" H6- PATIENT CODE 3191- SECONDARY SURGICAL INTERVENTION- LENS EXCHANGE CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL DEBRIS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.00/5.0/091 | OD |  | 10/09/2019 | 2019 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9272143&pc=QCB | 2023826-2019-02041 | 2019-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -7.50/6.0/086 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY THE SURGEON PLANS TO EXCHANGE THE LENS WITH A LONGER LENGTH LENS. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL DEBRIS IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -7.50/6.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. REPORTER STATES "PATIENT IS HAPPY." H6: PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION- LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.50/6.0/086 | OS |  | 10/09/2019 | 2019 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9441174&pc=QCB | 2023826-2019-02353 |  | Injury | STAAR SURGICAL COMPANY | 2019-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.5/+6.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. REPORTEDLY THE LENS IS "SMALL AND TURNING." IT IS UNKNOWN AS TO WHETHER THIS LENS REMAINS IMPLANTED OR NOT; ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5: CORRECTED TO: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.5/+6.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. REPORTEDLY THE LENS IS "SMALL AND TURNING." IT IS UNKNOWN AS TO WHETHER THIS LENS HAS BEEN EXPLANTED OR NOT; ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. H6: PATIENT CODE 2692 CORRECTED TO PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION (LENS EXPLANT). CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17.5/+6.0/082 | NA |  | 11/14/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9441177&pc=QCB | 2023826-2019-02354 |  | Injury | STAAR SURGICAL COMPANY | 2019-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+6.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. REPORTEDLY THE LENS IS "SMALL AND TURNING." IT IS UNKNOWN AS TO WHETHER THIS LENS REMAINS IMPLANTED OR NOT; ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5 - CORRECTED TO: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+6.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. REPORTEDLY THE LENS IS "SMALL AND TURNING. IT IS UNKNOWN AS TO WHETHER THIS LENS HAS BEEN EXPLANTED OR NOT; ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. H6 - PATIENT CODE 2692 CORRECTED TO PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION (LENS EXPLANT). CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+6.0/091 | NA |  | 11/14/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9483557&pc=QCB | 2023826-2019-02395 | 2019-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -4.50/5.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). REPORTEDLY "PATIENT NEEDS EXPLANT DUE TO EXCESS VAULTING AT AROUND 1400UM." PER THE REPORTER ON (B)(6) 2019 PATIENT "IS CURRENTLY EMMETROPIC." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B4 - CORRECTED TO 17-NOV-2019 IN INITIAL MDR. H6 - RESULT CODE 114 CORRECTED TO RESULT CODE 213 IN INITIAL MDR. (B)(4). Manufacturer Narrative: G4 - CORRECTED TO 17-NOV-2019 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: B5: UPDATED TO: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -4.50/5.5/089 (SPHERE / CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. EXCESSIVE VAULT OF 1500UM, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, PERIPHERAL ANTERIOR SYNECHIAE, AND IRIDOCORNEAL CONTACT WERE REPORTED. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. REPORTEDLY THE EYE IS "HEALTHY." THE CAUSE OF THIS EVENT IS REPORTED AS THE DEVICE. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGICAL INTERVENTION - LENS EXCHANGE). H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES; PERIPHERAL ANTERIOR SYNECHIAE, IRIDOCORNEAL CONTACT). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -4.50/5.5/089 | OS |  | 11/17/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9483562&pc=QCB | 2023826-2019-02396 | 2019-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -3.00/5.5/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). REPORTEDLY "PATIENT NEEDS EXPLANT DUE TO EXCESS VAULTING AT AROUND 1400UM." PER THE REPORTER ON (B)(6) 2019 PATIENT "IS CURRENTLY EMMETROPIC." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B4 - CORRECTED TO 17-NOV-2019 IN INITIAL MDR. H4 - CORRECTED TO 23-SEP-2016 IN INITIAL MDR. H6 - RESULT CODE 114 CORRECTED TO RESULT CODE 213 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: G4 - CORRECTED TO (B)(6) 2019 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: UPDATED TO: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -3.00/5.5/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. EXCESSIVE VAULT OF 1500UM, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, PERIPHERAL ANTERIOR SYNECHIAE, AND IRIDOCORNEAL CONTACT WERE REPORTED. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. REPORTEDLY THE EYE IS "HEALTHY." THE CAUSE OF THIS EVENT IS REPORTED AS THE DEVICE. PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGICAL INTERVENTION - LENS EXCHANGE) PATIENT CODE 3191: NO CODE AVAILABLE (SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES; PERIPHERAL ANTERIOR SYNECHIAE, IRIDOCORNEAL CONTACT). CLAIM# (B)(4). Manufacturer Narrative: SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -3.00/5.5/098 | OD |  | 11/17/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9528009&pc=QCB | 2023826-2019-02496 | 2020-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Known Impact Or Consequence To Patient; No Code Available | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, - 12.50/+1.5/029 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS REPORTED AS HAVING HIGH VAULTING AND WILL BE EXPLANTED AT A LATER DATE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THE LENS WAS IMPLANTED ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO UNKNOWN/PATIENT RELATED FACTOR. POST-OP PATIENT CONDITION - IMPROVED VA/IMPROVE PHOTOPHOBIA. THE LENS WAS DISCARDED. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED H6 - PATIENT CODE: 3191 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | - 12.50/+1.5/029 | OS |  | 12/11/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9528016&pc=QCB | 2023826-2019-02497 | 2019-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Known Impact Or Consequence To Patient; No Code Available | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 13.50/+4.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING HIGH VAULTING AND WILL BE EXPLANTED/EXCHANGED FOR A SHORTER LENS. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H6: DEVICE CODE: 1583 SHOULD BE CORRECTED TO 3189. B5: LENS IS REPORTED TO BE EXCHANGED WITH A SHORTER LENGTH LENS, NO REASON FOR EXCHANGE WAS REPORTED. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. POSTOPERATIVE REPORT INDICATED "PATIENT IS HAPPY" AND "PERFECT, IMPROVED VISION OVER PRE-OP". H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - THE LENS WAS REPORTED AS HAVING HIGH VAULTING SHOULD BE CORRECTED TO THE LENS WAS REPORTED TO BE TOO LONG. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | - 13.50/+4.0/092 | OS |  | 12/10/2019 | 2019 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9595583&pc=QCB | 2023826-2020-00123 | 2019-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.00/1.5/033 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: DATE OF EVENT IS (B)(6) 2020. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.00/1.5/033 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO LOW VUALT. THE EXCHANGE RESOLVED THE PROBLEM. DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HATIC TORN AND DEBRIS ON THE LENS. PATIENT CODE:3191 SECONDARY SURGERY (EXCHANGE). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.00/1.5/033 | OS |  | 10/29/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9664304&pc=QCB | 2023826-2020-00232 | 2020-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). PER THE REPORTER ON (B)(6) 2020, "LENS NEEDS TO BE EXCHANGED FOR A LARGER SIZE." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM TMICL12.1, -14.50/4.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LOW VAULT WITH ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2019, THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS REMOVED AND REPLACED FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS "TOO SMALL LENS SIZE.". H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE, REPOSITIONING. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: G4 - DATE RECEIVED BY MANUFACTURER: 31/JAN/2020 SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MEDWATCH REPORT #1. H6 - DEVICE CODE: 2923 SHOULD HAVE BEEN INCLUDED. D10 - RETURNED TO MANUFACTURER: 12/FEB/2020 SHOULD BE CORRECTED TO 13/FEB/2020. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: B5/H10 - INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019 SHOULD BE CORRECTED TO (B)(6) 2019 IN SUPPLEMENTAL MEDWATCH REPORT #1. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). (B)(4). | -14.50/4.0/092 | OS |  | 01/08/2020 | 2020 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9684571&pc=QCB | 2023826-2020-00264 | 2019-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/079 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. PATIENT IS HAPPY. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED CORRECTED DATA: B3: DATE OF EVENT: (B)(6) 2019 CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/+1.5/079 | OD |  | 01/19/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9689269&pc=QCB | 2023826-2020-00283 | 2020-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -7.50/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, LENS EXCHANGE TO OCCUR. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "ON 26-FEB-2020, THE SURGEON OPERATED ON THE LEFT EYE TO PERFORM ICL EXCHANGE." H6- PATIENT CODE 3191: LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B4- "(B)(6) 2020" SHOULD BE CORRECTED TO "(B)(6) 2020" IN THE INITIAL MDR. G4-"16-JAN-2020" SHOULD BE CORRECTED TO "15-JAN-2020" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1- ADVERSE EVENT SHOULD BE MARKED IN MDR SUPPLEMENTAL #2. B2- REQUIRED INTERVENTION SHOULD BE MARKED IN MDR SUPPLEMENTAL #2. D7- EXPLANT DATE: (B)(6) 2020 SHOULD BE IN MDR SUPPLEMENTAL #2. H1- SERIOUS INJURY SHOULD BE IN MDR SUPPLEMENTAL #2. CLAIM#(B)(4). Manufacturer Narrative: ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.50/1.5/089 | OS |  | 01/15/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9717739&pc=QCB | 2023826-2020-00329 | 2020-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -1.00/2.0/020 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT THE INITIAL LENS WAS IMPLANTED ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SMALLER LENGTH LENS AND THE PROBLEM RESOLVED. "PATIENT IS HAPPY". H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - RESULT CODE: 114 SHOULD BE CORRECTED TO 213, IN INITIAL MEDWATCH REPORT. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -1.00/2.0/020 | OD |  | 01/28/2020 | 2020 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9717740&pc=QCB | 2023826-2020-00330 | 2020-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -1.50/1.50/149 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H6 - RESULT CODE: 114 SHOULD BE CORRECTED TO 213, IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE INITIAL LENS WAS IMPLANTED ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED, AND A REPLACEMENT LENS WAS IMPLANTED INTO THE PATIENT'S EYE BUT WAS IMPLANTED UPSIDE DOWN. DURING REMOVAL OF THE LENS, THE LENS TORE. THIS OCCURRED INTRAOPERATIVELY. ANOTHER SHORTER LENGTH REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. "PATIENT IS HAPPY". H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -1.50/1.50/149 | OS |  | 01/28/2020 | 2020 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9811296&pc=QCB | 2023826-2020-00447 | 2020-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.50/+3.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. LOW VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.50/+3.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM, PATIENT IS HAPPY. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXCHANGE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.50/+3.0/081 | OS |  | 02/13/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9811419&pc=QCB | 2023826-2020-00446 | 2020-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.50/+3.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. LOW VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.50/+3.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM, PATIENT IS HAPPY. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXCHANGE. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.50/+3.5/093 | OD |  | 02/13/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9814229&pc=QCB | 2023826-2020-00471 | 2020-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -7.50/+0.5/042 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.50/+0.5/042 | OD |  | 02/17/2020 | 2020 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9814241&pc=QCB | 2023826-2020-00470 | 2020-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -7.00/+0.5/103 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.00/+0.5/103 | OS |  | 02/17/2020 | 2020 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9855436&pc=QCB | 2023826-2020-00536 | 2020-02-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient; No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -8.50/1.0/067 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REPORTEDLY "PATIENT IS UNHAPPY WITH HER VISION." PATIENT WAS GIVEN AN RX OF ALPHAGAN P OU QHS. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA:) B5- IN INITIAL MDR STATEMENT "THE PATIENT WAS GIVEN AN RX FOR ALPHAGAN" IS NOT APPLICABLE. THE "RX" IS REFERENCE TO EYEGLASS PRESCRIPTION PROVIDED BY THE OPTOMETRIST. PATIENT IS CONCURRENTLY TAKING ALPHAGAN OU QHS. H6: PATIENT CODE 3191: RESIDUAL REFRACTION ERROR. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.50/1.0/067 | OD |  | 02/26/2020 | 2020 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9880942&pc=QCB | 2023826-2020-00665 | 2020-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.5/+4.0/102 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING AND REMAINS IMPLANTED. CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2020. THE LENS EXCHANGE RESOLVED THE PROBLEM. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: B5 - "-13.5/+4.0/102 (SPHERE/CYLINDER/AXIS)" CORRECTED TO "-13.5/+4.0/107 (SPHERE/CYLINDER/AXIS)". H6 - DEVICE CODE 4117 IS NO LONGER APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/+4.0/102 | OD |  | 03/04/2020 | 2020 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9880973&pc=QCB | 2023826-2020-00666 | 2020-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.00/+1.5/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE REPORTER INDICATED THE LENS WAS IMPLANTED INTO THE WRONG EYE, THE LENS FOR THE LEFT EYE (OS) WAS IMPLANTED INTO THE RIGHT EYE (OD). THE LENS WILL BE EXPLANTED AND TO DATE REMAINS IMPLANTED. CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE LENS WAS EXCHANGED ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. H6 - PATIENT CODE 2692 UPDATED TO PATIENT CODE 3191 (SECONDARY SURGICAL INTERVENTION - LENS EXCHANGE). H6 - DEVICE CODE 3189 UPDATED TO DEVICE CODE 2993. CLAIM# (B)(4). Manufacturer Narrative: B5 - 'THE LENS WAS EXCHANGED ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED' UPDATED TO: 'THE LENS WAS EXPLANTED FROM THE OD ON (B)(6) 2020 AND THE CORRECT LENS WAS IMPLANTED INTO THE OS. THE PROBLEM WAS RESOLVED.' CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.00/+1.5/073 | OD |  | 03/04/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9889401&pc=QCB | 2023826-2020-00585 | 2019-11-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -08.50/0.5/144 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC DEFORMED . CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA : H6-PATIENT CODE 3191 IN INITIAL MDR IS NOT APPLICABLE AND IS CORRECTED TO 2692. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -08.50/0.5/144 | OD |  | 03/03/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9937637&pc=QCB | 2023826-2020-00707 | 2019-08-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - 'SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES WAS ALSO REPORTED.' SHOULD BE ADDED TO INITIAL MDR. H6 - PATIENT CODE 2692 CORRECT TO PATIENT CODE 3191 (NO CODE AVAILABLE - SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES) IN INITIAL MDR. H6 - DEVICE CODE 1494 (OFF-LABEL USE - AGE < 21 YEARS OLD) SHOULD BE ADDED TO INITIAL MDR. CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/1.0/090 | OD |  | 03/11/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9938269&pc=QCB | 2023826-2020-00815 | 2019-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS HAD LOW VAULTING AND REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -11.50/2.5/090 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. D7- (B)(6) 2020. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. METHOD CODE 4117 IN INITIAL MDR IS NOT APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTIFUGE VIAL; DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/+2.5/090 | OS |  | 03/18/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10057316&pc=QCB | 2023826-2020-01071 | 2020-04-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -8.00/1.0/044 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, "SURGEON IMPLANTED LENS INTO THE WRONG EYE" OF PATIENT. "EXCHANGE IS PLANNED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT. B2- REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICE). B5- ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. D7-EXPLANTED DATE: (B)(6) 2020. H6-PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.00/1.0/044 | OD |  | 04/27/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10112279&pc=QCB | 2023826-2020-01151 | 2020-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, WAS RECEIVED DAMAGED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00/1.5/089 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. NAME AND ADDRESS: (B)(6). TELEPHONE: (B)(6). CLAIM # (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H3 - CORRECTION TO PREVIOUS LENS EVALUATION- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND ONE HAPTIC BENT. (B)(4). | -09.00/1.5/089 | OS |  | 05/11/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8452984&pc=QCB | 2023826-2019-00483 | 2019-02-27 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-26 | QCB | IMPLANTABLE COLLAMER LENS | Ejection Problem | No Patient Involvement | P030016 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS ATTEMPTING TO IMPLANT A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -15.5/+3.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE FOOTPLATES OF THE LENS GOT STUCK IN THE INJECTOR. THERE WAS NO PATIENT CONTACT WITH THE LENS. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED; "ALL WENT WELL WITH PATIENT INTERACTIONS". Manufacturer Narrative: LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -15.5/+3.0/089 | OD | Physician | 03/01/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8453013&pc=QCB | 2023826-2019-00487 | 2019-02-27 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-26 | QCB | IMPLANTABLE COLLAMER LENS | Ejection Problem | No Patient Involvement | P030016 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS ATTEMPTING TO IMPLANT A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -15.5/+3.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE FOOTPLATES OF THE LENS GOT STUCK IN THE INJECTOR. THERE WAS NO PATIENT CONTACT WITH THE LENS. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED; "ALL WENT WELL WITH PATIENT INTERACTIONS". Manufacturer Narrative: LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -15.5/+3.0/089 | OS | Physician | 03/01/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8480532&pc=QCB | 2023826-2019-00511 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-04-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | Event Description: THE REPORTER INDICATED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -8.50/+2.5/090 (SPHERE/CYLINDER/AXIS), WAS DAMAGED DURING LOADING AND THERE WAS NO PATIENT CONTACT OR INJURY. THE BACKUP LENS WAS USED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO THE LENS. THE REPORTER DID NOT PROVIDE ANY ADDITIONAL INFORMATION. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.50/+2.5/090 | NA | Other Health Care Professional | 03/07/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8569319&pc=QCB | 2023826-2019-00723 | 2019-03-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | Event Description: THE REPORTER INDICATED THAT 13.2MM, VTICM5_13.2, -13.00/1.0/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING LOADING. THERE WAS NO PATIENT CONTACT. A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTIFUGE VIAL. CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND FOREIGN MATERIAL ON LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.00/1.0/89 | NA |  | 04/03/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8692893&pc=QCB | 2023826-2019-01062 | 2019-05-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Patient Involvement | P030016 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS LOADING A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -12.0/+4.0/070 (SPHERE/CYLINDER/AXIS) IT TORE. THIS OCCURRED ON (B)(6) 2019 AND THERE WAS NO PATIENT CONTACT WITH THIS LENS. A BACK UP LENS WAS THEN USED. THE SURGEON DOES NOT KNOW THE CAUSE OF THE EVENT. Manufacturer Narrative: CORRECTED TO (B)(6) 2019 IN INITIAL MDR. CORRECTED TO (B)(6) 2019 IN INITIAL MDR. CORRECTED TO DEVICE EVALUATION: THE LENS WAS RETURNED STUCK IN THE CARTRIDGE. VISUAL INSPECTION OF THE LENS FOUND THE LENS RETURNED IN THE CARTRIDGE. THE CARTRIDGE WAS RETURNED IN THE LENS CASE WITH CLEAR SURGICAL RESIDUE. VISUAL INSPECTION OF THE CARTRIDGE FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE CARTRIDGE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED STUCK IN THE CARTRIDGE AND IN THE LENS CASE. THERE WAS CLEAR SURGICAL RESIDUE ON THE CARTRIDGE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE DEVICE. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.0/+4.0/070 | NA |  | 05/15/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8692895&pc=QCB | 2023826-2019-01061 | 2019-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Patient Involvement | P030016 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS LOADING A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -9.0/+4.0/108 (SPHERE/CYLINDER/AXIS) IT BROKE. THIS OCCURRED ON (B)(6) 2019 AND THERE WAS NO PATIENT CONTACT WITH THIS LENS. THE SURGEON DOES NOT KNOW THE CAUSE OF THE EVENT. Manufacturer Narrative: DEVICE EVALUATION: "RESIDUE ON THE LENS" CORRECTED TO "RESIDUE ON THE CARTRIDGE". Manufacturer Narrative: DATE OF REPORT CORRECTED TO 15-MAY-2019 IN INITIAL MDR. DATE RECEIVED BY MANUFACTURER CORRECTED TO 15-MAY-2019 IN INITIAL MDR. CORRECTED TO - DEVICE EVALUATION: THE LENS WAS RETURNED STUCK IN THE CARTRIDGE. VISUAL INSPECTION OF THE LENS FOUND THE LENS STUCK IN THE CARTRIDGE. THE CARTRIDGE WAS RETURNED IN THE LENS CASE WITH CLEAR SURGICAL RESIDUE. VISUAL INSPECTION OF THE CARTRIDGE FOUND NO VISIBLE DAMAGE TO THE CARTRIDGE AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED STUCK IN THE CARTRIDGE AND IN THE LENS CASE. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE DEVICE. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.0/+4.0/108 | NA |  | 05/15/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8841796&pc=QCB | 2023826-2019-01371 | 2019-06-27 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -10.0/1.5/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE WHILE LOADING. THERE WAS NO PATIENT CONTACT. ANOTHER LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A TEAR IN ONE OF THE HAPTIC. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.0/1.5/070 | NA | Non-Healthcare Professional | 07/09/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8895274&pc=QCB | 2023826-2019-01422 | 2019-07-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | Event Description: THE REPORTER INDICATED THE SURGEON REPORTED HE NOTICED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+3.0/091 (SPHERE/CYLINDER/AXIS) WAS TORN/BROKEN WHEN OPENING THE VIAL AND GRABBING THE LENS WITH FORCEPS. THERE WAS NO PATIENT CONTACT WITH THE LENS. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND A PIECE OF THE LENS OPTIC AND ONE HAPTIC TORN OFF AND MISSING. (B)(4). Manufacturer Narrative: UDI # CORRECTED FROM UNK TO N/A. INITIAL REPORTER NAME AND ADDRESS: INCORRECT DOCTOR NAME REPORTED IN INITIAL MDR CORRECTED. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. IMPLANTED: NA. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+3.0/091 | NA |  | 07/15/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8902740&pc=QCB | 2023826-2019-01460 | 2019-07-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -11.0/1.5/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS NOTED TO BE TORN IN THE INJECTOR. THERE WAS NO PATIENT CONTACT OF INJURY AND A BACKUP LENS WAS IMPLANTED. A STAAR INJECTOR SYSTEM WAS USED BUT MODEL AND LOT NUMBERS ARE UNKNOWN. THE REPORTER STATED " I THINK THE LENS TEAR WAS DUE TO THE FOAM TIP PLUNGER NOT SEATING TIGHTLY INTO THE INJECTOR". Manufacturer Narrative: (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/1.5/108 | NA |  | 07/17/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8912823&pc=QCB | 2023826-2019-01502 | 2019-07-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Patient Involvement | P030016 | Event Description: THE REPORTER INDICATED THAT AFTER OPENING THE VIAL AND PLACING ONTO THE CARTRIDGE, THE SURGEON NOTED THAT A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/069 (SPHERE/CYLINDER/AXIS) HAD A "SCRATCH" ON IT. THIS OCCURRED ON (B)(6) 2019 AND THERE WAS NO PATIENT CONTACT WITH THIS LENS. A BACK UP LENS WAS THEN USED. Manufacturer Narrative: LENS WORK ORDER SEARCH - NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A CARTRIDGE CASE. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: NO PATIENT INVOLVEMENT. (B)(4). | -10.0/+1.0/069 | NA |  | 07/29/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8962472&pc=QCB | 2023826-2019-01586 | 2019-08-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | Event Description: THE REPORTER INDICATED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, - 8.50/+2.5/087 (SPHERE/CYLINDER/AXIS), TORE DURING THE LOADING PROCESS, WHILE BEING INSERTED INTO THE CARTRIDGE. THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE STATEMENT IS CORRECTED TO; " LENS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. CARTRIDGE RETURNED IN LENS BOX WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO DEVICE AND RESIDUE ON CARTRIDGE. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. A CARTRIDGE WAS RETURNED AND THERE WAS NO VISIBLE DAMAGE TO THE DEVICE. ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | - 8.50/+2.5/087 | NA |  | 08/07/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9176558&pc=QCB | 2023826-2019-01931 | 2019-07-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | Event Description: THE REPORTER INDICATED THAT THE DOCTOR LOADED A 12.6MM VTICMO12.6, -15.00/1.0/179 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS ACCORDING TO THE INSTRUCTION MANUAL, BUT FOUND LENS DAMAGED WHEN DELIVERING THE LENS INTO THE NOZZLE. THERE WAS NO PATIENT CONTACT. SURGEON DISCONTINUED THE SURGERY BECAUSE HE HAD NO ALTERNATIVE LENS AND PLANS TO IMPLANT A NEW LENS IN THE FUTURE. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN WITH RESIDUE ON THE LENS AND TWO PIECES OF THE LENS WERE RETURNED. MSI-PF INJECTOR WAS RETURNED AND DEVICE EVALUATION FOUND CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO DEVICE WITH RESIDUE ON THE INJECTOR, CARTRIDGE WAS NOT RETURNED. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.00/1.0/179 | NA |  | 09/12/2019 | 2019 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9257733&pc=QCB | 2023826-2019-01963 | 2019-09-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | Event Description: THE REPORTER INDICATED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/081 (SPHERE/CYLINDER/AXIS), WAS DAMAGED DURING LOADING. THERE WAS NO PATIENT CONTACT. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK RACE: UNK. IMPLANTED: NA. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/+2.0/081 | NA |  | 10/03/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9338152&pc=QCB | 2023826-2019-02180 | 2019-08-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.5/2.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE BEFORE/DURING LOADING. THERE WAS NO PATIENT CONTACT. THE REPORTER STATED THAT THE DAMAGE WAS NOTED DURING PULLING TO THE TIP OF THE CARTRIDGE. RUSHED LOADING WAS REPORTED FOR CAUSE OF EVENT, THE REPORTER STATED, "THE DOCTOR IS EXPERIENCED AND ITS NOT COMMON FOR HER TO MESS UP THE LOADING OF LENS". ON (B)(6) 2019 A REPLACEMENT LENS WAS IMPLANTED, AND IT RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/2.5/90 | NA |  | 10/24/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9386207&pc=QCB | 2023826-2019-02243 | 2019-07-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | Event Description: THE REPORTER INDICATED THE SURGEON OPENED THE VIAL OF A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/097 (SPHERE/CYLINDER/AXIS), AND NOTED THE LENS WAS DAMAGED BEFORE REMOVING THE LENS FROM THE VIAL. THERE WAS NO PATIENT CONTACT AND THE SURGERY WAS NOT STARTED. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. IMPLANTED: NA. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.50/+1.0/097 | NA |  | 11/03/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9498043&pc=QCB | 2023826-2019-02406 | 2019-11-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | Event Description: THE REPORTER INDICATED THE SURGEON WAS LOADING A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/087 (SPHERE/CYLINDER/AXIS) INTO THE NOZZLE AND THE LENS WAS DAMAGED. THERE WAS NO PATIENT CONTACT. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.0/+2.0/087 | NA |  | 11/19/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9577753&pc=QCB | 2023826-2020-00084 | 2019-12-19 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | Event Description: THE REPORTER INDICATED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, - 10.00/+4.0/091 (SPHERE/CYLINDER/AXIS), WAS TORN WHILE LOADING AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/+4.0/091 | NA | Other Health Care Professional | 12/20/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9612926&pc=QCB | 2023826-2020-00149 | 2019-10-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Patient Involvement | P030016 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS LOADING A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/092 (SPHERE/CYLINDER/AXIS) IT TORE. THIS OCCURRED ON (B)(6) 2019 AND THERE WAS NO PATIENT CONTACT WITH THIS LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN AND A BACK-UP LENS WAS SUCCESSFULLY IMPLANTED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.5/+1.0/092 | NA |  | 12/26/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9684576&pc=QCB | 2023826-2020-00260 | 2019-12-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Material Opacification; Manufacturing, Packaging or Shipping Problem | No Patient Involvement | N/A | Event Description: THE REPORTER INDICATED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.00/+1.5/082 (SPHERE/CYLINDER/AXIS), WAS REMOVED FROM THE VIAL AND WAS NOTED TO BE OPAQUE. THE LENS WAS NOT USED AND THERE WAS NO PATIENT CONTACT. THE SURGEON IMPLANTED A LONGER LENS THAT WAS AVAILABLE, AND THE PROBLEM WAS RESOLVED. THE PATIENTS EYE IS STABLE. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: "THE REPORTER STATED THAT WHEN THE SURGEON OPENED THE LENS HE SAW SOMETHING LIKE BUBBLE ON THE LENS SURFACE." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: LENS WAS RETURNED IN LIQUID WITH CLEAR SURGICAL RESIDUE AND MOISTURE IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6- DEVICE HISTORY REVIEW- (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CORRECTED DATA: H6- DEVICE CODE: 2975 IS NOT APPLICABLE IN PREVIOUS MDR CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.00/+1.5/082 | NA |  | 01/15/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9706459&pc=QCB | 2023826-2020-00302 | 2019-12-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.50/1.5/068 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS DAMAGED DURING LOADING PROCEDURE. THERE WAS NO PATIENT CONTACT WITH THE DEVICE. THE REPORTER STATED " DOCTOR ALWAYS PREPARE THE LOADING OF THE LENS BEFORE SHE OPENS THE EYE." IN THE REPORTERS OPINION THE DEVICE WAS THE CAUSE OF THIS EVENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER STATED THAT THE LENS WAS DAMAGED DURING LOADING. THE SURGEON DAMAGED THE LENS WITH THE FORCEPS. H6- CONCLUSION CODE: 4315 TO 61. CLAIM #: (B)(4). Manufacturer Narrative: CORRECTED DATA: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.50/1.5/068 | NA |  | 01/15/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9706706&pc=QCB | 2023826-2020-00301 | 2020-01-06 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | Event Description: THE REPORTER INDICATED THAT THE NURSE NOTICED THAT A 12.1MM, VTICMO12.1,-12.5/+1.5/132 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS TORN WHEN SHE OPENED THE LENS VIAL. THE LENS WAS NOT USED. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - METHOD CODE 3331, DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALAUTION: LENS WAS IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.5/+1.5/132 | NA |  | 01/16/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9722922&pc=QCB | 2023826-2020-00336 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-02-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -11.0/2.5/064 (SPHERE/CYLUNDER/AXIS), IMPLANTABLE COLLAMER LENS TORE WHILE LOADING. THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/2.5/064 | NA | Other Health Care Professional | 01/22/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781716&pc=QCB | 2023826-2020-00395 | 2019-10-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Patient Involvement | NA | Event Description: THE REPORTER INDICATED THAT, WHILE ATTEMPTING TO IMPLANT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), IT WAS DISCOVERED TO BE TORN. REPORTEDLY, THE LENS TORE DURING LOADING. THIS OCCURRED ON (B)(6) 2019. THERE WAS NO PATIENT CONTACT WITH THE LENS. ON (B)(6) 2019 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY, WITH CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH DEBRIS ON THE LENS. CLAIM#: (B)(4).. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.0/+2.0/083 | OS |  | 02/10/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781725&pc=QCB | 2023826-2020-00402 | 2019-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Patient Involvement | NA | Event Description: THE REPORTER INDICATED THAT AFTER THE SURGEON REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/072 (SPHERE/CYLINDER/AXIS) FROM THE VIAL, IT WAS DISCOVERED THAT THE LENS WAS TORN/BROKEN. THIS OCCURRED ON (B)(6) 2019. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE LENS DID NOT TOUCH THE PATIENT'S EYE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. RETURNED PRODUCT EVALUATION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/+2.5/072 | NA |  | 02/10/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9819167&pc=QCB | 2023826-2020-00450 | 2020-02-11 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Patient Involvement | NA | Event Description: THE REPORTER INDICATED THAT WHILE ATTEMPTING TO INJECT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5/+0.5/125 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE, IT WAS DISCOVERED THAT THE LENS WAS TORN/BROKEN. THIS OCCURRED ON (B)(6) 2020. THE CAUSE OF THE EVENT IS REPORTED AS "DEVICE??" THE LENS DID NOT TOUCH THE PATIENT'S EYE AND THERE WAS NO PATIENT INJURY. REPORTEDLY, THE LENS WAS STUCK TO THE PLUNGER. THE SURGEON TRIED TO REMOVE AND RELOAD BUT IT WAS "CUT." AN ALTERNATE LENS WAS SUCCESSFULLY PLANTED ON (B)(6) 2020. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE WITH MOISTURE. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN. CLAIM# : (B)(4). Manufacturer Narrative: CORRECTED DATA : D10- LENS RETURNED TO MANUFACTURER IN PREVIOUS MDR SHOULD BE 02-APR-2020 CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/+0.5/125 | OS |  | 02/12/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9880944&pc=QCB | 2023826-2020-00654 | 2020-03-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Material Deformation | No Patient Involvement | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON NOTED THE EDGE OF A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, - 7.00/+1.5/092 (SPHERE/CYLINDER/AXIS), WAS NOT SMOOTH AND DID NOT USE. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS USED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. THE LENS WAS VIEWED UNDER 10X, THE TORN HAPTIC WAS NOTED BUT NO OTHER OBSERVATIONS. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | - 7.00/+1.5/092 | NA |  | 03/04/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910965&pc=QCB | 2023826-2020-00680 | 2019-12-18 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Patient Involvement | NA | Event Description: THE REPORTER INDICATED THAT, WHILE ATTEMPTING TO IMPLANT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), IT WAS TORN. THIS OCCURRED ON (B)(6) 2019. THERE WAS NO PATIENT CONTACT WITH THE LENS. ON (B)(6) 2019 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/+2.5/090 | OD |  | 03/06/2020 | 2020 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9924474&pc=QCB | 2023826-2020-00835 | 2020-03-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.0/1.0/062 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS FOUND DAMAGE BEFORE/DURING LOADING THE LENS. A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/1.0/062 | NA |  | 03/10/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9965164&pc=QCB | 2023826-2020-00896 | 2020-02-20 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO LOAD A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50/+3.5/087 (SPHERE/CYLINDER/AXIS), INTO THE CARTRIDGE AND IT WAS NOTED THE LENS WAS DAMAGED. THE SURGEON FELT THE LENS WAS TORN DURING MANIPULATING USING THE FORCEP. THERE WAS NO PATIENT CONTACT. THE LENS WAS EXCHANGED FOR THE BACKUP LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.50/+3.5/087 | NA |  | 03/24/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9988428&pc=QCB | 2023826-2020-00952 | 2020-03-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | Event Description: THE REPORTER INDICATED THE SURGEON WAS LOADING A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.0/+1.5/006 (SPHERE/CYLINDER/AXIS), AND FOUND THE LENS DAMAGED. THERE WAS NO PATIENT CONTACT. THE SURGEON IMPLANTED A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5. IT WAS REPORTED CAUSE OF THE EVENT WAS USER ERROR. SHOULD HAVE BEEN INCLUDED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.0/+1.5/006 | NA |  | 04/03/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10116211&pc=QCB | 2023826-2020-01158 | 2020-01-16 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-03 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Patient Involvement | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON ATTEMPTED TO IMPLANT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/+1.5/019 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS BECAME STUCK IN THE CARTRIDGE/INJECTION SYSTEM. THIS OCCURRED ON (B)(6) 2020. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED INTRAOPERATIVELY AND THE PROBLEM WAS RESOLVED. THERE WAS NO PATIENT CONTACT WITH THE LENS AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.0/+1.5/019 | OD | Physician | 05/11/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10135940&pc=QCB | 2023826-2020-01168 | 2020-05-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Material Integrity Problem; No Apparent Adverse Event | No Patient Involvement | NA | Event Description: THE REPORTER INDICATED THE SURGEON FOUND A "LINE" ON THE LENS SURFACE OF A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+2.0/081 (SPHERE/CYLINDER/AXIS). THE SURGEON LOADED THE LENS BUT WAS WORRIED ABOUT THE SURGERY EFFECT AND VISUAL QUALITY OF THE LENS, SO PUT LENS BACK IN THE VIAL. THE SURGEON NOTED THE LENS WAS CRACKED IN CENTER OF LENS. THE LENS WAS NOT USED AND THERE WAS NO PATIENT CONTACT. THE REPLACEMENT LENS WAS IMPLANTED THE NEXT DAY AND THE PROBLEM WAS RESOLVED. THE PATIENTS CURRENT CONDITION IS GREAT. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.00/+2.0/081 | NA |  | 05/13/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10189346&pc=QCB | 2023826-2020-01280 | 2019-06-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | Event Description: THE REPORTER INDICATED THE SURGEON WAS LOADING A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.00/+2.5/098 (SPHERE/CYLINDER/AXIS), IN THE CARTRIDGE AND THE LENS TORE. THERE WAS NO PATIENT CONTACT. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.00/+2.5/098 | NA |  | 06/01/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10242559&pc=QCB | 2023826-2020-01433 | 2020-06-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | Event Description: THE REPORTER INDICATED THE SURGEON REPORTED WHILE REMOVING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+3.0/110 (SPHERE/CYLINDER/AXIS), FROM THE VIAL, IT WAS NOTED THE LENS WAS DAMAGED/SCRATCHES. THE LENS WAS NOT USED AND THERE WAS NO PATIENT CONTACT. THE REPORTER INDICATED THE LENS WAS RECEIVED IN GOOD CONDITION. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH DEBRIS ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.50/+3.0/110 | NA |  | 06/18/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10112184&pc=QCB | 2023826-2020-01149 | 2018-04-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement; No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.50/+1.0/089 (SPHERE/CYLINDER/AXIS), WAS TORN DURING INJECTION INTO THE PATIENT'S LEFT EYE (OS) AND THERE WAS NO PATIENT CONTACT OR INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS, ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: B5- THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -7.50/+1.0/089 (SPHERE/CYLINDER/AXIS) GOT STUCK IN THE CARTRIDGE OR INJECTION SYSTEM ON 26/04/2018. THERE WAS NO PATIENT CONTACT OR INJURY. REPORTEDLY "THE LENS WAS DAMAGED DURING THE INJECTION." ON (B)(6) 2018 A REPLACEMENT LENS OF THE SAME MODEL AND LENGTH WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.50/+1.0/089 | OS | Physician | 05/11/2020 | 2020 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8574021&pc=QCB | 2023826-2019-00722 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement; No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -14.0/1.5/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS FOUND TO BE BROKEN AND WITHOUT ASTIGMATISM LINES. NO PATIENT CONTACT. Manufacturer Narrative: THE DAMAGE (BREAK) WAS NOTICED WHEN THE LENS WAS REMOVED FROM THE LENS VIAL AND LOADED. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A TRANSPARENT BAG. VISUAL INSPECTION FOUND THAT A HAPTIC WAS BROKEN AND CONFIRMED THAT THE LENS HAS ASTIGMATISM MARKERS. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.0/1.5/070 | NA |  | 04/03/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966509&pc=QCB | 2023826-2019-01628 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Patient Involvement; No Known Impact Or Consequence To Patient | P030016 | Event Description: ON (B)(6) 2019 A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -16.0/+1.0/074 (SPHERE/CYLINDER/AXIS) WAS RETURNED TO STAAR SURGICAL. IT WAS RETURNED STILL INSIDE OF THE LENS CARTRIDGE AND THE BOX HAD "RESIDUE ON LENS" WRITTEN ON IT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: EVENT UPDATED INFORMATION: ON 06-SEP-2019 IT WAS DISCOVERED THAT THIS EVENT IS NOT MDR-REPORTABLE. HOWEVER, THE EVENT HAD ALREADY BEEN REPORTED. REPORTER STATES THAT THE SURGEON BELIEVES THE FORCEPS USED TO GRASP THE LENS HAD RUST ON THEM. THE SURGEON THEN DID NOT WANT TO RISK USING THE LENS. NO PATIENT CONTACT OCCURRED. CLAIM# (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED STILL INSIDE THE CARTRIDGE, INSIDE THE LENS BOX. VISUAL INSPECTION FOUND CLEAR SURGICAL RESIDUE ON THE LENS. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.0/+1.0/074 | NA |  | 08/07/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9597959&pc=QCB | 2023826-2020-00129 | 2019-11-14 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement; No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -08.50/2.0/105 (SPHERE/CYLINDER/AXIS), INTRAOCULAR LENS TORE DURING LOADING ON (B)(6) 2019. THERE WAS NO PATIENT CONTACT. A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -08.50/2.0/105 | NA | Physician | 12/26/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10225159&pc=QCB | 2023826-2020-01385 | 2020-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement; No Known Impact Or Consequence To Patient | NA | Event Description: A LENS TEAR/BREAK BEFORE LOADING ON A 12.6MM, VTICMO12.6, -13.5/3.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS REPORTED. IT WAS REPORTED THAT THE DAMAGE WAS NOTED WHILE REMOVING FROM VIAL. THE LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) AND REMOVED WITHIN THE SAME SURGERY. ON (B)(6) 2020 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT IS UNKNOWN . ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DAT: B4-DATE OF THIS REPORT IN INITIAL MDR SHOULD BE CORRECTED TO (B)(6) 2020 G4- DATE RECEIVED BY MANUFACTURER IN INITIAL MDR SHOULD BE CORRECTED TO (B)(6) 2020 CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION : H3: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION : B5- THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.5/3.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE/BROKE WHILE REMOVING FROM VIAL ON (B)(6) 2020. THE LENS WAS NOT IMPLANTED. THERE WAS NO PATIENT CONTACT. ON (B)(6) 2020 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. D6-IMPLANT DATE (B)(6) 2020 IN INITIAL MDR IS NOT APPLICABLE. D7-EXPLANTED DATE (B)(6) 2020 IN INITIAL MDR IS NOT APPLICABLE. CORRECTED DATA H6-PATIENT CODE 2692 IN INITIAL MDR IS CORRECTED TO 2645. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/3.5/089 | OS |  | 06/05/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8816879&pc=QCB | 2023826-2019-01218 | 2018-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Not Applicable | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -12.0/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/2.0/088 | OS |  | 06/24/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8816886&pc=QCB | 2023826-2019-01217 | 2019-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Not Applicable | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -14.0/2.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.0/2.0/079 | OD |  | 06/24/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8816888&pc=QCB | 2023826-2019-01222 | 2019-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Not Applicable | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICMO12.6, -9.0/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.0/2.0/089 | OS |  | 06/26/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8816887&pc=QCB | 2023826-2019-01219 | 2019-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Not Applicable; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICMO12.6, -8.0/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PATIENT CODE:3189 SHOULD BE CORRECTED TO 3191 IN INITAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS: FOAM TIP PLUNGER MODEL-FTP; LOT#-UNK SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) SHOULD HAVE BEEN INLCUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.0/2.5/090 | OS |  | 06/27/2019 | 2019 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8816890&pc=QCB | 2023826-2019-01221 | 2019-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Not Applicable; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICMO12.6, -7.5/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PATIENT CODE: 3189 SHOULD BE CORRECTED TO 3191 IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS: FOAM TIP PLUNGER MODEL-FTP; LOT#-UNK SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. PATIENT CODE 1494: OFF LABEL-USE (PATIENT'S ACD <3.00MM) SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.5/2.5/092 | OD |  | 06/27/2019 | 2019 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12073178&pc=QCB | 2023826-2021-02032 |  | Injury | STAAR SURGICAL COMPANY | 2021-06-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Pain | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.50/+2.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT IS EXPERIENCING PAIN OVER THE INNER EYE FOR NEARLY 10 MONTHS. " UNEVENTFUL SURGERY, GOOD VISION, GOOD POSITIONING AND GOOD VAULTING, CLEAR CORNEA, CLEAN LENS, NO PIGMENT DISPERSION" WAS REPORTED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/+2.0/092 | OD | Physician | 05/28/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12073179&pc=QCB | 2023826-2021-02031 |  | Injury | STAAR SURGICAL COMPANY | 2021-06-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Pain | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.50/+2.5/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT IS EXPERIENCING PAIN OVER THE INNER EYE FOR NEARLY 10 MONTHS. "UNEVENTFUL SURGERY, GOOD VISION, GOOD POSITIONING AND GOOD VAULTING, CLEAR CORNEA, CLEAN LENS, NO PIGMENT DISPERSION" WAS REPORTED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.50/+2.5/088 | OS | Physician | 05/28/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14620821&pc=QCB | 2023826-2022-01794 | 2022-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -05.00/+1.0/029 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT WAS COMPLAINING OF PAIN, POOR ABILITY TO READ AND LOW VAULT WAS REPORTED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5 - THE LENS WAS EXPLANTED ON (B)(6) 2022 AND PROBLEM RESOLVED. REPORTEDLY, "ICL HAS BEEN EXPLANTED, PATIENT GOES BACK TO USING REFRACTIVE APPLIANCES. CAUSE OF EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). | -05.00/+1.0/029 | OD | Physician | 09/06/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14639450&pc=QCB | 2023826-2022-01795 | 2022-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00/+1.0/161 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT WAS COMPLAINING OF PAIN, POOR ABILITY TO READ AND LOW VAULT WAS REPORTED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5 - THE LENS WAS EXPLANTED ON (B)(6) 2022 AND PROBLEM RESOLVED. REPORTEDLY, "ICL HAS BEEN EXPLANTED, PATIENT GOES BACK TO USING REFRACTIVE APPLIANCES. CAUSE OF EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -08.00/+1.0/161 | OS | Physician | 09/06/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17704775&pc=QCB | 2023826-2023-03953 | 2023-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.0/+2.0/180 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON EXPLANTED THE LENS ON (B)(6) 2023, DUE TO THE PATIENT HAVING DISSATISFACTION AND FOREIGN BODY SENSATION. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - PATIENT DISSATISACTION, FOREIGN SENSATION. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5- PATIENT ALSO HAD PAIN. CLAIM#: (B)(4). | -6.0/+2.0/180 | OS | Physician | 08/29/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652119&pc=QCB | 2023826-2019-00915 | 2019-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Pain; Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.0/+2.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES LOW REFRACTIVE SURPRISE, REFRACTIVE CHANGE OVER TIME, BLURRED VISION, PAIN, AND "MINOR TEMPORALLY ICL SHIFT." CORNEA IS CLEAR, NO INJUNCTIVAL INJECTION, IOP AT 14MMHG AT LAST VISIT. PATIENT CLAIMS NO HISTORY OF EYE INJURY. Manufacturer Narrative: PRODUCT PROBLEM: SHOULD BE CORRECTED TO ADVERSE EVENT IN INITIAL MEDWATCH REPORT. REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) SHOULD BE MARKED IN INITIAL MEDWATCH REPORT. DATE OF THIS REPORT: SHOULD BE CORRECTED TO 02/MAY/2019 IN INITIAL MEDWATCH REPORT. DATE RECEIVED BY MANUFACTURER: SHOULD BE CORRECTED TO 02/MAY/2019 IN INITIAL MEDWATCH REPORT. POST-OP PATIENT UCVA WAS RECORDED AS 20/25 AND THE VAULT WAS 458UM. REPOSITIONING WAS CONDUCTED ON A UNSPECIFIED DATE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+2.0/093 | OS |  | 05/02/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10454797&pc=QCB | 2023826-2020-01864 | 2020-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Pain; Blurred Vision; No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -10.00/1.5/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. PUPIL BLOCK WITH ELEVATED IOP (37MMHG) OBSERVED ON (B)(6) 2020. ADDITIONAL TREATMENT OF ENLARGEMENT OF PI AND A YAG WAS NECESSARY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS NON PATENT PI. REPORTEDLY, PATIENT CONDITION "IMPROVED - NO COMPLAINTS." Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS ALSO REPORTED THAT THE PATIENT EXPERIENCED BLURRY VISION WITH EXTREME PAIN. H6 - DEVICE CODE: 2993 SHOULD BE REMOVED AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/1.5/111 | OD |  | 08/03/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8525151&pc=QCB | 2023826-2019-00599 | 2019-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Pain; No Code Available | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -9.0/+4.0/070 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2019 AND PROBLEM WAS RESOLVED. THE PATIENT EXPERIENCED PAIN AND THE PAIN WAS RESOLVED WITH ICL EXCHANGE. THE CAUSE OF THE EVENT WAS DUE TO SIZING NOMOGRAM. THE PATIENT HAD LASIK OU ON (B)(6) 2019 TO CORRECT RESIDUAL ASTIGMATISM. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.0/+4.0/070 | OD |  | 03/18/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8525163&pc=QCB | 2023826-2019-00598 | 2019-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Pain; No Code Available | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -8.0/+4.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS IMPLANTED ON 25-FEB-2019 AND EXPLANTED ON 27-MAR-2019. THE PATIENT EXPERIENCED PAIN AND THE PAIN WAS RESOLVED AFTER ICL EXCHANGE.. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT HAD LASIK OU ON 14-MAY-2019 TO CORRECT RESIDUAL ASTIGMATISM. THE CAUSE OF THE EVENT WAS DUE TO SIZING NOMOGRAM. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.0/+4.0/094 | OS |  | 03/18/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10135938&pc=QCB | 2023826-2020-01166 | 2019-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Pain; No Code Available | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -9.00/+2.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) IN 2019. THE REPORTER HAS INDICATED THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS FOR A SHORTER LENS DUE TO THE LENS HAS VAULTED FORWARD AND POSSIBLY RUBBING AGAINST THE IRIS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5:THE LENS WAS IMPLANTED ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION AND INTERMITTENT PAIN IN THE DARK. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. H6 - PATIENT CODE: 3191 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; PIGMENT DISPERSION. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.00/+2.5/094 | OS |  | 05/18/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11478865&pc=QCB | 2023826-2021-00636 | 2021-02-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00/+1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE PUPIL WAS UNREACTIVE (FIXED) (BLOCKED PUPIL). THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE PATIENT IS NOW 8/10 IN THE LEFT EYE AND THE PUPIL IS ALMOST PERFECT BUT STILL NOT 100%. THE PATIENT IS BEING MONITORED. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. IMPLANT DATE: UNK THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT.CLAIM # (B)(4). | -08.00/+1.5/094 | OS |  | 02/17/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13311393&pc=QCB | 2023826-2022-00176 |  | Injury | STAAR SURGICAL COMPANY | 2022-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Pupillary Block; Blurred Vision |  | Event Description: THE REPORTER INDICATED THAT 12.6MM TMICL12.6; -7.00/+3.50/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021; AS A REPLACEMENT LENS. IT WAS NOTED THE PATIENT EXPERIENCED PUPIL BLOCK ON THE REPLACEMENT LENS WHICH WAS CAUSED FROM THE INITIAL IMPLANTED LENSES ANGLE CLOSURE. THE PATIENT THEN FELT VISION WASN'T GOOD AND STILL HAD A BLOCK PUPIL WITH RESIDUAL RX RIGHT EYE -0.25; +1.25; 145. THE SURGEON DECIDED TO PERFORM PRK SURGERY. IT WAS INDICATED THE PATIENT'S VISION IS STILL BLURRY. SEE MDR# 2023826-2022-00174 FOR INITIAL IMPLANT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER INDICATED THAT 12.6MM TMICL12.6; -7.00/+3.50/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021; AS A REPLACEMENT LENS. IT WAS NOTED THE PATIENT EXPERIENCED PUPIL BLOCK ON THE REPLACEMENT LENS WHICH WAS CAUSED FROM THE INITIAL IMPLANTED LENSES ANGLE CLOSURE. THE PATIENT THEN FELT VISION WASN'T GOOD AND STILL HAD A BLOCK PUPIL WITH RESIDUAL RX RIGHT EYE -0.25; +1.25; 145. THE SURGEON DECIDED TO PERFORM PRK SURGERY. IT WAS INDICATED THE PATIENT'S VISION IS STILL BLURRY. SEE MDR# 2023826-2022-00174 FOR INITIAL IMPLANT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -7.00/+3.50/105 | OD | Other Health Care Professional | 01/04/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16739723&pc=QCB | 2023826-2023-01370 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Red Eye(s); Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -9.00/1.0/071 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. IT WAS INDICATED THAT EVERYONE WAS NORMAL FOR THE 1ST FOUR DAYS AFTER SURGERY BUT ON THE 5TH DAY THE EYE WAS DILATED; WAS PRESCRIBED MEDICATION (PILOCARPINE); BUT PIGMENT DISPERSION STARTED AND RED EYE AND SERVER EYE PAIN. A DIFFERENT SURGEON HAD REPOSITIONED THE LENS BUT THIS DID NOT RESOLVE THE PROBLEM. A 3RD DOCTOR HAD PRESCRIBED ALPHAGAN P. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - DILATED PUPIL, PIGMENT DISPERSION. TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -9.00/1.0/071 | OD | Physician | 03/22/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13502410&pc=QCB | 2023826-2022-00410 | 2021-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Retinal Detachment; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.50/+1.5/123 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE REPORTER INDICATED THE REFRACTION WAS NOT STABLE AND THE LENS REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2021-02786. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.50/+1.5/123 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE REPORTER INDICATED THE REFRACTION WAS NOT STABLE AND THE LENS REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2021-02786. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE CAUSE OF THE REPORTED POSTOPERATIVE UNSTABLE REFRACTION WAS DUE TO A RETINAL DETACHMENT WHICH WILL BE TREATED AT THE MOMENT. THE RETINAL DETACHMENT WAS NOT CAUSED BY THE EVO VISIAN ICL. IT SEEMED TO BE CAUSED BY THE FURTHER HIGH MYOPIA OF THE PATIENT EYE. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: UNK. GMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: UNK. GMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -16.50/+1.5/123 | OD | Physician | 01/11/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8586209&pc=QCB | 2023826-2019-00734 | 2019-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Retinal Detachment; No Code Available | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A TORIC IMPLANTABLE COLLAMER LENS, IN (B)(6) 2019. THE REPORTER INDICATED THERE WAS A RETINAL DETACHMENT AND THE SURGERY WAS UNEVENTFUL. THE SURGEON INDICATED THE PATIENT WAS A VERY HIGH MYOPIC, WHICH WAS HER RISK FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/169 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENTS LEFT EYE (OS) ON 03/26/2019. THE PATIENT HAD A RETINAL DETACHMENT AND THE RETINA WAS REPAIRED. A SCLERAL BUCKLE WAS PLACED, WITH 20% GAS FILL AT THAT VISIT. THE LENS REMAINS IMPLANTED AND THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. THE PATIENT IS DOING WELL. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, RETINAL REPAIR, SCLERAL BUCKLE IN PLACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). | -10.5/+1.0/169 | OS |  | 04/08/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15099175&pc=QCB | 2023826-2022-02430 | 2022-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Swelling/ Edema | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -15.0/1.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND SWELLING; IT WAS NOTED "THE PATIENT CONTINUED TO FEEL THE WHOLE SWELLING; WHICH WAS NOT TOLERATED AFTER SEVERAL CONSUTATION". ON (B)(6) 2022 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT. LENS RETURNED IS NOT MODEL STATED IN THE CLAIM, THEREFORE DIMENSIONAL INSPECTION IS NOT REQUIRED. CLAIM#: (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -15.0/1.0/093 | OD | Physician | 06/27/2022 | 2022 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11448805&pc=QCB | 2023826-2021-00610 |  | Injury | STAAR SURGICAL COMPANY | 2021-03-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Swelling/ Edema; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/072 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS EDEMA, PACHYMETRY AT 600 MICRONS, WITH LIMITED VISION 20/25 AND REFRACTION -0.5. THE DOCTOR BELIEVES THIS SHOULD IMPROVE WITH TIME. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+1.5/072 | OS | Physician | 02/12/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13085403&pc=QCB | 2023826-2021-04924 | 2021-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS) | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/+3.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. IT WAS REPORTED TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS OBSERVED. STEROIDS WERE PRESCRIBED. AT THE POST-OP VISIT ON (B)(6) 2021, THE PROBLEM WAS RESOLVED AND STEROIDS WERE NO LONGER PRESCRIBED. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/+3.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. IT WAS REPORTED TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS OBSERVED. STEROIDS WERE PRESCRIBED. AT THE POST-OP VISIT ON (B)(6) 2021, THE PROBLEM WAS RESOLVED AND STEROIDS WERE NO LONGER PRESCRIBED. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: NA. PMA/510/K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: NA. PMA/510/K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). | -14.50/+3.0/090 | OS | Physician | 11/30/2021 | 2021 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13182687&pc=QCB | 2023826-2022-00035 |  | Injury | STAAR SURGICAL COMPANY | 2022-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS) | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/106 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS TASS WAS OBSERVED WHICH SUBSIDED WITH MEDICAL TREATMENT. THE LENS IS STABLE IN THE EYE. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). | -10.0/+2.0/106 | OS | Physician | 12/13/2021 | 2021 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13958873&pc=QCB | 2023826-2022-00953 | 2021-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | Toxic Anterior Segment Syndrome (TASS) |  | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -10.5/2.5/177 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO TOXIC ANTERIOR SEGMENT SYNDROME (TASS). DIAGNOSTIC PERFORMED: INTENSE INFLAMMATION WITHIN THE ANTERIOR CHAMBER. TREATMENT PERFORMED: ANTIMICROBIAL TREATMENT AND ANTI-INFLAMMATORY TREATMENT WERE GIVEN. THE REPLACEMENT LENS RESOLVED THE PROBLEM. REPORTER STATES "IT IS DIFFICULT TO DETERMINE THE CAUSE. Manufacturer Narrative: WEIGHT - RACE:UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). | -10.5/2.5/177 | OS | Physician | 03/07/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14190878&pc=QCB | 2023826-2022-01293 | 2022-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS) | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -18.0/+1.0/084 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTS TASS-DIAGNOSTICS WERE NOT PERFORMED AND MEDICATION WAS ADMINISTERED. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR: THE DEVICE DID NOT FAIL TO PERFORM AS EXPECTED. THE SURGEON ALSO REPORTS, "ALLERGIES TO LENSES." Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: H6-INVESTIGATION CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW-BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: B5 - CAUSE OF THE EVENT WAS PATIENT-RELATED FACTOR, NOT USER ERROR. CLAIM# (B)(4). | -18.0/+1.0/084 | OD | Physician | 03/28/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15153793&pc=QCB | 2023826-2022-02512 | 2022-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS) | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -06.00/+1.0/009 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE REPORTER INDICATED TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS FOUND AFTER SURGERY. TREATMENT WAS PERFORMED - ANTERIOR CHAMBER LAVAGE AND KENACORT CONJUNCTIVAL INJECTION. RECOVERY TREND WITH STEROIDS TREATMENT. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | -06.00/+1.0/009 | OD | Physician | 07/07/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15681653&pc=QCB | 2023826-2022-03832 | 2022-05-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS) | P030017 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -6.0/2.0/075 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS OBSERVED. NO DIAGNOSTICS WERE PERFORMED. CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. LENS REMAINS IMPLANTED. LINDELON WAS PRESCRIBED. REPORTEDLY, "AFTER 2 MONTHS PRESCRIPTION OF LINDELON, SHE IS ON THE MEND." PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. STATEMENT SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -6.0/2.0/075 | OS | Physician | 09/28/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15681668&pc=QCB | 2023826-2022-03833 | 2022-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS) | P030017 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -7.5/1.0/179 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS OBSERVED. NO DIAGNOSTICS WERE PERFORMED. CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. LENS REMAINS IMPLANTED. LINDELON WAS PRESCRIBED. REPORTEDLY, "AFTER 1 MONTHS PRESCRIPTION OF LINDELON, SHE IS ON THE MEND." PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. STATEMENT SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT#: (B)(4) IS ASSOCIATED WITH THE PATIENT'S FELLOW EYE (OS). (B)(4). | -7.5/1.0/179 | OD | Physician | 09/29/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15681720&pc=QCB | 2023826-2022-03834 | 2022-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS) | P030017 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -7.5/1.0/179 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS OBSERVED. NO DIAGNOSTICS WERE PERFORMED. CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. LENS REMAINS IMPLANTED. LINDELON WAS PRESCRIBED. REPORTEDLY, "AFTER 1 MONTHS PRESCRIPTION OF LINDELON, SHE IS ON THE MEND." PROBLEM WAS RESOLVED. Manufacturer Narrative: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT #(B)(4) IS ASSOCIATED WITH THE PATIENT'S FELLOW EYE (OD). CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. - STATEMENT SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). | -7.5/1.0/179 | OS | Physician | 09/29/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15765123&pc=QCB | 2023826-2022-03998 | 2022-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | Toxic Anterior Segment Syndrome (TASS) | P030016 | Event Description: THE REPORTER STATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -10.50/1.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) (B)(6) 2022. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS REPORTED. THE LENS REMAINS IMPLANTED. STEROID EYE DROPS AND ORAL MEDICATION WERE PRESCRIBED. REPORTER STATES THE PROBLEM HAS NOT RESOLVED AND PATIENT OUTCOME STATUS IS UNKNOWN. CAUSE REPORTED AS DEVICE. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | -10.50/1.0/089 | OD | Physician | 10/13/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15765127&pc=QCB | 2023826-2022-03999 | 2022-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | Toxic Anterior Segment Syndrome (TASS) | P030016 | Event Description: THE REPORTER STATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -10.00/3.0/082 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) (B)(6) 2022. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS REPORTED. THE LENS REMAINS IMPLANTED. STEROID EYE DROPS AND ORAL MEDICATION WERE PRESCRIBED. REPORTER STATES THE PROBLEM HAS NOT RESOLVED AND PATIENT OUTCOME STATUS IS UNKNOWN. CAUSE REPORTED AS DEVICE. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | -10.00/3.0/082 | OS | Physician | 10/13/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17515441&pc=QCB | 2023826-2023-03338 | 2023-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | Uveitis | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -6.0/1.5/106 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. REPORTEDLY THE SURGERY WAS UNEVENTFUL, GOOD VISUAL RESULT ACHIEVED. POST OPERATIVELY MAXIDEX (STEROID DROPS) UNABLE TO STOP USE AS ONCE CEASED UVEITIS FLARE UP OCCURS. LENS REMAIN IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6-HEALTH IMPACT 4644: MAXIDEX DROPS. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -6.0/1.5/106 | OD | Physician | 08/01/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17515558&pc=QCB | 2023826-2023-03339 | 2023-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | Uveitis | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -6.0/2.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. REPORTEDLY THE SURGERY WAS UNEVENTFUL, GOOD VISUAL RESULT ACHIEVED. POST OPERATIVELY MAXIDEX (STEROID DROPS) UNABLE TO STOP USE AS ONCE CEASED UVEITIS FLARE UP OCCURS. LENS REMAIN IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. . Manufacturer Narrative: ADDITIONAL INFORMATION: H6-HEALTH IMPACT 4644: MAXIDEX DROPS. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -6.0/2.0/090 | OS | Physician | 08/01/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17777436&pc=QCB | 2023826-2023-04124 | 2023-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Uveitis | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -14.00/+1.00/133 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS THE PATIENT HAS UVEITIS WHICH WAS NOTED (B)(6) 2023. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -14.00/+1.00/133 | OS | Physician | 09/04/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18164007&pc=QCB | 2023826-2023-05151 |  | Injury | STAAR SURGICAL COMPANY | 2023-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Uveitis | P030016 | Event Description: INFLAMMATORY REACTION 9 DAYS AFTER IMPLANTABLE COLLAMER LENS IMPLANTATION WAS REPORTED. RPORTER IS UNCERTAIN OF THE ETIOLOGY BUT IS TREATING IT AS UVEITITS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). | -7.0/2.0/072 | OS | Physician | 10/20/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15849035&pc=QCB | 2023826-2022-04266 | 2022-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Uveitis; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -10.0/3.0/096 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ENDOPHTHALMITIS WAS REPORTED. DIAGNOSTICS WA PERFORME 'UVEITIS'. TREATMENT WAS PERFORMED: MINOCIN CAP 50MG; PREDNISOLONE 5MG; AXADIN CAP 150MG; VIGAMOX EYE DROPS 0.5%. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS BUT DIFFERENT AXIS. THE REPLACEMENT LENS RESOLVED THE PROBLEM. "IMPROVEMENT IN SYMPTOMS AFTER REPLACEMENT." Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -10.0/3.0/096 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. IN THE COMPLAINT QUESTIONNAIRE, THE BOX MARKED ENDOPHTHALMITIS WAS TICKED AND THE DIAGNOSTICS PERFORMED WAS DESCRIBED AS 'UVEITIS'. LATER CLARIFICATION FROM THE REPORTER STATED THAT ENDOPHTHALMITIS WAS NOT DIAGNOSED NOR WAS THERE ANY EVIDENCE OF THIS. AT THIS TIME, MEDICAL MANAGEMENT IN THE FORM OF ANTIBIOTICS (TOPICAL AND SYSTEMIC) WITH SYSTEMIC STEROIDS WITH GASTRIC PROTECTION WAS PRESCRIBED. FOLLOWING THIS, THE REPORTER NOTED A "SIGNIFICANT IMPROVEMENT IN SYMPTOMS". HOWEVER SUBSEQUENT COMMUNICATION STATED THAT "INFLAMMATION OCCURRED A WEEK AFTER WHEN STEROIDS DISCONTINUED". THREE WEEKS POST-IMPLANTATION THE LENS WAS EXCHANGED FOR A NEW LENS OF SAME PARAMETERS WHICH RESOLVED THE ISSUE. H65: CLINICAL CODE: 4581:ENDOPHTHALMITIS IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: CLINICAL CODE: (B)(4):ENDOPHTHALMITIS HEALTH IMPACT : 4644:MINOCIN CAP 50MG; PREDNISOLONE 5MG; AXADIN CAP 150MG; VIGAMOX EYE DROPS 0.5%. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.0/3.0/096 | OD | Physician | 10/27/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930479&pc=QCB | 2023826-2020-02833 | 2020-09-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Uveitis; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXPLANTED, DIAGNOSTICS AND TREATMENT PERFORMED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. UVEITIS; PUPIL MEMBRANE WAS CLOSED, IRIS SYNECHIA WAS REPORTED. TREATMENT OF HORMONES; NON STEROIDAL ANTI-INFLAMMATORY DRUGS, MYDRIASIS TREATMENT, TREATMENT IMPROVED, REPEATED ATTACKS, AFTER CONSULTATION DECIDED TO TAKE OUT THE LENS. ON (B)(6) 2020 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. H3- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. H6 CLINICAL CODE: 4581: PUPIL MEMBRANE CLOSED; IRIS SYNECHIA. HEALTH IMPACT: 4625: MYDRIASIS TREATMENT. HEALTH IMPACT : 4644: HORMONES; NON-STEROIDAL ANTI- INFLAMMATORY DRUGS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.0/+2.0/089 | OD |  | 11/05/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16010278&pc=QCB | 2023826-2022-04648 | 2022-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Adverse Event Without Identified Device or Use Problem | Uveitis; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -1.50/+3.5/078 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE SURGEON REPORTS UVEITIS AFTER SURGERY WITHOUT PREVIOUS HISTORY; DEPOSITS IN ICL THAT DECREASE VISUAL ACUITY; PATIENT IS TREATED WITH CORTICOSTEROIDS; AND THERE IS PIGMENT DISPERSION. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN; AND DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CORRECTED DATA: H6: MEDICAL DEVICE PROBLEM CODE-1494 SHOULD BE ADDED FOR CORRECTION. OFF-LABEL USAGE. CLAIM#(B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -1.50/+3.5/078 | OD | Physician | 11/21/2022 | 2022 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9706691&pc=QCB | 2023826-2020-00304 | 2020-01-10 | Injury | STAAR SURGICAL COMPANY | 2020-02-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Uveitis; Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.00/1.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. PATIENT WAS DIAGNOSED WITH UVEITIS ON (B)(6) 2020 (VISUAL ACUITY 6/4.5-2) AND WAS PRESCRIBED PRED FORTE. AT PATIENT'S (B)(6) 2020 FOLLOW-UP, BLURRY VISION WAS OBSERVED (VISUAL ACUITY: 6/7.5-1, PINHOLE: 6/6-1). PATIENT WAS PRESCRIBED CIPROFLOXACIN, DEXAMETHASONE, CYCLOPENTOLATE, CRAVIT. PER REPORTER ON (B)(6) 2020, "PATIENT HAVE FULLY RECOVERED FROM THE UVEITIS." LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.00/1.0/075 | OS | Physician | 01/21/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930493&pc=QCB | 2023826-2020-02834 | 2020-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Uveitis; No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED, DIAGNOSTICS AND TREATMENT PERFORMED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. UVEITIS WAS REPORTED. TREATMENT OF HORMONES; NON STEROIDAL ANTI-INFLAMMATORY DRUGS, MYDRIASIS TREATMENT, TREATMENT IMPROVED, REPEATED ATTACKS, AFTER CONSULTATION DECIDED TO TAKE OUT THE LENS. ON (B)(6) 2020 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. D9- (B)(6) 2020. HEALTH IMPACT: 4625: MYDRIASIS TREATMENT. HEALTH IMPACT : 4644: HORMONES; NON-STEROIDAL ANTI- INFLAMMATORY DRUGS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.5/+2.5/090 | OS |  | 11/05/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9941512&pc=QCB | 2023826-2020-00870 | 2020-01-31 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances | NA | Event Description: THE RE4PORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.5/1.0/90 9 (SPHERE/CYLINDER/AXIS),IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT IS REPORTED TO BE EXPERIENCING "DOUBLE VISION THAT CAN BE SEEN UP AND DOWN IN THE DARK PLACES". THE SYMPTOMS DO NOT OCCUR IN BRIGHT PLACES. THE SURGEON REPORTED THAT THE "PATIENT HAS NORMAL IOP, NORMAL VAULT, NORMAL OCULAR BODY AND NO COMPLICATIONS." THE REPORTER STATED " THE DOCTOR DOES NOT INTEND TO POINT OUT ANY PRODUCT DEFECTS." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTE TO THE COMPLAINT ISSUE. (B)(4). | -7.5/1.0/90 9 | OS | Physician Assistant | 03/12/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9941534&pc=QCB | 2023826-2020-00871 | 2020-01-31 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2,-11.0/1.5/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT IS REPORTED TO BE EXPERIENCING "DOUBLE VISION THAT CAN BE SEEN UP AND DOWN IN THE DARK PLACES". THE SYMPTOMS DO NOT OCCUR IN BRIGHT PLACES. THE SURGEON REPORTED THAT THE "PATIENT HAS NORMAL IOP, NORMAL VAULT, NORMAL OCULAR BODY AND NO COMPLICATIONS." THE REPORTER STATED " THE DOCTOR DOES NOT INTEND TO POINT OUT ANY PRODUCT DEFECTS." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTE TO THE COMPLAINT ISSUE. (B)(4). | -11.0/1.5/91 | OD | Physician Assistant | 03/12/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10740194&pc=QCB | 2023826-2020-02482 | 2020-09-20 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Visual Disturbances | NA#S-P#NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -06.00/+2.0/058 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THE PATIENT WAS UNHAPPY WITH VISION (RESIDUAL MYOPIA). THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -06.00/+2.0/058 | OS | Physician | 10/05/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10740224&pc=QCB | 2023826-2020-02481 | 2020-09-20 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Visual Disturbances | NA#S-P#NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50/+2.5/110 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THE PATIENT WAS UNHAPPY WITH VISION (RESIDUAL MYOPIA). THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -06.50/+2.5/110 | OD | Physician | 10/05/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098621&pc=QCB | 2023826-2020-03250 | 2020-10-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Optical Problem | Visual Disturbances | N/A | Event Description: THE REPORTER STATES THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. REPORTEDLY, THE PATIENT IS EXPERIENCING LIGHTNING/FLASHES WHILE HE LIES IN BED AND TURNS HIS HEAD; STANDING UP THE VISION IS FINE. RETINAL DIAGNOSTIC SHOWS NO CHANGES AND THE SURGEON THINKS THIS COULD BE A "STRESS SYMPTOMATIC." Manufacturer Narrative: B5-UPDATED INFORMATION: THE HALOS HAVE BECOME LESS AND LESS. THE PATIENT IS HAPPY. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.5/+2.0/089 | OS |  | 12/04/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238074&pc=QCB | 2023826-2021-00173 | 2021-06-01 | Injury | STAAR SURGICAL COMPANY | 2021-01-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6, -11.50/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED AND THE LENS WAS REPLACED IN A SECOND SURGERY WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR AND DEVICE. REPORTEDLY, "WTW MEASUREMENT CALLED FOR 12.6MM LENS(STAAR OPTIMIZED). UNCLEAR IF THERE WAS A AN ABNORMAL SULCUS ANATOMY CAUSING EXCESS VAULT. PER REPORTER, "VAULT IS NOW EXCELLENT, PATIENT NOTICES SIGNIFICANTLY REDUCED GLARE." Manufacturer Narrative: CLAIM# (B)(4). | -11.50/2.0/095 | OS | Physician | 01/07/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238075&pc=QCB | 2023826-2021-00213 | 2021-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6, -11.50/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, "WE ARE PLANNING TO EXCHANGE LENS IN MARCH." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: REPORTEDLY, "WE EXCHANGED HER LEFT EYE A COUPLE OF MONTHS AGO. SIGNIFICANTLY IMPROVED HER OVERALL VISION, BUT ESPECIALLY HER NIGHT VISION. SHE HAD A WTW OF 11.5 WTW OU AND I PLACED A 12.6 LENS IN EACH EYE. 300% CT VAULT OU." CLAIM (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/2.0/095 | OD | Physician | 01/25/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11411793&pc=QCB | 2023826-2021-00568 | 2021-01-15 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -6.00/5.0/045 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LOW VAULT WITHOUT ROTATION WAS OBSERVED AND THE PATIENT EXPERIENCED DIFFICULTY WITH ACCOMMODATION FOR NEAR OBJECTS, THE REPORTER STATED " SX REPORTS DIFFICULTY WITH ACCOMMODATION POST ICL IMPLANTATION IN THE CONTEXT OF GOOD VISION AND REFRACTION DURING 1 DAY AND 1 WEEK POST OP REVIEW. BETTER WHEN OBJECTS ARE HELD FURTHER AWAY." LENS REMAINS IMPLANTED FOR NOW THE DOCTOR WILL MONITOR THE EYE. THE REPORTER STATED " PATIENTS CURRENT CONDITION IS FINE, ONLY ISSUE IS DIFFICULTY FOR READING NEAR SOMETIMES. VA/ EYE PRESSURE IS GOOD. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT-RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.00/5.0/045 | OD | Physician | 02/02/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558804&pc=QCB | 2023826-2021-00706 | 2020-11-17 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -4.5/+1.0/052 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, AFTER IMPLANT, THE PATIENT FELT UNCOMFORTABLE WITH THE GLARE WHEN DRIVING AT NIGHT AND SO THE ICL WAS REMOVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -4.5/+1.0/052 | OS | Physician | 02/25/2021 | 2021 | 55
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563271&pc=QCB | 2023826-2021-00759 |  | Injury | STAAR SURGICAL COMPANY | 2021-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -16.00/4.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. DIPLOPIA AND ASTIGMATISM WAS OBSERVED, WITH POSSIBLE ROTATION. THE LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2021, LENS ROTATION IS PLANNED. WILL PROVIDE FURTHER INFORMATION ONCE A RESOLUTION IS REACHED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -16.0/4.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OD) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, ORIGINAL LENS TOO LARGE, SMALLER REPLACED LENS ROTATED. MFR # 2023826-2021-00652. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.00/4.0/079 | OS |  | 03/10/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563284&pc=QCB | 2023826-2021-00757 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -16.00/4.0/129 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 201. DIPLOPIA AND ASTIGMATISM WAS OBSERVED, WITH POSSIBLE ROTATION. THE LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2021, LENS ROTATION IS PLANNED. WILL PROVIDE FURTHER INFORMATION ONCE A RESOLUTION IS REACHED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -16.00/4.0/129 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 201. DIPLOPIA AND ASTIGMATISM WAS OBSERVED, WITH POSSIBLE ROTATION. THE LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2021, "DIPLOPIA AND ASTIGMATISM HAVE RESOLVED." CORRECTED INFORMATION: B1 - REPORTED AS ADVERSE EVENT - CORRECT TO PRODUCTION PROBLEM. B2 - REPORTED AS REQUIRED INTERVENT TO PREVENT PERMANENT IMPAIRMENT/DAMAGE - NO OUTCOMES ATTRIBUTED TO ADVERSE EVENT. H1 - REPORTED AS SERIOUS INJURY - CORRECT TO MALFUNCTION. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO [OR#] SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -16.00/4.0/129 | OD |  | 03/10/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11820848&pc=QCB | 2023826-2021-01410 | 2021-05-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | Event Description: THE REPORTER/PATIENT INDICATED THAT A 13.2MM, TMICL13.2, -11.00/1.5/053 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, "GLARE, VISUAL QUALITY IS BAD AND NOT IMPROVING." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER/PATIENT INDICATED THAT A 13.2MM TMICL13.2 -11.00/1.5/053 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY "GLARE VISUAL QUALITY IS BAD AND NOT IMPROVING." GLARE/HALOES REPORTED. LENS REMAINS IMPLANTED. PROBLEM NOT RESOLVED. STATUS OF THE EYE IS QUIET - WELL HEALED. REPORTEDLY THE CAUSE OF THIS EVENT IS DUE TO THE DEVICE. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.00/1.5/053 | OD | Non-Healthcare Professional | 04/23/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11820914&pc=QCB | 2023826-2021-01411 | 2021-05-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | Event Description: THE REPORTER/PATIENT INDICATED THAT A 13.2MM, TMICL13.2, -10.50/1.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REPORTEDLY, "GLARE, VISUAL QUALITY IS BAD AND NOT IMPROVING." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER/PATIENT INDICATED THAT A 13.2MM TMICL13.2 -10.50/1.5/085 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REPORTEDLY, "GLARE, VISUAL QUALITY IS BAD AND NOT IMPROVING." LENS REMAINS IMPLANTED. PROBLEM NOT RESOLVED. STATUS OF THE EYE IS QUIET - WELL HEALED. REPORTEDLY THE CAUSE OF THIS EVENT IS DUE TO THE DEVICE. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.50/1.5/085 | OS | Non-Healthcare Professional | 04/23/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12281705&pc=QCB | 2023826-2021-03069 | 2021-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; Device Dislodged or Dislocated | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -13.50/4.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LENS WAS MISALIGNED BY 15 DEGREES. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR, MOST LIKELY CAUSE IS THAT INK AXIS MARKING WERE OFF TARGET FOR THE PROCEDURE. REPORTEDLY, LENS REPOSITIONING IS PLANNED. WILL ADVISE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -13.50/4.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LENS WAS MISALIGNED BY 15 DEGREES. LENS REPOSITIONED ON (B)(6) 2021 DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR, MOST LIKELY CAUSE IS THAT INK AXIS MARKING WERE OFF TARGET FOR THE PROCEDURE. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.50/4.0/090 | OS |  | 07/14/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12478533&pc=QCB | 2023826-2021-03579 | 2021-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00/+2.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND GLARE (THE PATIENTS PUPIL WAS 5MM). THE PUPIL SIZE WAS STILL 5MM FOR HALF A MONTH AFTER SURGERY, SO THE SURGEON DECIDED NOT TO IMPLANT ANOTHER ICL LENS FOR THE TIME BEING. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.00/+2.5/089 | OD | Physician | 08/27/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12479095&pc=QCB | 2023826-2021-03592 | 2021-06-27 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.00/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE REPORTER STATED THAT THE PATIENT DEVELOPED DIPLOPIA WITH VERY LOW VAULT AND LENS ROTATED. LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.00/2.0/094 | OD | Physician | 08/19/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12479109&pc=QCB | 2023826-2021-03593 | 2021-06-30 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.50/2.5/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE REPORTER STATED THAT THE PATIENT DEVELOPED DIPLOPIA WITH VERY LOW VAULT AND LENS ROTATED. LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.50/2.5/077 | OS | Physician | 08/19/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12558495&pc=QCB | 2023826-2021-03810 | 2021-05-21 | Injury | STAAR SURGICAL COMPANY | 2021-09-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE SURGEON REPORTS GLARE/HALOS AND PILOCARPINE ADMINISTRATION. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM# (B)(4). | -8.5/+1.0/082 | OS | Physician | 09/08/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725161&pc=QCB | 2023826-2021-04177 | 2020-08-05 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -15.5/3.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND POWER LENS DUE TO GLARE/HALOES. THIS EXCHANGE DID NOT RESOLVE THE PROBLEM. THE REPORTER STATED "GLARE/ABERATIONS IMPROVED BUT STILL PRESENT. PLAN ON DOING LASIK FOR RESIDUAL REFRACTIVE ERROR." IN THE REPORTERS OPINION " DEVICE AND PATIENT-RELATED FACTOR" WAS THE CAUSE OF THIS EVENT. FOR CAUSE DETAILS THE REPORTER STATED " DEVICE CAUSES MILD ABERATIONS- PT VERY SENSITIVE TO ABERATIONS". Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -15.5/3.5/105 | OD | Physician | 09/29/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725173&pc=QCB | 2023826-2021-04176 | 2020-08-05 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -15.0/4.0/089 (SPHERE / CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND POWER LENS DUE TO GLARE/ HALOES. THIS EXCHANGE DID NOT RESOLVE THE PROBLEM. THE REPORTER STATED "GLARE/ ABERRATIONS IMPROVED BUT STILL PRESENT. PLAN ON DOING LASIK FOR RESIDUAL REFRACTIVE ERROR." IN THE REPORTERS OPINION "DEVICE AND PATIENT-RELATED FACTOR" WAS THE CAUSE OF THIS EVENT. FOR CAUSE DETAILS THE REPORTER STATED "DEVICE CAUSES GLARE/ABERRATIONS- PT VERY SENSITIVE TO GLARE/ABERRATIONS". Manufacturer Narrative: (B)(4). | -15.0/4.0/089 | OS | Physician | 09/29/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13057182&pc=QCB | 2023826-2021-04829 | 2021-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/070 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED OF NIGHT GLARE, WHICH AFFECTED HIS NORMAL LIFE AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/070 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED OF NIGHT GLARE, WHICH AFFECTED HIS NORMAL LIFE AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(6). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(6). | -09.50/+1.5/070 | OS | Physician | 11/22/2021 | 2021 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13057258&pc=QCB | 2023826-2021-04828 | 2021-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/111 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED OF NIGHT GLARE, WHICH AFFECTED HIS NORMAL LIFE. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.00/+1.5/111 | OD | Physician | 11/22/2021 | 2021 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082455&pc=QCB | 2023826-2021-04974 | 2021-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0/+2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXPLANTED DUE TO "UNACCEPTABLE GLARE AFTER IMPLANTATION OF LENS." THE GLARE DISAPPEARED AFTER THE REMOVAL OF THE LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0/+2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXPLANTED DUE TO "UNACCEPTABLE GLARE AFTER IMPLANTATION OF LENS." THE GLARE DISAPPEARED AFTER THE REMOVAL OF THE LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -16.0/+2.5/091 | OD | Physician | 12/01/2021 | 2021 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13362150&pc=QCB | 2023826-2022-00347 | 2021-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -4/+6/55 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "PATIENT SAW GLARE".Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -4/+6/55 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "PATIENT SAW GLARE". Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -4/+6/55 | OD | Physician | 01/21/2022 | 2022 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14093007&pc=QCB | 2023826-2022-01195 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Visual Disturbances |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM TMICL_12.1; -9.00/1.0/180 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS THE PATIENT WANTED TO TRY FOR MONOVISION; IN WHICH SHE WAS UNHAPPY WITH THE DEGREE OF MONOVISION AND WOULD PREFER TO HAVE AN EXCHANGE FOR A POWER CLOSER TO HER DISTANCE REQUIREMENT. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). | -9.00/1.0/180 | NA | Other Health Care Professional | 03/17/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14706970&pc=QCB | 2023826-2022-01749 | 2022-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED A 13.2MM VTICMO13.2 -13.5/3.0/112 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REPORTEDLY GLARES/HALOES OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2022 WHICH RESOLVED THE PROBLEM. CAUSE IS DUE TO THE DEVICE - THE PATIENT FELT THAT THE POSTOPERATIVE HALO WAS BECOMING MORE SERIOUS AND COULD NOT BEAR IT AND ASKED TO REMOVE IT. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. (B)(4). | -13.5/3.0/112 | OD | Physician | 05/16/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14706973&pc=QCB | 2023826-2022-01748 | 2022-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED A 13.2MM VTICMO13.2 -14.5/5.0/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS OF WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REPORTEDLY GLARES/HALOES OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2022 WHICH RESOLVED THE PROBLEM. CAUSE IS DUE TO THE DEVICE - THE PATIENT FELT THAT THE POSTOPERATIVE HALO WAS BECOMING MORE SERIOUS AND COULD NOT BEAR IT AND ASKED TO REMOVE IT. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | -14.5/5.0/086 | OS | Physician | 05/16/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15003548&pc=QCB | 2023826-2022-02198 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/154 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6). POST-OP, THE UCVA WAS GOOD BUT THERE WAS GLARE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK .EXPLANT DATE: NA. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/+1.5/154 | OD | Other Health Care Professional | 06/16/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15567406&pc=QCB | 2023826-2022-03564 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 13.2MM, VTICM5_13.2, -9.50/4.0/91. (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) AS A REPLACEMENT LENS AFTER THE INITIAL LENS WAS FOUND DAMAGE BEFORE LOADING. IT WAS REPORTED THAT POST ICL IMPLANTATION THE PATIENT EXPERIENCED GLARE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION THE GLARE RESOLVED. LENS REMAINS IMPLANTED. CLAIM # (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.50/4.0/91 | OD | Physician | 09/14/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16000576&pc=QCB | 2023826-2022-04596 | 2022-11-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.5/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE REPORTER STATES THAT THE PATIENT HAS AN USTABLE DISTANCE VISION. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). Manufacturer Narrative: B5 - REPORTEDLY, PROBLEM IS RESOLVED. "THE PATIENT HAS GOTTEN USED TO THE SITUATION AND DOES NOT WANT ANY FURTHER INTERVENTIONS." CLAIM# (B)(4). | -9.5/2.5/093 | OD | Physician | 11/17/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16040405&pc=QCB | 2023826-2022-04702 | 2021-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Optical Problem | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -3.5/6.0/084 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2021, THE LENS WAS REMOVED. REPORTEDLY, "EXCESSIVE GLARE AND HALO, PATIENT CAN'T TOLERATE. CAUSE IS REPORTED AS THE DEVICE AND PATIENT RELATED FACTOR. REPORTEDLY, STATUS OF THE EYE IS "RETURNED TO PRE-OP STATUS." Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | -3.5/6.0/084 | OS | Physician | 11/24/2022 | 2022 | 44 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16296735&pc=QCB | 2023826-2023-00431 |  | Injury | STAAR SURGICAL COMPANY | 2023-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -15.0/2.00/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). LOW VAULT WITH ROTATION WAS OBSERVED AND THE PATIENT EXPERIENCED DECREASED VISION AND HALO GHOST IMAGE. THE REPORTER STATES THAT ICL ROTATION WAS PERFORMED IN JUNE 2020 AND THAT HE NOW PLANS TO EXCHANGE FOR LONGER LENGTH LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). IMPLANT DATE: UNKNOWN. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS. REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -15.0/2.00/100 | OS | Physician | 01/04/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16296738&pc=QCB | 2023826-2023-00429 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -10.00/5.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN (B)(6) 2019. LOW VAULT WITH ROTATION WAS OBSERVED AND THE PATIENT EXPERIENCED SEVER GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS DUE TO USER ERROR. THE REPORTER ALSO STATES THAT THE DEVICE FAILED TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATES THAT A LENS WITH WRONG LENGTH WAS ORDERED AND THAT THE CAUSE OF ROTATION IS LOW VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(6). IMPLANT DATE: (B)(6) 2019. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -10.00/5.0/095 | OD | Physician | 01/04/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16296741&pc=QCB | 2023826-2023-00433 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -7.5/1.0/64 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) IN (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED AND THE PATIENT EXPERIENCED SEVER GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. FOR CAUSE DETAILS THE REPORTER STATES "HIGH VAULT WITH UNKNOWN CAUSE". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: IMPLANT DATE: (B)(6) 2022. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.5/1.0/64 | OS | Physician | 01/04/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16296747&pc=QCB | 2023826-2023-00430 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -10.00/4.5/77(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) IN (B)(6) 2019. LOW VAULT WITH ROTATION WAS OBSERVED AND THE PATIENT EXPERIENCED SEVER GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS DUE TO USER ERROR. THE REPORTER ALSO STATES THAT THE DEVICE FAILED TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATES THAT A LENS WITH WRONG LENGTH WAS ORDERED AND THAT THE CAUSE OF ROTATION IS LOW VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). IMPLANT DATE: (B)(6) 2019. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -10.00/4.5/77 | OS | Physician | 01/04/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16296766&pc=QCB | 2023826-2023-00428 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.50/3.0/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) IN (B)(6) 2021. LOW VAULT WITH ROTATION WAS OBSERVED AND THE PATIENT EXPERIENCED SEVER GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS DUE TO USER ERROR. THE REPORTER ALSO STATES THAT THE DEVICE FAILED TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATES "CAUSE OF ROTATION IS LOW VAULT". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). IMPLANT DATE: (B)(6) 2021. TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1: PRODUCT PROBLEM SHOULD BE CORRECTED TO ADVERSE EVENT. B2: OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS SHOULD BE ADDED FOR CORRECTION. H3: SERIOUS INJURY SHOULD BE ADDED FOR CORRECTION. CLAIM# (B)(4). | -9.50/3.0/068 | OS | Physician | 01/04/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16296833&pc=QCB | 2023826-2023-00432 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -7.5/1.5/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN SEPTEMBER 2022. EXCESSIVE VAULT WAS OBSERVED AND THE PATIENT EXPERIENCED SEVER GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. FOR CAUSE DETAILS THE REPORTER STATES "HIGH VAULT WITH UNKNOWN CAUSE". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: PATIENT IDENTIFIER:(B)(6). IMPLANT DATE: SEPTEMBER 2022 TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7, -7.5/1.5/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) IN (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED AND THE PATIENT EXPERIENCED SEVERE GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. FOR CAUSE DETAILS THE REPORTER STATES "HIGH VAULT WITH UNKNOWN CAUSE". H6: MEDICAL DEVICE PROBLEM CODE 1494: OFF-LABEL USE: AGE>45 YEARS OLD AT THE TIME OF IMPLANTATION. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7, -7.5/1.5/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED AND THE PATIENT EXPERIENCED SEVERE GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. FOR CAUSE DETAILS THE REPORTER STATES "HIGH VAULT WITH UNKNOWN CAUSE". CLAIM# (B)(4). | -7.5/1.5/94 | OD | Physician | 01/04/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16742900&pc=QCB | 2023826-2023-01256 | 2023-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.50/+0.5/177 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023. DUE TO LOW VAULTING AND THE PATIENT EXPERIENCED DOUBLE VISION AND GLARE AT NIGHT. THE PATIENT WAS PRESCRIBED MEDICATION BUT THE SYMPTOMS DID NOT IMPROVE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS PATIENT-RELATED FACTOR AND THE DEVICE DID FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS RETURNED IN PIECES AND UNABLE TO PERFORM THE DIMENSIONAL INSPECTION DUE TO THE DAMAGED LENS. CLAIM # (B)(4). | -11.50/+0.5/177 | OD | Physician | 03/16/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16877245&pc=QCB | 2023826-2023-01587 | 2023-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50/+1.0/107 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6)2022. THE LENS WAS EXPLANTED ON (B)(6)2023 DUE TO THE PATIENT EXPERIENCED GLARE AND THE PROBLEM WAS RESOLVED. THE PATIENT COULD NOT ACCEPT THE GLARE AND REQUESTED THE LENS BE REMOVED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4). | -08.50/+1.0/107 | OD | Physician | 04/10/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17273274&pc=QCB | 2023826-2023-02719 | 2023-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -18.0/2.0/115 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED SEVERE DOUBLE VISION. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL/POWER LENS, AND LIMBAL RELAXING INCISION WAS PERFORMED TO CORRECT ITS ASTIGMATISM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -18.0/2.0/115 | OD | Physician | 07/03/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17302685&pc=QCB | 2023826-2023-02829 | 2023-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50/+3.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4). | -08.50/+3.0/087 | OD | Physician | 06/19/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17322117&pc=QCB | 2023826-2023-02830 | 2023-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -06.50/+2.5/075 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -06.50/+2.5/075 | OS | Physician | 06/19/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18279670&pc=QCB | 2023826-2023-05479 | 2023-11-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS OF -2.0/0.5/1 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND REFRACTIVE CHANGE OVERTIME ARE OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -2.0/0.5/1 | OD | Physician | 11/22/2023 | 2023 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18279671&pc=QCB | 2023826-2023-05480 | 2023-11-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS OF -2.0/0.5/69 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6). LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND REFRACTIVE CHANGE OVERTIME ARE OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# 732815. | -2.0/0.5/69 | OS | Physician | 11/22/2023 | 2023 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18456573&pc=QCB | 2023826-2024-00038 | 2023-08-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.5/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT AND GLARE/HALOS WERE OBSERVED. CAUSE OF THE EVENT IS OTHER. Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | -9.0/1.5/91 | OS | Physician | 12/28/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18456592&pc=QCB | 2023826-2024-00039 | 2023-08-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -10.0/1.5/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT AND GLARE/HALOS WERE OBSERVED. CAUSE OF THE EVENT IS OTHER. Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | -10.0/1.5/91 | OD | Physician | 12/28/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13684401&pc=QCB | 2023826-2022-00551 | 2021-12-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6, -9.50/1.0/84 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, GLARE, AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, GLARE REDUCED A LOT, VISION BETTER PUPIL BACK TO NORMAL. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.50/1.0/84 | OD | Physician | 02/15/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10852592&pc=QCB | 2023826-2020-02686 | 2020-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Discomfort | NA#S-P#NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50/+1.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT WAS UNCOMFORTABLE AND EXPERIENCED VISION FATIGUE. THE SURGEON WILL NOT IMPLANT ANOTHER ICL LENS. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETUNRED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -06.50/+1.0/087 | OD |  | 10/21/2020 | 2020 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10938663&pc=QCB | 2023826-2020-02843 |  | Injury | STAAR SUGICAL COMPANY | 2020-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -13.50/3.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE ON (B)(6) 2019. THE PATIENT CONTINUED TO STRUGGLE WITH GLARE AND HALO SINCE INITIAL SURGERY, POSSIBLE LENS EXCHANGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -13.50/3.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND GLARE/HALOES. THE REPORTER STATED " PUPIL SIZE GREATER OD > OS DUE TO EXCESSIVE VAULT CAUSING GLARE/HALO OD. " THIS EXCHANGE RESOLVED THE PROBLEM, "GLARE/HALO COMPLAINT HAS RESOLVED WITH 12.6MM ICL IMPROVING VAULT." CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.50/3.0/094 | OD |  | 11/02/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8322676&pc=QCB | 2023826-2019-00229 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM, VTICMO13.7, -12.0/4.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018 AND THE PATIENT IS COMPLAINING OF GLARE/HALOES. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. IT WAS REPORTED THAT REFRACTIVE TREATMENT WAS PERFORMED. ON PATIENTS LAST VISIT DATE: (B)(6) 2019, STATUS OF THE EYE WAS REPORTED AS "NORMAL", BCVA: 20/25-2, UCVA: 20/30 AND IOP: 12 MMHG. THE FOLLOWING WAS REPORTED AS ADDITIONAL INFORMATION: " PATIENT EXPERIENCING RESIDUAL HYPEROPIA AFTER SURGERY, ALSO COMPLAINING OF GLARE & HALOES". ADDITIONAL INFORMATION HAS BEEN REQUESTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM, VTICMO13.7, -12.0/4.0/075 (SPHERE/CYLINDE/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018 AND THE PATIENT IS COMPLAINING OF GLARE/HALOES. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. ON PATIENT'S LAST VISIT DATE: (B)(6) 2019, STATUS OF THE EYE WAS REPORTED AS "NORMAL", BCVA 20/25-2, UCVA:20/30 AND IOP: 12MMHG. THE FOLLOWING WAS REPORTED AS ADDITIONAL INFORMATION: "PATIENT EXPERIENCING RESIDUAL HYPEROPIA AFTER SURGERY, ALSO COMPLAINING OF GLARE AND HALOES". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.0/4.0/075 | OD |  | 01/09/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8544980&pc=QCB | 2023826-2019-00594 | 2018-03-07 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -1.00/2.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. IT WAS REPORTED THAT GLARE/HALOES WERE OBSERVED. CAUSE OF THE EVENT IS REPORTED AS CENTRAL FLOW HOLE. LENS REMAINS IMPLANTED. REPORTEDLY, LENS EXCHANGE IS PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -1.00/2.5/081 | OS | Physician | 03/26/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9323816&pc=QCB | 2023826-2019-02164 | 2018-09-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem; No Apparent Adverse Event | Visual Disturbances; Halo | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -0.50/2.0/174 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. GLARE/HALOES WERE OBSERVED. LENS REMAINS IMPLANTED. ON (B)(6) 2019 REPORTER INDICATED "THIS PATIENT HAS ALREADY EXCHANGED HIS LENS ON OS AND MIGHT NEED TO DO IT ON OD."ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - DEVICE CODE 3189 CORRECTED TO DEVICE CODE 3001. (B)(4). Manufacturer Narrative: H6 - RESULT CODE 213 CORRECTED TO RESULT CODE 114. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -0.50/2.0/174 | OD |  | 10/10/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9427016&pc=QCB | 2023826-2019-02296 | 2019-11-14 | Injury | STAAR SURGICAL COMPANY | 2019-12-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.00/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019 DUE TO GLARE/HALOES. PER REPORTER, "PATIENT EXPERIENCED CONCENTRIC CIRCLES OF LIGHT THRU LENS HOLD - INTOLERABLE." PATIENT WILL NOT BE PROCEEDING WITH ANOTHER LENS." Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/1.0/093 | OD | Administrator/Supervisor | 11/14/2019 | 2019 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9633749&pc=QCB | 2023826-2020-00161 | 2019-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.00/4.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. GLARES/HALOES WERE OBSERVED. ON (B)(6) 2020 THE LENS WAS REMOVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY AFTER EXPLANT (REMOVAL)." Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT IMPACT CODE 4629 SHOULD BE IN THE INITIAL MDR. H10- DEVICE EVALUATION: "RETURNED IN CARTRIDGE" SHOULD BE REMOVED FROM PREVIOUS MDR. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: THE LENS WAS RETURNED WITH CLEAR SURGICAL RESIDUE AND MOISTURE IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS RETURNED IN THE CARTRIDGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.00/4.5/092 | OS |  | 01/10/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10135948&pc=QCB | 2023826-2020-01171 | 2020-03-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+1.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT COMPLAINED OF GLARE/HALOS AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B1 - SHOULD BE CORRECTED TO MALFUNCTION. B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE SHOULD BE REMOVED AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.00/+1.0/093 | OD |  | 05/15/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10135999&pc=QCB | 2023826-2020-01172 | 2020-03-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+1.0/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT COMPLAINED OF GLARE/HALOS AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B1 - CORRECTED TO MALFUNCTION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.00/+1.0/085 | OS |  | 05/15/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10740529&pc=QCB | 2023826-2020-02488 | 2020-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo | NA#S-P#NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -15.50/+3.0/072 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS EXCESSIVE VAULTING, WITH GLARE AND HALOS. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. "PATIENT IS HAPPY" CORRECTED INFORMATION: B5 - (B)(6) 2020 IMPLANT DATE SHOULD BE CORRECTED TO (B)(6) 2020. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.50/+3.0/072 | OS |  | 10/09/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10740575&pc=QCB | 2023826-2020-02487 | 2020-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo | NA#S-P#NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.00/+4.5/107 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPORTED AS EXCESSIVE VAULTING, WITH GLARE AND HALOS. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5 - (B))(6) 2020 IMPLANT DATE SHOULD BE CORRECTED TO (B)(6) 2020. ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE LENS WAS EXPLANTED ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17.00/+4.5/107 | OD |  | 10/09/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984748&pc=QCB | 2023826-2020-02858 | 2020-11-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.00/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. GLARE/HALOES HAVE BEEN OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: B1 - REPORTED AS PRODUCT PROBLEM - UPDATE TO ADVERSE EVENT. B2 - UPDATE TO REQUIRD INTERVENTION TO PREVENT IMPAIREMENT/DAMAGE. B5 - NOTED LENS REMAINS IMPLANTED. LENS REMOVED ON (B)(6) 2020 AND THE PROBLEM RESOLVED. D6B = (B)(6) 2020. H6- HEALTH EFFECT IMPACT CODE - REPORTED AS 2199, UPDATED TO 4629. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON PRODUCT. VISUAL INSPECTION FOUND OPTIC AND HAPTIC TORN WITH FOREIGN MATERIAL ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.00/2.0/094 | OD |  | 11/16/2020 | 2020 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11206966&pc=QCB | 2023826-2021-00106 | 2020-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -10.50/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED ON (B)(6) 2020 DUE TO LOW VAULT AND GLARE/HALOES. THIS RESOLVED THE PROBLEM. REPORTEDLY, "THERE IS NO PLAN TO IMPLANT ANOTHER LENS." Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.50/2.0/089 | OD |  | 12/28/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238056&pc=QCB | 2023826-2021-00172 | 2020-12-11 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10/1.0/131 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). GLARE/HALOES WERE OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, "IRIS IS NOT FULLY JUMPING OVER OPTICAL ZONE. MESOPIC PUPIL MID-DILATED POST OP (5MM). EDGE OF THE PUPIL IS NOT IN THE WAY OF THE LENS AND THE PUPIL DOES CONSTRICT AND RELAX BUT DOES NOT FULLY CONSTRICT. PATIENT TO BE MONITORED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10/1.0/131 | OS | Physician | 01/07/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238066&pc=QCB | 2023826-2021-00171 | 2020-12-11 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/1.0/169 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). GLARE/HALOES WERE OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, "IRIS IS NOT FULLY JUMPING OVER OPTICAL ZONE. MESOPIC PUPIL MID-DILATED POST OP (5MM). EDGE OF THE PUPIL IS NOT IN THE WAY OF THE LENS AND THE PUPIL DOES CONSTRICT AND RELAX BUT DOES NOT FULLY CONSTRICT. PATIENT TO BE MONITORED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.50/1.0/169 | OD | Physician | 01/07/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11368705&pc=QCB | 2023826-2021-00458 | 2020-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+3.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED GLARE/HALOS AND THE PROBLEM WAS RESOLVED. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -09.50/+3.5/084 | OS |  | 01/24/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11368840&pc=QCB | 2023826-2021-00457 | 2020-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+4.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED GLARE/HALOS AND THE PROBLEM WAS RESOLVED. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND THE HAPTIC BENT AND DEFORMED. DIMENSIONAL FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.50/+4.5/089 | OD |  | 01/24/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11642581&pc=QCB | 2023826-2021-00943 | 2020-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -6.5/1.0/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE RIGHT EYE (OD) ON (B)(6) 2020. THE REPORTER STATED " ONE YEAR AGO I GOT THE EVO VISIAN ICL IMPLANTED IN MY EYES (I'M MYOPIC WITH ASTIGMATISM). THE SURGERY WENT WELL AND MY EYES HEALED WELL. I DON'T HAVE ANY HEALTH PROBLEMS, BUT DESPITE HAVING 100% EYESIGHT, I HAVE STRONG SIDE EFFECTS: I SEE LARGE CIRCLES OF LIGHT AT ANY TIME OF THE DAY WHEN I AM ILLUMINATED BY LAMPS OR THE SUN. THE LIGHT CIRCLES ARE EXTREMELY BRIGHT, LARGE AND DEPEND ON HOW THE LIGHT FALLS INTO THE EYE. SOME TUNNELS OR AVENUES OF STREET LAMPS TRIGGER SEVERAL CIRCLES OF LIGHT ONE AFTER THE OTHER, WHICH ARE GETTING BIGGER AND BIGGER. FROM THE SUNLIGHT I ADDITIONALLY SOMETIMES SEE A SMALL CIRCLE OF LIGHT IN THE MIDDLE, WHICH APPEARS IN LIGHT RAINBOW COLORS. IN DAWN AND IN POORLY LIT ROOMS I SEE HIGH-CONTRAST EDGES / CONTOURS TWICE, WHEREAT THEY BLUR INTO EACH OTHER. THE DOUBLE EDGES ARE NOTICEABLY LIGHTER AND BLURRY. EVEN WHEN I WATCH TV IN THE EVENING, THE IMAGES KEEP BLURRING. DUE TO THE RISKS OF A SURGERY, I DO NOT WANT TO HAVE THE LENSES REMOVED PREMATURELY, BUT I AM VERY UNHAPPY WITH MY CURRENT CONDITION AND I REGRET THE SURGERY VERY MUCH." ON (B)(6) 2021 THE REPORTER STATED " WE HAVE NOW STARTED AN ATTEMPT AT THERAPY WITH EYE DROPS AND I HOPE THAT MY SYMPTOMS CAN BE ALLEVIATED WITH THEM." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER UPDATED INFORMATION THE LENS WILL STAY IN THE. THE REPORTER STATED " THE LENS WILL STAY IN THE EYE, BECAUSE THE UCVA 20/16 IS PERFECT , PATIENT IS TRYING TO GET USED TO HALOS." CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -6.5/1.0/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY THE PATIENT EXPERIENCED GLARES; HALOES; BLURRY VISION AND LARGE CIRCLES OF LIGHT. ON (B)(6) 2021 THE REPORTER STATED "WE HAVE NOW STARTED AN ATTEMPT AT THERAPY WITH EYE DROPS AND I HOPE THAT MY SYMPTOMS CAN BE ALLEVIATED WITH THEM." AFTER SOME TIME REPORTS THAT "THE LENS WILL STAY IN THE EYE BECAUSE THE UCVA 20/16 IS PERFECT; PATIENT IS TRYING TO GET USED TO HALOS." LENS REMAINS IMPLANTED. CLAIM # (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -6.5/1.0/101 | OD | Physician | 03/10/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11642627&pc=QCB | 2023826-2021-00944 | 2020-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -6.5/1.0/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE LEFT EYE (OS) ON (B)(6) 2020. THE REPORTER STATED " ONE YEAR AGO I GOT THE EVO VISIAN ICL IMPLANTED IN MY EYES (I'M MYOPIC WITH ASTIGMATISM). THE SURGERY WENT WELL AND MY EYES HEALED WELL. I DON'T HAVE ANY HEALTH PROBLEMS, BUT DESPITE HAVING 100% EYESIGHT, I HAVE STRONG SIDE EFFECTS: I SEE LARGE CIRCLES OF LIGHT AT ANY TIME OF THE DAY WHEN I AM ILLUMINATED BY LAMPS OR THE SUN. THE LIGHT CIRCLES ARE EXTREMELY BRIGHT, LARGE AND DEPEND ON HOW THE LIGHT FALLS INTO THE EYE. SOME TUNNELS OR AVENUES OF STREET LAMPS TRIGGER SEVERAL CIRCLES OF LIGHT ONE AFTER THE OTHER, WHICH ARE GETTING BIGGER AND BIGGER. FROM THE SUNLIGHT I ADDITIONALLY SOMETIMES SEE A SMALL CIRCLE OF LIGHT IN THE MIDDLE, WHICH APPEARS IN LIGHT RAINBOW COLORS. IN DAWN AND IN POORLY LIT ROOMS I SEE HIGH-CONTRAST EDGES / CONTOURS TWICE, WHEREAT THEY BLUR INTO EACH OTHER. THE DOUBLE EDGES ARE NOTICEABLY LIGHTER AND BLURRY. EVEN WHEN I WATCH TV IN THE EVENING, THE IMAGES KEEP BLURRING. DUE TO THE RISKS OF A SURGERY, I DO NOT WANT TO HAVE THE LENSES REMOVED PREMATURELY, BUT I AM VERY UNHAPPY WITH MY CURRENT CONDITION AND I REGRET THE SURGERY VERY MUCH." ON 06 APR 2021 THE REPORTER STATED " WE HAVE NOW STARTED AN ATTEMPT AT THERAPY WITH EYE DROPS AND I HOPE THAT MY SYMPTOMS CAN BE ALLEVIATED WITH THEM." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER UPDATED INFORMATION THE LENS WILL STAY IN THE. THE REPORTER STATED " THE LENS WILL STAY IN THE EYE, BECAUSE THE UCVA 20/16 IS PERFECT , PATIENT IS TRYING TO GET USED TO HALOS." CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -6.5/1.0/069 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE LEFT EYE (OS) ON (B)(6) 2020. REPORTEDLY THE PATIENT EXPERIENCED GLARES; HALOES; BLURRY VISION AND LARGE CIRCLES OF LIGHT. ON 06/04/2021 THE REPORTER STATED "WE HAVE NOW STARTED AN ATTEMPT AT THERAPY WITH EYE DROPS AND I HOPE THAT MY SYMPTOMS CAN BE ALLEVIATED WITH THEM." AFTER SOME TIME REPORTS THAT "THE LENS WILL STAY IN THE EYE BECAUSE THE UCVA 20/16 IS PERFECT; PATIENT IS TRYING TO GET USED TO HALOS." LENS REMAINS IMPLANTED. CLAIM # (B)(4). Manufacturer Narrative: SEX-RACE:UNK. DATE OF EVENT:UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -6.5/1.0/069 | OS | Physician | 03/10/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651557&pc=QCB | 2023826-2021-00883 | 2020-11-17 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -4.50/1.5/143 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED ON (B)(6) 2020 DUE TO GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. PER THE REPORTER ON (B)(6) 2021, NO PLANS TO IMPLANT ANOTHER LENS, UCVA OF THIS PATIENT WAS GREAT AFTER THE SURGERY, BUT THE PATIENT COULDN'T STAND THE HALO AND GLARE, PATIENT INSISTED TO TAKE OUT THE LENS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4), ACD <3.0MM, OVER 45 YRS OF AGE AT DATE OF IMPLANT. (B)(4). | -4.50/1.5/143 | OD | Physician | 03/21/2021 | 2021 | 55
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721921&pc=QCB | 2023826-2021-01040 | 2019-10-01 | Injury | STAAR SURGICAL COMPANY | 2021-04-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/2.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.00/2.5/081 | OD | Physician | 04/01/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11778986&pc=QCB | 2023826-2021-01325 | 2021-02-23 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.50/1.0/139 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT HAD SERIOUS NIGHTTIME DAZZLES AFTER SURGERY AND REQUESTED LENS BE REMOVED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.50/1.0/139 | OD | Physician | 04/15/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11779037&pc=QCB | 2023826-2021-01326 | 2021-02-23 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.00/1.00/051 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT HAD SERIOUS NIGHTTIME DAZZLES AFTER SURGERY AND REQUESTED LENS BE REMOVED. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -16.00/1.00/051 | OS | Physician | 04/15/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11820836&pc=QCB | 2023826-2021-01405 | 2018-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.50/1.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS,INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. GLARE/HALOS WERE OBSERVED. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- "THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -8.50/1.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, LENS NATURALLY SITS VERY SLIGHTLY SUPERIOR/TEMPORAL. EYE DROPS (ALPHAGAN) DO NOT WORK. LENS REMAINS IMPLANTED. STATUS OF EYE; "ALL WITHIN NORMAL LIMITS." H6: HEALTH IMPACT CODE- 2199 IN INITIAL MDR IS CORRECTED TO 4644. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.50/1.0/092 | OS |  | 04/20/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222777&pc=QCB | 2023826-2021-02889 | 2021-02-10 | Injury | STAAR SURGICAL COMPANY | 2021-07-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -6.0/6.0/104, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THIS EXPLANT RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS PATIENT RELATED FACTOR AND IT WAS UNKNOWN IF THE DEVICE FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.0/6.0/104 | OD | Physician | 06/24/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222778&pc=QCB | 2023826-2021-02888 | 2020-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Visual Disturbances; Halo | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.0/1.5/115, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED LOW VAULT, LENS/DISLOCATION/SUBLUXATION AND GLARE/HALOES. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. THE REPORTER STATED " LENS INITIALLY HAD A SMALL VAULT AND IS OFF-CENTERED AS A RESULT. IT IS CURRENTLY OFF-CENTERED TEMPORALLY INFERIORLY DISLOCATION. AXIAL POSITION AND OVER-OVER REFRACTION THEREFORE ARE NOT RELIABLE. PATIENT IS DETECTING HALOS AND GLARE DUE TO DISLOCATION." LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H6: WORK ORDER SEARCH - NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). Manufacturer Narrative: B5: 12.6MM VTICM5 -13.0/1.5/104 (SPHERE/CYLINDER/AXIS). CLAIM# (B)(4). Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. B5 - "THE LENS NO LONGER REMAINS IMPLANTED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL/ LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION, LENS DISLOCATION/SUBLUXATION, AND GLARE HALOES. THE PROBLEM WAS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4629 SHOULD BE ADDED TO THE INITAL MDR. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.0/1.5/115 | OS | Physician | 06/24/2021 | 2021 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12587386&pc=QCB | 2023826-2021-03898 | 2021-08-10 | Injury | STAAR SURGICAL COMPANY | 2021-10-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.5/1.0/090, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE FOUND. CLAIM # (B)(4). | -10.5/1.0/090 | OS | Physician | 09/06/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12587379&pc=QCB | 2023826-2021-03897 | 2021-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.0/1.0/075, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5- CORRECTION: INITIAL LENS WAS IMPLANTED (B)(6) 2021. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K : THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -13.0/1.0/075 | OD | Physician | 09/06/2021 | 2021 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12593044&pc=QCB | 2023826-2021-03919 | 2021-09-01 | Injury | STAAR SURGICAL COMPANY | 2021-10-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.0/+1.0/012, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. THE REPORTER STATED " THE PATIENT HAS AN UCVA OF 20/20 WITH ICLS BUT MASSIVE ISSUES WITH HALOS AND GLARE AT NIGHT. SO THE SURGEON EXPLANTED THE ICL AND NOW ALL IS FINE." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE : UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE FOUND. CLAIM # (B)(4). | -11.0/+1.0/012 | OD | Physician | 09/07/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12593615&pc=QCB | 2023826-2021-03920 | 2021-09-01 | Injury | STAAR SURGICAL COMPANY | 2021-10-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -8.0/+0.5/033, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. THE REPORTER STATED " THE PATIENT HAS AN UCVA OF 20/20 WITH ICLS BUT MASSIVE ISSUES WITH HALOS AND GLARE AT NIGHT. SO THE SURGEON EXPLANTED THE ICL AND NOW ALL IS FINE." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE FOUND. CLAIM # (B)(4). | -8.0/+0.5/033 | OS | Physician | 09/07/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12791947&pc=QCB | 2023826-2021-04321 | 2021-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER STATED " 2 WEEKS POST-OP ICL PATIENT HAS COMPLAINTS OF HALO". PER UPDATED INFORMATION THE REPORTER STATED THAT THE PATIENT IS DOING QUITE WELL, WITH GLARE THAT CAN BE MANAGED WITH BRIMONIDINE. LENS REMIANS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. DATE IMPLANTED: UNK. DEVICE MANUFACTURING DATE: UNK, NO SERIAL NUMBER REPORTED. (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -6.0/1.0/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. PATIENT COMPLAINTS OF HALO/GLARE POST ICL IMPLANTATION. PER (B)(6) 2021 EMAIL FROM REPORTED THE REPORTER STATES "PATIENT IS ACTUALLY DOING QUITE WELL, WITH GLARE THAT HE CAN MANAGE WITH BRIMONIDINE". PATIENT-RELATED FACTOR AND DEVICE WAS REPORTED FOR CAUSE OF EVENT. HOWEVER, IT WAS REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. D4: (B)(6) 2023. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -6.0/1.0/102 | OD | Physician | 10/11/2021 | 2021 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12872823&pc=QCB | 2023826-2021-04493 | 2021-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/011 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED (THE LENS WAS IMPLANTED VERTICALLY). THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHINICITY , RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -13.00/+1.0/011 | OD | Physician | 10/28/2021 | 2021 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12872831&pc=QCB | 2023826-2021-04494 | 2021-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+3.5/079 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED GLARE/ HALOES AND THE PROBLEM WAS RESOLVED. ANOTHER ICL LENS WAS NOT IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.50/+3.5/079 | OS | Physician | 10/28/2021 | 2021 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14169183&pc=QCB | 2023826-2022-01222 | 2021-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.5/1.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE AND HALOES AND THE PATIENT FEELING UNCOMFORTABLE AND COULD NOT TOLERATE. AFTER REPEATED COMMUNICATION, THE PATIENT STILL INSISTED ON REMOVING THE LENS. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -11.5/1.0/091 | OS | Physician | 03/21/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14169248&pc=QCB | 2023826-2022-01221 | 2021-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.0/1.0/015 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND THE PATIENT FEELING UNCOMFORTABLE AND COULD NOT TOLERATE. AFTER REPEATED COMMUNICATION, THE PATIENT STILL INSISTED ON REMOVING THE LENS. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.0/1.0/015 | OD | Physician | 03/21/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14734140&pc=QCB | 2023826-2022-01850 | 2022-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.50/+1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO GLARE/HALOES. THE LENS WAS REPLACED WITH A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM #: (B)(4). | -08.50/+1.5/093 | OS | Physician | 05/19/2022 | 2022 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15003801&pc=QCB | 2023826-2022-02199 | 2022-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+1.5/004 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO PATIENT COMPLAINED OF GLARE/HALOS. THE LENS WAS REPLACED WITH AN IOL LENS AND THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT: UNK. PATIENT ETHNICITY: UNK. PATIENT RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -18.00/+1.5/004 | OD | Physician | 06/14/2022 | 2022 | 59 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15591507&pc=QCB | 2023826-2022-03594 | 2022-08-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM13.2, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT, GALRE AND HALOES, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- "THE REPORTER INDICATED THAT A 13.2MM VTICM13.2" SHOULD BE CORRECTED TO "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2." CLAIM# (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 13.2MM VTICM13.2 -8.50/1.0/130 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT, GLARE AND HALOES. LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM#: (B)(4). | -8.50/1.0/130 | OD | Physician | 09/13/2022 | 2022 | 46 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16030530&pc=QCB | 2023826-2022-04434 | 2022-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS -9.5/1.0/080 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND GLARE/HALOES. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTD TO BE PATIENT-RELATED FACTOR. HOWEVER, THE REPORTER STATED THAT THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -9.5/1.0/080 | OD | Physician | 11/22/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16130089&pc=QCB | 2023826-2023-00094 | 2022-12-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -01.00/+1.5/087, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT COMPLAINED OF GLARE/HALOES AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -01.00/+1.5/087 | OD | Physician | 12/14/2022 | 2022 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301533&pc=QCB | 2023826-2023-00382 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.00/5.5/047 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. EXCESSIVE VAULT WAS OBSERVED WITH SEVERE GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. PATIENT HAS A PRE-EXISTING CONDITION OF KERATOCONUS. CAUSE REPORTED AS UNKNOWN. "JUST IN CASE OF KERATOCONUS HIGH VAULT IS COMMON EVENT AND THIS CASES AS HISTORY OF KERRARING (INTRA CORNEAL RING)." Manufacturer Narrative: IMPLANT DATE: (B)(6) 2021. DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -11.00/5.5/047 | OD | Physician | 01/04/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301607&pc=QCB | 2023826-2023-00383 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -75/5.0/065 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. EXCESSIVE VAULT WAS OBSERVED WITH SEVERE GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. PATIENT HAS A PRE-EXISTING CONDITION OF KERATOCONUS. CAUSE REPORTED AS UNKNOWN. "JUST IN CASE OF KERATOCONUS HIGH VAULT IS COMMON EVENT AND THIS CASES AS HISTORY OF KERRARING (INTRA CORNEAL RING)." Manufacturer Narrative: IMPLANT DATE: (B)(6) 2021. DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -75/5.0/065 | OS | Physician | 01/04/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301634&pc=QCB | 2023826-2023-00384 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -14.0/2.5/065 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WITH ROTATION WAS OBSERVED WITH SEVERE GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. PATIENT HAS A PRE-EXISTING CONDITION OF KERATOCONUS. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - CORRECTED TO : "THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -14.0/2.5/068 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WITH ROTATION WAS OBSERVED WITH SEVERE GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. PATIENT HAS A PRE-EXISTING CONDITION OF KERATOCONUS. CAUSE REPORTED AS UNKNOWN." IN THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: IMPLANT DATE: (B)(6) 2020. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH. NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | -14.0/2.5/065 | OD | Physician | 01/04/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301672&pc=QCB | 2023826-2023-00385 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -14.0/2.5/065 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WITH ROTATION WAS OBSERVED WITH SEVERE GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. PATIENT HAS A PRE-EXISTING CONDITION OF KERATOCONUS. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: IMPLANT DATE: (B)(6) 2021. DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). (B)(4). Manufacturer Narrative: B1 - CORRECTED TO: ADVERSE EVENT IN THE INITIAL MDR. B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/ DAMAGE SHOULD BE ADDED TO THE INITIAL MDR. H1 - CORRECTED TO : SERIOUS INJURY IN THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: B1 - CORRECTED TO: ADVERSE EVENT IN THE INITIAL MDR. B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/ DAMAGE SHOULD BE REMOVED FROM MFR REPORT #: 2023826-2023-00385-SUPP1. H1 - CORRECTED TO : MALUNCTION IN THE INITIAL MDR. CLAIM # (B)(4). | -14.0/2.5/065 | OD | Physician | 01/04/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16506881&pc=QCB | 2023826-2023-00625 | 2022-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.00/+2.5/168 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO THE PATIENT EXPERIENCED GLARE/HALOES AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.00/+2.5/168 | OS | Physician | 02/08/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16595253&pc=QCB | 2023826-2023-00996 | 2023-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS OF -14.0/1.0/051 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 , THE LENS WAS REMOVED AND REPLACED FOR A DIFFERENT MODEL, LONGER LENGTH LENS DUE TO LOW VAULT, LENS OPACITY, AND GLARE HALOES. THE CAUSE OF THE EVENT WAS REPORTED AS OTHER. REPORTEDLY, "PATIENT STILL WITH LENS OPACITY SUPRISINGLY, BUT BETTER VAULT" Manufacturer Narrative: CORRECTED DATA: H6- 1494: OFF-LABEL USAGE- PRE-EXISTING CONDITION OF PROGRESSIVE MYOPIA. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). | -14.0/1.0/051 | OS | Physician | 02/21/2023 | 2023 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17434410&pc=QCB | 2023826-2023-03186 | 2023-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6, -14.5/4.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON 2(B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT STRONGLY REQUESTED REMOVAL AND NO REPLACEMENT. Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK, H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.5/4.0/087 | OS | Physician | 07/11/2023 | 2023 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17440152&pc=QCB | 2023826-2023-03185 | 2023-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6, -13.0/3.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT STRONGLY REQUESTED REMOVAL AND NO REPLACEMENT. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.0/3.5/094 | OD | Physician | 07/11/2023 | 2023 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17495490&pc=QCB | 2023826-2023-03323 | 2023-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -4.0/0.5/174 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SPHERICAL LENS DUE TO GLARE/HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: "2140" SHOULD BE REMOVED AND "2227" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "1401" SHOULD BE REMOVED AND "3001" SHOULD BE ADDED. (B)(4) Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -4.0/0.5/174 | OD | Physician | 07/22/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733767&pc=QCB | 2023826-2023-04036 | 2023-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.00/ 3.0/114 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.00/ 3.0/114 | OD | Physician | 08/18/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733799&pc=QCB | 2023826-2023-04037 | 2023-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -7.50/2.0/145 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -7.50/2.0/145 | OS | Physician | 08/18/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17882968&pc=QCB | 2023826-2023-04447 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -16.00/1.0/120 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS AN EXCHANGE LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. HALOS/GLARE WITH NIGHTTIME DRIVING ARE REPORTED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -16.00/1.0/120 | OD | Other Health Care Professional | 09/07/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17961829&pc=QCB | 2023826-2023-04557 | 2023-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE/HALOS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.00/+1.0/094 | OD | Physician | 09/27/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18067593&pc=QCB | 2023826-2023-04969 | 2023-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.5/2.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.5/2.0/090 | OD | Physician | 11/01/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18075196&pc=QCB | 2023826-2023-04725 | 2023-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6. -11.5/1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO NIGHTTIME GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS OTHER: NIGHTTIME GLARE. REPORTEDLY, PATIENT REQUEST REMOVAL OF LENS. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.5/1.5/091 | OD | Physician | 10/16/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18075197&pc=QCB | 2023826-2023-04726 | 2023-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6. -12.0/2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO NIGHTTIME GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS OTHER: NIGHTTIME GLARE. REPORTEDLY, PATIENT REQUEST REMOVAL OF LENS. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -12.0/2.5/091 | OS | Physician | 10/16/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18080758&pc=QCB | 2023826-2023-04992 | 2023-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13/2.5/071 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13/2.5/071 | OS | Physician | 11/01/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18117390&pc=QCB | 2023826-2023-04761 | 2021-09-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -05.50/+4.0/077 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE/HALOS AND WAS NOT SATISFIED WITH RESULTS. THE PROBLEM WAS RESOLVED. POST-OP THE EYE WAS NORMAL AND WAS WITHOUT SIGNS AND SYMPTOMS. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -05.50/+4.0/077 | OD | Physician | 10/19/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136579&pc=QCB | 2023826-2023-04762 | 2021-09-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -05.50/+5.5/103 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE/HALOS AND WAS NOT SATISFIED WITH RESULTS. THE PROBLEM WAS RESOLVED. POST-OP THE EYE WAS NORMAL AND WAS WITHOUT SIGNS AND SYMPTOMS. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -05.50/+5.5/103 | OS | Physician | 10/19/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18440045&pc=QCB | 2023826-2024-00048 | 2023-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -9.00/1.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT WITH GLARE/HALOES WAS REPORTED. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#:(B)(4). | -9.00/1.0/095 | OD | Other Health Care Professional | 12/05/2023 | 2023 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18440457&pc=QCB | 2023826-2024-00049 | 2023-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -8.00/1.0//055 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT WITH GLARE/HALOES WAS REPORTED. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -8.00/1.0//055 | OS | Other Health Care Professional | 12/05/2023 | 2023 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11129006&pc=QCB | 2023826-2021-00086 | 2020-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.50/1.0/010(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.50/1.0/010 | OS |  | 12/11/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11129010&pc=QCB | 2023826-2021-00085 | 2020-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.00/1.0/170(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTFIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.00/1.0/170 | OD |  | 12/11/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11368769&pc=QCB | 2023826-2021-00471 | 2020-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.00/+2.0/089 | OD |  | 01/24/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11386930&pc=QCB | 2023826-2021-00544 | 2020-12-22 | Injury | STAAR SURGICAL COMPANY | 2021-02-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.00/1.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF ICA, AND GLARE/HALOES. THIS EXCHANGE RESOLVED THE PROBLEM. "NORMAL" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. FOR ADDITIONAL INFORMATION "OS- VAULT 1000 OBSERVED" WAS REPORTED. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. DATE OF EVENT: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -10.00/1.0/082 | OD | Physician | 01/29/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11697117&pc=QCB | 2023826-2021-01089 | 2021-03-19 | Injury | STAAR SURGICAL COMPANY | 2021-04-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.0/+1.5/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, UNREACTIVE (FIXED) PUPIL, AND GLARE/HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN: "UNREACTIVE PUPIL AND GLARE HALO DUE TO HIGH VAULT WERE CONFIRMED. THE LENS SIZE DID NOT FIT THE PATIENT." Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH EFFECT CLINICAL CODE: (B)(4). CLAIM# (B)(4). | -10.0/+1.5/078 | OS | Physician | 03/24/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11697124&pc=QCB | 2023826-2021-01088 | 2021-03-19 | Injury | STAAR SURGICAL COMPANY | 2021-04-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.0/+1.0/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, UNREACTIVE (FIXED) PUPIL, AND GLARE/HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN: "UNREACTIVE PUPIL AND GLARE HALO DUE TO HIGH VAULT WERE CONFIRMED. THE LENS SIZE DID NOT FIT THE PATIENT." Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+1.0/080 | OD | Physician | 03/24/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119230&pc=QCB | 2023826-2021-02331 | 2021-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS THEN EXPLANTED AND THEN REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, GLARE/HALOS, BLURRED VISION, AND SPARKLES. THE EYE "LOOKS GOOD" NOW. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: SPHERE/CYLINDER/AXIS OF THE IMPLANTED LENS IS -0.5/+1.5/049. THE LENS WAS IMPLANTED ON (B)(6) 2021. THE PATIENT CAN SEE CENTRAL RINGS. ON (B)(6) 2021 THE LENS WAS EXCHANGED. THE PROBLEM WAS NOT RESOLVED. SEE MFR REP# 2023826-2021-02969 FOR THE REPLACEMENT LENS. H6-HEALTH EFFECT CLINICAL CODE 4581 (SEES CENTRAL RINGS). H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. DATE OF EVENT - UNK. SERIAL NUMBER, EXPIRATION DATE: UNK. IMPLANT, EXPLANT DATE - UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MFG DATE - UNK. CLAIM# (B)(4). | -0.5/+1.5/049 | OS |  | 06/10/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12669520&pc=QCB | 2023826-2021-04044 | 2021-09-24 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD). THREE MONTHS AFTER SURGERY, THE PATIENT REPORTED DYSPHOTOPSIAS AT NIGHT, WITH MESOPIC PUPIL (6MM). THE PATIENT REPORTED GLARE/HALOS. THE PATIENT WAS TREATED WITH ALPHAGAN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. IMPLANT DATE: UNK. EXPLANT DATE: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.00/+1.5/092 | OD | Physician | 09/27/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15296711&pc=QCB | 2023826-2022-02793 | 2022-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -05.00/+0.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE EYE WAS QUIET. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: EXPIRATION DATE: 31-OCT-2024 (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4), SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). | -05.00/+0.5/090 | OD | Physician | 08/03/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15307873&pc=QCB | 2023826-2022-02794 | 2022-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -05.00/+0.5/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE EYE WAS QUIET. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4), SHOULD BE ADDED TO INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. EXPIRATION DATE: 30-APR-2025. HEALTH IMPACT- CLINICAL CODE: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -05.00/+0.5/098 | OS | Physician | 08/03/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15476840&pc=QCB | 2023826-2022-03279 | 2022-08-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -18/1.00/106 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON . THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; ANGLE CLOSURE AND GLARE/HALOES. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "ERROR ON LENS DIAMETER CALCULATION. ANGLE CLOSURE ON ONE CUADRANT (NASAL)". Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -18/1.00/106 | OS | Physician | 08/26/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16939032&pc=QCB | 2023826-2023-01654 | 2023-04-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.0/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, GLARE/HALOES, AND EXCESSIVE VAULT HAS BEEN OBSERVED. LENS REMAINS IMPLANTED. PATIENT REMAINS UNDER OBSERVATION. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. REPORTEDLY, PATIENTS ANATOMY MAY BE ABNORMAL OR A MEASUREMENT ERROR. Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -7.00/1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND GLARES/HALOES. THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "NARROW ANGLE, HIGH VAULT, NO OTHER SYMPTOMS, MONITORING FOR NOW". ADDITIONAL INFORMATION PROVIDED: "WOULD LIKE TO REQUEST A CONSULTATION WITH AN MD TO DISCUSS THE SCENARIO AND TO RECEIVE MEDICAL PERSPECTIVE ON HIS PLANT TO OBSERVE." THE CAUSE OF THE EVENT IS A PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. CAUSE DETAILS PROVIDED: "HIGH VAULT OD DESPITE ACCURATE PREOP MEASUREMENTS, PATIENTS ANATOMY MAY BE ABNORMAL; OR A MEASUREMENT ERROR. " CLAIM# (B)(4). | -7.0/1.0/091 | OD | Physician | 04/28/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16939036&pc=QCB | 2023826-2023-01655 | 2023-04-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -10.5/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, GLARE/HALOES, AND EXCESSIVE VAULT HAS BEEN OBSERVED. LENS REMAINS IMPLANTED. PATIENT REMAINS UNDER OBSERVATION. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. REPORTEDLY, PATIENTS ANATOMY MAY BE ABNORMAL OR A MEASUREMENT ERROR. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -12.0/+1.5/114 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND GLARES/HALOES. THE LENS REMAINS IMPLANTED, AND THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "NARROW ANGLE, HIGH VAULT, NO OTHER SYMPTOMS, MONITORING FOR NOW". ADDITIONAL INFORMATION PROVIDED: "WOULD LIKE TO REQUEST A CONSULTATION WITH AN MD TO DISCUSS THE SCENARIO AND TO RECEIVE MEDICAL PERSPECTIVE ON HIS PLANT TO OBSERVE." THE CAUSE OF THE EVENT IS A PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. CAUSE DETAILS PROVIDED: "HIGH VAULT OS DESPITE ACCURATE PREOP MEASUREMENTS, PATIENTS ANATOMY MAY BE ABNORMAL; OR A MEASUREMENT ERROR." CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -10.50/1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) HE PATIENT EXPERIENCED AN EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND GLARES/HALOES. THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "NARROW ANGLE, HIGH VAULT, NO OTHER SYMPTOMS, MONITORING FOR NOW". ADDITIONAL INFORMATION PROVIDED: "WOULD LIKE TO REQUEST A CONSULTATION WITH AN MD TO DISCUSS THE SCENARIO AND TO RECEIVE MEDICAL PERSPECTIVE ON HIS PLANT TO OBSERVE." THE CAUSE OF THE EVENT IS A PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. CAUSE DETAILS PROVIDED: "HIGH VAULT OS DESPITE ACCURATE PREOP MEASUREMENTS, PATIENTS ANATOMY MAY BE ABNORMAL; OR A MEASUREMENT ERROR. CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/1.0/090 | OS | Physician | 04/27/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17138291&pc=QCB | 2023826-2023-02315 | 2023-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+1.5/046 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING, UNREACTIVE (FIXED) PUPIL AND GLARE/HALOES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THERE WERE NO COMPLAINTS, EYE WAS CALM AND THE PUPIL WAS REACTIVE. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - UNREACTIVE (FIXED) PUPIL. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: A PI WAS PERFORMED ON (B)(6) 2023 AND WAS THE FIRST PI (NO NEED OF BEFORE PI - MYOPIC LENS). H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - ADDITION OF PI. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "4581- WIDE PUPIL " SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER-18.0/+1.5/46 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, UNREACTIVE (FIXED) PUPIL, 'WIDE' PUPIL, AND GLARES/HALOES. ON (B)(6) 2023 PI WAS PERFORMED. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL AND POWER, SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE: "NO COMPLAINTS, CALM EYE, REACTIVE PUPIL (STATUS VERY GOOD)". THE CAUSE OF THE EVENT WAS THE DEVICE. CAUSE DETAILS PROVIDED: "TO BIG SIZE OF THE LENS UNUSUAL STRUCTURE OF THE EYE". CLAIM# (B)(4). | -18.00/+1.5/046 | OS | Physician | 05/24/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11887128&pc=QCB | 2023826-2021-01641 | 2021-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50/+2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING ROTATED AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED TOGETHER WITH THE PATIENT, THE DOCTOR DECIDED NOT TO PERFORM A SURGICAL INTERVENTION. ALL IS GOOD AND THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: IN INITIAL MDR THE WORK ORDER SEARCH RESULTS SHOULD BE CORRECTED TO: NO SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. H6 MEDICAL DEVICE CODE- 2923 CODE IN INITIAL MDR IS NOT APPLICABLE. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: AT A POST-OP EXAMINATION ON (B)(6) 2021, THE SURGEON DETECTED AN EXCESSIVE VAULTING. THERE WAS PIGMENT DISPERSION AND PATIENT COMPLAINED OF GLARE/HALOS. H6: HEALTH IMPACT CLINICAL CODE: 4581 PIGMENT DISPERSION. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THAT ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, PIGMENT DISPERSION, AND GLARE/HALOS. THE PROBLEM WAS RESOLVED AND THE "PATIENT IS HAPPY". CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. | -10.50/+2.0/091 | OS | Physician | 04/30/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12756627&pc=QCB | 2023826-2021-04227 | 2021-09-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; Discomfort | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7,-12.5/2.5/101, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. PATIENT EXPERIENCED EXCESSIVE VAULT AND GLARE/HALOES, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN, CAUSE DETAILS " LENS IS TO BIG FOR PATIENT" WAS REPORTED. PER UPDATED INFORMATION THE REPORTER STATED " PATIENT IS JUST DISCOMFORT WITH THE VISUAL ACUITY IOP IS FINE." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: B5- "PATIENT ALSO EXPERIENCED DISCOMFORT". CLAIM# (B)(4). | -12.5/2.5/101 | OS | Physician | 10/05/2021 | 2021 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17057307&pc=QCB | 2023826-2023-02178 | 2023-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -10.0/4.5/176 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -10.0/4.5/176 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT, GLARES, AND HALOES. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -10.0/4.5/176 | OS | Physician | 05/08/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17091344&pc=QCB | 2023826-2023-02233 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/044 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR DIFFERENT MODEL, SHORTER LENGTH LENS, DUE TO EXCESSIVE VAULT AND GLARE/HALOES. Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- IMPACT CODE: "2227" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -12.0/1.5/044 | OD | Physician | 05/23/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8519001&pc=QCB | 2023826-2019-00612 | 2019-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.5/+3.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WASEXPLANTED DUE TO GLARE/HALOS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+3.0/090 | OS |  | 03/22/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8519013&pc=QCB | 2023826-2019-00608 | 2019-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+3.5/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+3.5/095 | OD |  | 03/22/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552494&pc=QCB | 2023826-2019-00659 | 2019-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THE PROBLEM WAS RESOLVED AFTER EXPLANT. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+1.0/090 | OS |  | 03/29/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8746386&pc=QCB | 2023826-2019-01163 | 2019-05-04 | Injury | STAAR SURGICAL COMPANY | 2019-06-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.0/+2.5/096 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO THE PATIENT COULD NOT TOLERATE GLARE AND HALOS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.0/+2.5/096 | OD | Physician | 06/04/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8895522&pc=QCB | 2023826-2019-01453 | 2019-03-02 | Injury | STAAR SURGICAL COMPANY | 2019-08-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.0/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. IT WAS REPORTED THAT AFTER EXPLANTING THE LENS THE PROBLEM WAS RESOLVED, PATIENT CONDITION WAS STABLE. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.0/1.0/090 | OS | Physician | 07/16/2019 | 2019 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966508&pc=QCB | 2023826-2019-01647 | 2019-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS EXCESSIVE VAULT, GLARE/HALOS, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA) AND IRIS PROLAPSE. THE CAUSE OF THE EVENT IS REPORTED AS "NOMOGRAM ERROR." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: DATE OF EVENT CORRECTED TO (B)(6) 2019 IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: EVENT DESCRIPTION CORRECTED TO STATE THAT A 13.7MM VTICM5 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.00/6.0/031 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS FOUND WITHIN ASSOCIATED LOTS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -3.00/6.0/031 | OS |  | 08/09/2019 | 2019 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966519&pc=QCB | 2023826-2019-01646 | 2019-05-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS EXCESSIVE VAULT, GLARE/HALOS, AND NEGATIVE DYSPHOTOPSIA. THE CAUSE OF THE EVENT IS REPORTED AS "NOMOGRAM ERROR." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPES WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# : (B)(4). Manufacturer Narrative: CORRECTED TO STATE THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.50/1.0/101 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). CLAIM# : (B)(4). Manufacturer Narrative: DATE OF EVENT: CORRECTED TO (B)(6) 2019 IN INTIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: INITIAL MDR SHOULD HAVE INCLUDED PATIENT CODE 3191: NEGATIVE DYSPHOTOPSIA. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. COMMON DEVICE NAME: PRODUCT CODE: UNK. MODEL #, IMPLANTED, EXPLANTED: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. DEVICE MANUFACTURE DATE: UNK. (B)(4). | -7.50/1.0/101 | OS |  | 08/09/2019 | 2019 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8969999&pc=QCB | 2023826-2019-01662 | 2018-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -0.50/2.0/179 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. HALOS AND "LIGHTSPRED" WERE OBSERVED. LENS REMAINS IMPLANTED. REPORTER INDICATED THAT "THE SURGEON DELAYED POSSIBLE LENS EXCHANGE DUE TO VERY GOOD REFRACTIVE RESULT. PATIENT WOULD LIKE TO EXCHANGE LENS, STILL HAVING DAILY PROBLEMS WITH GLARE." PER THE REPORTER ON (B)(6) 2019 EXCHANGE WILL NOT HAPPEN UNTIL END OF (B)(6). Manufacturer Narrative: PATIENT GENDER: UNK SHOULD BE REMOVED FROM INITIAL MEDWATCH REPORT. ADVERSE EVENT. REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL BUT SAME POWER/SIZE LENS AND THE PROBLEM WAS RESOLVED. PATIENT STATUS WAS REPORTED AS "GOOD". EXPLANT DATE: (B)(6) 2019. TYPE OF REPORTABLE EVENT: SERIOUS. PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -0.50/2.0/179 | OS |  | 08/13/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9115540&pc=QCB | 2023826-2019-01765 | 2019-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -9.00/1.0/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. GLARE/HALOES WERE OBSERVED AND THE LENS WAS REMOVED ON (B)(6) 2019. REPORTEDLY "LENS WAS REMOVED DUE TO STRONG HALOES. PATIENT WEARS GLASSES AGAIN." Manufacturer Narrative: CORRECTED DATA: PATIENT CODE 2227 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. ADDITIONAL INFORMATION: DATE OF EVENT UPDATED TO (B)(6) 2019 IN INITIAL MDR. PER UPDATED INFORMATION, THE CAUSE OF THE EVENT IS DUE TO A PATIENT RELATED FACTOR. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL STUCK TO A PIECE OF PAPER WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS, BROWN RESIDUE ON THE LENS AND THE LENS RETURNED STUCK TO A PIECE OF PAPER WITH INK. THE LENS HAS BLUISH/ PURPLE TINT POSSIBLY FROM THE INK ON THE PAPER. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.00/1.0/103 | OS |  | 09/09/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9115541&pc=QCB | 2023826-2019-01764 | 2019-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.00/1.0/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. GLARE/HALOES WAS OBSERVED AND THE LENS WAS REMOVED ON (B)(6) 2019. REPORTEDLY "LENS WAS REMOVED DUE TO STRONG HALOES. PATIENT WEARS GLASSES AGAIN." Manufacturer Narrative: PATIENT CODE 2227 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. DATE OF EVENT: UPDATED TO (B)(6) 2019 IN INITIAL MDR. EVENT DESCRIPTION: PER UPDATED INFORMATION, THE CAUSE OF THE EVENT IS DUE TO A PATIENT RELATED FACTOR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND RED RESIDUE ON THE LENS AND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.00/1.0/102 | OD |  | 09/09/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9203032&pc=QCB | 2023826-2019-01953 | 2019-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -08.00/1.0/133 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. GLARE/HALOES WERE OBSERVED. THE LENS REMAINS IMPLANTED. REPORTEDLY SURGEON NOTED "HALOS/GLARE PRESENT BOTH WITH SMALL AND DILATED PUPIL. RELATED IN MY OPINION TO CENTRAL HOLE." CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR AND DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 3191: NO CODE AVAILABLE, SMALL AND DILATED PUPIL. RESULT CODE 114 CORRECTED TO RESULT CODE 213. PER UPDATED INFORMATION, THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL, LONGER LENGTH, DIFFERENT POWER LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. CLAIM#(B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED INA MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: RESULT CODE 114 SHOULD BE ADDED TO THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -08.00/1.0/133 | OD |  | 09/27/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9252703&pc=QCB | 2023826-2019-01960 |  | Injury | STAAR SURGICAL COMPANY | 2019-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/+0.5/170 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING AND THE PUPIL SIZE WAS LARGE/REACTIVE. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL BE MONITORED FOR 2 MONTHS. Manufacturer Narrative: CORRECTED DATA: EVENT - "THE PATIENT EXPERIENCES GLARE/HALOES DURING DAY AND NIGHT TIME." SHOULD BE ADDED IN THE INITIAL MDR." CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANTED: NA THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.0/+0.5/170 | OD |  | 09/30/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9252792&pc=QCB | 2023826-2019-01961 |  | Injury | STAAR SURGICAL COMPANY | 2019-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+1.0/015 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING AND THE PUPIL SIZE WAS LARGE/REACTIVE. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL BE MONITORED FOR 2 MONTHS. Manufacturer Narrative: CORRECTED DATA: "THE PATIENT EXPERIENCES GLARE/HALOES DURING DAY AND NIGHT TIME." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.5/+1.0/015 | OS |  | 09/30/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9492810&pc=QCB | 2023826-2019-02418 | 2019-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -12.00/2.0/091 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. IT WAS REPORTED THAT THE EXCHANGED LENS RESOLVED THE PROBLEM BUT "STILL A LITTLE HIGH VAULT". Manufacturer Narrative: EVENT: "GLARE AND HALOS WERE OBSERVED." SHOULD BE ADDED IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | -12.00/2.0/091 | OD |  | 11/20/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9492867&pc=QCB | 2023826-2019-02419 | 2019-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -12.00/2.0/089 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. IT WAS REPORTED THAT THE EXCHANGED LENS RESOLVED THE PROBLEM BUT "STILL A LITTLE HIGH VAULT". ABNORMAL SULCUS TO SULCUS WIDTH WAS REPORTED AS CAUSE OF EVENT. Manufacturer Narrative: CORRECTED DATA: B5- " GLARE AND HALOS WERE OBSERVED." SHOULD BE ADDED IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. (B)(4). | -12.00/2.0/089 | OS |  | 11/20/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9498138&pc=QCB | 2023826-2019-02410 | 2019-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARES AND HALOS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO SIZING VARIATION IN OUTCOME. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+1.0/094 | OS |  | 11/20/2019 | 2019 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9633747&pc=QCB | 2023826-2020-00160 | 2019-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.50/4.5/083 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. GLARES/HALOES WERE OBSERVED. ON (B)(6) 2020 THE LENS WAS REMOVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY AFTER EXPLANT (REMOVAL)." Manufacturer Narrative: ADDITIONAL INFORMATION: H3: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- PATIENT CODE 3191 SHOULD BE IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.50/4.5/083 | OD |  | 01/10/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9881022&pc=QCB | 2023826-2020-00659 | 2019-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED GLARE AND HALOS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.5/+1.5/084 | OD |  | 03/03/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9916667&pc=QCB | 2023826-2020-00735 | 2018-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.5/+3.5/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND GLARE/HALOS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND RESIDUE AND DEBRIS ON THE LENS AND LENS WAS RETURNED IN PIECES - UNABLE TO EVALUATE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.5/+3.5/082 | OD |  | 03/11/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9916671&pc=QCB | 2023826-2020-00736 | 2018-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+3.5/105 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND GLARE/HALOS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+3.5/105 | OS |  | 03/11/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10105254&pc=QCB | 2023826-2020-01079 | 2019-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -12.00/1.5/119 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO GLARE/HALOES AND LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/1.5/119 | OS |  | 05/11/2020 | 2020 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138815&pc=QCB | 2023826-2020-01192 | 2020-05-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem; No Apparent Adverse Event | Visual Disturbances; Halo; No Code Available | N/A | Event Description: REPORTER INDICATED AFTER IMPLANTABLE COLLAMER LENS SURGERY. STARBURST AND FLOATERS ARE BEING OBSERVED. REPORTEDLY, HIGH REFRACTIVE POWER (AROUND -16 D) AND PATIENT NOTED STARBURST, FLOATERS AND HALOS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.50/1.5/039 (SPERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT HAS EXPERIENCED STARBURST, FLOATERS, AND GLARE/HALOES FOR THE PAST TWO YEARS. THE PATIENT HAS BEEN REGULARLY VISITING THE SURGEON SINCE IMPLANTATION. THE LENS REMAINS IMPLANTED. H6 - DEVICE CODE 3189 IS NOT APPLICABLE IN INITIAL MDR H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PRECLUDE PERMANENT IMPAIRMENT/DAMAGE IN INITIAL MDR. B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.50/1.5/039 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT HAS EXPERIENCED STARBURST, FLOATERS, AND GLARE/HALOES FOR THE PAST TWO YEARS. THE CLINIC SUGGESTED TO USE PILO (MEDICATION) FOR A MONTH (DURING NIGHT TIME). THE PATIENT WAS FINE FOR 3 DAYS BUT STARTED HAVING THE SAME PROBLEM BEYOND THAT. THE PATIENT HAS BEEN REGULARLY VISITING THE SURGEON SINCE IMPLANTATION. THE LENS REMAINS IMPLANTED. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL AND SUPPLEMENTAL MDR. H6 - HEALTH EFFECT IMPACT CODE 4644 - PILO MEDS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17.50/1.5/039 | OS |  | 05/22/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10184788&pc=QCB | 2023826-2020-01309 | 2020-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.00/2.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS REMOVED DUE TO GLARE/HALOS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/2.0/091 | OD |  | 06/02/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10184802&pc=QCB | 2023826-2020-01310 | 2020-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.50/3.0/096 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS REMOVED DUE TO GLARE/HALOS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED DRY IN A MICO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.50/3.0/096 | OS |  | 06/02/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10491985&pc=QCB | 2023826-2020-01969 | 2020-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+1.5/108 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, SMALL ANTERIOR CHAMBER AND GLARE/HALOS. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "LENS WAS EXCHANGED ON (B)(6) 2020 WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM." H3- DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH DEBRIS ON THE LENS. H6- DEVICE IMPACT CODE 4629- LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: "H6- DEVICE HISTORY RECORD- BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK WORD(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE." SHOULD HAVE BEEN IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -7.0/+1.5/108 | OD |  | 08/10/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10585633&pc=QCB | 2023826-2020-02222 | 2020-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Visual Disturbances; No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REPORTEDLY, "PATIENT STILL HAVING SOME GHOSTING ON THE LETTERS ON THE CHART NO MATTER WHERE THE AXIS IS TURNED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -5.00/2.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THERE WAS LENS ROTATION NOT ASSOCIATED WITH LOW VAULT WITH SOME "GHOSTING OF LETTERS ON THE CHART NO MATTER WHERE THE AXIS IS TURNED", AND ON (B)(6) 2020 THE LENS WAS REPOSITIONED AND THE PROBLEM WAS RESOLVED. POSTOPERATIVE REPORT INDICATED THAT "NORMAL POST OP COURSE. PATIENT STILL HAS A SUTURE OS AT AXIS 85". CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. G3 - DATE IN INITIAL MEDWATCH REPORT SHOULD BE CORRECTED TO 28-AUG-2020. H6 - HEALTH EFFECT CLINICAL CODE: 2140 SHOULD BE CORRECTED TO 4582. H6 - HEALTH EFFECT IMPACT CODE: 4625 - ADDITIONAL SUTURE. H6 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). (B)(4). | -5.00/2.0/089 | OS |  | 08/28/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10823660&pc=QCB | 2023826-2020-02592 | 2020-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Visual Disturbances; No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.50/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: IT WAS LATER CLARIFIED THAT THE PATIENT ALSO EXPERIENCED REFRACTIVE SURPRISE WHICH WAS ALSO RESOLVED WITH THE NOTED LENS EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: H6: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. SHOULD HAVE PREVIOUSLY BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -14.50/2.5/088 | OD |  | 10/20/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14439778&pc=QCB | 2023826-2022-01596 | 2022-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Unintended Movement; No Apparent Adverse Event | Visual Disturbances; No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER STATES THAT A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS -8.0/+2.0/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. REPORTEDLY, THE LENS FLIPPED AND WAS REMOVED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5: THE REPORTER STATED THAT THE SURGEON IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY A 12.6MM VICM5_12.6 -8.00/+2.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 DUE TO LENS BEING IMPLANTED UPSIDE DOWN. THIS RESOLVED THE PROBLEM. THE STATUS OF EYE IS REPORTED AS " BLURRY VISION AT NIGHT, PATIENT SEE HALOES AND GLARE AT NIGHT." CAUSE OF THE EVENT IS REPORTED AS "DEVICE". H6: DEVICE CODE: 1494: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM# (B)(4). Manufacturer Narrative: B5: THE REPORTER STATED THAT THE SURGEON IMPLANTED AND FLIPPED A 12.6MM VICM5_12.6 -8.00/+2.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 DUE TO THE LENS BEING IMPLANTED UPSIDE DOWN. ON THE SAME DAY BUT DIFFERENT SURGERY, THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. THE STATUS OF EYE IS REPORTED AS " BLURRY VISION AT NIGHT, PATIENT SEE HALOES AND GLARE AT NIGHT." CAUSE OF THE EVENT IS REPORTED AS "DEVICE". CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION- UNK. IMPLANT DATE, EXPLANT DATE-UNK. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -8.0/+2.0/088 | OD | Physician | 08/09/2022 | 2022 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14887046&pc=QCB | 2023826-2022-02097 | 2022-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.00/+4.50/84 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -14.00/+4.50/84 | OD | Physician | 06/03/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14889364&pc=QCB | 2023826-2022-02098 | 2022-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.50/+4.50/81 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.50/+4.50/81 | OS | Physician | 06/03/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8688274&pc=QCB | 2023826-2019-01077 | 2019-04-01 | Injury | STAAR SURGICAL COMPANY | 2019-06-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICM5_13.2, -10/0.5/75 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO A PATIENT RELATED FACTOR. IT WAS REPORTED THAT THE PROBLEM RESOLVED AFTER THE EXPLANT, PATIENTS CURRENT CONDITION IS WELL AND HAPPY. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10/0.5/75 | OS | Physician | 05/16/2019 | 2019 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8688384&pc=QCB | 2023826-2019-01078 | 2019-04-01 | Injury | STAAR SURGICAL COMPANY | 2019-06-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICM5_12.6, -8.0/0.5/81 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO A PATIENT RELATED FACTOR. IT WAS REPORTED THAT THE PROBLEM RESOLVED AFTER THE EXPLANT, PATIENTS CURRENT CONDITION IS WELL. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.0/0.5/81 | OD | Physician | 05/16/2019 | 2019 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10053827&pc=QCB | 2023826-2020-01098 | 2019-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Visual Disturbances; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-18.0/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO A PATIENT RELATED ISSUE AND THE PROBLEM RESOLVED. THE REPORTER STATED " THE PATIENT WAS UNABLE TO ADAPT TO THE GLARE AFTER THE OPERATION AND ASKED TO REMOVE THE LENS AND CANCEL THE OPERATION." Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- RESULT CODE IN INITIAL MDR SHOULD BE 213. H6- RESULT CODE 114 SHOULD BE IN PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.0/2.0/090 | OD |  | 04/16/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138800&pc=QCB | 2023826-2020-01190 | 2020-03-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.00/2.0/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT WITH ROTATION WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020 AND (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. IN THE REPORTER'S OPINION, THE CAUSE OF THE EVENT IS ATTRIBUTED TO THE DEVICE. "LENS SIZING IS SMALL, CAUSING LENS TO ROTATE." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: B5 - IT WAS REPORTED PATIENT ALSO EXPERIENCED DOUBLE VISION WHICH WAS RESOLVED WITH THE LENS EXCHANGE. THIS SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/2.0/112 | OS |  | 05/21/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10369229&pc=QCB | 2023826-2020-01678 |  | Injury | STAAR SURGICAL COMPANY | 2020-08-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.50/3.0/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED HIGH GLARE IMMEDIATELY AFTER REMOVING THE DRESSING ONE WEEK AFTER SURGERY AND IS STILL SUFFERING FROM HIGH GLARE. THE REPORTER STATED "ANNOYING GLARE ESPECIALLY AT NIGHT, ESPECIALLY WHEN DRIVING. PATIENT DECLARES AN IMPROVEMENT DURING THE INSTALLATION OF THE PILOCARPINE BUT UPON STOPPING THE GLARE RESUMES."ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.50/3.0/100 | OD | Physician | 07/09/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12628804&pc=QCB | 2023826-2021-03986 | 2021-04-29 | Injury | STAAR SURGICAL COMPANY | 2021-10-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Impairment | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2M VTICMO13.2 IMPLANTABLE COLLAMER LENS, -0.50/+6.0/160 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED PERMANENT LOSS OF BCVA. THE LENS WAS REPLACED ON (B)(6) 2021 WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR (THE WRONG REFRACTION WAS TYPED INTO THE ICL FORM). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. G4: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -0.50/+6.0/160 | OD | Physician | 09/14/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12628810&pc=QCB | 2023826-2021-03990 | 2021-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Impairment | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/+1.0/090, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED ABOUT HIS VISION AND IT WAS FOUND THE WRONG LENS WAS IMPLANTED IN ERROR. THE LENS WAS REPLACED WITH THE CORRECT LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: AFTER THE LENS WAS REPLACED WITH THE CORRECT LENS, THE PROBLEM WAS RESOLVED. THERE WERE NO SYMPTOMS AND NO PATIENT COMPLAINT. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. CORRECTED DATA: B3: 13-SEP-2021. CLAIM#: (B)(4). Manufacturer Narrative: AGE: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -14.50/+1.0/090 | OD | Physician | 09/17/2021 | 2021 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14365144&pc=QCB | 2023826-2022-01479 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Impairment; No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+2.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS REPORTED AS LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+2.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2012. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE CHANGE OVER TIME. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. CLAIM# (B)(4). | -09.50/+2.5/086 | OS | Physician | 04/11/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17200399&pc=QCB | 2023826-2023-02519 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-06-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Vitreous Floaters | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED, A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT COMPLAINED OF VITREOUS FLOATERS POST-OP. AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED, BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS PATIENT-RELATED FACTOR AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED (PRE-EXISTING CONDITION OF FLOATERS). THE LENS REMAINED IMPLANTED AND THE EYE IS HEALTHY. CLAIM # (B)(4). | -11.00/+1.0/085 | OS | Other Health Care Professional | 06/07/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16873135&pc=MTA | 2023826-2023-01664 | 2023-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Patient Device Interaction Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTED COLLAMER LENS INTO THE PATIENT'S EYE ON (B)(6) 2023. IT WAS INDICATED THAT THE LENS WAS DISPLACED THROUGH THE PUPIL. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6: HEALTH EFFECT- CLINICAL CODE: 4581- "IRIS ATROPHY." B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.50/3.0/141 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS DISLOCATION AND IRIS ATROPHY. REPORTEDLY "HALF OF THE LENS DISPLACED THROUGH THE PUPIL" AND "PLANNING TO REPOSITION". THE LENS WAS REPOSITIONED AND THIS RESOLVED THE PROBLEM. STATUS OF THE EYE: "BACK TO NORMAL". THE CAUSE OF THE EVENT: PATIENT RELATED FACTOR - "PATIENT INADVERTENTLY RUBBED EYE". CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- IMPACT CODE: "4628" SHOULD BE ADDED. H6- TYPE OF INVESTIGATION CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -9.50/3.0/141 | OD | Physician | 04/12/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17056304&pc=MTA | 2023826-2023-02129 | 2023-05-10 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | UNK | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). LOW VAULT WAS OBSERVED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.0/053 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. LOW VAULT WAS OBSERVED. PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS UNKNOWN. B7: PREVIOUS CORNEAL OR REFRACTIVE SURGERY. OTHER: PREVIOUS POSTERIOR CHAMBER LENS IMPLANTATION. H6: HEALTH IMPACT- CLINICAL CODE:4581 - REFRACTIVE CHANGE OVERTIME. H6: WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. Manufacturer Narrative: A4-A6:UNK; B3: UNK; D4 (EXPERIATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK; H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM # (B)(4). | -10.0/1.0/053 | OD | Physician | 05/12/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4551528&pc=MTA | 2023826-2015-00205 | 2014-12-25 | Injury | STAAR SURGICAL COMPANY | 2015-02-26 | MTA | VISIAN TICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND AN REFRACTIVE SURPRISE WAS NOTED. PATIENT EXPERIENCING BLURRED VISION. THE LENS WAS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S BUT NOT MARKETED IN THE U.S. (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH, DEVICE HISTORY RECORD. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: ACCORDING TO THE SURGEON, THE LENS WAS ALIGNED AT 20 DEGREES CCW, WHICH IS 5 DEGREES OFF-AXIS. THE PATIENT EXPERIENCED BLURRED VISION AND THE SURGEON DECIDED TO PERFORM A SECONDARY INTERVENTION TO ROTATE THE LENS 31 DEGREES ACCORDING TO ABERROMETRY TESTING. THIS RESOLVED THE PROBLEM. SEVERAL FACTORS MAY BE INVOLVED IN A POST-OPERATIVE REFRACTIVE SURPRISE, INCLUDING INAPPROPRIATE REFRACTIVE SURGICAL PLANNING, A WRONG LENS CALCULATION/SELECTION, PREOPERATIVE INACCURATE DETERMINATION OF HE EYE'S MEASUREMENT TO DETERMINE THE POWER OF THE ICL, INDUCED REFRACTIVE CHANGE BY INCISIONS, UNSTABLE CORNEA AND/OR REFRACTION PRIOR TO IMPLANTATION, CL-INDUCED WARPAGE, WRONG LENS, UPSIDE-DOWN LENS, MISALIGNED LENS, ROTATION, TILTING, ECT. THERE IS NOT ENOUGH CLINICAL DATA TO DETERMINE THE CAUSE OF THE EVENT. CONCLUSION: BASED ON THE COMPLAINT INFORMATION, LENS WORK ORDER SEARCH AND MEDICAL REVIEW, WE WERE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. THE LENS REMAINS IMPLANTED. (B)(4). | -9.5/+4/08 | OD | Physician | 01/29/2015 | 2015 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6037536&pc=MTA | 2023826-2016-01482 | 2016-09-23 | Injury | STAAR SURGICAL COMPANY | 2016-10-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS - -18/1/145 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS REMAINS IMPLANTED. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE SURGEON DECIDED TO MAKE A LASIK TOUCH-UP FOR BOTH EYES TO CORRECT REFRACTIVE SURPRISE. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18/1/145 | OD | Physician | 09/23/2016 | 2016 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6231187&pc=MTA | 2023826-2017-00030 |  | Injury | STAAR SURGICAL COMPANY | 2017-01-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/+2.5/110 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -15.00/+2.5/110 | OD | Physician | 08/01/2016 | 2016 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6074352&pc=MTA | 2023826-2016-00909 | 2016-06-03 | Injury | STAAR SURGICAL COMPANY | 2016-11-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -04.00/+5.5/093 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -04.00/+5.5/093 | OD | Physician | 06/07/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4789550&pc=MTA | 2023826-2015-00578 |  | Injury | STAAR SURGICAL COMPANY | 2015-05-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE REPORTER INDICATED THE PATIENT WAS HAVING DIFFICULTIES WITH NEAR VISION (BLURRED). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4) IMPLANTED. DEVICE EVALUATED BY MANUFACTURER? NO: LENS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: BASED ON THE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -6.5/+1.5/089 | OD | Other | 04/28/2015 | 2015 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4789592&pc=MTA | 2023826-2015-00579 |  | Injury | STAAR SURGICAL COMPANY | 2015-05-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE REPORTER INDICATED THERE WAS A RESIDUAL HYPEROPIC ERROR AND THE PATIENT WAS HAVING DIFFICULTIES WITH NEAR VISION (BLURRED). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4): (OTHER): LENS IMPLANTED. EVALUATION CODES: METHOD -(OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (OTHER): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS IMPLANTED. Manufacturer Narrative: BASED ON THE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -9.5/+3.0/086 | OS | Other | 04/28/2015 | 2015 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7213981&pc=MTA | 2023826-2018-00135 | 2017-11-01 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-23 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | Blurred Vision | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.5/+2.0/134 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE SURGEON REPORTS LENS ROTATION AND BLURRED VISION. ADDITIONALLY, THE SURGEON STATES THE LENS IS "ON A DEAD VERTICAL POSITION." REPORTEDLY, LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -12.5/+2.0/134 | OD | Physician | 12/29/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7136399&pc=MTA | 2023826-2017-01940 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.5/+3.0/X090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ONE YEAR LATER, THE REPORTER INDICATED A REFRACTIVE SURPRISE. THE PATIENT IS HAPPY AND THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -12.5/3.0/090 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ONE YEAR LATER THE REPORTER INDICATED LENS ROTATION AND REFRACTIVE SURPRISE. THE PATIENT IS HAPPY AND THE LENS REMAINS IMPLANTED IN THE PATIENT'S EYE. (B)(4). Manufacturer Narrative: CORRECTED DATA: THE IMPLANTATION DATE IN THE PREVIOUS SUPPLEMENTAL MDR IS CORRECTED FROM "(B)(6) 2017" TO "(B)(6) 2016". CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.5/+3.0/090 | OD | Physician | 11/28/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6326555&pc=MTA | 2023826-2017-00153 | 2015-02-10 | Injury | STAAR SURGICAL COMPANY | 2017-02-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -4.5/+6.0/90 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. ON PATIENT'S LAST VISIT ON (B)(6) 2015, THE PATIENT'S BEST CORRECTED VISUAL ACUITY WAS 20/20. THE STATUS OF THE EYE WAS A "QUIET EYE" AND "GOOD VAULTING". Manufacturer Narrative: ADDITIONAL DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. THE PATIENT HAS AN ANTERIOR CHAMBER DEPTH OF 2.93MM. THIS LENS MODEL IS INDICATED FOR USE WITH AN ANTERIOR CHAMBER DEPTH (ACD) OF EQUAL OR GREATER THAN 3.0 MM, AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE. CLAIM #(B)(4). LENS NOT RETURNED. | -4.5/+6.0/90 | OS | Physician | 01/17/2017 | 2017 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6429909&pc=MTA | 2023826-2017-00487 | 2015-12-10 | Injury | STAAR SURGICAL COMPANY | 2017-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM13.2, DIOPTER -17.0/+6.0/062 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE AND BLURRY VISION. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED. | -17.0/+6.0/062 | OS | Physician | 06/15/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6501471&pc=MTA | 2023826-2017-00653 | 2016-04-11 | Injury | STAAR SURGICAL COMPANY | 2017-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM12.6, DIOPTER -3.5/3.5/031 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: IN THE INITIAL REPORT, THE LENS MODEL IS DESCRIBED AS A VTICM12.6. THE CORRECT MODEL IS VTICMO12.6. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. OFF-LABEL USE: ACD < 3.0MM (2.87MM), AGE (B)(6) AT TIME OF IMPLANT. (B)(4). | -3.5/3.5/031 | OS | Physician | 03/20/2017 | 2017 | 47
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6564060&pc=MTA | 2023826-2017-00804 | 2016-10-19 | Injury | STAAR SURGICAL COMPANY | 2017-05-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO 13.2, DIOPTER -18.0/+2.5/X178 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED. | -18/2.5/178 | OD | Physician | 04/12/2017 | 2017 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6393933&pc=MTA | 2023826-2017-00368 | 2017-01-02 | Injury | STAAR SURGICAL COMPANY AG | 2017-03-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -0.50/+3.50/168 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAIN IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE NOT EVALUATED BY MANUFACTURE BECAUSE LENS NOT RETURNED. (B)(4). | -0.50/+3.50/168 | OD | Physician | 02/09/2017 | 2017 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7274885&pc=MTA | 2023826-2018-00224 | 2017-02-02 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8/4/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) . THE SURGEON NOTED REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION IS AVAILABLE AS OF THE TIME OF MDR SUBMISSION. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.00/4.00/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON NOTED REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. CORRECTED FROM (B)(6) 2017 TO UNK. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: CORRECTION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS; -8.00/4.00/100 (SPHERE/CYLINDER/AXIS); INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8/4/100 | OD | Physician | 01/17/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7275256&pc=MTA | 2023826-2018-00225 | 2017-02-02 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.50/+3.00/87 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) . THE SURGEON NOTED REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION IS AVAILABLE AS OF THE TIME OF MDR SUBMISSION. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.50/+3.00/87 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON NOTED REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. CORRECTED FROM (B)(6) 2017 TO UNK CLAIM# (B)(4). Manufacturer Narrative: CORRECTION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS; -6.50/+3.00/087 (SPHERE/CYLINDER/AXIS); INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.50/+3.00/87 | OS | Physician | 01/17/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6545668&pc=MTA | 2023826-2017-00764 | 2017-04-04 | Injury | STAAR SURGICAL COMPANY | 2017-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -17.00/250/125 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED, HOWEVER A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: PREVIOUS: LENS REMAINS IMPLANTED. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A SAME SIZE, SIMILAR DIOPTER LENS. UCVA IS 20/20, VAULT IS ~500UM. THE PATIENT IS HAPPY, PROBLEM IS RESOLVED. METHOD CODE, PREVIOUS:(NO TESTING METHODS PERFORMED). UPDATED: (DEVICE FROM SAME LOT EVALUATED). RESULT CODE, PREVIOUS: (OPERATIONAL PROBLEM). UPDATED: (NO FAILURE DETECTED). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQUID IN A LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. DIMENSIONAL INSPECTION FOUND LENS IS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PREVIOUS: LENS DIOPTER -17.00/250/125; CORRECTED: LENS DIOPTER -17.00/2.5/125. DEVICE EVALUATION, PREVIOUS: STATED THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. CORRECTED: NO CLEAR RESIDUE/DEBRIS WAS FOUND ON THE PRODUCT. (B)(4). Manufacturer Narrative: PREVIOUS: LENS REMAINS IMPLANTED. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A SAME SIZE, SIMILAR DIOPTER LENS. UCVA IS 20/20, VAULT IS ~500UM. THE PATIENT IS HAPPY, PROBLEM IS RESOLVED. METHOD CODE, PREVIOUS:(NO TESTING METHODS PERFORMED). UPDATED: (DEVICE FROM SAME LOT EVALUATED). RESULT CODE, PREVIOUS: (OPERATIONAL PROBLEM). UPDATED: (NO FAILURE DETECTED). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -17.00/250/125 | OS | Physician | 04/05/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6680940&pc=MTA | 2023826-2017-01087 | 2017-05-11 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.5/76 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DAY OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. ACCORDING TO THE CUSTOMER, THE SURGEON DOES NOT HAVE A PLAN TO EXCHANGE THE LENS, AND WILL FOLLOW-UP THE PATIENT'S CONDITION WITHIN THE NEXT FEW MONTHS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.5/+2.5/76 | OS | Physician | 06/09/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6941344&pc=MTA | 2023826-2017-01558 | 2017-07-15 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-11 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.5/+3.0/091 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. A REFRACTIVE SURPRISE WAS NOTED BY THE SURGEON DURING THE POST-OP VISIT, DATE OF (B)(6) 2017. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.5/+3.0/091 | OD | Physician | 09/13/2017 | 2017 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6952440&pc=MTA | 2023826-2017-01566 | 2017-08-03 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -9.5/+3.0/076 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. Manufacturer Narrative: CORRECTED DATA: LENS REMAINS IMPLANTED. DOCTOR WANTS TO OBSERVE FOR A FEW MONTHS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT UNDER THE AGE OF 21 YEARS OLD. (B)(4). | -9.5/+3.0/076 | OD | Physician | 09/19/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6938985&pc=MTA | 2023826-2017-01554 | 2017-08-28 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICM5_12.6, -7.5/+3.0/96 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.5/+3.0/96 | OS | Physician | 09/13/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6970586&pc=MTA | 2023826-2017-01595 | 2017-09-25 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_13.2 , -11.5/+1/009 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE DOCTOR HAS DECIDED TO WAIT ANOTHER FEW MONTHS TO SEE WHAT HAS HAPPENED WITH THE REFRACTION. LENS IS STILL CURRENTLY IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | -11.5/+1/009 | OS | Physician | 09/25/2017 | 2017 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6970590&pc=MTA | 2023826-2017-01568 | 2017-09-25 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_13.2 , -11/1.5/151 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | -11/1.5/151 | OD | Physician | 09/25/2017 | 2017 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7097940&pc=MTA | 2023826-2017-01835 | 2017-09-27 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -8.50/+2.00/080 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED IN THE EYE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.50/+2.00/080 | OS | Physician | 11/08/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7097943&pc=MTA | 2023826-2017-01836 | 2017-09-28 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -10.50/+3.50/097 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED IN THE EYE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.50/+3.50/097 | OD | Physician | 11/08/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6987323&pc=MTA | 2023826-2017-01633 | 2017-10-02 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -3.0/4.5/X070 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE SURGEON IS PLANNING TO EXCHANGE THE LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER STATED THAT THE LENS WILL BE ROTATED IN (B)(6) 2018. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -3.0/4.5/070 | OS | Physician | 10/03/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7146626&pc=MTA | 2023826-2017-01941 | 2017-10-19 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -8.0/+2.0/157 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE WITH THE POST-OP MANIFEST REFRACTION, SPH. -0.1 AMD CYL. -1.75. THE SURGEON IS PLANNING TO PERFORM LASIK FOR ADDITIONAL CORRECTION. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.0/+2.0/157 | OD | Physician | 11/27/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7205784&pc=MTA | 2023826-2018-00103 | 2017-12-04 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7, -10.5/+6/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED IN THE EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/+6/098 | OD | Physician | 12/20/2017 | 2017 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6395325&pc=MTA | 2023826-2017-00359 |  | Injury | STAAR SURGICAL COMPANY | 2017-03-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.00/+6.0/111 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DAY OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. THIS CASE IS CURRENTLY UNDER MEDICAL CONSULTATION. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.00/+6.0/111 | OD | Physician | 02/08/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6429982&pc=MTA | 2023826-2017-00475 |  | Injury | STAAR SURGICAL COMPANY | 2017-03-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.5/+2/167 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. DOCTOR IS MONITORING THE PATIENT STATUS. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CONCLUSION CODE: THE REPORTER CONFIRMED THAT THE DOCTOR MADE A MISTAKE WHEN WRITING LENS POWER AND PRE-OP K VALUES IN THE OS IN THE OCOS FORM. (B)(4). Manufacturer Narrative: CONCLUSION CODE: THE REPORTER CONFIRMED THAT THE DOCTOR MADE A MISTAKE WHEN WRITING LENS POWER AND PRE-OP K VALUES IN THE OS IN THE OCOS FORM. (B)(4). Manufacturer Narrative: CONCLUSION CODE: THE REPORTER CONFIRMED THAT THE DOCTOR MADE A MISTAKE WHEN WRITING LENS POWER AND PRE-OP K VALUES IN THE OS IN THE OCOS FORM. (B)(4). | -17.5/+2/167 | OS | Physician | 02/23/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336540&pc=MTA | 2023826-2018-00411 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2, IMPLANTABLE COLLAMER LENS, -13.50/+2.0/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD). CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.50/+2.0/080 | OD | Physician | 02/12/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7398188&pc=MTA | 2023826-2018-00497 | 2018-02-20 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -3.50/5.5/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. IT WAS NOTICED THE THE "REFRACTIVE IS STILL THE SAME AS BEFORE DURING SURGERY" AND THERE IS ALMOST NO CORRECTION. LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, OF -3.50/5.5/107 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. IT WAS NOTICED THAT THE "REFRACTIVE IS STILL THE SAME AS BEFORE DURING SURGERY" AND THERE IS ALMOST NO CORRECTION. LENS REMAINS IMPLANTED IN THE PATIENT'S EYE. FROM THE (B)(4) CALCULATION FORM WE CAN NOTICE THAT THE LEFT EYE (OS) VALUES WERE USED FOR CALCULATION OF RIGHT EYE (OD) LENS REQUIREMENT. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | -3.50/5.5/107 | OD | Physician | 03/06/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7398192&pc=MTA | 2023826-2018-00498 | 2018-02-20 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -4.00/3.0/119 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. IT WAS NOTICED THE "REFRACTIVE IS STILL THE SAME AS BEFORE DURING SURGERY" AND THERE IS ALMOST NO CORRECTION. LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -4.00/ 3.0/119 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. IT WAS NOTICED THAT THE "REFRACTIVE IS STILL THE SAME AS BEFORE DURING SURGERY" AND THERE IS ALMOST NO CORRECTION. LENS REMAINS IMPLANTED IN THE PATIENT'S EYE. FROM THE (B)(4) CALCULATION FORM WE CAN NOTICE THAT THE RIGHT EYE (OD) VALUES WERE USED FOR CALCULATION OF LEFT EYE (OS) LENS REQUIREMENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -4.00/3.0/119 | OS | Physician | 03/06/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6108230&pc=MTA | 2023826-2016-01651 | 2016-10-14 | Injury | STAAR SURGICAL COMPANY | 2016-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5 3.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/074 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CONCLUSION CODE: PREVIOUSLY, IT WAS REPORTED THAT THE CONCLUSION CODE. THIS IS INCORRECT. THE CORRECT CONCLUSION CODE WAS REPORTED IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THE CORRECT MODEL NUMBER AND LENGTH OF THE LENS IS VTICM5 13.2 IMPLANTABLE COLLAMER LENS. REPORT SOURCE: USER FACILITY TO DISTRIBUTOR - PREVIOUS SELECTION 'USER FACILITY' IS INCORRECT, THE CORRECT SELECTION SHOULD BE 'DISTRIBUTOR.' (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.00/+1.5/074 | OD | Physician | 10/17/2016 | 2016 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6223864&pc=MTA | 2023826-2017-00021 | 2016-12-07 | Injury | STAAR SURGICAL COMPANY | 2017-01-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+2.0/066 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -12.50/+2.0/066 | OD | Physician | 12/07/2016 | 2016
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7089425&pc=MTA | 2023826-2017-01830 | 2017-11-07 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGERY IMPLANTED A 13.2 MM VTICM5_13.2, -4.5/+4.5/075 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AS WELL AS REFRACTIVE SURPRISE WAS NOTICED. THE LENS IS PLANNED TO BE REPOSITIONED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -4.5/+4.5/075 | OS | Physician | 11/07/2017 | 2017 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6227635&pc=MTA | 2023826-2017-00016 | 2016-11-18 | Injury | STAAR SURGICAL COMPANY | 2017-01-05 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Insufficient Information | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+3.0/78 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. AFTER IMPLANTATION, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT. Manufacturer Narrative: ADDITIONAL DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED. | -14.0/+3.0/78 | OS | Physician | 12/07/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7201639&pc=MTA | 2023826-2018-00086 | 2017-11-16 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -1.5/+1.50/079 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE AND EXCESSIVE VAULT. IN ADDITION, THE ICL WAS IMPLANTED "IN A PIGGY-BACK CONDITION." REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE U.S. CONCLUSION CODE: OFF-LABEL USE [OVER (B)(6) YEARS OF AGE AT TIME OF IMPLANT ((B)(6) YRS), PREVIOUS CORNEAL OR REFRACTIVE SURGERY]. (B)(4) | -1.5/+1.50/079 | OD | Physician | 12/20/2017 | 2017 | 72
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336748&pc=MTA | 2023826-2018-00377 | 2018-02-01 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.0/+2.5/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+2.5/088 | OS | Physician | 02/16/2018 | 2018 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8052129&pc=MTA | 2023826-2018-01618 | 2018-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -2.50/6.0/168 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND LOSS OF BCVA; THIS EXPLANT RESOLVED THE PROBLEM. IT WAS REPORTED THAT THE PATIENT IS DOING FINE AFTER EXPLANTATION. Manufacturer Narrative: CORRECTED DATA: "KERATOCONUS" NOT INCLUDED IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: (B)(4). LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. UNABLE TO PERFORM DIMENSIONAL INSPECTION SINCE HAPTIC IS MISSING. (B)(4). Manufacturer Narrative: CORRECTED DATA: RESULT CODE CORRECTED TO: "114" ONLY IN THE SUPPLEMENTAL MDR#2, PREVIOUS CODE NOT APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4). | -2.50/6.0/168 | OD |  | 10/24/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7051045&pc=MTA | 2023826-2017-01755 | 2017-09-30 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/099 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD LOW VAULTING AND LENS ROTATION. THE PATIENT COMPLAINED OF BLURRY VISION. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER STATED THAT THE EXCHANGE IS LIKELY TO BE IN (B)(6) 2018. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. REPORT SOURCE: USER FACILITY IS INCORRECT, SHOULD BE DISTRIBUTOR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.0/099 | OS | Physician | 10/30/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6877224&pc=MTA | 2023826-2017-01462 | 2016-02-18 | Injury | STAAR SURGICAL COMPANY | 2017-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -13.5/+1.5/080 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE PATIENT BEGAN EXPERIENCING GLARE/HALOES AND BLURRED VISION. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR THE SAME SIZE, SIMILAR DIOPTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PATIENT POST-OP BCVA WAS 20/25 AND UCVA WAS 20/25 ON (B)(6) 2016. PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/+1.5/080 | OS | Physician | 08/22/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135239&pc=MTA | 2023826-2017-01925 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -5.5/+5.5/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE AND STATES "IN SPITE OF NORMAL VAULT POST-OP DATA" HE BELIEVES THE CURRENT LENS IS LARGER THAN NEEDED. THE SURGEON NOTES THE OPTIC PART OF THE LENS IN THE INFERIROR SIGMENT IS "A BIT OBTRUSIVE." NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: CLAIM #: (B)(4). Manufacturer Narrative: CORRECTED DATA: THE SURGEON NOTES THAT AFTER THE DILATION, "THE OPTIC PART OF THE LENS IN THE INFERIOR SEGMENT IS A BIT PROTRUSIVE". CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -5.5/+5.5/082 | OD | Physician | 11/27/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135248&pc=MTA | 2023826-2017-01926 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.0/+6.0/079 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE AND STATES "IN SPITE OF NORMAL VAULT POST-OP DATA" HE BELIEVES THE CURRENT LENS IS LARGER THAN NEEDED. THE SURGEON NOTES THE OPTIC PART OF THE LENS IN THE INFERIOR SIGMENT IS "A BIT OBTRUSIVE." NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: CORRECTED DATA: THE SURGEON NOTES THAT AFTER THE DILATION, "THE OPTIC PART OF THE LENS IN THE INFERIOR SEGMENT IS A BIT PROTRUSIVE". CLAIM # (B)(4). Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -7.0/+6.0/079 | OS | Physician | 11/27/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6227640&pc=MTA | 2023826-2017-00017 | 2016-12-07 | Injury | STAAR SURGICAL COMPANY | 2017-01-05 | MTA | IMPLANTABLE COLLAMER LENS | Insufficient Information | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/2.0/152 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT. ADDITIONALLY, THE COMPLAINT FORM STATES, "CUSTOMER MIXED UP THE CALCULATION DATA BETWEEN OD AND OS." Manufacturer Narrative: PATIENT WEIGHT-UNK. PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED. | -11.5/2.0/152 | OS | Physician | 12/07/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6242804&pc=MTA | 2023826-2017-00041 | 2016-12-10 | Injury | STAAR SURGICAL COMPANY | 2017-01-11 | MTA | IMPLANTABLE COLLAMER LENS | Insufficient Information | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+3.0/82 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT AMD THEREFORE, HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: PATIENT WEIGHT-UNK. PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED. | -10.5/+3.0/82 | OD | Physician | 12/13/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6242817&pc=MTA | 2023826-2017-00039 | 2016-12-10 | Injury | STAAR SURGICAL COMPANY | 2017-01-11 | MTA | IMPLANTABLE COLLAMER LENS | Insufficient Information | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+3.0/81 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT AND THEREFORE, HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: ADDITIONAL DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTIONAL DATA: CORRECT OUTCOME ATTRIBUTED TO ADVERSE EVENT IS OTHER SERIOUS (DEVICE REMAINS IMPLANTED). (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED. | -10.5/+3.0/81 | OS | Physician | 12/13/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7783494&pc=MTA | 2023826-2018-01202 | 2018-07-18 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.5/+4.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, THE PATIENT IS NOT SATISFIED WITH THE VISION BECAUSE IT IS DIFFERENT FROM THE EXPECTED REFRACTION. Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.5/+4.0/094 | OS | Physician | 07/20/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8071556&pc=MTA | 2023826-2018-01682 | 2018-09-04 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -4.00/4.0/155 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED IN THE EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -4.00/4.0/155 | OS | Physician | 10/24/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8090071&pc=MTA | 2023826-2018-01702 | 2018-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.0/+4.0/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: EVENT - THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.0/4.0/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2018. THE SURGEON NOTES REFRACTIVE SURPRISE AND POST OP DATA SHOWS LOSS OF VISUAL ACUITY. THE LES REMAINS IMPLANTED WITH LASER TREATMENT PLANNED FOR THE FUTURE. DEVICE CODE 4110: LENS WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINT TYPES EVENTS. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+4.0/072 | OS |  | 10/26/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8090052&pc=MTA | 2023826-2018-01684 | 2018-10-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -3.5/+6.0/102 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE SURGEON REPORTS BLURRED VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: DATE OF EVENT CORRECTED TO: "UNKNOWN". CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -3.5/+6.0/102 | OD |  | 10/24/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372545&pc=MTA | 2023826-2019-00314 | 2018-10-23 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -4.5/+6.0/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED AND THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -4.5/+6.0/080 | OD | Physician | 01/31/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8227420&pc=MTA | 2023826-2019-00048 | 2018-11-10 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -4.5/+4.5/178 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -4.5/+4.5/178 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS REPORTED. THE LENS REMAINS IMPLANTED. CORRECTED TO: "PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB". (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -4.5/+4.5/178 | OD | Physician | 12/12/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8241988&pc=MTA | 2023826-2019-00086 | 2018-12-17 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -14.00/4.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SUPRISE WAS OBSERVED ON (B)(6) 2018. REMOVAL OF THE LENS AND EXCHANGE IS PLANNED ONCE NEW LENS IS CALCULATED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.00/4.0/105 | OS | Physician | 12/18/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8403806&pc=MTA | 2023826-2019-00366 | 2019-01-12 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -7.00/6.0/020 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTER STATES, "WE DON'T KNOW WHY SHE STILL HAS A REFRACTIVE CYLINDER ERROR ALTHOUGH THE LENS IS NOT ROTATED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -7.00/6.0/020 | OS | Physician | 02/05/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372275&pc=MTA | 2023826-2019-00302 | 2019-01-19 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.00/+6.0/086 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO ASTIGMATIC AXIS ERROR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.00/+6.0/086 | OD | Physician | 01/28/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11206969&pc=MTA | 2023826-2021-00112 | 2020-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS). REFRACTIVE SURPRISE, LENS ROTATION NOT ASSOCIATED TO LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.00/3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2021 THE LENS WAS REPOSITIONED AND THE THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY". H6 - HEALTH EFFECT IMPACT CODE REPORTED AS 2199 - UPDATED/CHANGED TO 4629 H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: DB6 DATE: LENS NOT EXPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.00/3.0/093 | OS |  | 12/30/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17218627&pc=MTA | 2023826-2023-02556 | 2023-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -13.50/1.5/062 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE IS REPORTED AS UNKNOWN. CLAIM# (B)(4) Manufacturer Narrative: H6 - INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4) | -13.50/1.5/062 | OS | Physician | 05/30/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7783466&pc=MTA | 2023826-2018-01203 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-08-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.5/+4.5/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. Manufacturer Narrative: CORRECTED DATA: DATE OF REPORT SHOULD BE CORRECTED TO 19-JUL-2018 IN THE ORIGINAL MDR. DATE RECEIVED SHOULD BE CORRECTED TO 19-JUL-2018 IN THE ORIGINAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.5/+4.5/100 | OD | Physician | 07/19/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7908155&pc=MTA | 2023826-2018-01426 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -3.00/+3.0/080 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -3.00/+3.0/080 | OS | Physician | 08/27/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7908160&pc=MTA | 2023826-2018-01427 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -1.50/+3.0/X044 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S., (B)(4). | -1.50/+3.0/044 | OS | Physician | 08/27/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7916208&pc=MTA | 2023826-2018-01451 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-09-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.0/4.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED, AS OF THE DATE OF MDR REPORTING NO INTERVENTION HAS BEEN PERFORMED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.0/4.5/094 | NA | Physician | 08/29/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7928342&pc=MTA | 2023826-2018-01467 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-10-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM013.2 IMPLANTABLE COLLAMER LENS, -3.00/+6.0/X092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -3.00/+6.0/092 | OS | Physician | 09/28/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8143302&pc=MTA | 2023826-2018-01787 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-12-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM013.2 IMPLANTABLE COLLAMER LENS, -5.0/+2.0/166 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE REPORTER INDICATED THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME AND THE LENS REMAINS IMPLANTED. THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS REMAINS IMPLANTED. THE PATIENTS CURRENT CONDITION IS VISUAL ACUITY IS NOT AS GOOD AS IT SHOULD BE BUT THE PATIENT CONDITION IS OK. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.0/+2.0/166 | OD |  | 11/08/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13459217&pc=MTA | 2023826-2022-00392 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-02-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+0.5/075 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS BLURRED VISION. REPORTEDLY, "PREPARE FOR FOLLOW-UP OBSERVATION AND VISUAL FUNCTION TRAINING." Manufacturer Narrative: WEIGHT, RACE, ETHNICITY:UNK. SEX: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -8.5/+0.5/075 | OS | Physician | 01/12/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15440641&pc=MTA | 2023826-2022-03146 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-09-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | Event Description: THE REPORTER INDICATED THAT A PATIENT WHO HAD AN IMPLANTABLE COLLAMER LENS IMPLANTED ON (B)(6) 2022 EXPERIENCED A -6D MYOPIC SHIFT AT HIS 1 DAY POST-OP VISIT. PER UPDATED INFORMATION THE REPORTER STATES " THE PATIENT RETURNED FOR HIS 1 WK POST OP CHECK TODAY WITH PLANO REFRACTION AND PERFECT VISION. THERE IS NO ISSUE AT THIS TIME. I BELIEVE THE 1 DAY MYOPIC SHIFT SEEN ON EXAM WAS DUE TO USE OF MIOSTAT DURING SURGERY AND THIS EFFECT HAS RESOLVED". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. DEVICE MANUFACTURING DATE: UNK, NO SERIAL NUMBER REPORTED. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A PATIENT WHO HAD A 13.2MM VTICMO13.2 -16.50/1.5/099 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS IMPLANTED ON (B)(6) 2022 EXPERIENCED A 6D MYOPIC SHIFT AT HIS 1 DAY POST-OP VISIT. PER UPDATED INFORMATION THE REPORTER STATES " THE PATIENT RETURNED FOR HIS 1 WK POST OP CHECK TODAY WITH PLANO REFRACTION AND PERFECT VISION. THERE IS NO ISSUE AT THIS TIME. I BELIEVE THE 1 DAY MYOPIC SHIFT SEEN ON EXAM WAS DUE TO USE OF MIOSTAT DURING SURGERY AND THIS EFFECT HAS RESOLVED". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -16.50/1.5/099 | NA | Physician | 08/17/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16010741&pc=MTA | 2023826-2022-04480 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-12-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -10.5/2.0/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). LENS ROTATION AND REFRACTIVE CHANGE WAS OBSERVED. Manufacturer Narrative: CORRECTED DATA: B1: ADVERSE EVENT SHOULD BE CORRECTED TO PRODUCT PROBLEM. B2: OUTCOMES ATTRIBUTED TO ADVERSE EVENT- SHOULD BE OMITTED FOR CORRECTION. CLAIM#(B)(4). Manufacturer Narrative: WEIGHT AND RACE:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -10.5/2.0/073 | OD | Physician | 11/21/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16010748&pc=MTA | 2023826-2022-04486 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-12-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -10.5/2.0/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). LENS ROTATION AND REFRACTIVE CHANGE WAS OBSERVED. Manufacturer Narrative: CORRECTED DATA: B1: ADVERSE EVENT SHOULD BE CORRECTED TO PRODUCT PROBLEM. B2: OUTCOMES ATTRIBUTED TO ADVERSE EVENT- SHOULD BE OMITTED FOR CORRECTION. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: CORRECTION LENS DIOPTER IS -11.0/1.0/118 . CLAIM#: (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). | -10.5/2.0/073 | OS | Physician | 11/21/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12920600&pc=MTA | 2023826-2021-04619 | 2021-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS REFRACTIVE SURPRISE AND LENS DISLOCATION. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: RECEIVED CQ ON (B)(6) 2022: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS WITH DIOPTER -17.50/+3.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS (13.7MM VTICMO13.7 DIOPTER -17.0/+3.0/105) ON (B)(6) 2022 DUE TO REFRACTIVE SURPRISE AND LENS DISLOCATION/SUBLUXATION. PROBLEM RESOLVED. REPORTEDLY, "PATIENT IS HAPPY!" THE CAUSE OF THIS EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: A LENS WORK ORDER SEARCH WAS PERFORMED. CLAIM#: (B)(4) IS REGARDING EXCESSIVE VAULTING WHICH IS UNRELATED TO THIS COMPLAINT. NO SIMILAR COMPLAINTS FOUND. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO-UNK. DATE OF EVENT-UNK. PRODUCT CODE UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK/UNK ENTRY. PRODUCT INFO, IMPLANT DATE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE-UNK. (B)(4). | -17.50/+3.0/093 | OS | Physician | 11/04/2021 | 2021 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372531&pc=MTA | 2023826-2019-00316 | 2018-10-26 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Improper or Incorrect Procedure or Method | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -3.5/+5.5/079 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED AND THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -3.5/+5.5/079 | OS | Physician | 01/31/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824424&pc=MTA | 2023826-2018-01262 | 2018-06-11 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -3.00/+6.0/015 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND THERE WAS REDUCTION OF BEST CORRECTED VISUAL ACUITY. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: (B)(4).. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -3.00/+6.0/015 | OS | Physician | 07/29/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7765095&pc=MTA | 2023826-2018-01173 | 2018-07-09 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-08 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -4.0/+2.5/076 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -4.0/+2.5/076 | OS | Physician | 07/17/2018 | 2018 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8351772&pc=MTA | 2023826-2019-00274 | 2019-01-08 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.50/4.0/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. IT WAS REPORTED THAT REFRACTIVE SUPRISE WAS OBSERVED AND SURGEON HAS NOT YET MADE THE DECISION WITH PATIENT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.50/4.0/93 | OD | Physician | 01/22/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8354792&pc=MTA | 2023826-2019-00275 | 2019-01-08 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.50/4.0/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. IT WAS REPORTED THAT REFRACTIVE SURPRISE WAS OBSERVED AND SURGEON HAS NOT YET MADE THE DECISION WITH PATIENT. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/4.0/90 | OS | Physician | 01/22/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6284774&pc=MTA | 2023826-2017-00156 | 2017-01-03 | Injury | STAAR SURGICAL COMPANY | 2017-01-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Use of Device Problem | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS,-4.0/+1.5/17 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LEFT EYE WAS IMPLANTED WITH THE CALCULATED RIGHT EYE (OD) LENS, RESULTING IN REFRACTIVE SURPRISE. REPOSITIONING SURGERY IS PLANNED FOR (B)(6) 2017. Manufacturer Narrative: LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2017 AND THE OUTCOME WAS GOOD. THE PATIENT'S LAST VISIT WAS ON (B)(6) 2017 WITH UNCORRECTED VISUAL ACUITY (UCVA) 20/25. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -4.0/+1.5/17 | OS | Physician | 01/03/2017 | 2017 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6284878&pc=MTA | 2023826-2017-00155 | 2017-01-03 | Injury | STAAR SURGICAL COMPANY | 2017-01-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Use of Device Problem | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.0/+1.5/172 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE RIGHT EYE (OD) WAS IMPLANTED WITH THE CALCULATED LEFT EYE (OS) LENS, RESULTING IN REFRACTIVE SURPRISE. REPOSITIONING SURGERY IS PLANNED FOR (B)(6) 2017. Manufacturer Narrative: LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2017 AND THE OUTCOME WAS GOOD. THE PATIENT'S LAST VISIT WAS ON (B)(6) 2017 WITH UNCORRECTED VISUAL ACUITY (UCVA) 20/25. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: THE CUSTOMER REPORTED THAT THE SURGEON INADVERTENTLY IMPLANTED THE LENS INTERCHANGED BETWEEN EYES. CLAIM # (B)(4). Manufacturer Narrative: LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2017 AND THE OUTCOME WAS GOOD. THE PATIENT'S LAST VISIT WAS ON (B)(6) 2017 WITH UNCORRECTED VISUAL ACUITY (UCVA) 20/25. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: THE CUSTOMER REPORTED THAT THE SURGEON INADVERTENTLY IMPLANTED THE LENS INTERCHANGED BETWEEN EYES. CLAIM # (B)(4). | -4.0/+1.5/172 | OD | Physician | 01/03/2017 | 2017 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6508705&pc=MTA | 2023826-2017-00682 | 2017-02-09 | Injury | STAAR SURGICAL COMPANY | 2017-04-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Use of Device Problem | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.00/+2.00/180 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE CUSTOMER STATED THAT THE DR. MADE A MISTAKE IN CALCULATION OF THE OPTIC. AT THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE SURGEON PERFORMED THE LIMBAL INCISION TO REDUCE ASTIGMATISM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.00/+2.00/180 | OS | Physician | 03/23/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6711503&pc=MTA | 2023826-2017-01175 | 2017-04-17 | Injury | STAAR SURGICAL COMPANY | 2017-07-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; Discomfort; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, DIOPTER -15.5/+2.0/110 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULT, REFRACTIVE SURPRISE, AND UNCOMFORTABLE VISION. THE LENS WAS REPOSITIONED THREE TIMES ON (B)(6) 2015, (B)(6) 2016, AND (B)(6) 2017 BUT UNSUCCESSFUL, THE LENS WAS ROTATED AGAIN. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED, THE LENS IS IN POSITION AT DESIRED AXIS. AS OF THE DATE OF MDR SUBMISSION THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN US BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.5/+2.0/110 | OD | Physician | 06/22/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135234&pc=MTA | 2023826-2017-01923 | 2017-10-26 | Injury | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; Halo; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -4.5/+2.5/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE, BLURRED VISION, GLARE/HALOES. REPORTEDLY, THE LENS WAS REPOSITIONED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: THE REPORTER STATED THE LENS WAS EXPLANTED ON (B)(6) 2018. (B)(4) Manufacturer Narrative: THE REPORTER STATES THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2018. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -4.5/+2.5/100 | OD | Physician | 11/27/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5525079&pc=MTA | 2023826-2016-00333 | 2015-11-09 | Injury | STAAR SURGICAL COMPANY | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5/+1.5/130 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO A REFRACTIVE SURPRISE, BLURRED VISION AND GLARES/HALOS. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | -7.5/+1.5/130 | OD | Physician | 02/27/2016 | 2016 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6559295&pc=MTA | 2023826-2017-00792 | 2017-04-13 | Injury | STAAR SURGICAL COMPANY | 2017-05-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/090 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE REPORTER INDICATED AT ONE DAY POST-OP, THE LENS HAD AN EXCESSIVE VAULT, REFRACTIVE SURPRISE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE SURGEON REPOSITIONED THE LENS ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PATIENT EXPERIENCED GLARE/HALOES. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+1.0/090 | OD | Physician | 04/14/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4919768&pc=MTA | 2023826-2015-00838 | 2015-02-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-16 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND EXCESSIVE VAULT WAS OBSERVED. PATIENT EXPERIENCED BLURRED VISION AND GLARE/HALOS. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2015 AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. BASED ON THE COMPLAINT INFORMATION AND LENS WORK ORDER SEARCH, WE WERE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4). | -15.0/+1.5/130 | OD | Physician | 06/16/2015 | 2015 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5312180&pc=MTA | 2023826-2015-01675 | 2015-06-18 | Injury | STAAR SURGICAL COMPANY | 2015-12-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/006 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE, HALOS AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH PERFORMED AND FOUND NO SIMILAR COMPLAINTS. LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO,LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -13.0/+1.5/006 | OD | Physician | 11/26/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5685199&pc=MTA | 2023826-2016-00756 | 2016-04-14 | Injury | STAAR SURGICAL COMPANY | 2016-05-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/008 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT AND A REFRACTIVE SURPRISE. THE PATIENT'S BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE AND FOREIGN MATERIAL - FIBERS PRESENT ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS IS IN SPECIFICATIONS. CORRECTIONAL DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4). | -18.0/+1.5/008 | OS | Physician | 05/03/2016 | 2016 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5831916&pc=MTA | 2023826-2016-00948 |  | Injury | STAAR SURGICAL COMPANY | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+3.5/113 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT AND THE PATIENT EXPERIENCED BLURRED VISION, GLARE AND HALOS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: METHOD: (DEVICE FROM SAME LOT EVALUATED); WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). Manufacturer Narrative: DATE OF EVENT: UNK. DATE OF EXPLANT: UNK. DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). | -7.0/+3.5/113 | OS | Physician | 06/29/2016 | 2016
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6216613&pc=MTA | 2023826-2016-01850 | 2016-10-09 | Injury | STAAR SURGICAL COMPANY | 2016-12-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50/+5.00/054 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. AFTER TWO WEEKS, THE LENS WAS REPOSITIONED. ADDITIONALLY, ON (B)(6) 2016 THE LENS WAS EXPLANTED DUE TO LOW VAULTING, ROTATION, GLARE/HALOES, BLURRED VISION, AND LOSS OF BCVA (BEST CORRECTED VISUAL ACUITY). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE MANUFACTURE DATE: PREVIOUS DATE WAS INCORRECT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4), SECONDARY SURGICAL INTERVENTION, REPOSITIONING, LOSS OF BCVA, LENS EXPLANT. (B)(4). LENS NOT RETURNED. | -08.50/+5.00/054 | OS | Physician | 12/02/2016 | 2016 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5540150&pc=MTA | 2023826-2016-00375 | 2016-01-15 | Injury | STAAR SURGICAL COMPANY | 2016-03-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+2.0/076 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, LENS DISLOCATION/SUBLUXATION AND ROTATION. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOS AND BLURRED VISION. THE LENS WAS REPOSITIONED AND THE INCISION WAS ENLARGED. THE LENS WAS EXCHANGED FOR A LONGER LENS BUT THAT LENS ALSO TILTED AND SUBLUXED - SEE MFR REPORT # 2023826-2016-00376. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/80 AND THE PATIENT EXPERIENCED BLUR/HALOS AND DIPLOPIA. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAPTIC SCRATCHED; DIMENSIONAL INSPECTION FOUND THE LENS IS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DATE OF BIRTH - UNK. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -16.50/+2.0/076 | OS | Physician | 03/03/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5047717&pc=MTA | 2023826-2015-01115 | 2014-11-21 | Injury | STAAR SURGICAL COMPANY | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -15.5/+4.0/94 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON 11/10/2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, GLARE/HALOS AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4) NEW INCISION. SECONDARY SURGERY. GLARE. LENS EXPLANTED. METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). LENS NOT RETURNED. | -15.5/+4.0/94 | OD | Physician | 08/04/2015 | 2015 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5540385&pc=MTA | 2023826-2016-00376 |  | Injury | STAAR SURGICAL COMPANY | 2016-03-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.50/+2.5/075 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS DISLOCATED/SUBLUXED. THE LENS REMAINS IMPLANTED. THIS WAS THE REPLACEMENT LENS - SEE MFR REPORT # 2023826-2016-00375 FOR THE FIRST LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/80 AND THE PATIENT EXPERIENCED BLUR/HALOS AND DIPLOPIA. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. METHOD: WORK ORDER SEARCH. EVALUATION RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -16.50/+2.5/075 | OS | Physician | 03/03/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5432678&pc=MTA | 2023826-2016-00150 | 2015-08-25 | Injury | STAAR SURGICAL COMPANY | 2016-02-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; Loss of Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/052 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOSS OF BEST-CORRECTED VISUAL ACUITY (POST-OP BCVA - 20/25). THE LENS WAS EXCHANGED FOR A SHORTER LENS (VICMO12.6, -11.0 DIOPTER) AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN AND SCRATCHES. (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -10.5/+1.5/052 | OS | Physician | 12/22/2015 | 2015 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5085796&pc=MTA | 2023826-2015-01195 | 2015-07-27 | Injury | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Loss of Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.50/+5.0/100 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD AN EXCESSIVE VAULT AND THE PATIENT EXPERIENCED LOSS OF BEST CORRECTED VISUAL ACUITY. THE PATIENT ALSO EXPERIENCED PIGMENT DISPERSION, UNREACTIVE PUPIL (FIXED), BLURRED VISION AND INFLAMMATION. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: BASED ON THE ABOVE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT DATE OF BIRTH - UNK. PT WEIGHT: UNK. EXPLANT DATE: UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | -8.50/+5.0/100 | OD | Physician | 08/20/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5671131&pc=MTA | 2023826-2016-00738 | 2016-01-23 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Loss of Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND LOSS OF BEST-CORRECTED VISUAL ACUITY (POST-OP BCVA - 20/20). THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/16. Manufacturer Narrative: THE LENS DESCRIPTION IS -18.0/+2.0/X087 (SPHERE/ CYLINDER/ AXIS). PATIENT'S LAST VISIT WAS ON (B)(6) 2016 WITH UNCORRECTED VISUAL ACUITY (UCVA) 20/40. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT OPTIC AND HAPTIC WERE TORN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | -18.0/+2.0/090 | OS | Physician | 04/21/2016 | 2016 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5867161&pc=MTA | 2023826-2016-01048 | 2016-02-12 | Injury | STAAR SURGICAL COMPANY | 2016-08-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Torn Material; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Loss of Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/097 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS TORE DURING DELIVERY INTO THE EYE. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT, LENS ROTATION/DISLOCATION. THE PATIENT EXPERIENCED LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS. THE EXCHANGED LENS RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE PATIENT'S POST-OP (ON (B)(6) 2016) UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. CORRECTED DATA: THE CORRECT REPORT SOURCE IS DISTRIBUTOR, NOT USER FACILITY. (B)(4). Manufacturer Narrative: (B)(4). | -9.5/+1.5/097 | OD | Physician | 07/11/2016 | 2016 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5834798&pc=MTA | 2023826-2016-00996 | 2016-06-24 | Injury | STAAR SURGICAL COMPANY | 2016-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Loss of Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.00/+5.0/090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, REFRACTIVE SURPRISE AND LENS DISLOCATION/SUBLUXATION. THE PATIENT EXPERIENCED LOSS OF BEST-CORRECTED VISUAL ACUITY AND BLURRED VISION. THE REPORTER INDICATED THE LENS ROTATION WAS DUE TO WEAKNESS OF ZONULES. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE FOREIGN MATERIAL ON LENS SURFACE - FIBERS. DIMENSIONAL INSPECTION FOUND THE LENS IS IN SPECIFICATIONS. THE PATIENT IS UNDER THE AGE OF 21 YEARS. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | -6.00/+5.0/090 | OS | Physician | 06/30/2016 | 2016 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5834802&pc=MTA | 2023826-2016-00997 | 2016-06-24 | Injury | STAAR SURGICAL COMPANY | 2016-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Loss of Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.50/+5.0/125 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, REFRACTIVE SURPRISE AND LENS DISLOCATION/SUBLUXATION. THE PATIENT EXPERIENCED LOSS OF BEST-CORRECTED VISUAL ACUITY AND BLURRED VISION. THE REPORTER INDICATED THE LENS ROTATION WAS DUE TO WEAKNESS OF ZONULES. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (FIBERS). DIMENSIONAL INSPECTION FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION: PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS OLD. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | -5.50/+5.0/125 | OD | Physician | 06/30/2016 | 2016 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6001950&pc=MTA | 2023826-2016-01394 | 2016-08-10 | Injury | STAAR SURGICAL COMPANY | 2016-10-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Use of Device Problem; Insufficient Information | Blurred Vision; Loss of Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+2.0/001 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOSS OF BEST CORRECTED VISUAL ACUITY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/SIZE LENS BUT DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS, BUT FOREIGN MATERIAL ON LENS SURFACE (NOT POSSIBLE TO SPECIFY). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT. DATE OF BIRTH AND WEIGHT: - UNK. (B)(4). | -9.0/+2.0/001 | OD | Physician | 09/12/2016 | 2016 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15341923&pc=MTA | 2023826-2022-03007 | 2022-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. IT HAS BEEN REPORTED THE LENS HAS ROTATED. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.50/+2.5/171 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION AND BLURRED VISION. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "SHORT DIAMETER". H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -6.50/+2.5/171 | OD | Other Health Care Professional | 08/04/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14402698&pc=MTA | 2023826-2022-01529 | 2022-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER STATES AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON REPORTS AN OVERSIZED ICL. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. B3 - DATE OF EVENT (B)(6) 2022 SHOULD BE ADDED TO THE INITIAL MDR. B5 - "THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS, -17.00/2.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE LOW VAULT WITH ROTATION, LENS "DISLOCATION"/SUBLUXATION, AND BLURRED VISION. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/ LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER MARKED THE CAUSE OF THE EVENT AS THE DEVICE." SHOULD BE ADDED TO THE INITIAL MDR. D6A - IMPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4627 SHOULD BE ADDED TO INITIAL MDR. H6 - HEALTH EFFECT - CLINICAL CODE: 2137 SHOULD BE ADDED TO INITIAL MDR. H6 - HEALTH EFFECT - DEVICE CODE: 1401 -2923 SHOULD BE ADDED TO INITIAL MDR. H6 - HEALTH EFFECT - TYPE IF INVESTIGATION CODE: 4110-41111 :" LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT." SHOULD BE ADDED TO INITIAL MDR. CLAIM #: (B)(4). Manufacturer Narrative: (B)(4). | -17/2/89 | OD | Physician | 04/18/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8113371&pc=MTA | 2023826-2018-01753 | 2016-05-02 | Injury | STAAR SURGICAL COMPANY AG | 2018-11-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -5.00/+5.0/150 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOSS OF BEST-CORRECTED VISUAL ACUITY (BCVA) AND BLURRED VISION. THE SURGEON HAS NOT DECIDED AT THIS TIME IF ANOTHER ICL LENS WILL BE IMPLANTED. THE REPORTER INDICATED THERE WAS NO CATARACT, NO ICL ROTATION AND NO PROGRESSION OF KERATOCONUS. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(6) BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.00/+5.0/150 | OD | Physician | 10/29/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6426150&pc=MTA | 2023826-2017-00460 | 2017-02-28 | Injury | STAAR SURGICAL COMPANY | 2017-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+2.0/X009 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. AT ONE DAY POST-OP THERE WAS A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED - THE REPORTER INDICATED THE LENS WAS REPOSITIONED ON (B)(6) 2017. THE PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/20 AND THE ISSUE WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.5/+2.0/009 | OS | Physician | 02/22/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4635312&pc=MTA | 2023826-2015-00382 | 2014-11-10 | Injury | STAAR SURGICAL COMPANY | 2015-03-26 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THERE WAS AN INCIDENT OF REFRACTIVE SURPRISE AS WELL AS LENS DISLOCATION/SUBLUXATION, AND BLURRED VISION. THE LENS WAS REPOSITIONED AND REMAINS IMPLANTED. Manufacturer Narrative: THIS DEVICE IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. EXPLANT DATE: N/A. (B)(4). DEVICE EVALUATED BY THE MANUFACTURER? NO. LENS REMAINS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS IMPLANTED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT):UPON REVIEW OF THE DEVICE HISTORY RECORD, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON THE ABOVE INVESTIGATION, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -10.50/+4.5/178 | OS | Other | 03/09/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4600864&pc=MTA | 2023826-2015-00312 | 2015-02-10 | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THE LENS HAD A REFRACTIVE SURPRISE AND THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND BLURRED VISION. THE LENS REMAINS IMPLANTED. THE PATIENT WAS TREATED FOR CHOROIDAL EFFUSION (SUSPECT) WITH TOPIC CORTICOSTEROIDS, THE ANGLES WERE COMING BACK TO PREVIOUS AND THERE WAS GOOD EVOLUTION OF THE PATIENT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (REFRACTIVE SURPRISE); (CHOROIDAL EFFUSION-SUSPECT); BLURRED VISION. (LENS IMPLANTED). DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS IMPLANTED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -1.50/+6.0/060 | OS | Other | 02/12/2015 | 2015 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6089928&pc=MTA | 2023826-2016-01601 | 2016-03-11 | Injury | STAAR SURGICAL COMPANY AG | 2016-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -14.5/+6.0/086 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2009. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(6). (B)(4). ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS HAIR. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | -14.5/+6.0/086 | OS | Physician | 10/11/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5720709&pc=MTA | 2023826-2016-00373 |  | Injury | STAAR SURGICAL COMPANY | 2016-06-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/177 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD A REFRACTIVE SURPRISE AND WAS REPOSITIONED ON (B)(6) 2015. THE LENS REMAINS IMPLANTED AND IS PLANNED TO EXCHANGE WITH ANOTHER LENS. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4). | -8.5/+1.0/177 | OD | Physician | 03/24/2016 | 2016
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5509447&pc=MTA | 2023826-2016-00286 | 2015-11-25 | Injury | STAAR SURGICAL COMPANY | 2016-03-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+2.5/075 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE AND BLURRED VISION. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN. (B)(4). NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF BLURRED VISION AND THE LENS EXCHANGE. (B)(4). THERE WAS NO BREAK IN THE MATERIAL DUE TO MATERIAL INTEGRITY. (B)(4). Manufacturer Narrative: DEVICE EVALUATION NARRATIVE WAS INCORRECT. VISUAL INSPECTION DID NOT FIND THE OPTIC TORN. VISUAL INSPECTION FOUND THE HAPTIC TORN. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -15.0/+2.5/075 | OS | Physician | 02/16/2016 | 2016 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6918200&pc=MTA | 2023826-2017-01534 | 2016-04-19 | Injury | STAAR SURGICAL COMPANY | 2017-10-05 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 DIOPTER -15.0/+4.0/091 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE SURGEON NOTED LENS ROTATION, REFRACTIVE CHANGE OVER TIME, AND LENS REPOSITIONING. THE CLAIM REPORT STATES THERE IS A PLAN TO EXCHANGE THE LENS ("RE-IMPLANTATION"), AND THE SAME REPORT STATES THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ANY FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: LENS REMAINS IMPLANTED; PLAN TO EXCHANGE. LENS WAS EXCHANGED IN (B)(6) 2016 (EXACT DATE NOT KNOWN) FOR A LARGER, DIFFERENT DIOPTER LENS. NO ADDITIONAL INFORMATION IS AVAILABLE. (B)(4). Manufacturer Narrative: PREVIOUS: LENS WAS EXCHANGED IN (B)(6) 2016 (EXACT DATE NOT KNOWN) FOR A LARGER, DIFFERENT DIOPTER LENS. NO ADDITIONAL INFORMATION IS AVAILABLE. - UPDATED: THE LENS WAS EXCHANGED FOR A LARGER, DIFFERENT DIOPTER LENS ON (B)(6) 2016. THE LENS EXCHANGE RESOLVED THE PROBLEM. EXPLANTED DATE: IT IS CORRECTED FROM (B)(6) 2016 TO (B)(6) 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. PATIENT CODE (B)(4) (NO CODE AVAILABLE): SECONDARY SURGERY, LENS REPOSITIONING. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.0/+4.0/091 | OD | Physician | 09/07/2017 | 2017 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6609744&pc=MTA | 2023826-2017-00917 | 2016-04-24 | Injury | STAAR SURGICAL COMPANY | 2017-06-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+3.5/080 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE SURPRISE WERE NOTED. THE LENS WAS REPOSITIONED ON (B)(6) 2016, BUT THIS DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXCHANGED ON (B)(6) 2016 FOR ANOTHER LENS OF A DIFFERENT MODEL/DIOPTER AND THE PROBLEM WAS RESOLVED. PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 AND UNCORRECTED VISUAL ACUITY WAS 20/100 ON (B)(6) 2016. Manufacturer Narrative: THIS DEVICE IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND THE HAPTIC TORN/BROKEN. (B)(4). CONCLUSION CODE AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS UNDER 21 YEARS OF AGE AT THE TIME OF IMPLANTATION. (B)(4). | -17.0/+3.5/080 | OS | Physician | 06/20/2016 | 2016 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6618987&pc=MTA | 2023826-2017-00910 | 2016-08-03 | Injury | STAAR SURGICAL COMPANY | 2017-06-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/169 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, BLURRED VISION. THE LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2016, BUT UNSUCCESSFUL. ON (B)(6) 2016 THE LENS WAS EXCHANGED. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BENT AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+2.5/169 | OD | Physician | 05/09/2017 | 2017 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7144234&pc=MTA | 2023826-2017-01947 | 2017-06-01 | Injury | STAAR SURGICAL COMPANY | 2017-12-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/+2.0/117 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS REPORTED AS HAVING ROTATED (NOT ASSOCIATED TO A LOW VAULT) AND THERE WAS A REFRACTIVE SURPRISE. THE PATIENT EXPERIENCED POOR VISUAL ACUITY. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS SUCCESSFULLY REPOSITIONED AND REMAINS IMPLANTED. AT FIRST DAY POST-OP, THE PATIENT HAD ENLARGED PUPILS BUT IS MUCH HAPPIER. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.50/+2.0/117 | OS | Physician | 12/04/2017 | 2017 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6970600&pc=MTA | 2023826-2017-01597 | 2017-07-16 | Injury | STAAR SURGICAL COMPANY | 2017-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1, -13/+4.0/154 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, IT WAS NOTICED THAT THE LENS ROTATED. ON (B)(6) 2017, THE SURGEON REDIALED THE LEFT EYE. THE PATIENT WAS HAPPY FOR A WEEK, THEN THE LENS ROTATED AGAIN ON (B)(6) 2017. AS OF (B)(6) 2017, THE VAULT IS 410 MICRONS, BCVA IS 20/80 AND UCVA IS LESS THAN 20/200. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THE SURGEON STATED THAT PATIENT DIDN`T WANT EXCHANGE FOR A NON-TORIC ICL. (B)(4). Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2017. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13/+4.0/154 | OS | Physician | 09/26/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7089423&pc=MTA | 2023826-2017-01829 | 2017-11-07 | Injury | STAAR SURGICAL COMPANY | 2017-12-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -5.0/+5/086 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND REFRACTIVE SURPRISE WERE NOTICED. THE LENS WAS THEN REPOSITIONED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.0/+5/086 | OD | Physician | 11/07/2017 | 2017 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7180692&pc=MTA | 2023826-2018-00039 | 2017-11-30 | Injury | STAAR SURGICAL COMPANY | 2018-01-10 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.0/+3.5/062 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE SURPRISE, LENS ROTATION, BLURRED VISION, AND "RESIDUAL POWER." Manufacturer Narrative: THE LENS WAS REPOSITIONED AND REMAINS IMPLANTED. "PATIENT IS COMFORTABLE WITH 6/6 VISION." (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.0/+3.5/062 | OS | Physician | 12/14/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7180665&pc=MTA | 2023826-2018-00035 | 2017-12-03 | Injury | STAAR SURGICAL COMPANY | 2018-01-10 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.0/+6.0/099 (SPHERE/ CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS REPOSITIONED. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT IS REPORTED. LENS REPOSITIONING DID NOT RESOLVE THE PROBLEM. SURGEON STATES THAT THE LENS ROTATED POST-OPERATIVELY. LENS REMAINS IMPLANTED. Manufacturer Narrative: THE SURGEON ALSO NOTICED REFRACTIVE SURPRISE IN THE PATIENT'S EYE. (B)(4) Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -4.0/+6.0/099 | OS | Physician | 12/14/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7136378&pc=MTA | 2023826-2017-01938 | 2016-09-20 | Injury | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -4.5/+6.0/006 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING ROTATED AND WAS REPOSITIONED ON (B)(6) 2017. AT A POST-OP VISIT ON (B)(6) 2017, THE LENS HAD ROTATED AND WAS REPOSITIONED AGAIN. THE LENS WAS NOT STABLE AND TO DATE, REMAINS IMPLANTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -4.5/6.0/006 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING ROTATED AND WAS REPOSITIONED ON (B)(6) 2017. AT A POST-OP VISIT ON (B)(6) 2017, IT WAS REPORTED THAT THE LENS WAS NOT STABLE AND "ROTATED AGAIN BUT LESS". THE LENS REMAINS IMPLANTED IN THE PATIENT'S EYE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -4.5/+6.0/006 | OS | Physician | 11/28/2017 | 2017 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6701661&pc=MTA | 2023826-2017-01156 | 2012-01-11 | Injury | STAAR SURGICAL COMPANY AG | 2017-07-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM13.2, DIOPTER -12.0/+1.0/91 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, AND REFRACTIVE CHANGE OVERTIME. THE CUSTOMER STATED CURRENTLY THE LENS REMAINS IMPLANTED, AND THE DR. PLANS TO EXCHANGE THE LENS WITH HIGHER REFRACTIVE POWER TO RESOLVED THE PROBLEM FOR THE PATIENT. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS -12.0/+1.0/91 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED FOR SAME MODEL DIFFERENT DIOPTER LENS. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/25. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: SHOULD HAVE BEEN RECORDED. FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: FUNCTIONAL INSPECTION WAS PERFORMED BUT DID NOT FIND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.0/+1.0/91 | OS | Physician | 06/14/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6307761&pc=MTA | 2023826-2017-00166 | 2012-06-13 | Injury | STAAR SURGICAL COMPANY AG | 2017-02-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50/1.5/093 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE PATIENT HAS EXPERIENCED REFRACTIVE CHANGE OVERTIME. AS OF THE DATE OF THIS MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED FOR ANOTHER LENS ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED OR MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). LENS NOT RETURNED. | -15.50/1.5/093 | OD | Physician | 01/10/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6428720&pc=MTA | 2023826-2017-00464 | 2015-11-11 | Injury | STAAR SURGICAL COMPANY AG | 2017-03-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS, -10.5/+3.0/105 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO REFRACTIVE SURPRISE AND LENS DISLOCATION/SUBLUXATION. THE SURGEON ATTEMPTED TO REPOSITION THE LENS BUT THE RESULT WAS NOT STABLE. THE LENS WAS THEN EXCHANGED FOR A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED. THE SURGEON IS PLANNING TO PERFORM CORNEAL REFRACTIVE SURGERY TO REDUCE THE ERROR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED WAS RETURNED IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC TORN, LENS HAPTIC BROKEN, AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (DRIED, DISCOLORED MATERIAL). (B)(4). | -10.5/+3.0/105 | OS | Physician | 07/12/2016 | 2016 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6442677&pc=MTA | 2023826-2017-00516 | 2016-05-07 | Injury | STAAR SURGICAL COMPANY | 2017-03-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/79 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2017 THE LENS WAS EXCHANGED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/79 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR ANOTHER LENS OF A DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY ON (B)(6) 2017 WAS 20/25. DEVICE EVALUATION: THE DEVICE WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/+2.0/79 | OD | Physician | 02/28/2017 | 2017 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6393825&pc=MTA | 2023826-2017-00364 | 2016-07-18 | Injury | STAAR SURGICAL COMPANY | 2017-03-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.00/+0.5/86 DIOPTER IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR THE SAME LENS MODEL WITH A DIOPTER -7/ +0.5/84 DIOPTER AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION-LENS WAS RETURNED IN A SMALL VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON IT. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: CORRECTED DATA: CONCLUSION CODE NOT APPLICABLE, PATIENT'S ACD IS ABOVE 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). CONCLUSION: REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. PATIENT'S ANTERIOR ENDOTHELIUM CHAMBER DEPTH IS BELOW 3.0MM FOR VICL/VTICL. (B)(4). | -8.00/+0.5/86 | OD | Physician | 02/09/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6429971&pc=MTA | 2023826-2017-00470 | 2016-08-08 | Injury | STAAR SURGICAL COMPANY AG | 2017-03-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -3.0/+1.5/31 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, AND ENLARGEMENT OF INCISION SURGERY. ON (B)(6) 2016, THE LENS WAS EXPLANTED. THE LENS WAS THEN EXCHANGED WITH THE SAME SIZE, AND SIMILAR DIOPTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED.Event Description: THERE WAS NO ENLARGEMENT OF THE INCISION ON (B)(6) 2016. Manufacturer Narrative: (B)(4). Manufacturer Narrative: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT AND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THERE WAS NO ENLARGEMENT OF INCISION ON 08/08/2016 AS PREVIOUSLY REPORTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION: OFF LABEL USE (MORE THAN 45 YEARS OF AGE AT THE TIME OF LENS IMPLANT). (B)(4). | -3.0/+1.5/31 | OD | Physician | 02/24/2017 | 2017 | 57
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6381058&pc=MTA | 2023826-2017-00320 | 2016-11-07 | Injury | STAAR SURGICAL COMPANY | 2017-03-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, DIOPTER -11.50/+3.00/0.87 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ALSO ON (B)(6) 2016, THE SURGEON RE-POSITIONED THE LENS. ON (B)(6) 2016 THE PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (PREVIOUS): REPOSITIONING WAS PERFORMED ON (B)(6) 2016. (CORRECTED) REPOSITIONING WAS PERFORMED ON (B)(6) 2016. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED. | -11.50/+3.00/0.87 | OS | Physician | 02/06/2017 | 2017 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6414543&pc=MTA | 2023826-2017-00433 | 2017-01-11 | Injury | STAAR SURGICAL COMPANY | 2017-03-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/3.0/087 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE DOCTOR ALSO STATED THE PATIENT EXPERIENCED RESIDUAL MIXED ASTIGMATISM, AND BELIVES THAT THE LENS MAY BE OVERROTATED. AS OF THE DAY OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE (B)(4) NO CODE AVAILABLE (RESIDUAL MIXED ASTIGMATISM). RESULT CODE: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.00/3.0/087 | OS | Physician | 02/20/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6395339&pc=MTA | 2023826-2017-00362 | 2017-02-06 | Injury | STAAR SURGICAL COMPANY | 2017-03-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.50/+2.0/76 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SAME SIZE AND SIMILAR DIOPTER DUE TO THE PATIENT EXPERIENCING REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND THICK, STICKY, CLEAR RESIDUE ON PRODUCT. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.50/+2.0/76 | OD | Physician | 02/09/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6487511&pc=MTA | 2023826-2017-00629 | 2017-02-28 | Injury | STAAR SURGICAL COMPANY | 2017-04-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, DIOPTER -11.50/+0.5/X013 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR SAME SIZE, SIMILAR DIOPTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION THE LENS HAS NOT YET BEEN RETURNED. Manufacturer Narrative: PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS.(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | -11.50/+0.5/013 | OS | Physician | 03/14/2017 | 2017 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6473540&pc=MTA | 2023826-2017-00593 | 2017-03-09 | Injury | STAAR SURGICAL COMPANY | 2017-04-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, DIOPTER -11.50/+3.0/X085 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -11.50/+3.0/085 | OS | Physician | 03/10/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6667615&pc=MTA | 2023826-2017-01034 | 2017-04-12 | Injury | STAAR SURGICAL COMPANY | 2017-06-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICM5_12.1, DIOPTER -2.0/+3.5/005 IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT THEN BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SAME SIZE, DIFFERENT DIOPTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. LENS HAS BEEN RETURNED AND EVALUATED. LENS WAS RETURNED IN LIQUID, IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. UPDATED BASED ON PRODUCT EVALUATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | -2.0/+3.5/005 | OS | Physician | 05/30/2017 | 2017 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6701659&pc=MTA | 2023826-2017-01166 | 2017-05-09 | Injury | STAAR SURGICAL COMPANY | 2017-07-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.5/83 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. THIS CASE IS UNDER MEDICAL CONSULT. Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.0/+2.5/83 | OS | Physician | 06/19/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6657557&pc=MTA | 2023826-2017-01018 | 2017-05-24 | Injury | STAAR SURGICAL COMPANY | 2017-06-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/2.5/20 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SAME SIZE AND DIFFERENT DIOPTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR THE UNITS WITHIN THE SAME LOT. (B)(4). | -12.5/2.5/20 | OD | Physician | 05/25/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6674235&pc=MTA | 2023826-2017-01075 | 2017-05-24 | Injury | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.0/+4.0/080 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT ALSO EXPERIENCES HYPEROPIA. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -7.0/+4.0/080 | OD | Physician | 06/05/2017 | 2017 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6725634&pc=MTA | 2023826-2017-01205 | 2017-06-23 | Injury | STAAR SURGICAL COMPANY | 2017-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, DIOPTER -13.00/+3.0/X096 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS REPOSITIONED ON (B)(6) 2017 DUE TO REFRACTIVE SURPRISE. SURGEON PLANS TO PERFORM AN ADDITIONAL LENS ROTATION. Manufacturer Narrative: STATES THAT AN ADDITIONAL LENS ROTATION IS PLANNED. UPDATED/CORRECTED: REPORTEDLY, THE LENS WAS REPOSITIONED ON (B)(6) 2017. AS OF (B)(6) 2017 THE "PATIENT IS HAPPY AFTER REPOSITIONING." THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.00/+3.0/096 | OD | Physician | 06/22/2017 | 2017 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6725726&pc=MTA | 2023826-2017-01206 | 2017-06-23 | Injury | STAAR SURGICAL COMPANY | 2017-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, DIOPTER -13.50/+3.50/X192 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS REPOSITIONED ON (B)(6) 2017 DUE TO REFRACTIVE SURPRISE. SURGEON WANTS TO KNOW THE AMOUNT OF REQUIRED ROTATION FOR PLANO REFRACTION. Manufacturer Narrative: THE LENS WAS REPOSITIONED ON (B)(6) 2017. THE LENS ROTATED BACK AGAIN AFTER REPOSITIONING, SO THE SURGEON PLANS TO MAKE "BIOPTIC LVC" TO "CORRECT RX." (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.50/+3.50/192 | OS | Physician | 06/26/2017 | 2017 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6711486&pc=MTA | 2023826-2017-01182 | 2017-06-27 | Injury | STAAR SURGICAL COMPANY | 2017-07-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, DIOPTER -0.5/+6/109 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, MISALIGNMENT. AS OF THE DATE OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. ATTEMPTS TO CONTACT TO THE CUSTOMER FOR ADDITIONAL INFORMATION HAVE BEEN ON GOING. Manufacturer Narrative: THE IMPLANTED VTICM5_12.6 LENS WAS REPOSITIONED ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THE SURGEON DECIDED NOT TO DO A SECOND REPOSITIONING, THE PATIENT IS HAPPY WITH THE SLIDE CYLINDER AT THE MOMENT. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/22. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN US BUT NOT MARKETED IN THE U.S. (B)(4). | -0.5/+6/109 | OD | Physician | 06/27/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6970595&pc=MTA | 2023826-2017-01603 | 2017-07-11 | Injury | STAAR SURGICAL COMPANY | 2017-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -1.5/1.5/175 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THE LENS WAS THEN EXCHANGED FOR A SIMILAR LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL AND THERE WAS PRESENCE OF CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TO BE BROKEN. METHOD CODE: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -1.5/1.5/175 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. A SIMILAR LENS WAS THEN IMPLANTED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. THE DATE OF EXPLANT IS CORRECTED FROM (B)(6) 2017 TO (B)(6) 2017. THE DEVICE MANUFACTURED DATE IS CORRECTED FROM 25-JUN-2015 TO 26-JUN-2015. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -1.5/1.5/175 | OD | Physician | 07/27/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050900&pc=MTA | 2023826-2017-01707 | 2017-07-27 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -5.50/+5.0/070 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SIMILAR LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL AND HAD CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.50/+5.0/070 | OS | Physician | 10/24/2017 | 2017 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7335039&pc=MTA | 2023826-2018-00340 | 2017-07-28 | Injury | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -9.50/+4.0/85 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2017. ON (B)(6) 2018, THE LENS WAS REMOVED DUE PATIENT'S DISSATISFACTION WITH HIS VISUAL ACUITY WHICH WAS 20/30 AFTER THE SURGERY. HIS PRE-OP VISUAL ACUITY IS 20/20. THE LENS WAS EXCHANGED FOR LONGER DIFFERENT LENS. THE PATIENT CURRENT VISUAL ACUITY IS 20/25. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID INSIDE A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND HAD PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. SECONDARY SURGERY, LENS EXCHANGE. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.50/+4.0/85 | OD | Physician | 02/13/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6952444&pc=MTA | 2023826-2017-01564 | 2017-08-03 | Injury | STAAR SURGICAL COMPANY | 2017-10-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -8.5/+1.0/087 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SIMILAR LENS. THE PATIENT'S UCVA IS 20/30 AS OF (B)(6) 2017 AND MANIFEST REFRACTION IS +0.50/-1.25/30 AND 20/20 AS OF (B)(6) 2017. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL AND HAD CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/+1.0/087 | OD | Physician | 09/19/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6855452&pc=MTA | 2023826-2017-01420 | 2017-08-14 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -8.50/4.0/91 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. ADDITIONAL REFRACTIVE TREATMENT (ENCHANCEMENT) WAS DONE. THE SURGEON DECIDED TO WAIT A FEW WEEKS AFTER THE SURGERY TO ALLOW THE EYE AND CORNEA TO HEAL AND GET STABLE. THE SURGEON DECIDED THAT NO EXCHANGE IS NEEDED, BUT WILL PERFORM A LASER TOUCH UP IF NECESSARY. Manufacturer Narrative: ADDITIONAL REFRACTIVE TREATMENT (ENHANCEMENT) WAS NOT PERFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/4.0/91 | OD | Physician | 08/14/2017 | 2017 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6912388&pc=MTA | 2023826-2017-01496 | 2017-08-30 | Injury | STAAR SURGICAL COMPANY | 2017-10-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -7.0/+1.5/111 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: LENS EXCHANGE WAS PERFORMED ON (B)(6) 2017. PATIENT VISUAL ACUITY WAS NOT CORRECTED COMPLETELY AFTER EXCHANGE. THE DOCTOR HAS DECIDED TO PERFORM LASIK FOR ADDITIONAL CORRECTION. PATIENT POST-OP, PRE-LASIK, MR IS SPH. -0.1 AND CYL. -1.75. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.0/+1.5/111 | OD | Physician | 09/06/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6941354&pc=MTA | 2023826-2017-01559 | 2017-09-15 | Injury | STAAR SURGICAL COMPANY | 2017-10-11 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.5/+2.0/079 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. A REFRACTIVE SURPRISE WAS NOTED BY THE SURGEON, INCIDENT DATE REPORTED AS (B)(6) 2017. THE SURGEON IS CONSIDERING LENS EXCHANGE OR ROTATION. Manufacturer Narrative: UPDATED INFORMATION: LENS REPOSITIONING PERFORMED. SURGEON STATES, "VISUS IS NOW 1.25." THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). [ANTERIOR ENDOTHELIUM CHAMBER DEPTH LESS THAN 3.0MM. (B)(4). | -6.5/+2.0/079 | OD | Physician | 09/15/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135226&pc=MTA | 2023826-2017-01920 | 2017-10-19 | Injury | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/180 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED. NO ADDITIONAL INFORMATION IS AVAILABLE AS OF THE TIME OF MDR SUBMISSION. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT CURRENTLY IS NOT PLANNED TO HAVE ANY ADDITIONAL SURGERY. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT, LENS REPOSITIONING). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.0/+2.0/180 | OS | Physician | 11/24/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050082&pc=MTA | 2023826-2017-01757 | 2017-10-24 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 DIOPTER -0.50/+5.5/089 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE, BLURRED VISION, AND LENS RE-ALIGNMENT. REPORTEDLY, THE LENS WAS RE-ALIGNED ON (B)(6) 2017. THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ANY FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: DATE OF THIS REPORT: IT IS CORRECTED FROM "26-OCT-2017" TO " 25-OCT-2017". THE DIOPTER VALUE OF THE INTIAL LENS IMPLANTED IS CORRECTED FROM "-0.50/+5.5/089 SPHERE/ CYLINDER/ AXIS" TO " -0.5/+5.5/084 SPHERE/ CYLINDER/ AXIS". DATE RECEIVED BY MANUFACTURER: IT IS CORRECTED FROM "26-OCT-2017" TO " 25-OCT-2017". (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -0.50/+5.5/089 | OS | Physician | 10/25/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050099&pc=MTA | 2023826-2017-01756 | 2017-10-24 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 DIOPTER -1.50/+4.5/089 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE, BLURRED VISION, AND LENS RE-ALIGNMENT. REPORTEDLY, THE LENS WAS RE-ALIGNED ON (B)(6) 2017. THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ANY FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: DATE OF THIS REPORT: IT IS CORRECTED FROM " 26-OCT-2017" TO "25-OCT-2017" IN THE INITIAL MDR. DATE RECEIVED BY MANUFACTURER: IT IS CORRECTED FROM "26-OCT-2017" TO " 25-OCT-2017" IN THE INITIAL MDR. (B)(4) Manufacturer Narrative: THE DIOPTER VALUE OF THE INITIAL LENS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) IS CORRECTED FROM "-1.50/ +4.5/ 089 SPHERE/ CYLINDER/ AXIS" TO "-2.0/ +4.5/ 072 SPHERE/ CYLINDER/ AXIS". (B)(4) Manufacturer Narrative: (B)(4). THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -1.50/+4.5/089 | OD | Physician | 10/25/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7156010&pc=MTA | 2023826-2017-01970 | 2017-12-05 | Injury | STAAR SURGICAL COMPANY | 2017-12-29 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -3.0/+4.5/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION IS AVAILABLE AS OF THE TIME OF MDR SUBMISSION. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS). DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -3.0/+4.5/086 | OS | Physician | 12/06/2017 | 2017 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7484262&pc=MTA | 2023826-2018-00640 | 2018-02-21 | Injury | STAAR SURGICAL COMPANY | 2018-05-03 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -6.5/+6.0/152 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), DECREASED VISUAL ACUITY, BLURRED VISION, ENLARGEMENT OF PI OR ADDITION OF NEW PI PERFORMED VIA YAG, AND ADMINISTRATION OF TOPICAL 2% PILOCARPINE. THE STATUS OF THE EYE IS LISTED AS SUPERIOR IRIDECTOMY, NARROW ANTERIOR CHAMBER. ADDITIONAL INFORMATION-FIRST DAY POST-OP THE VISUAL ACUITY WITHOUT CORRECTION WAS "ALMOST 20/20," BUT "A FEW DAYS AFTER" IT WAS 20/50. POST-YAG LASER AND TOPICAL PILOCARPINE THE VA WAS 20/30 BUT "TODAY" IS 20/80 WITHOUT CORRECTION. THE LAST VISIT DATE IS LISTED AS (B)(6) 2018. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE SURGEON ALSO NOTICED LOSS OF BCVA. "PELLUCID DEGENERATION, PREVIOUS CORNEAL SURGERY (NON- REFRACTIVE)" IS CORRECTED TO "PELLUCID DEGENERATION". (B)(4) Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED STATEMENT: ON (B)(4) 2019 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL, SHORTER LENGTH LENS. THE PROBLEM IS RESOLVED. (SECONDARY SURGERY, LENS EXCHANGE). CLAIM # (B)(4). Manufacturer Narrative: UPDATED/ADDITIONAL INFORMATION: ON (B)(6) 2019 THE SURGEON REPORTS THE PATIENT'S BCVA IS 20/30 IN HIS RIGHT (OD) EYE. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SIGNIFICANT REDUCTION OF ICA, ENLARGEMENT OR ADDITION OF PI-YAG, ADMINISTRATION OF TOPICAL 2% PILOCARPINE, SUPERIOR IRIDECTOMY, NARROW ANTERIOR CHAMBER). CONCLUSION CODE: OFF-LABEL USE [PREVIOUS HISTORY: PELLUCID DEGENERATION, CORNEAL SURGERY (NON-REFRACTIVE)] (B)(4). | -6.5/+6.0/152 | OD | Physician | 04/04/2018 | 2018 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7403701&pc=MTA | 2023826-2018-00514 | 2018-03-12 | Injury | STAAR SURGICAL COMPANY | 2018-04-06 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/050 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED ON (B)(6) 2018 WHICH DID NOT RESOLVE THE PROBLEM. REPORTEDLY THE LENS REMAINS IMPLANTED AND "VISION IS STILL BLURRY." Manufacturer Narrative: UPDATED/ADDITIONAL INFORMATION: THE LENS WAS REPOSITIONED AGAIN ON (B)(6) 2018. AFTER REPOSITIONING, THE PATIENT IS "VERY HAPPY WITH THE OUTCOME" AND "HE WAS READING WITH NO EFFORT AT ALL." THE PROBLEM IS RESOLVED. (B)(4) Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. PATIENT CODE : NO CODE AVAILABLE (SECONDARY SURGERY, LENS REPOSITIONING) LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.0/+1.5/050 | OS | Physician | 03/12/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594283&pc=MTA | 2023826-2018-00894 | 2018-05-21 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/+5.0/036 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED DUE TO REFRACTIVE SURPRISE. REPORTEDLY, THE SURGEON PERFORMED THE "WRONG CALCULATION." THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CONCLUSION CODE : USER ERROR (SURGEON PERFORMED WRONG CALCULATIONS) (B)(4). | -10.0/+5.0/036 | OD | Physician | 05/21/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6314688&pc=MTA | 2023826-2017-00158 |  | Injury | STAAR SURGICAL COMPANY | 2017-02-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+2.0/102 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AFTER THE SURGERY WITH AN UNCORRECTED VISUAL ACUITY (UCVA) OF 6/36 AND A VAULT VALUE OF 250 MICROMETERS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: THE LENS WAS EXPLANTED AND EXCHANGED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. SECONDARY SURGICAL INTERVENTION, EXPLANT, LENS EXCHANGE. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN A LENS CASE/ VIAL AND HAD CLEAR SURGICAL RESIDUE ON LENS SURFACE. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -15.0/+2.0/102 | OS | Physician | 01/16/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6673303&pc=MTA | 2023826-2017-01052 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -17.5/+2.0/X174 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ASTIGMATISM WAS NOTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUS REPORT STATEMENT "SINGLE-USE DEVICE WAS REPROCESSED AND REUSED ON A PATIENT" IS INCORRECT. (B)(4). Manufacturer Narrative: CORRECTED DATA: TYPE OF REPORTABLE EVENT: "MALFUNCTION" INSTEAD OF "SERIOUS INJURY". (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17.5/+2.0/174 | OD | Physician | 09/04/2016 | 2016
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6674075&pc=MTA | 2023826-2017-01054 |  | Injury | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -7.0/+6.0/96 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. BLURRY VISION WAS REPORTED. THE LENS WAS THEN ROTATED 10 DEGREES ANTI CLOACKWISE ON (B)(6) 2017 WHICH IMPROVED UNAIDED VISION FROM 0.7 TO 0.9. Manufacturer Narrative: CORRECTED DATA: PREVIOUS REPORT STATEMENT "SINGLE-USE DEVICE WAS REPROCESSED AND REUSED ON A PATIENT" IS INCORRECT. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). LENS NOT RETURNED. | -7.0/+6.0/96 | OD | Physician | 06/03/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6795668&pc=MTA | 2023826-2017-01318 |  | Injury | STAAR SURGICAL COMPANY | 2017-08-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6)2016 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS, AND THE PROBLEM WAS RESOLVED. THE DR. STATED THAT WHEN HE ORDERING THE PRIMARY LENS, THE RIGHT EYE ASTIGMATISM READINGS HAVE BEEN MIXED WITH LEFT EYE CYLINDER. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS PRODUCT TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PMA 510(K):_THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: THE SURGEON STATED THAT WHEN ORDERING PRIMARY ICL, THE RIGHT EYE ASTIGMATISM READINGS HAVE BEEN MIXED WITH LEFT EYE CYLINDER. (B)(4). | -13.5/+2.0/080 | OD | Physician | 07/18/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7201657&pc=MTA | 2023826-2018-00093 |  | Injury | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.5/+4.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: UPDATED INFORMATION: THE SURGEON HAS DECIDED TO TREAT THE PATIENT USING PRK-REFRACTIVE LASER TREATMENT. THE LENS WILL REMAIN IMPLANTED. (B)(4) Manufacturer Narrative: PMA 510(K): -THIS PRODUCT IS NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | -9.5/+4.0/094 | OD | Physician | 12/22/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7201662&pc=MTA | 2023826-2018-00095 |  | Injury | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.5/+4.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: UPDATED INFORMATION: REPORTEDLY, THE SURGEON HAS DECIDED TO TREAT THE PATIENT USING PRK-REFRACTIVE LASER TREATMENT. THE LENS WILL REMAIN IMPLANTED. (B)(4) Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | -7.5/+4.5/094 | OS | Physician | 12/22/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7419742&pc=MTA | 2023826-2018-00526 |  | Injury | STAAR SURGICAL COMPANY | 2018-04-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -3.50/+6.0/172 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -3.50/+6.0/172 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. (B)(4). AGE IS (B)(6) YEARS, NOT (B)(6) YEARS AS PREVIOUSLY REPORTED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE IN SPECIFICATION. CONCLUSION: - THIS LENS MODEL IS CONTRAINDICATED IN THE PRESENCE OF KERATOCONUS AND PREVIOUS CORNEAL SURGERY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -3.50/+6.0/172 | OD | Physician | 03/12/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7557296&pc=MTA | 2023826-2018-00803 |  | Injury | STAAR SURGICAL COMPANY | 2018-05-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED PATIENT HAD A 12.5MM, TICM125V4, -16.0/+2.5/019 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IMPLANTED INTO THE LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT RECENTLY CAME BACK WITH REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADD'L INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: THE IMPLANTED DATE IS CORRECTED FROM "(B)(6) 2017" TO "(B)(6) 2007". ADDITIONAL DATA: THE REPORTER INDICATED THAT THE PATIENT HAD A 12.5MM TICM145V4, -16.0/+2.5/019 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS IMPLANTED INTO THE LEFT EYE (OS) ON (B)(6) 2007. THE PATIENT RECENTLY HAD REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED ON (B)(6) 2018. THE PROBLEM WAS RESOLVED. SECONDARY SURGICAL INTERVENTION, EXPLANT AND NEW LENS IMPLANTED LATER. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FIBERS ON LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.0/+2.5/019 | OS | Physician | 05/02/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4903447&pc=MTA | 2023826-2015-00807 | 2015-03-12 | Injury | STAAR SURGICAL COMPANY | 2015-07-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011 AND REFRACTIVE CHANGE WAS NOTED ON (B)(6) 2015, DUE TO MYOPIA PROGRESSION. THE LENS REMAINS IMPLANTED AND THE SURGEON PLANS TO EXCHANGE FOR A DIFFERENT POWER. THE PATIENT'S LAST BCVA WAS 20/25. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4) REFRACTIVE SURPRISE. CONCLUSIONS: BASED ON THE COMPLAINT INFORMATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). LENS NOT YET EXCHANGED. | -15.5/+2.0/176 | OS | Physician | 06/09/2015 | 2015
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6117561&pc=MTA | 2023826-2016-01637 | 2016-10-26 | Injury | STAAR SURGICAL COMPANY | 2016-11-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/046 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGE FOR THE SAME MODEL, SIZE BUT DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT LENS DIOPTER MEASUREMENTS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION WAS ALSO PERFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED FOR THE SAME MODEL, SIZE, DIFFERENT DIOPTER LENS. (B)(4): BLURRED VISION/REFRACTIVE SURPRISE. (B)(4). LENS NOT RETURNED. | -18.0/+1.5/046 | OD | Physician | 10/26/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6682110&pc=MTA | 2023826-2017-01090 | 2017-05-30 | Injury | STAAR SURGICAL COMPANY | 2017-06-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, DIOPTER -1.5/+3.0/090 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. APPROXIMATELY TWO MONTHS AFTER RECOVERY, PATIENT WAS HIT IN THE EYE BY A CHILD. THE PATIENT'S VISION WORSENED AND PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. PHYSICIAN REPOSITIONED THE LENS TWICE WHICH DID NOT RESOLVE THE ISSUE. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED. SURGEON PLANS TO DO ANOTHER REPOSITION AFTER CONSULTING WITH THE MEDICAL AFFAIRS TEM Manufacturer Narrative: ON (B)(6) 2017 THE LENS WAS REPOSITIONED ACCORDING TO RE-CALCULATION RESULTS. THE PATIENT IS HAPPY, THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4) CONCLUSION CODE:OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.82MM)]. (B)(4). DEVICE REMAINS IMPLANTED. | -1.5/+3.0/090 | OD | Physician | 05/31/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5250502&pc=MTA | 2023826-2015-01535 | 2015-01-01 | Injury | STAAR SURGICAL COMPANY | 2015-11-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/1.0/X075 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. REFRACTIVE SURPRISE WAS NOTED IN (B)(6) 2015. THE ICL WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SIMILAR LENS MODEL.THIS RESOLVED THE PROBLEM. POST-OP, UCVA WAS 20/20 SEE MFR. #2023826-2015-01534 FOR RIGHT EYE. Manufacturer Narrative: (B)(4). CORRECTED DATA: BEST ESTIMATE IS (B)(6) 2015 ((B)(6) 2015). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -11.00/1.0/X075 (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -11.00/1.0/075 | OS | Physician | 10/28/2015 | 2015 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5325949&pc=MTA | 2023826-2015-01712 | 2015-05-19 | Injury | STAAR SURGICAL COMPANY | 2015-12-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/063 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) . DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. UNK EVENT PROBLEM AND EVALUATION CODES: PATIENT: NEW INCISION, SECONDARY SURGERY, REFRACTIVE SURPRISE. DEVICE: EXPLANTED. METHOD: WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -13.5/+1.5/063 | OD | Physician | 11/16/2015 | 2015 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5654713&pc=MTA | 2023826-2016-00656 | 2015-11-13 | Injury | STAAR SURGICAL COMPANY | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -2.0/+1.5/122 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS WAS RETURNED IN LIQUID IN THE LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS IS WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE DIOPTER MEASUREMENT IS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -2.0/+1.5/122 | OD | Physician | 05/05/2016 | 2016 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5596424&pc=MTA | 2023826-2016-00489 | 2015-12-30 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -14.5/+5.0/076 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2009. THE LENS HAD A REFRACTIVE CHANGE OVERTIME AND THE PLAN IS TO EXPLANT AND EXCHANGE THE LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PATIENT HAD LENS EXCHANGE, AND PROBLEM WAS RESOLVED. PATIENT IS SLIGHTLY MONOVISION. PREVIOUS REPORT SOURCE, USER FACILITY, WAS INCORRECT. CORRECT INFORMATION IS DISTRIBUTOR. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4). | -14.5/+5.0/076 | OD | Physician | 01/15/2016 | 2016 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5596425&pc=MTA | 2023826-2016-00490 | 2015-12-30 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -15.5/+5.5/095 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2009. THE LENS HAD A REFRACTIVE CHANGE OVERTIME AND THE PLAN IS TO EXPLANT AND EXCHANGE THE LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR ANOTHER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: (B)(4). DEVICE DISCARDED BY USER. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). | -15.5/+5.5/095 | OS | Physician | 01/15/2016 | 2016 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5917266&pc=MTA | 2023826-2016-01182 | 2015-12-31 | Injury | STAAR SURGICAL COMPANY | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -15.0/+4.5/118 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/14. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DATE RECEIVED BY MANUFACTURER: DATE ON FOLLOW-UP #1 IS INCORRECT. CORRECT DATE WAS 08/29/2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED. | -15.0/+4.5/118 | OS | Physician | 08/02/2016 | 2016 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5324056&pc=MTA | 2023826-2015-01717 | 2016-01-20 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.0/+3.5/026 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD A REFRACTIVE SURPRISE AND THE PATIENT'S VISION WAS BLURRED, WITH SHADOW. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED WITH ANOTHER LENS (ICL), SAME MODEL, LENGTH AND -7.5 DIOPTER, THIS RESOLVED THE PROBLEM. THE PATIENT WAS DOING WELL WITH BEST CORRECTED VISUAL ACUITY (BCVA) 6/6. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL (HAIR,FIBERS) ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -5.0/+3.5/026 | OS | Physician | 12/01/2015 | 2015 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6089924&pc=MTA | 2023826-2016-01600 | 2016-03-11 | Injury | STAAR SURGICAL COMPANY AG | 2016-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -15.5/+6.0/092 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2009. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PATIENT'S LAST VISIT WAS ON (B)(6) 2016 WITH UNCORRECTED VISUAL ACUITY (UCVA) 20/16. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4). | -15.5/+6.0/092 | OD | Physician | 10/11/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5730070&pc=MTA | 2023826-2016-00818 | 2016-04-05 | Injury | STAAR SURGICAL COMPANY | 2016-06-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+4.5/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULT: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: BLURRED VISION WAS REPORTED. LENS REPOSITIONING WAS PERFORMED AND THE PROBLEM WAS RESOLVED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS REMAINS IMPLANTED. (B)(4). | -9.5/+4.5/085 | OS | Physician | 05/12/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6074344&pc=MTA | 2023826-2016-00908 | 2016-06-03 | Injury | STAAR SURGICAL COMPANY | 2016-11-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -03.50/+5.0/113 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. Manufacturer Narrative: ON (B)(6) 2016, THE REPORTER STATED THAT THE PATIENT EXPERIENCED REFRACTIVE SURPRISE ASSOCIATED WITH LENS MISALIGNMENT. FIVE MONTHS AFTER SURGERY, THE SURGEON DECIDED TO PERFORM A SECONDARY INTERVENTION TO REPOSITION THE LENS. TICL WAS STILL FOUND TO BE OFF-AXIS AFTER REPOSITIONING AND THE SURGEON DECIDED TO REMOVE THE LENS ON (B)(6) 2016 AND EXCHANGE THE LENS FOR LONGER ONE WITH THE SAME POWER. THIS SOLVED THE PROBLEM. DEVICE EVALUATION - RETURNED PRODUCT EVALUATION FOUND THE LENS WAS RETURNED DRY AND BENT IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS IS IN SPECIFICATIONS. WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT WITHIN ASSOCIATED LOTS WAS FOUND. (B)(4). REPORT SOURCE: "USER FACILITY" TO "DISTRIBUTOR". (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -03.50/+5.0/113 | OS | Physician | 06/07/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5799981&pc=MTA | 2023826-2016-00931 | 2016-06-13 | Injury | STAAR SURGICAL COMPANY | 2016-07-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+4.0/042 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER INDICATED THE LENS HAD A REFRACTIVE SURPRISE AND REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). | -8.5/+4.0/042 | OS | Physician | 06/16/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5972072&pc=MTA | 2023826-2016-01271 | 2016-08-21 | Injury | STAAR SURGICAL COMPANY | 2016-09-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+2.0/009 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED INFORMATION: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). | -7.0/+2.0/009 | OS | Physician | 08/24/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5972074&pc=MTA | 2023826-2016-01270 | 2016-08-22 | Injury | STAAR SURGICAL COMPANY | 2016-09-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+2.0/177 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). | -9.0/+2.0/177 | OD | Physician | 08/24/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6059254&pc=MTA | 2023826-2016-01537 | 2016-09-19 | Injury | STAAR SURGICAL COMPANY | 2016-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+6.0/078 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+6.0/078 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. THE LENS WAS EXCHANGED ON (B)(6) 2017 FOR ANOTHER LENS OF THE SAME MODEL BUT DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). METHOD: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | -6.5/+6.0/078 | OD | Physician | 10/06/2016 | 2016 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6223834&pc=MTA | 2023826-2017-00013 | 2016-10-10 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -16.00/+3.0/089 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2010. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVER TIME. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EXPLANT DATE: N/A. (B)(4). LENS NOT RETURNED. | -16.00/+3.0/089 | OS | Physician | 12/07/2016 | 2016 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5867090&pc=MTA | 2023826-2016-01047 |  | Injury | STAAR SURGICAL COMPANY | 2016-08-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.5/+3.0/082 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A REFRACDTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). | -12.5/+3.0/082 | OD | Physician | 07/11/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978094&pc=MTA | 2023826-2017-01614 | 2017-02-08 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -6.5/+1.5/X063 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS REPOSITIONED. A REFRACTIVE SURPRISE FORM WAS COMPLETED BY THE SURGEON. SURGEON STATES PATIENT IS "NOT HAPPY" AND HIS BCVA IS "IMPROVING WITH REFRACTION." SURGEON IS INQUIRING AS TO WHETHER REALIGNMENT OR LENS EXCHANGE SHOULD BE PERFORMED. Manufacturer Narrative: ADDITIONAL ITEMS NOTED BY SURGEON: LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. (B)(4). ADDITIONAL CONCLUSION CODE FOR OFF-LABEL USE ADDED (PREVIOUS CORNEAL/REFRACTIVE SURGERY). (B)(4). Manufacturer Narrative: UPDATED INFORMATION: ACCORDING TO CLAIMANT THE LENS WAS REPOSITIONED AN ADDITIONAL TIME ON (B)(6) 2018. AT A FOLLOW-UP VISIT THE SURGEON NOTES THE PATIENT'S VISION IS NOT IMPROVING. THE SURGEON STATES IT'S UNKNOWN IF THE LENS HAS ROTATED AGAIN BUT IT'S NOT GIVING A GOOD REFRACTIVE OUTCOME. THE SURGEON BELIEVES THE ICL MAY BE SMALL. REPORTEDLY, THE PATIENT IS UNHAPPY AND HAS DECIDED TO REMOVE THE LENS. (B)(4) Manufacturer Narrative: EXPLANT DATE: "UNK" IS CORRECTED TO "N/A". (B)(4) Manufacturer Narrative: UPDATED INFORMATION: LENS WAS EXPLANTED AND A NEW LENS WAS IMPLANTED ON (B)(6) 2018. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | -6.5/+1.5/063 | OS | Physician | 09/28/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6927297&pc=MTA | 2023826-2017-01530 | 2017-07-01 | Injury | STAAR SURGICAL COMPANY | 2017-10-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -9.5/+3.0/175 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD ROTATED (NOT ASSOCIATED WITH LOW VAULTING), THERE WAS LOSS OF BEST-CORRECTED VISUAL ACUITY AND THE SURGEON NOTED A REFRACTIVE SURPRISE. THE LENS WAS ALSO NOTED TO HAVE TORN DURING INSERTION. THE LENS WAS REPOSITIONED ON (B)(6) 2017 BUT HAD ROTATED AGAIN, WITH LOSS OF BEST-CORRECTED VISUAL ACUITY AFTER REPOSITIONING INTERVENTION. THE LENS REMAINS IMPLANTED AND THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/60. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE DOCTOR HAS REMOVED THE LENS AFTER MAKING SURE THAT THE TORN HAPTIC AFFECTS THE STABILITY OF LENS IN SULCUS. THE DOCTOR WILL RE-EVALUATE IN A MONTH AFTER REMOVAL. WORK ORDER SEARCH: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THE LENS WAS REMOVED ON (B)(6) 2017. THE REPORTER STATED THAT THE DOCTOR IS EXPECTED TO RE-ASSESS THE PRE-OP REFRACTION AND MEASUREMENTS. THIS WAS SCHEDULED TO HAPPEN IN (B)(6) BUT MAYBE PATIENT DID NOT COME BACK YET. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+3.0/175 | OD | Physician | 09/11/2017 | 2017 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6940698&pc=MTA | 2023826-2017-01493 | 2017-08-28 | Injury | STAAR SURGICAL COMPANY | 2017-10-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICM5_12.6, -7.0/+2.5/88 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. ON (B)(6) 2017, THE LENS WAS REPOSITIONED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.0/+2.5/88 | OD | Physician | 09/13/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6955554&pc=MTA | 2023826-2017-01578 | 2017-09-19 | Injury | STAAR SURGICAL COMPANY | 2017-10-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0/+0.5/X135 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTED A REFRACTIVE SURPRISE. CURRENTLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: PATIENT IDENTIFIER IS CORRECTED FROM "(B)(6)" TO "(B)(6)" IN THE INITIAL. THE SURGEON NOTED A REFRACTIVE SURPRISE CAUSED DUE TO "WRONG ROTATION OF LENS". (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: "CURRENTLY, THE LENS REMAINS IMPLANTED." ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SAME SIZE, SIMILAR POWER LENS. REPORTEDLY, "THE OUTCOME IS VERY GOOD. UCVA 20/20." (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.0/+0.5/135 | OD | Physician | 09/21/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7487562&pc=MTA | 2023826-2018-00673 | 2018-04-11 | Injury | STAAR SURGICAL COMPANY | 2018-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+3.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD ROTATED (NOT ASSOCIATED TO A LOW VAULT) AND THERE WAS A REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED 3 TIMES BUT THE PROBLEM WAS NOT RESOLVED. THE PATIENT HAD LASER TOUCH UP ON (B)(6) 2017. THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT AND EXCHANGE FOR A LONGER LENS. Manufacturer Narrative: THE REPORTER STATED THE LENS WAS EXCHANGED ON (B)(6) 2018 FOR A LONGER LENS. THE EXCHANGE RESOLVED THE PROBLEM. (B)(4). CORRECTED INFORMATION: INJECTION MODEL SYSTEM: FTP; INJECTOR MODEL: MSI-PF; CARTRIDGE MODEL: SFC-45 NOT IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, LENS REPOSITIONING. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.50/+3.5/094 | OD | Physician | 04/16/2018 | 2018 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594290&pc=MTA | 2023826-2018-00890 | 2018-05-10 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -18.0/+2.5/062 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES REFRACTIVE SURPRISE AND LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. ADDITIONALLY, THE SURGEON STATES THE LENS WAS REPOSITIONED ON (B)(6) 2018; EACH TIME THE LENS WAS REPOSITIONED IT ROTATED AT A 90 DEGREE ANGLE AGAIN. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS REPOSITIONING). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.0/+2.5/062 | OS | Physician | 05/17/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978090&pc=MTA | 2023826-2017-01613 |  | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.0/+2.5/X096 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS REPOSITIONED. A REFRACTIVE SURPRISE FORM WAS COMPLETED BY THE SURGEON. SURGEON STATES PATIENT IS "NOT HAPPY" AND HIS BCVA IS "IMPROVING WITH REFRACTION." SURGEON IS INQUIRING AS TO WHETHER REALIGNMENT OR LENS EXCHANGE SHOULD BE PERFORMED. Manufacturer Narrative: ADDITIONAL ITEMS REPORTED BY SURGEON: LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, REFRACTIVE SURPRISE. (B)(4). Manufacturer Narrative: THE DIOPTER VALUE OF THE INITIAL LENS, VTICMO13.2 IMPLANTED IS CORRECTED FROM -5.0/+2.5/096 TO -6.0/+2.5/096 DIOPTER. (B)(4). Manufacturer Narrative: UPDATED INFORMATION: ACCORDING TO CLAIMANT THE LENS WAS REPOSITIONED AN ADDITIONAL TIME ON (B)(6) 2018. AT A FOLLOW-UP VISIT THE SURGEON NOTES THE PATIENT'S VISION IS NOT IMPROVING. THE SURGEON STATES IT'S UNKNOWN IF THE LENS HAS ROTATED AGAIN BUT IT'S NOT GIVING A GOOD REFRACTIVE OUTCOME. THE SURGEON BELIEVES THE ICL MAY BE SMALL. REPORTEDLY, THE PATIENT IS UNHAPPY AND HAS DECIDED TO REMOVE THE LENS. (B)(4) Manufacturer Narrative: EXPLANT DATE: "UNK" IS CORRECTED TO "N/A".(B)(4) Manufacturer Narrative: UPDATED INFORMATION: LENS WAS EXPLANTED AND A NEW LENS WAS IMPLANTED ON (B)(6) 2018. THE PROBLEM WAS RESOLVED (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S.. SECONDARY SURGERY, LENS REPOSITIONING. KERATOCONUS EYE. (B)(4). | -5.0/+2.5/096 | OD | Physician | 09/28/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7020606&pc=MTA | 2023826-2017-01695 | 2017-02-10 | Injury | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Insufficient Information | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/002 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS (12.1 MM) BUT DIFFERENT DIOPTER (-7.0/+1.0/012) AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+1.0/002 | OD | Physician | 10/15/2017 | 2017 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7020855&pc=MTA | 2023826-2017-01696 | 2017-02-11 | Injury | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Insufficient Information | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/180 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS (12.1 MM) BUT DIFFERENT DIOPTER (-7.0/+1.0/012) AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+1.0/180 | OS | Physician | 10/15/2017 | 2017 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7144255&pc=MTA | 2023826-2017-01912 | 2017-07-10 | Injury | STAAR SURGICAL COMPANY | 2017-12-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Insufficient Information | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+5.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL (12.6MM) BUT DIFFERENT DIOPTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR SAME MODEL/DIFFERENT DIOPTER LENS. PATIENT CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4). | -13.5/+5.0/095 | OD | Physician | 12/03/2017 | 2017 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5521849&pc=MTA | 2023826-2016-00325 | 2016-02-13 | Injury | STAAR SURGICAL COMPANY | 2016-03-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/083 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THE LENS ROTATED AND HAD A REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED AND REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS ALSO HAD REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND WAS EXCHANGED FOR A LONGER LENS, SAME MODEL AND SIMILAR DIOPTER. THE PATIENT IS DOING FINE. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (PARTICULATES) ON LENS SURFACE. (B)(4). THE DATES ON THE INITIAL WERE INCORRECTLY REPORTED AS (B)(6) 2016. THE CORRECT DATE IS (B)(6) 2016. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -13.0/+2.5/083 | OS | Physician | 02/22/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6832781&pc=MTA | 2023826-2017-01389 | 2013-10-08 | Injury | STAAR SURGICAL COMPANY | 2017-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -9.5/+1.5/102 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE PATIENT BEGAN EXPERIENCING LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING AND REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. PATIENT IS HAPPY AND ALL IS FINE. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS) AND THE LENS HAS A CURVED SHAPE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/+1.5/102 | OD | Physician | 08/08/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6832815&pc=MTA | 2023826-2017-01390 | 2013-10-08 | Injury | STAAR SURGICAL COMPANY | 2017-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -13.5/+2.5/078 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE PATIENT BEGAN EXPERIENCING LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING AND REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY AND EVERYTHING IS FINE. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS) AND THE LENS HAS A CURVED SHAPE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/+2.5/078 | OS | Physician | 08/08/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6365387&pc=MTA | 2023826-2017-00289 | 2015-12-10 | Injury | STAAR SURGICAL COMPANY | 2017-02-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/2.0/170 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING WITH REFRACTIVE SURPRISE, PIGMENT DISPERSION, AND OTHER UNKNOWN ADVERSE EVENT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A CENTRIFUGE TUBE, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED AND DEBRIS ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4). | -12.5/2.0/170 | OD | Physician | 02/01/2017 | 2017 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6645753&pc=MTA | 2023826-2017-00988 | 2016-02-14 | Injury | STAAR SURGICAL COMPANY | 2017-06-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -11.5/+2/107 IMPLANTABLE COLLAMER LENS (ICL), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT BEGAN TO EXPERIENCE LOW VAULT, GLARES/HALOES, DECENTRALIZATION, AND "FUNNY VISION." THE LENS WAS REPOSITIONED.TO DATE, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUSLY STATED THE LENS REMAINS IMPLANTED. THE CURRENT LENS HAS BEEN EXCHANGED FOR A LONGER, SIMILAR DIOPTER LENS ON (B)(6) 2017. THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND A TEAR IN THE LENS OPTIC AND PIECE OF HAPTIC MISSING FROM THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE REMAINS IMPLANTED. | -11.5/+2/107 | OD | Physician | 05/22/2017 | 2017 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6703130&pc=MTA | 2023826-2017-01160 | 2016-03-08 | Injury | STAAR SURGICAL COMPANY | 2017-07-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -10.0/+3.0/110 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT BEGAN EXPERIENCING LOW VAULT AND REFRACTIVE SURPRISE. THE LENS IS PLANNED TO BE EXCHANGED BUT CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2017. THE PROBLEM WAS RESOLVED AND "ALL ARE HAPPY". (B)(4) Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.0/+3.0/110 | OD | Physician | 06/13/2017 | 2017 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6703148&pc=MTA | 2023826-2017-01161 | 2016-03-08 | Injury | STAAR SURGICAL COMPANY | 2017-07-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -10.5/+2.0/087 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT BEGAN EXPERIENCING LOW VAULT (NO ROTATION) AND REFRACTIVE SURPRISE. THE LENS IS PLANNED TO BE EXCHANGED BUT CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2017. THE PROBLEM WAS RESOLVED AND "ALL ARE HAPPY". (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/+2.0/087 | OS | Physician | 06/13/2017 | 2017 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438818&pc=MTA | 2023826-2018-00562 | 2016-06-15 | Injury | STAAR SURGICAL COMPANY | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.5/+1.5/052 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE SURGEON NOTES REFRACTIVE CHANGE OVER TIME AND LOW VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED STATEMENT ON 12-SEP-2018 THE LENS WAS EXCHANGED WITH A LONGER LENS. THE PROBLEM IS RESOLVED. THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.5/+1.5/052 | OS | Physician | 03/21/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6443489&pc=MTA | 2023826-2017-00512 | 2016-09-26 | Injury | STAAR SURGICAL COMPANY | 2017-03-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.0/+2.0/X089 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, REFRACTIVE SURPRISE, AND SHALLOW ANTERIOR CHAMBER. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DRY, CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4). CONCLUSION CODE: OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT ((B)(6)), ACD < 3.0MM (2.8MM)] (B)(4). | -6.0/+2.0/089 | OD | Physician | 02/27/2017 | 2017 | 51
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6443524&pc=MTA | 2023826-2017-00513 | 2016-09-26 | Injury | STAAR SURGICAL COMPANY | 2017-03-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/X087 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS AND SIMILAR DIOPTER DUE TO EXCESSIVE VAULT, REFRACTIVE SURPRISE, AND SHALLOW ANTERIOR CHAMBER. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DRY, CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4). | -8.0/+1.5/087 | OS | Physician | 02/27/2017 | 2017 | 51
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6842513&pc=MTA | 2023826-2017-01337 | 2016-10-07 | Injury | STAAR SURGICAL COMPANY | 2017-09-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12/1/78 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2017 DUE TO EXCESSIVE VAULT, REFRACTIVE SURPRISE, REFRACTIVE CHANGE OVER TIME, AND HYPEROPIA SIDE-EFFECT AFTER A WEEK. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP STATED THAT THE BEST CORRECTED VISUAL ACUITY (BCVA) IS 20/13 AND UNCORRECTED VISUAL ACUITY (UCVA) IS 20/16. Manufacturer Narrative: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE US BUT NOT MARKETED IN THE US. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). | -12/1/78 | OD | Physician | 08/07/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6404238&pc=MTA | 2023826-2017-00392 | 2017-01-31 | Injury | STAAR SURGICAL COMPANY | 2017-03-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18/5/009 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, REFRACTIVE SURPRISE, AND BLURRED VISION. AS THE DATE MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: UPDATED/ADDITIONAL INFORMATION-THE PATIENT COMPLAINED OF BLURRED VISION, "AS IF SEEING EVERYTHING THROUGH WATER." THE SURGEON SUSPECTED ISSUES MAY BE RELATED TO A VERY DIFFICULT PERSONAL PERIOD THE PATIENT IS GOING THROUGH. THE LENS WAS EXPLANTED ON (B)(6) 2017. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WAS REPORTED FOR UNITS IN THE SAME LOT. (B)(4). | -18/5/009 | OS | Physician | 02/13/2017 | 2017 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6415314&pc=MTA | 2023826-2017-00434 | 2017-01-31 | Injury | STAAR SURGICAL COMPANY | 2017-03-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+4.5/144 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, FLUCTUATIONS IN VISUAL ACUITY, AND ACCOMMODATIVE SPASM. ON (B)(6) 2017 A SECONDARY SURGERY FOR IRIDECTOMY WAS PERFORMED, AND THE PATIENT STARTED PILOCARPINE TREATMENT. BASED ON SURGEON'S COMMENTS DOCTOR IS MONITORING THE PATIENT'S STATUS. AT THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). RESULT: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. CONCLUSION: OFF LABEL USE (KERATOCONUS EYE). (B)(4). | -12.5/+4.5/144 | OD | Physician | 02/21/2017 | 2017 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6832836&pc=MTA | 2023826-2017-01391 | 2017-04-14 | Injury | STAAR SURGICAL COMPANY | 2017-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -8.0/+1.0/087 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ON (B)(6) 2017, THE SURGEON REPORTED LOW VAULT AND REFRACTIVE CHANGE OVER TIME. Manufacturer Narrative: CORRECTED DATA: THE DATE OF IMPLANTATION WAS CORRECTED FROM (B)(6) 2015 TO (B)(6) 2015. CLAIM# (B)(4). Manufacturer Narrative: PLANNED EXCHANGE IS CANCELLED AS THE PATIENT IS PREGNANT. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 -8.0/+1.0/087 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE SURGEON NOTICED LOW VAULT AND REFRACTIVE CHANGE OVER TIME ON (B)(6) 2017. THE LENS REMAINS IMPLANTED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: CORRECTED DATA: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES); WAS NOT SELECTED IN PREVIOUS MDR. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 OF -8.0/+1.0/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON 03/06/2015. SURGEON REPORTED LOW VAULT AND REFRACTIVE CHANGE OVER TIME. SURGERY WAS CANCELED IN 2017 DUE TO PATIENT BEING PREGNANT. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LARGER LENGTH LENS AND PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+1.0/087 | OD | Physician | 08/08/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6832866&pc=MTA | 2023826-2017-01392 | 2017-04-14 | Injury | STAAR SURGICAL COMPANY | 2017-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -7.5/+1.0/079 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ON (B)(6) 2017, THE SURGEON REPORTED LOW VAULT AND REFRACTIVE CHANGE OVER TIME. Manufacturer Narrative: CORRECTED DATA: THE DATE OF IMPLANTATION IS CORRECTED FROM (B)(6) 2015 TO (B)(6) 2015. CLAIM# (B)(4). Manufacturer Narrative: PLANNED EXCHANGE IS CANCELLED AS THE PATIENT IS PREGNANT. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 -7.5/+1.0/079 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE SURGEON NOTICED LOW VAULT AND REFRACTIVE CHANGE OVER TIME ON (B)(6) 2017. THE LENS REMAINS IMPLANTED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 OF -7.5/+1.0/079 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(4) 2015. SURGEON REPORTED LOW VAULT AND REFRACTIVE CHANGE OVER TIME. SURGERY WAS CANCELED IN 2017 DUE TO PATIENT BEING PREGNANT. ON (B)(6)2019 THE LENS WAS EXCHANGED FOR A LARGER LENGTH LENS AND PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | -7.5/+1.0/079 | OS | Physician | 08/08/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6847031&pc=MTA | 2023826-2017-01408 | 2017-05-23 | Injury | STAAR SURGICAL COMPANY | 2017-09-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -13.0/2.0/090 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION, BLURRED VISION, AND SHALLOW AC. A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: THE LENS EXCHANGE FOR A SHORTER LENS RESOLVED THE PROBLEM. CORRECTED DATA: THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION, BLURRED VISION, AND SHALLOW AC. THE LENS WAS EXCHANGED ON (B)(6) 2017 FOR A SHORTER LENS. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.0/2.0/090 | OD | Physician | 08/10/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7139572&pc=MTA | 2023826-2017-01930 | 2017-09-27 | Injury | STAAR SURGICAL COMPANY | 2017-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/87 (SPHERE/CYLINDER/AXIS) , IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCE LOW VAULT AND REFRACTIVE SURPRISE. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. WORK ORDER SEARCH - NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/+2.5/87 | OS | Physician | 11/22/2017 | 2017 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7139744&pc=MTA | 2023826-2017-01929 | 2017-09-27 | Injury | STAAR SURGICAL COMPANY | 2017-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0/+2.5/91 (SPHERE/CYLINDER/AXIS) , IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCE LOW VAULT AND REFRACTIVE SURPRISE. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON 11 APR 2018 DUE TO LOW VAULTING AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. WORK ORDER SEARCH- NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.0/+2.5/91 | OD | Physician | 11/22/2017 | 2017 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7359037&pc=MTA | 2023826-2018-00338 | 2018-02-07 | Injury | STAAR SURGICAL COMPANY | 2018-03-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICM5_13.7, -12.00/+1.5/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED. A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS . VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: CORRECTED DATA: "ADVERSE EVENT" IS CORRECTED TO "PRODUCT PROBLEM". "OTHER SERIOUS" IS REMOVED FROM THE INITIAL MDR. TYPE OF REPORTABLE EVENT: IT IS CORRECTED FROM "SERIOUS INJURY" TO "MALFUNCTION". CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | -12.00/+1.5/102 | OD | Physician | 02/20/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7639247&pc=MTA | 2023826-2018-00978 | 2018-04-03 | Injury | STAAR SURGICAL COMPANY | 2018-06-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICM5_12.1, -11.0/+2.0/81 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION - LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION FOUND RESIDUE AND DEBRIS ON LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/+2.0/81 | OS | Physician | 05/29/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7696918&pc=MTA | 2023826-2018-01031 | 2018-06-19 | Injury | STAAR SURGICAL COMPANY | 2018-07-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+3.5/066 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES EXCESSIVE VAULTING AND REFRACTIVE SURPRISE. THE CAUSE IS LISTED AS A PATIENT-RELATED FACTOR AND STATES THE LENS WAS DIFFICULT TO MANIPULATE INTO THE CORRECT POSITION BECAUSE THE "PATIENT WAS VERY NERVOUS AND HECTIC." THE SURGEON ALSO STATES THE SULCUS SIZE DOES NOT CORRELATE WITH THE WHITE TO WHITE MEASUREMENT. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED/CORRECTED INFORMATION: ON (B)(4) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM IS RESOLVED AND THE "PATIENT IS HAPPY WITH 20/20 VISION." (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.5/+3.5/066 | OS | Physician | 06/20/2018 | 2018 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6060685&pc=MTA | 2023826-2016-01533 |  | Injury | STAAR SURGICAL COMPANY | 2016-10-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+3.0/092 DIOPTER, IN THE PATIENT'S RIGHT EYE. THE PATIENT EXPERIENCED EXCESSIVE VAULTING OF THE LENS, A SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, BLURRED VISION, AND UNREACTIVE PUPIL. THE LENS WAS REPOSITIONED, AND THE ANTERIOR CHAMBER WAS WASHED AND IRRIGATED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2 VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+3.0/092 DIOPTER, IN THE PATIENT'S RIGHT EYE. THE PATIENT EXPERIENCED EXCESSIVE VAULTING OF THE LENS, A SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, BLURRED VISION, UNREACTIVE PUPIL, AND IRIS PUSHED TOWARDS THE CORNEA. THE LENS WAS REPOSITIONED AND THE ANTERIOR CHAMBER WAS WASHED AND IRRIGATED/EVACUATED OF REMAINING VISCO/FLUIDS. (B)(4). CORRECTED DATA: EXPLANTED DATE: N/A. (B)(6). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -13.00/+3.0/092 | OD | Physician | 10/12/2016 | 2016 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5199927&pc=MTA | 2023826-2015-01454 | 2014-07-05 | Injury | STAAR SURGICAL COMPANY | 2015-11-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER -5.5/+2.0/X081 DIOPTER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LOW VAULT, LENS ROTATION, BLURRED VISION AND REFRACTIVE CHANGE WAS NOTED (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 AND WAS EXCHANGED FOR A LONGER LENS LENGTH. THE PROBLEM WAS RESOLVED. PATIENT'S LAST VISIT ON (B)(6) 2014 - UCVA 20/32. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -5.5/+2.0/X081. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -5.5/+2.0/081 | OD | Physician | 08/31/2015 | 2015 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4806370&pc=MTA | 2023826-2015-00588 | 2014-08-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) AND NOTED EXCESSIVE VAULT, WITH A REFRACTIVE SURPRISE/REFRACTIVE CHANGE OVERTIME, IN THE POST-OPERATIVE PERIOD. THE PATIENT EXPERIENCED BLURRED VISION. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. (B)(4). LENS REMAINS IMPLANTED. | -10.0/+5.0/011 | OS | Physician | 03/09/2015 | 2015 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4740720&pc=MTA | 2023826-2015-00489 | 2014-10-23 | Injury | STAAR SURGICAL COMPANY | 2015-04-30 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WERE NOTED, PATIENT EXPERIENCED A REDUCTION OF IRIDO-CORNEAL ANGLES AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITH A DIFFERENT DIOPTER ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST KNOWN UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -3.5/+1/50 | OS | Physician | 03/31/2015 | 2015 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5193772&pc=MTA | 2023826-2015-01444 | 2015-10-01 | Injury | STAAR SURGICAL COMPANY | 2015-11-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/112 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND A REFRACTIVE SURPRISE. THE SURGEON IS PLANNING ON EXPLANTING THE LENS. THE LENS REMAINS IMPLANTED.Event Description: THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A LONGER LENS WITH A DIFFERENT DIOPTER. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A LONGER LENS WITH A DIFFERENT DIOPTER. THIS EXCHANGE RESOLVED THE PROBLEM. (B)(4). AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.80MM AT THE TIME OF IMPLANTATION. CLAIM #(B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, PATIENT AGE BELOW 21 OR OVER 45 YEARS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. EXPLANT DATE: UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. (B)(4) METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) LENS NOT RETURNED. | -8.5/+1.5/112 | OD | Physician | 10/02/2015 | 2015 | 49
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5464177&pc=MTA | 2023826-2016-00200 | 2016-01-29 | Injury | STAAR SURGICAL COMPANY | 2016-02-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+0.5/075 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD AN EXCESSIVE VAULT AND A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -9.50/+0.5/075 | OS | Physician | 02/05/2016 | 2016 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6089943&pc=MTA | 2023826-2016-01603 | 2016-08-01 | Injury | STAAR SURGICAL COMPANY AG | 2016-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+4.0/093 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, LENS ROTATION AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/23.Event Description: THE PROBLEM WAS RESOLVED AFTER LENS EXCHANGE. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: CONCLUSION CODE: (OFF-LABEL USE): PATIENT ACD WAS BELOW 3.00 MM. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN CASE/VIAL; SURGICAL RESIDUE WAS CLEAR; VISUAL INSPECTION FOUND HAPTIC BROKEN; DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND WITHIN SPECIFICATIONS. (B)(4). | -9.0/+4.0/093 | OD | Physician | 10/12/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6247436&pc=MTA | 2023826-2017-00065 | 2016-08-03 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -11.0/+2.5/098, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO THE PATIENT REPORTING LOW VAULT, ROTATION OF THE LENS, AND REFRACTIVE CHANGE OVER TIME. THE SURGEON STATED THAT THE REFRACTIVE CHANGE OVER TIME IS DUE TO THE LOW VAULT AND ICL ROTATION. THE LENS HAS BEEN RETURNED, BUT HAS NOT YET BEEN EVALUATED. Manufacturer Narrative: THE LENS HAS BEEN RETURNED AND HAS BEEN EVALUATED. (B)(4). THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC DEFORMED AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (HAIRS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.5/098 | OD | Physician | 12/16/2016 | 2016 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7245220&pc=MTA | 2023826-2018-00205 | 2017-01-17 | Injury | STAAR SURGICAL COMPANY | 2018-02-06 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.5/+1.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE CHANGE OVER TIME, LOW VAULT, AND LENS ROTATION. REPORTEDLY, LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. CORRECTED/UPDATED INFORMATION: ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS. THE PROBLEM IS RESOLVED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.5/+1.0/087 | OD | Physician | 01/16/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7245224&pc=MTA | 2023826-2018-00207 | 2017-01-17 | Injury | STAAR SURGICAL COMPANY | 2018-02-06 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.0/+1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE CHANGE OVER TIME, LOW VAULT, AND LENS ROTATION. REPORTEDLY, LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: PREVIOUSLY STATED THE LENS REMAINS IMPLANTED. CORRECTED/UPDATED INFORMATION: ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM IS RESOLVED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH WHITE LIQUID ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.0/+1.5/091 | OS | Physician | 01/16/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7268474&pc=MTA | 2023826-2018-00194 | 2017-01-31 | Injury | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -12.5/+1.5/65 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. EXCESSIVE VAULT, LENS ROTATION, AND REFRACTIVE CHANGE OVER TIME WAS NOTICED. A SURGERY FOR A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 AND A NEW LONGER LENGTH LENS WAS IMPLANTED LATER ON (B)(6) 2018. THE PROBLEM WAS RESOLVED. EXPLANT DATE: (B)(6) 2018. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: CORRECTED DATA: IT IS CORRECTED FROM "ADVERSE EVENT" TO "PRODUCT PROBLEM" IN THE INITIAL MDR. "EXCESSIVE VAULT" IS CORRECTED TO "LOW VAULT" IN THE INITIAL MDR. TYPE OF REPORTABLE EVENT: IT IS CORRECTED FROM "SERIOUS INJURY" TO "MALFUNCTION" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.5/+1.5/65 | OS | Physician | 01/16/2018 | 2018 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6673242&pc=MTA | 2023826-2017-01044 | 2017-05-03 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/092 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE. THE DR. MENTIONED THAT THE LENS IS DECENTERED IN PUPIL TEMPORALLY, HE WONDERED IF A TEMPORAL HAPTIC COULD BE DAMAGED/FOLDED. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. ACCORDING TO THE DR., THE PATIENT IS MANAGING THE PROBLEM WITH CONTACT LENSES, AND THE DR. ALSO PLANNED TO EXCHANGE THE LENS. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/+1.5/092 | OD | Physician | 06/01/2017 | 2017 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6595410&pc=MTA | 2023826-2017-00892 | 2017-05-04 | Injury | STAAR SURGICAL COMPANY | 2017-05-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, DIOPTER -7.00/+6.0/118 IMPLANTABLE COLLAMER LENS (ICL), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT BEGAN TO EXPERIENCE LOW VAULT, REFRACTIVE SURPRISE, AND LENS ROTATION. THE LENS WAS REPOSITIONED (B)(6) 2017.TO DATE, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUS: LENS REMAINS IMPLANTED. CURRENT: ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A LONGER LENS. VAULT IS NORMAL, PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE REMAINS IMPLANTED. | -7.00/+6.0/118 | OD | Physician | 05/05/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6865845&pc=MTA | 2023826-2017-01453 | 2017-05-20 | Injury | STAAR SURGICAL COMPANY | 2017-09-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -9.0/+3.5/095 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED BECAUSE LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE WERE NOTED BY THE SURGEON. THE REPLACEMENT LENS WAS OF LONGER LENGTH AND SIMILAR DIOPTER. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL AND THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND THE HAPTIC BENT AND DEFORMED. (B)(4). DEVICE MANUFACTURER DATE: IT IS CORRECTED TO "27-MAY-2016". (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+3.5/095 | OS | Physician | 08/18/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7169880&pc=MTA | 2023826-2018-00012 | 2017-11-25 | Injury | STAAR SURGICAL COMPANY | 2018-01-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE PATIENT EXPERIENCED BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE VAULT IS ADEQUATE AND PATIENT IS EXTREMELY HAPPY. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT; THE LENS MISSING A PIECE OF ITS HAPTIC AND PRESENCE OF RESIDUE AND DEBRIS ON SURFACE OF THE LENS. METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S (B)(4). (B)(4). | -9.0/+1.0/100 | OD | Physician | 12/18/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7169893&pc=MTA | 2023826-2018-00013 | 2017-11-25 | Injury | STAAR SURGICAL COMPANY | 2018-01-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE PATIENT EXPERIENCED BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE VAULT IS ADEQUATE AND PATIENT IS EXTREMELY HAPPY. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S (B)(4). (B)(4). | -8.5/+1.0/093 | OS | Physician | 12/18/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7481819&pc=MTA | 2023826-2018-00628 | 2018-02-27 | Injury | STAAR SURGICAL COMPANY | 2018-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -11.0/2.0/082 (SPHERE/AXIS/SPHERE), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED ON (B)(46 2018 DUE TO LOW VAULTING, LENS ROTATION, AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND BENT ALONG WITH PRESENCE OF RESIDUE ON THE LENS SURFACE. THE LENS WAS ALSO MISSING THE TORN PIECE OF ITS HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/2.0/082 | OS | Physician | 04/05/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7639140&pc=MTA | 2023826-2018-00977 | 2018-04-03 | Injury | STAAR SURGICAL COMPANY | 2018-06-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -11.0/+2.5/86 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/+2.5/86 | OD | Physician | 05/29/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6589138&pc=MTA | 2023826-2017-00867 |  | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | INPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: ON (B)(6) 2017 THE REPORTER STATED THAT HE NEEDED TO RETURN A LENS THAT WAS "REMOVED FROM THE EYE DUE TO SIZING ISSUES." ATTEMPTS TO RETRIEVE ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: UPDATED/CORRECTED INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, DIOPTER -16.0/+4.5/115 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT BEGAN EXPERIENCING LOW VAULT, LENS ROTATION, BLURRED VISION. THE LENS WAS REPOSITIONED. THE LENS WAS EXCHANGED ON (B)(6) 2017. THE PROBLEM IS RESOLVED. ADDITIONAL INFORMATION ADDED. LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH SOME MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: PREVIOUS: 26-APR-2017. _CORRECTED: 25-APR-2017. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -16.0/+4.5/115 | OD | Physician | 04/25/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7114673&pc=MTA | 2023826-2017-01860 |  | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -16.00/+1.50/89 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING THAT CAUSE LENS ROTATION TWICE AND REFRACTIVE CHANGE OVER TIME. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.00/+1.50/89 | OD | Physician | 11/14/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6769268&pc=MTA | 2023826-2017-01284 |  | Injury | STAAR SURGICAL COMPANY | 2017-08-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Insufficient Information | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON INPLANTED AN ICL IMPLANTABLE COLLAMER LENS, AND THERE WAS A REFRACTIVE SURPRISE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+3.0/180 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING, LENS ROTATION AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED. THE PATIENT STILL HAS SOME REFRACTION LEFT, BUT THE SURGEON DECIDED NOT TO ROTATE THE LENS. THE SURGEON IS PLANNING TO DO A LASER TOUCH-UP. (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.50/+3.0/180 | OD | Physician | 07/14/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6675761&pc=MTA | 2023826-2017-01083 | 2016-12-26 | Injury | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.50/+4.5/062 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND REFRACTIVE SURPRISE. THE SURGEON STATED THAT THE LENS SEEMS NOT TO BE ROTATED AT SLIT LAMP, AND REFRACTION WAS CHANGING THE FIRST WEEKS BUT IT IS STABLE SINCE 2 MONTHS. THE DOCTOR PLANS TO EXCHANGE THE LENS. Manufacturer Narrative: THE DOCTOR PLANS TO EXCHANGE THE LENS. UPDATED ON 21-AUG-2017, THE LENS WAS EXCHANGED FOR A SHORTER LENS, SIMILAR DIOPTER. THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQUID IN A MICRO CENTRIFUGE VIAL AND HAD CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF DEBRIS ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS DEVICE IS NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4)- NO CODE AVAILABLE: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES METHOD CODE: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4). | -4.50/+4.5/062 | OS | Physician | 06/06/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4621347&pc=MTA | 2023826-2015-00356 | 2014-11-04 | Injury | STAAR SURGICAL COMPANY | 2015-03-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE WITH AN INCIDENT OF REFRACTIVE SURPRISE. AN EXCHANGE OF THE LENS IS PLANNED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4): EVALUATION: METHOD: WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS NOT RETURNED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE IN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE LENS WAS EXPLANTED 5 MONTHS AFTER THE INITIAL SURGERY. THE LENS WAS EXCHANGED FOR A SMALLER LENS WITH DIFFERENT POWER TO ADJUST THE ERROR AND THE PROBLEM WAS RESOLVED. EXPLANTED: (B)(6) 2015. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE IN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | -2/+2.5/180 | OD | Other | 02/20/2015 | 2015 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5865428&pc=MTA | 2023826-2016-01033 | 2015-12-10 | Injury | STAAR SURGICAL COMPANY | 2016-08-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+3.0/080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND A REFRACTIVE SURPRISE AND REMAINS IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WEIGHT: UNK. OUTCOMES ATTRIBUTED TO ADVERSE EVENTS: UNK. EXPLANT DATE: N/A. (B)(4). LENS REMAINS IMPLANTED. | -10.0/+3.0/080 | OD | Physician | 07/08/2016 | 2016 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5838759&pc=MTA | 2023826-2016-00993 | 2016-06-09 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.5/3.0/014 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL WITH PARTICULATES ON LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: PREVIOUSLY REPORTED EVENT DESCRIPTION IS INCORRECT. CORRECT EVENT DESCRIPTION IS THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -5.5/3.0/014 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXCHANGED ON (B)(6) 2016 DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | -5.5/3.0/014 | OS | Physician | 06/30/2016 | 2016 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135217&pc=MTA | 2023826-2017-01919 | 2012-02-21 | Injury | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -17.0/+2.5/097 (SPHERE/CYLINDAR/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. REPORTEDLY, THE SURGEON NOTES EXCESSIVE VAULT AND REFRACTIVE CHANGE OVER TIME. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: REPORTEDLY, WHEN THE SURGEON EXAMINED THE EYE PRIOR TO PLANNED LENS EXCHANGE SURGERY, IT WAS DISCOVERED THAT THE LENS HAD ROTATED. THE SURGEON REPOSITIONED THE LENS AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.0/+2.5/097 | OS | Physician | 11/24/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594258&pc=MTA | 2023826-2018-00888 | 2018-03-20 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -5.5/+1.0/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON NOTES LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2018. THE EXCHANGE WAS SUCCESSFUL. PROBLEM RESOLVED. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND VISUAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: CORRECTED DATA: IT IS CORRECTED TO (B)(6) 2018 IN THE SUPPLEMENTAL MDR#1. DEVICE EVALUATION CORRECTED TO: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.5/+1.0/070 | OD | Physician | 05/16/2018 | 2018 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4635334&pc=MTA | 2023826-2015-00381 | 2014-11-17 | Injury | STAAR SURGICAL COMPANY | 2015-03-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Device Dislodged or Dislocated; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LOW VAULT WAS NOTED AS WELL AS LENS DISLOCATION/SUBLUXATION, REFRACTIVE CHANGE OVER TIME, AND BLURRED VISION. THE LENS WAS EXPLANTED ON (B)(6) 2014. THE PATIENT'S LAST BVCA WAS 20/60 AS OF (B)(6) 2014. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY THE MANUFACTURER? NO. LENS WAS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: (B)(4). MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THE LENS DESCRIPTION IS -10.0/+3.0/063 (SPHERE/CYLINDER/AXIS). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CONCLUSION: AT THE TIME OF IMPLANTATION, PATIENT WAS (B)(6) (I.E. UNDER THE AGE OF 21 YEARS) AND HAD PRE-EXISTING PROGRESSIVE MYOPIA OR HYPEROPIA, BOTH OF WHICH ARE CONTRAINDICATED AS PER THE DFU OF THIS LENS MODEL. (B)(4). | -10.0/+3.0/063 | OS | Other | 03/09/2015 | 2015 | 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5678984&pc=MTA | 2023826-2016-00743 | 2015-11-18 | Injury | STAAR SURGICAL COMPANY | 2016-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+2.5/066 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD A LOW VAULT, ROTATION, AND A REFRACTIVE SURPRISE, AND WAS REPOSITIONED. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BROKEN AND FOREIGN MATERIAL (PARTICULATES) ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION: AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.84MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: LENS NOT RETURNED. (B)(4). | -15.5/+2.5/066 | OS | Physician | 05/09/2016 | 2016 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4634764&pc=MTA | 2023826-2015-00380 | 2014-11-17 | Injury | STAAR SURGICAL COMPANY | 2015-03-26 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LOW VAULT WAS NOTED AS WELL AS LENS DISLOCATION/SUBLUXATION, REFRACTIVE CHANGE OVER TIME, AND BLURRED VISION. THE LENS WAS REPOSITIONED. THE LENS WAS EXPLANTED ON (B)(6) 2014. THE PATIENT'S LAST BVCA WAS 20/60 AS OF (B)(6) 2014. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). NOT RETURNED. Manufacturer Narrative: PER MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). | -11.00/+3.0/106 | OD | Other | 03/09/2015 | 2015 | 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5968182&pc=MTA | 2023826-2016-01309 | 2016-07-30 | Injury | STAAR SURGICAL COMPANY | 2016-09-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+4.5/110 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, ROTATION AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (SMALL FIBERS). (B)(4). METHOD CODE: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA (B)(4). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -11.5/+4.5/110 | OD | Physician | 08/22/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5018700&pc=MTA | 2023826-2015-01065 | 2015-07-21 | Injury | STAAR SURGICAL COMPANY | 2015-08-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/+1.50X45 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT COMPLAINT OF BLURRED VISION. ON EXAMINATION, THE PATIENT'S UCVA: 6/15 WITH -0.50/+1.75X115. A REFRACTIVE SURPRISE WAS REPORTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: UPON REVIEW OF THE DEVICE HISTORY RECORD, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON THE INVESTIGATION, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. EVALUATION CONCLUSION CODE(S) BASED ON THE COMPLAINT HISTORY, DEVICE HISTORY RECORD REVIEW AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - (OTHER): LENS REMAINS IMPLANTED. EVENT PROBLEM AND EVALUATION CODES: PATIENT CODE: (B)(4)- NO CODE AVAILABLE (REFRACTIVE SURPRISE), UNLABELED DEVICE CODE: (B)(4)- NO INFORMATION. METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED.(B)(4). OTHER TEXT : LENS REMAINS IMPLANTED | -12.50/+1.5045 | OS | Physician | 07/21/2015 | 2015 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6817519&pc=MTA | 2023826-2017-01367 | 2017-07-28 | Injury | STAAR SURGICAL COMPANY | 2017-08-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/170 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS REPOSITIONED ON (B)(6) 2017. THE REPORTER INDICATED THE REASON FOR THE PROBLEM WAS NOT A ROTATION OF THE LENS CAUSED THROUGH A SIZING PROBLEM, IT WAS THE SURGEON WHO IMPLANTED THE LENS 5º CW (CLOCKWISE) INSTEAD OF 5º CCW(COUNTER CLOCKWISE). THE PATIENT'S CURRENT CONDITION IS UCVA - 20/20, AND PATIENT IS HAPPY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.5/170 | OD | Physician | 07/28/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6589758&pc=MTA | 2023826-2017-00773 |  | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AND REMOVED A PHAKIC IMPLANTABLE COLLAMER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/+2.0/118 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE HAPTIC BENT. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. DIMENSIONAL AND FUNCTIONAL TESTING WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. (B)(4). Manufacturer Narrative: (B)(4). THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 +1, THE EYE IS QUIET. DUANE'S RETRACTION SYNDROME. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/+2.0/118 | OD | Physician | 04/27/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8164819&pc=MTA | 2023826-2018-01818 | 2017-11-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -16.00/1.0/058 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENTS IS RPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTION INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.00/1.0/058 | OS |  | 11/20/2018 | 2018 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7700430&pc=MTA | 2023826-2018-01066 | 2018-05-31 | Injury | STAAR SURGICAL COMPANY | 2018-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14/1.5/24 (SPHERE/CYLIBDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND HAPTIC TORN/BROKEN/DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14/1.5/24 | OD | Physician | 06/19/2018 | 2018 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824451&pc=MTA | 2023826-2018-01256 | 2018-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-08-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+2.5/171 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT EXPERIENCING A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME MODEL BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS REPORT IS BEING SUBMITTED PER FDA REQUEST. THIS REPORT WAS PREVIOUSLY FILED WITH INCORRECT MFR NUMBER. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -8.0/+2.5/171 | OS | Physician | 07/30/2018 | 2018 | 17 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7696916&pc=MTA | 2023826-2018-01035 | 2018-06-18 | Injury | STAAR SURGICAL COMPANY | 2018-07-17 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/178 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY THE LENS REMAINS IMPLANTED. THE SURGEON STATES THERE IS A "SIGN ERROR ON ASTIGMATISM" THAT OCCURRED DURING LENS CALCULATION. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. CORRECTED STATEMENT: ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR THE SAME SIZE, DIFFERENT SPHERE/CYLINDER/AXIS LENS. THE PROBLEM IS RESOLVED. THE STATUS OF THE EYE IS REPORTED AS "PERFECT". (B)(4). Manufacturer Narrative: PREVIOUS MDR SHOULD INCLUDE THE FOLLOWING STATEMENT: "USE/HANDLING AND ADDITIONAL PROCESSING IMPACTS LENS DIMENSIONS/FUNCTIONS; RETURNED LENS MAY NOT MEET ORIGINAL SPECIFICATIONS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE CYLINDER AND AXIS TO BE WITHIN SPECIFICATIONS BUT FOUND THE SPHERE TO BE 0.07D OFF SPECIFICATIONS. FUNCTIONAL INSPECTION AT THE TIME OF FINAL INSPECTION PRIOR TO PRODUCT RELEASE FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.0/+1.0/178 | OS | Physician | 06/21/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7765576&pc=MTA | 2023826-2018-01175 | 2018-06-21 | Injury | STAAR SURGICAL COMPANY | 2018-08-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -16.00/4.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SUPRISE WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.00/4.5/90 | OS | Physician | 07/09/2018 | 2018 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7961068&pc=MTA | 2023826-2018-01509 | 2018-06-21 | Injury | STAAR SURGICAL COMPANY | 2018-10-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A VTICMO12.6,-18.00/3.0/148 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018 WITH AN OBSERVATION OF REFRACTIVE SURPRISE. REPOSITIONING OF THE LENS WAS CONDUCTED AND THE PROBLEM WAS RESOLVED. THE REPORTER STATED THAT 'THE PATIENT IS ABSOLUTELY FINE' AND THAT 'VN IS 6/9'. IT WAS REPORTED THAT THE CAUSE OF EVENT WAS DUE TO A USER ERROR, SURGEON ADMITTED TO IMPLEMENTING INCORRECT DATA IN OCOS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.00/3.0/148 | OS | Physician | 08/23/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7914678&pc=MTA | 2023826-2018-01450 | 2018-06-28 | Injury | STAAR SURGICAL COMPANY | 2018-09-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.0/+3.0/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.0/+3.0/94 | OD | Physician | 08/29/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7906882&pc=MTA | 2023826-2018-01446 | 2018-07-23 | Injury | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.0/+1.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE IS LISTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.0/+1.5/074 | OS | Physician | 08/31/2018 | 2018 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7906886&pc=MTA | 2023826-2018-01443 | 2018-07-23 | Injury | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, IMPLANTABLE COLLAMER LENS, -6.5/+1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE IS LISTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN VIAL WITH CLEAR SURGICAL DEBRIS/RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -6.5/+1.5/090 | OD | Physician | 08/31/2018 | 2018 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7860735&pc=MTA | 2023826-2018-01248 | 2018-08-14 | Injury | STAAR SURGICAL COMPANY | 2018-09-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.00/+10/67 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. REFRACTIVE SUPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED WITH EXCHANGE PLANNED. PATIENT IS FINE. Manufacturer Narrative: DIOPTER CORRECTED TO "-9.00/+1.0/167" IN THE INITIAL MDR. THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -9.00/+1.0/167 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED DUE TO REFRACTIVE SURPRISE. THE REPORTER STATED THAT "THE EXCHANGE WENT WELL AND THE PATIENT IS HAPPY". THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.00/+10/67 | OD | Physician | 08/14/2018 | 2018 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7860748&pc=MTA | 2023826-2018-01249 | 2018-08-14 | Injury | STAAR SURGICAL COMPANY | 2018-09-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.00/+10/67 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SUPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED WITH EXCHANGE PLANNED. PATIENT IS FINE. Manufacturer Narrative: CORRECTED DATA: DIOPTER CORRECTED TO: "-8.00/0.5/002" IN THE INITIAL MDR. EXPIRATION DATE CORRECTED TO: 30-APR-2021 IN THE INITIAL MDR. MANUFACTURE DATE CORRECTED TO: 11-MAY-2018 IN THE INITIAL MDR. ADDITIONAL DATA: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -8.00/0.5/002 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED DUE TO REFRACTIVE SURPRISE. THE REPORTER STATED THAT "THE EXCHANGE WENT WELL AND THE PATIENT IS HAPPY". THE PROBLEM WAS RESOLVED. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.00/+10/67 | OS | Physician | 08/14/2018 | 2018 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8113360&pc=MTA | 2023826-2018-01756 | 2018-09-06 | Injury | STAAR SURGICAL COMPANY | 2018-11-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -0.50/+1.5/065 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC BENT AND DEFORMED. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. DIMENSIONAL AND REFRACTIVE (FUNCTIONAL) VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -0.50/+1.5/065 | OS | Physician | 10/29/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8227438&pc=MTA | 2023826-2019-00059 | 2018-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/103 (SPHERE/CYLINDAR/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE SURGEON REPORTS REFRACTIVE CHANGE OVER TIME AND "VISUAL POWER PROBLEM ON CLOSE DISTANCES." REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: PROBLEM IS RESOLVED WITH THE EXCHANGE; "ALL IS FINE!". (B)(4). Manufacturer Narrative: ADDED APPLICABLE OUTCOME/ PREVIOUSLY STATED THE LENS REMAINS IMPLANTED. UPDATED INFORMATION: ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL, SPHERE, CYLINDER, AXIS LENS. REPORTEDLY, POST-OP RESULTS ARE NOT CURRENTLY AVAILABLE. DISTRIBUTOR CHECKED IN ERROR. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION EVENT: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2013. THE SURGEON REPORTS REFRACTIVE CHANGE OVER TIME AND "VISUAL POWER PROBLEM ON CLOSE DISTANCES." THE LENS WAS EXPLANTED ON A LATER DATE. ADDITIONAL INFORMATION REGARDING THE PATIENT STATUS AND IF PROBLEM WAS RESOLVED HAS BEEN REQUEST. IF ADDITIONAL DETAILS ARE AVAILABLE A SUPPLEMENTAL MDR WILL BE SUBMITTED. DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN AND RESIDUE ON LENS SURFACE. PATIENT CODE- (B)(4). CLAIM# COMPLAINT-(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/+1.0/103 | OS |  | 12/14/2018 | 2018 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7975584&pc=MTA | 2023826-2018-01531 | 2018-09-19 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_13.7, -8.50/1.0/076 (SPHERE/CYLINDER/AXIS) IMPLANTABLE LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED AND THE LENS WAS REMOVED ON (B)(6) 2018. A REPLACEMENT LENS (DIFFERENT LENGTH AND DIOPTER) WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. IT WAS REPORTED THAT THE CAUSE OF EVENT WAS USER ERROR, THE SURGEON USED THE WRONG RX DATA FOR ICL POWER CALCULATION. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/1.0/076 | OD | Physician | 09/19/2018 | 2018 | 49
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7975592&pc=MTA | 2023826-2018-01532 | 2018-09-19 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_13.7, -9.00/0.5/097 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED AND THE LENS WAS REMOVED ON (B)(6) 2018. A REPLACEMENT LENS (DIFFERENT LENGTH AND DIOPTER) WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. IT WAS REPORTED THAT THE CAUSE OF EVENT WAS USER ERROR, THE SURGEON USED THE WRONG RX DATA FOR ICL POWER CALCULATION. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/0.5/097 | OS | Physician | 09/19/2018 | 2018 | 49
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8229422&pc=MTA | 2023826-2019-00023 | 2018-09-20 | Injury | STAAR SURGICAL COMPANY | 2019-01-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.5/+3.0/101 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. | -17.5/+3.0/101 | OD | Physician | 12/09/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8268673&pc=MTA | 2023826-2019-00018 | 2018-10-20 | Injury | STAAR SURGICAL COMPANY | 2019-01-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.00/1.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED WITHIN THE SAME SURGERY DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.00/1.0/096 | OD | Physician | 12/25/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8270834&pc=MTA | 2023826-2019-00017 | 2018-10-20 | Injury | STAAR SURGICAL COMPANY | 2019-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.00/2.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SAME SIZE , "SIMILARE" (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. WITHIN THE SAME SURGERY. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.00/2.0/085 | OS | Physician | 12/25/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8089154&pc=MTA | 2023826-2018-01716 | 2018-10-22 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -2.50/5.5/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED. IT WAS REPORTED THAT THE SURGEON PLANS TO DO A PHOTOREFRACTIVE KERATECTOMY (PRK) TO RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED: IT WAS REPORTED THAT "AFTER PRK REFRACTION NOW IS -0,25/-0,25 WITH A BCVA OF 20/20". (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -2.50/5.5/099 | OS | Physician | 10/23/2018 | 2018 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8276381&pc=MTA | 2023826-2019-00156 | 2018-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -3.5/+1.0/170 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME MODEL/LENGTH LENS BUT DIFFERENT SPHERE/CYLINDER/AXIS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -3.5/+1.0/170 | OD |  | 12/31/2018 | 2018 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8166385&pc=MTA | 2023826-2018-01798 | 2018-11-03 | Injury | STAAR SURGICAL COMPANY | 2018-12-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/X166 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO USER ERROR (I.T. ERROR AT THE SURGEON'S OFFICE. THE PATIENTS REFRACTION DATA WAS INPUT INCORRECTLY AS PLUS CYLINDER INSTEAD OF A MINUS CYLINDER). THE PATIENT HAD RESIDUAL ASTIGMATISM POST-OPERATIVE AND A REFRACTIVE TREATMENT (LASIK) IS PLANNED TO CORRECT PROCEDURE. Manufacturer Narrative: THE REPORTER INDICATED THE PATIENT HAD LASIK SURGERY ON (B)(6) 2018 FOR CORRECTION AND THE PATIENT IS HAPPY WITH THE RESULTS. (B)(4). Manufacturer Narrative: PT INFO: UNK. EXPLANT DATE: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/+1.5/166 | OS | Physician | 11/15/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8149333&pc=MTA | 2023826-2018-01801 | 2018-11-08 | Injury | STAAR SURGICAL COMPANY | 2018-12-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -3.5/+1.5/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SAME SIZE, SIMILAR DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS NOT RESOLVED. REPORTEDLY, THE "PATIENT IS DISSATISFIED WITH THE MYOPIA REMAINING." THE CAUSE IS LISTED AS UNKNOWN. SEE MFR REPORT# 2023826-2018-01804 FOR CASE CONTINUATION. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FIBERS ON THE LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -3.5/+1.5/085 | OS | Physician | 11/19/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7891566&pc=MTA | 2023826-2018-01420 |  | Injury | STAAR SURGICAL COMPANY | 2018-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -6.00/+1.0/69 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER DUE TO OBSERVATION OF REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVERTIME. IT WAS REPORTED THAT THE DESIRED DEGREE OF MYOPIA WAS NEVER ACHEIVED AND OVER TIME IT DRIFTED EVEN LESS MYOPIC. Manufacturer Narrative: ADDITIONAL DATA: ADDITIONAL INFORMATION WAS RECEIVED. REPORTER INDICATED THAT THE EXCHANGED LENS RESOLVED THE PROBLEM. LAST REPORT STATES THAT VISION IS IMPROVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. THE LENS WAS RETURNED IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. | -6.00/+1.0/69 | OS | Other Health Care Professional | 08/21/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7918080&pc=MTA | 2023826-2018-01453 |  | Injury | STAAR SURGICAL COMPANY | 2018-09-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A VTICMO13.2, -8.50/5.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). REFRACTIVE SURPRISE WAS OBSERVED AND ROTATION/REPOSITIONING WAS PERFORMED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: ADDITIONAL INFORMATION: " THE DOCTOR DID THE REPOSITIONING LAST WEEK AND HE IS WAITING FOR PATIENT'S FOLLOW UP TO SEE THE RESULT". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: ADDITIONAL INFORMATION: " THE ICL WAS ROTATED ANTI-CLOCKWISE TO THE DESIRED AXIS. REFRACTION IN HER LAST FOLLOW UP VISIT (3 WEEKS FOLLOWING THE PROCEDURE) IS AS FOLLOWS: -0.25 -0.75 AXIS 70 WITH UCVA OF 20/25. PATIENT IS HAPPY". (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.50/5.5/091 | OS | Physician | 08/29/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8052105&pc=MTA | 2023826-2018-01626 | 2018-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -7.50/1.5/014 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SUPRISE WAS OBSERVED, THE SURGEON DECIDED TO DO A REPOSITIONING OF THE LENS (19 DEGREES CCW). POST SURGERY RESULTS ARE PENDING AT THE TIME OF THIS REPORT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. PATENT IS FINE AT THE MOMENT, STRUGGLING WITH RX DUE TO REFRACTIVE SUPRISE. IT WAS REPORTED THAT THE CAUSE OF EVENT WAS USER ERROR. SURGEON USED THE WRONG AXIS OF THE RX. Manufacturer Narrative: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 -7.50/1.5/014 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) -2018. REFRACTIVE SURPRISE WAS OBSERVED. IT WAS REPORTED THAT THE SURGEON USED THE WRONG AXIS OF THE RX. ON (B)(6) 2018, REPORTER STATED: "PATIENT CONDITIONS ARE NOT BAD, BUT THE UCVA IS ONLY 20/25 BECAUSE OF THE REFRACTIVE SURPRISE". ON (B)(6) 2018 IT WAS STATED: "PATIENT IS FINE-AT THE MOMENT, HE'S ONLY A BIT STRUGGLING WITH HIS RX ON OS (BECAUSE OF REFRACTIVE SURPRISE)". THE SURGEON PLANS TO DO A REPOSITIONING OF THE LENS (19 DEGREES CCW). IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. (B)(4). "THE ROTATION HAS NOT BEEN DONE DUE TIME PROBLEMS OF THE PATIENT, YET IT IS PLANNED FOR MID OF DECEMBER". IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE LENS WAS REPOSITIONED ON (B)(6) 2019, "STABLE REFRACTION AFTER REPOSITIONING." AND THE PROBLEM WAS RESOLVED. PATIENT CODE ADDITIONAL INFORMATION: 3191 - SECONDARY SURGICAL INTERVENTION, REPOSITIONING. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.50/1.5/014 | OS |  | 10/09/2018 | 2018 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7974926&pc=MTA | 2023826-2018-01528 | 2018-04-11 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/+2.0/081 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 AND (B)(6) 2018 THE LENS WAS RE-POSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND BLURRED VISION. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS. REPORTEDLY, THE LENS EXCHANGE RESOLVED THE PROBLEM. THE SURGEON STATES "SUSPECTED ZONULAR WEAKNESS THAT COULD CAUSED LENS ROTATION MULTIPLE TIMES." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+2.0/081 | OS | Physician | 09/21/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7974933&pc=MTA | 2023826-2018-01527 | 2018-04-11 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.0/+1.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 AND (B)(6) 2018 THE LENS WAS RE-POSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND BLURRED VISION. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS. REPORTEDLY, THE LENS EXCHANGE RESOLVED THE PROBLEM. THE SURGEON STATES "SUSPECTED ZONULAR WEAKNESS THAT COULD CAUSED LENS ROTATION MULTIPLE TIMES." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND A TEAR IN THE HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.5/086 | OD | Physician | 09/21/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7710867&pc=MTA | 2023826-2018-01071 | 2018-05-18 | Injury | STAAR SURGICAL COMPANY | 2018-07-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0/+6.0/007 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS REPORTED AS HAVING A REFRACTIVE SURPRISE AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER STATED THAT THE "DOCTOR DECIDED TO ROTATE THE LENS BUT AFTER SIX DAYS THE LENS RETURNED TO ITS INITIAL POSITION OBTAINING THE SAME REFRACTIVE SURPRISE. THE DOCTOR WILL ORDER THE SAME LENS WITH THE SAME POWER BUT A BIGGER SIZE BECAUSE THE VAULT AFTER THESE DAYS HAS DECREASED." (B)(4). Manufacturer Narrative: ADVERSE EVENT" NOT INCLUDED IN THE SUPPLEMENTAL MDR#2. 'REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICE)" NOT INCLUDED IN THE SUPPLEMENTAL MDR#2. (B)(4). Manufacturer Narrative: LENS WAS EXCHANGED FOR A SIMILAR POWER AND A LARGER SIZE ON (B)(6) 2018. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.0/+6.0/007 | OD | Physician | 06/25/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7851176&pc=MTA | 2023826-2018-01292 | 2018-07-24 | Injury | STAAR SURGICAL COMPANY | 2018-09-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1 MM, VTICM5_12.1, -13.00+/4.5/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2018, LENS REMAINS IMPLANTED, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00+/4.5/91 | OS | Physician | 08/07/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8182343&pc=MTA | 2023826-2018-01887 | 2018-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -0.50/+6.0/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED AND THE PROBLEM WAS RESOLVED. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE "LENS WAS NOT STABLE" ON (B)(6) 2019. ON (B)(6) 2019 THE REPORTER STATED, "THE SURGEON TRIED TO CORRECT THE REFRACTIVE SURPRISE WITH A REPOSITIONING OF BOTH LENSES. FOR THIS HE ROTATED THE MISALIGNED TICLS TO HE AXIS OF THE IODS. AFTER THAT (1DAY POSTOP) FIRST THE REFRACTIONS WERE GOOD AND EVERYTHING WAS FINE. AT THE NEXT EXAMINATION MID OF (B)(6) THE TICLS HAVE ROTATED AGAIN TO THE SAME MISALIGNED AXIS AS BEFORE. IN THIS POSITION THE LENSES WERE NOW STABLE AND SO THE SURGEON DECIDED TO MAKE A LASER TOUCH UP TO CORRECT THE RESIDUAL REFRACTION ON BOTH EYES". PATIENT CODE: 3191 SECONDARY SURGICAL INTERVENTION, LASER TOUCH UP. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: DEVICE CODE: (B)(4). NOT INCLUDED IN THE SUPPLEMENTAL MDR#2 CLAIM #(B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE PATIENT HAD DECIDED NOT TO HAVE A LVC PROCEDURE TO CORRECT CURRENT REFRACTIVE ERROR AFTER TICL IMPLANTATION. THE PLAN IS TO EXPLANT THE LENS. CORRECTED DATA: PATIENT CODE: 3191 - (SECONDARY SURGICAL INTERVENTION, LASER TOUCH UP) - PREVIOUS CODE NOT APPLICABLE, PATIENT DECIDED NOT TO HAVE LASER TOUCH UP (LVC). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -0.50/+6.0/098 | OS |  | 11/26/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8272716&pc=MTA | 2023826-2019-00152 |  | Injury | STAAR SURGICAL COMPANY | 2019-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.00/+2.0/091 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD ROTATED AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND LOSS OF BEST-CORRECTED VISUAL ACUITY (BCVA). THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED, THE LENS ROTATED AGAIN. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO LENS ROTATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.00/+2.0/091 | OD | Physician | 12/27/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6725096&pc=MTA | 2023826-2017-01217 | 2017-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2017-07-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -1.0/+6.0/85 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND LOSS OF BEST CORRECTED VISUAL ACUITY (BCVA). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: CORRECTED DATA: CONCLUSION: PREVIOUS CODE IS NOT APPLICABLE, THIS LENS MODEL IS CONTRAINDICATED IN THE PRESENCE OF ANY OF THE FOLLOWING CIRCUMSTANCES AND/OR CONDITIONS SUCH AS FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: IT WAS REPORTED ON (B)(6) 2019 THE LENS WAS EXPLANTED. (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -1.0/+6.0/85 | OS |  | 06/28/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7589617&pc=MTA | 2023826-2018-00878 | 2018-04-25 | Injury | STAAR SURGICAL COMPANY | 2018-06-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+4.0/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTION (FUNCTIONAL) INSPECTION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2018. THE LENS WAS EXCHANGED FOR A SAME MODEL/DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/+4.0/074 | OD | Physician | 05/14/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8227439&pc=MTA | 2023826-2019-00056 | 2009-03-13 | Injury | STAAR SURGICAL COMPANY | 2019-01-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/029 (SPHERE/CYLINDAR/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2009. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED INFORMATION: ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS. THE PROBLEM WAS RESOLVED. "PATIENT IS HAPPY AND EXCHANGE WENT WELL!" (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+1.0/029 | OS | Physician | 12/14/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7773875&pc=MTA | 2023826-2018-01187 | 2016-07-25 | Injury | STAAR SURGICAL COMPANY | 2018-08-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM, VTICMO13.2, -13.0/+5.5/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 REPOSITIONING WAS PERFORMED DUE TO MISALIGNMENT. ONE DAY AFTER REPOSITIONING UCVA/BCVA WAS 20/40 AND 20/20, IOP WAS 15 MMHG, VAULT- 902 MICRONS. RX WAS THE SAME AS BEFORE MANIPULATION. IT WAS REPORTED THAT THE SURGEON PLANS TO REPLACE THE ICL DUE TO THE RESIDUAL REFRACTIVE ERROR AND HIGH VAULT. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.0/+5.5/89 | OS | Physician | 07/11/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8021583&pc=MTA | 2023826-2018-01572 | 2018-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-10-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.50/+1.0/107 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS INTRAOPERATIVELY ON (B)(6) 2018. THE "VAULT WAS GREATLY IMPROVED" AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND BLUE RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.50/+1.0/107 | OD |  | 10/01/2018 | 2018 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7790898&pc=MTA | 2023826-2018-01237 | 2018-07-09 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_12.6,-5.5/+1.5/87 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. EXCESSIVE VAULT, REFRACTIVE SURPRISE, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS OBSERVED. THE LENS REMAINS IMPLANTED AND THE SURGEON IS NOT SURE IF THE LENS WILL BE EXPLANTED OR EXCHANGED, THE SURGEON WILL BE MONITORING OVER THE NEXT MONTHS. PATIENT'S CONDITION REMAINS THE SAME. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.5/+1.5/87 | OD | Physician | 07/17/2018 | 2018 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7765084&pc=MTA | 2023826-2018-01171 | 2018-07-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -3.0/+1.5/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON NOTES LOW VAULT AND REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A LENS VIAL DRY WITH CLEAR SURGICAL RESIDUE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "THE LENS WAS REMOVED ON AN UNKNOWN DATE. ON (B)(6) 2019 A LONGER LENGRTH REPLACEMENT LENS WAS IMPLANTED. THE REPLACEMENT LENS RESOLVED THE PROBLEM." SHOULD BE ADDED TO THE INITIAL MDR. H6- PATIENT CODE 3191: SECONDARY SURGERY; LENS EXPLANT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -3.0/+1.5/082 | OD |  | 07/17/2018 | 2018 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6711588&pc=MTA | 2023826-2017-01176 | 2015-09-07 | Injury | STAAR SURGICAL COMPANY | 2017-07-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -15.5/+2.0/65 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THE LENS WAS REPOSITIONED ON(B)(6)2015. THE PROBLEM WAS RESOLVED, THE LENS IS IN POSITION AT DESIRED AXIS. AS OF THE DATE OF MDR SUBMISSION THE LENS REMAINS IMPLANTED, AND WILL NOT RETURN. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO_13.2 -15.5/+2.0/065 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS REPOSITIONED ON (B)(6) 2015 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. ON (B)(6) 2018 IT WAS REPORTED THAT REFRACTIVE SURPRISE AND LOW VAULT WITH ROTATION WERE OBSERVED IN THE PATIENT; THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2018 WHICH RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BENT AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN US BUT NOT MARKETED IN THE U.S. (B)(4). | -15.5/+2.0/65 | OS | Physician | 06/22/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8194494&pc=MTA | 2023826-2018-01903 | 2018-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+3.5/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS REPOSITIONED. THE SURGEON REPORTS LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: EVENT- LENS WAS EXPLANTED. EXPLANT DAT- UNK. PATIENT CODE 3191: NO CODE AVAILABLE - SECONDARY SURGERY, EXPLANT. Manufacturer Narrative: PRODUCT EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM #: (B)(4). Manufacturer Narrative: DESCRIBE EVENT OR PROBLEM: ADDITIONAL INFORMATION: ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS. THE PROBLEM IS RESOLVED, THE PATIENT IS "VERY HAPPY!" DISTRIBUTOR PREVIOUSLY CHECKED IN ERROR. PATIENT CODE 3191 PREVIOUSLY STATED LENS EXPLANT UPDATED, 3191 SHOULD STATE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+3.5/097 | OD |  | 11/30/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8227441&pc=MTA | 2023826-2019-00057 | 2018-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.0/+5.0/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE DUE TO THE ROTATION OF THE LENS. ON (B)(6) 2018 THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. REPORTEDLY, THE LENS RE-ROTATED AFTER REPOSITIONING, THE SURGEON STATING "NON-STABLE SULCUS POSITION (TOO SHORT ICL)." Manufacturer Narrative: ADDITIONAL INFORMATION: REFER TO MFR REPORT # 2023826-2019-01063 FOR CLAIM# (B)(4). FOR FURTHER INVESTIGATION DETAILS. (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). ADDITIONAL INFORMATION: PROBLEM IS NOT RESOLVED DUE TO THERE STILL BEING REFRACTIVE SURPRISE. (B)(4). Manufacturer Narrative: PREVIOUSLY STATED LENS WAS EXPLANTED AT A LATER DATE. UPDATED INFORMATION: LENS WAS EXCHANGED ON 15-JAN-2019 WITH A LONGER, DIFFERENT MODEL LENS. PROBLEM IS RESOLVED. DISTRIBUTOR WAS CHECKED IN ERROR. PREVIOUSLY STATED LENS EXPLANT. UPDATED-LENS EXCHANGE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 -11.0/+5.0/080 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE DUE TO THE ROTATION OF THE LENS. ON (B)(6) 2018 THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. REPORTEDLY, THE LENS RE-ROTATED AFTER REPOSITIONING, THE SURGEON STATING "NON-STABLE SULCUS POSITION (TOO SHORT ICL)." THE LENS WAS EXPLANTED ON A LATER DATE. ADDITIONAL INFORMATION REGARDING THE PATIENT STATUS AND IF PROBLEM WAS RESOLVED HAS BEEN REQUEST. IF ADDITIONAL DETAILS ARE AVAILABLE A SUPPLEMENTAL MDR WILL BE SUBMITTED. DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN AND RESIDUE ON LENS SURFACE. PATIENT CODE- 3191: SECONDARY SURGICAL INTERVENTION, LENS EXPLANT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -11.0/+5.0/080 | OS |  | 12/14/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8172154&pc=MTA | 2023826-2018-01843 | 2018-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -6.5/+2.5/173 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES REFRACTIVE SURPRISE AND LENS DISLOCATION. ON (B)(6) 2018 THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. SURGEON ALSO STATES THE VAULT IS 200UM AND PLANS TO IMPLANT A LONGER LENS. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: ON (B)(6) 2019, THE LENS WAS REPOSITIONED A 2ND TIME. REPORTEDLY, THE "PATIENT AND SURGEON ARE HAPPY WITH THE RESULTS." THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.5/+2.5/173 | OS |  | 11/23/2018 | 2018 | 51
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8172170&pc=MTA | 2023826-2018-01851 | 2018-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -18.0/+3.0/071 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO LOW VAULT, LENS ROTATION, PIGMENT DISPERSION, AND BLURRED VISION. REPORTEDLY, THE ICL HAD ROTATED TO THE VERTICAL POSITION. IT IS NOTED THAT THE PATIENT DOES NOT WANT TO UNDERGO ANOTHER OPERATION. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE STATUS OF THE EYE IS NOTED AS "CALM" AFTER THE EXPLANT. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN/BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.0/+3.0/071 | OS |  | 11/23/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7488984&pc=MTA | 2023826-2018-00677 | 2018-04-16 | Injury | STAAR SURGICAL COMPANY | 2018-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/+1.0/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND REFRACTTIVE SURPRISE. THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT AND EXCHANGE FOR ANOTHER LENS. Manufacturer Narrative: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2018 AND EXCHANGED FOR A SHORTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. (B)(4). CORRECTED DATA: PRODUCT PROBLEM TO ADVERSE EVENT MALFUNCTION TO SERIOUS INJURY (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL VERIFICATION WAS PERFORMED ON THE LENS. INVESTIGATION FOUND LENS WAS WITHIN SPECIFICATION AT THE TIME OF RELEASE. RESULTS OF RETURNED LENS FUNCTIONAL EVALUATION ARE INCONCLUSIVE; LENS MAY HAVE BEEN MISSHAPEN DURING PERIOD LENS REMAINED IMPLANTED OR DUE TO HANDLING DURING EXPLANT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. | -14.50/+1.0/100 | OS | Physician | 04/16/2018 | 2018 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7201647&pc=MTA | 2023826-2018-00091 | 2017-07-01 | Injury | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Shelf Life Exceeded; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -4.50/+1.5/055 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2017. THE LENS WAS REPOSITIONED ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A LONGER LENS, THE SURGEON NOTING LOW VAULT, LENS ROTATION, AND LOSS OF BCVA. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING AND CLEAR RESIDUE ON LENS. (B)(4) Manufacturer Narrative: PATIENT INFORMATION: UNK. THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | -4.50/+1.5/055 | OS | Physician | 12/21/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5654727&pc=MTA | 2023826-2016-00659 | 2010-12-28 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; No Information; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -8.0/+2.0/092 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2010. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND ROTATION AND REMAINS IMPLANTED. IT IS PLANNED TO EXPLANT THE LENS. Manufacturer Narrative: THE SURGEON PERFORMED A SECONDARY INTERVENTION TO REPOSITION THE LENS TO THE ORIGINAL AXIS. AN EXCHANGE IS PLANNED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. OUTCOME ATTRIBUTED TO ADVERSE EVENT: UNK. EXPLANT DATE: N/A. DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. (B)(4). LENS REMAINS IMPLANTED. | -8.0/+2.0/092 | OS | Physician | 04/27/2016 | 2016 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8219629&pc=MTA | 2023826-2019-00045 | 2018-12-05 | Injury | STAAR SURGICAL COMPANY | 2019-01-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Known Impact Or Consequence To Patient; No Code Available; Missing Value Reason | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_13.2, -9.5/+1.0/003 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. LOW VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. IT WAS REPORTED THAT THE CAUSE OF EVENT WAS A PATIENT RELATED FACTOR. SURGEON PLANS TO EXCHANGE THE LENS IN (B)(6); IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: ADVERSE EVENT SHOULD BE MARKED IN PREVIOUS MDR. REQUIRED INTERVENTION SHOULD BE MARKED IN PREVIOUS MDR. REFRACTIVE SURPRISE WAS ALSO REPORTED. THE LENS WAS EXCHANGED FOR A LONGER LENGTH, STRONGER POWER LENS. SHOULD BE ADDED TO THE INITIAL MDR. DATE RECEIVED BY MFR: SHOULD BE CORRECTED TO "OCTOBER 14,2019" IN THE PREVIOUS MDR. (B)(4). Manufacturer Narrative: CORRECTED DATA: "LENS REMAINS IMPLANTED" SHOULD BE CORRECTED TO "LENS WAS EXCHANGED ON (B)(6) 2019 FOR A LONGER LENGTH LENS. THE PROBLEM WAS RESOLVED." IN THE INITIAL MDR. "THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR"SHOULD BE CORRECTED TO " THE CAUSE OF THE EVENT WAS USER ERROR." IN THE INITIAL MDR. ADDITIONAL INFORMATION: SECONDARY SURGICAL INTERVENTION - EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARK. ETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/+1.0/003 | OS | Physician | 12/06/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6361178&pc=MTA | 2023826-2017-00262 | 2017-01-02 | Injury | STAAR SURGICAL COMPANY | 2017-02-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; Overwear Syndrome | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15/0.5/149 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER STATED EXCESSIVE VAULT, GLARE/HALOES, AND PHOTOPHOBIA WAS OBSERVED ON (B)(6) 2017. THE LENS REMAINS IMPLANTED AS OF (B)(6) 2017. Manufacturer Narrative: CORRECTED DATA: EXCESSIVE VAULT, GLARES/HALOES, AND PHOTOPHOBIA WAS OBSERVED ON (B)(6) 2017. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WORK ORDER SEARCHES: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). LENS NOT RETURNED. | -15/0.5/149 | OS | Physician | 01/31/2017 | 2017 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8021086&pc=MTA | 2023826-2018-01581 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-10-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | N/A | Event Description: IT WAS REPORTED THAT A 13.2MM, VTICMO13.2, -16.00/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE ON (B)(6) 2018. REPORTER STATES "MY ISSUE IS THAT I CAN SEE GREAT OUTDOORS BUT SUFFER FROM PATCHES INDOORS OR UNDER NON/NATURAL LIGHT. I HAVE BEEN TO MY CHECKUPS AND THEY SAY ALL IS OK. BUT THE PATCHES ARE MAKING MY VISION INDOORS EXHAUSTIVE. WE HAVE HAD A SECOND OPINION FROM ANOTHER OPTICIAN AND THE GRAYISH AND BLURRED IMAGES IS CAUSED DUE TO THE LIGHT REFRACTION ON THE LENS ESPECIALLY IN LOW LIGHT WHEN THE PUPIL IS AT ITS BIGGEST. WE BELIEVE MY WIFE'S ISSUE IS BECAUSE SHE STILL HAS ASTIGMATISM. THIS IS THE CULPRIT OF THE BLURRY IMAGES ESPECIALLY IN LOW LIGHT CONDITIONS." TO DATE THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: PATIENT IS UNDER TREATMENT WITH BRIMONIDINE (TO REDUCE HALOES) AND WILL BE CLOSELY MONITORED. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.00/2.0/090 | OS | Non-Healthcare Professional | 09/30/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8021095&pc=MTA | 2023826-2018-01582 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-10-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | N/A | Event Description: IT WAS REPORTED THAT A 13.2MM, VTICMO13.2, -16.00/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE ON (B)(6) 2018. REPORTER STATES "MY ISSUE IS THAT I CAN SEE GREAT OUTDOORS BUT SUFFER FROM PATCHES INDOORS OR UNDER NON/NATURAL LIGHT. I HAVE BEEN TO MY CHECKUPS AND THEY SAY ALL IS OK. BUT THE PATCHES ARE MAKING MY VISION INDOORS EXHAUSTIVE. WE HAVE HAD A SECOND OPINION FROM ANOTHER OPTICIAN AND THE GRAYISH AND BLURRED IMAGES IS CAUSED DUE TO THE LIGHT REFRACTION ON THE LENS ESPECIALLY IN LOW LIGHT WHEN THE PUPIL IS AT ITS BIGGEST. WE BELIEVE MY WIFE'S ISSUE IS BECAUSE SHE STILL HAS ASTIGMATISM. THIS IS THE CULPRIT OF THE BLURRY IMAGES ESPECIALLY IN LOW LIGHT CONDITIONS." TO DATE THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT IS UNDER TREATMENT WITH BRIMONIDINE (TO REDUCE HALOES) AND WILL BE CLOSELY MONITORED. EVENT: "VTICMO13.2, -16.00/2.0/90" SHOULD BE CORRECTED TO "VTICMO13.2, -16.00/2.0/94" IN THE PREVIOUS MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.00/2.0/090 | OD |  | 09/30/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16210270&pc=MTA | 2023826-2023-00247 | 2021-11-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.50/+3.0/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON 06-SEP-2021. THE PATIENT EXPERIENCED LENS ROTATION; UNREACTIVE (FIXED) PUPIL; AND IMMEDIATELY AFTER SURGERY; GLARE/HALOES AND BLURRED VISION. LENS REMAINS IMPLANTED. THE PROBLEM IS NOT RESOLVED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: UNREACTIVE (FIXED) PUPIL. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.50/+3.0/89 | OD | Physician | 12/27/2022 | 2022 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8379600&pc=MTA | 2023826-2019-00351 | 2018-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Blurred Vision; Visual Disturbances; Halo; Discomfort; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.5/+4.5/039 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE, WITH GLARE/HALOS AND BLURRED VISION. THE PATIENT HAD DISCOMFORT WITH THEIR VISION IN COMPARISON WITH THE ICL LENS IN THE RIGHT EYE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PER UPDATED INFORMATION, THE LENS WAS EXPLANTED ON (B)(6) 2019. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, LENS EXPLANT. METHOD CODE 4117 NO LONGER APPLICABLE. INITIAL MDR SHOULD HAVE INCLUDED DEVICE CODE 1494: OFF-LABEL USE, PRE-EXISTING HISTORY OF KERATOCONUS AND CORNEAL CROSS-LINKING. INITIAL MDR SHOULD HAVE INCLUDED METHOD CODE 4111. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: B5 - PER UPDATED INFORMATION, THE PATIENT IS STABLE AND HAS RETURNED TO PREOPERATIVE REFRACTION. CLAIM# (B)(4). Manufacturer Narrative: B5 - PER UPDATED INFORMATION, THE LENS WAS EXPLANTED ON (B)(6) 2019. D7 - UPDATED TO (B)(6) 2019. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+4.5/039 | OS |  | 01/31/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6876813&pc=MTA | 2023826-2017-01461 | 2016-02-18 | Injury | STAAR SURGICAL COMPANY | 2017-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -14.5/+2.0/071 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE PATIENT BEGAN EXPERIENCING GLARE/HALOES AND BLURRED VISION. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A SHORTER AND SIMILAR DIOPTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DATE OF THIS REPORT: IT IS CORRECTED TO "(B)(6) 2017" FROM "(B)(6) 2017" IN THE INITIAL. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -14.5/+2.0/071 | OD | Physician | 08/16/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6752257&pc=MTA | 2023826-2017-01144 | 2016-12-24 | Injury | STAAR SURGICAL COMPANY | 2017-07-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -7.5/+1.0/92 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO GLARE/HALOES, BLURRED VISION, AND FOREIGN BODY SENSATION. THERE IS NO PLANS TO EXCHANGE FOR A NEW LENS. ON (B)(6) 2017, PATIENT'S BCVA WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.5/+1.0/92 | OS | Physician | 07/03/2017 | 2017 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6752269&pc=MTA | 2023826-2017-01143 | 2016-12-24 | Injury | STAAR SURGICAL COMPANY | 2017-07-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -9/+2/88 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO GLARE/HALOES, BLURRED VISION, AND FOREIGN BODY SENSATION. THERE IS NO PLANS TO EXCHANGE FOR A NEW LENS. ON (B)(6) 2017, PATIENT'S BCVA WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THE LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9/+2/88 | OD | Physician | 07/03/2017 | 2017 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135224&pc=MTA | 2023826-2017-01924 | 2017-10-26 | Injury | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -5.0/+4.0/096 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE, BLURRED VISION, GLARE/HALOES. REPORTEDLY, THE LENS WAS REPOSITIONED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THE REPORTER STATED THE LENS WAS EXPLANTED ON (B)(6) 2018. (B)(4) Manufacturer Narrative: THE REPORTER STATES THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2018. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGERY, LENS REPOSITIONING H6-LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.0/+4.0/096 | OS | Physician | 11/27/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320740&pc=MTA | 2023826-2018-00322 | 2016-12-01 | Injury | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; Visual Disturbances; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -4.5/+1.5/180 (SPHERE/ CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED. REPORTEDLY, THE PATIENT EXPERIENCED "PROBLEMS READING AND DOING COMPUTER WORK." THE SURGEON STATES THE EYE IS "CALM AND QUIET" AFTER THE EXCHANGE AND THAT THE EYE IS 0.5 MORE MYOPIC THAN EXPECTED. THE SURGEON ALSO REPORTS THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND HAPTIC BROKEN/BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND "PROBLEMS READING AND DOING COMPUTER WORK." (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -4.5/+1.5/180 | OD | Physician | 02/08/2018 | 2018 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6520730&pc=MTA | 2023826-2017-00707 | 2015-05-12 | Injury | STAAR SURGICAL COMPANY | 2017-04-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Visual Disturbances; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -7.0/+1.5/089 IMPLANTABLE COLLAMER LENS (ICL), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT BEGAN TO EXPERIENCE LOW VAULT, REFRACTIVE SURPRISE, AND REFRACTIVE CHANGE OVER TIME. THE SURGEON STATED THAT THE ICL IS ON THE AXIS. TO DATE, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: STATED LENS REMAINS IMPLANTED. LENS WAS EXCHANGED ON (B)(6) 2017 FOR A LARGER, DIFFERENT DIOPTER LENS. THE PROBLEM WAS RESOLVED. FINAL POST-OP RESULTS: VAULT 500UM, UCVA 20/20. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT, FOREIGN MATERIAL ON LENS SURFACE (CLEAR AND RED RESIDUE), AND A PIECE OF HAPTIC MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE REMAINS IMPLANTED. | -7.0/+1.5/089 | OD | Physician | 03/28/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6461375&pc=MTA | 2023826-2017-00573 | 2015-10-24 | Injury | STAAR SURGICAL COMPANY | 2017-04-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -16.5/+5.5/088 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO THE PATIENT EXPERIENCING LOW VAULT AND LENS ROTATION. IN ADDITION, THE PATIENT SUFFERED SIGNIFICANT VISION LOSS AND GLARE WHILE DRIVING AT NIGHT THE PROBLEM WAS RESOLVED WITH THE EXCHANGE. AS OF THE DATE OF MDR REPORTING THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.5/+5.5/088 | OD | Physician | 03/08/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6211496&pc=MTA | 2023826-2016-01841 | 2016-10-06 | Injury | STAAR SURGICAL COMPANY | 2016-12-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; Visual Impairment; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12/03/78 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO REFRACTIVE SURPRISE AND LOSS OF BCVA (BEST CORRECTED VISUAL ACUITY). THE LENS WAS EXCHANGED FOR ANOTHER LENS VTICM 13.2 MM -12/03/79 DIOPTER. THE PROBLEM WAS RESOLVED AFTER THE LENS WAS EXCHANGED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION AND FUNCTIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). THE LENS WAS EXCHANGED FOR A 13.2 MM VTICM5_13.2, NOT A VTICM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12/03/78 | OS | Physician | 11/28/2016 | 2016 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4629094&pc=MTA | 2023826-2015-00368 |  | Injury | STAAR SURGICAL COMPANY | 2015-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Visual Impairment; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS). THERE WAS A REFRACTIVE SURPRISE AND THE VISUAL ACUITY WAS NOT GOOD ENOUGH. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE FACILITY REPORTED THE LENS WAS IMPLANTED ON (B)(6) 2014. SEVEN MONTHS AFTER THE INITIAL SURGERY, THE SURGEON DECIDED TO REPOSITION THE LENS 30 CW TO THE TARGET POSITION OF 3 CCW. HOWEVER, THE LENS ROTATED AGAIN TO THE SAME POSITION. THE SURGEON IS PLANNING TO REMOVE AND REPLACE THE LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: LOW VAULT WAS REPORTED. LENS WAS EXPLANTED SOME TIME AFTER REPOSITIONING THE LENS. DEVICE EVALUATION: AN EVALUATION WAS PERFORMED; ADDITIONAL INVESTIGATION FOUND THE DEVICE WAS IN SPECIFICATION AT TIME OF RELEASE. (B)(4). AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM AS WELL AS PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, RUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY IS CONTRAINDICATED. (B)(4). Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS IMPLANTED. Manufacturer Narrative: EVALUATION METHOD: DEVICE HISTORY RECORD REVIEW RESULTS: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. SEVERAL FACTORS MAY BE INVOLVED IN A POST-OPERATIVE REFRACTIVE SURPRISE, INCLUDING INAPPROPRIATE REFRACTIVE SURGICAL PLANNING, A WRONG LENS CALCULATION/SELECTION, PREOPERATIVE INACCURATE DETERMINATION OF THE EYE'S MEASUREMENTS TO DETERMINE THE POWER OF THE ICL, INDUCED REFRACTIVE CHANGE BY INCISIONS, UNSTABLE CORNEA AND/OR REFRACTION PRIOR TO IMPLANTATION, CL-INDUCED WARPAGE, UPSIDE-DOWN LENS, WRONG LENS, MISALIGNED LENS, ROTATION, TILTING, ETC. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CLINICAL STUDIES HAVE SHOWN THAT TORIC ICL ROTATION OCCURS IN 0.5% OF PATIENTS WHERE A TICL HAS BEEN IMPLANTED. SEVERAL FACTORS MAY CONTRIBUTE TO THIS PHENOMENON (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -15.0/+2.5/87 | OS | Other | 03/02/2015 | 2015 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5495703&pc=MTA | 2023826-2016-00263 | 2015-11-05 | Injury | STAAR SURGICAL COMPANY | 2016-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Visual Impairment; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+3.5/118 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE LENS ROTATING AND THE PATIENT EXPERIENCED LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS EXCHANGED FOR A LONGER SPHERICAL LENS WITH ANOTHER DIOPTER AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -18.0/+3.5/118 | OD | Physician | 01/13/2016 | 2016 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5496053&pc=MTA | 2023826-2016-00264 | 2015-11-05 | Injury | STAAR SURGICAL COMPANY | 2016-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Visual Impairment; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+3.5/093 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE LENS ROTATING AND THE PATIENT EXPERIENCED LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS EXCHANGED FOR A LONGER SPHERICAL LENS WITH ANOTHER DIOPTER AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -18.0/+3.5/093 | OS | Physician | 01/13/2016 | 2016 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6522593&pc=MTA | 2023826-2017-00710 | 2017-03-20 | Injury | STAAR SURGICAL COMPANY | 2017-04-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, DIOPTER -11.0/1.0/091 IMPLANTABLE COLLAMER LENS (ICL), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT BEGAN TO EXPERIENCE LOW VAULT AND LENS OPACITY. TO DATE, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(6). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE REMAINS IMPLANTED. | -11.0/1.0/091 | OD | Physician | 03/28/2017 | 2017 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6534124&pc=MTA | 2023826-2017-00711 | 2017-03-20 | Injury | STAAR SURGICAL COMPANY | 2017-05-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, DIOPTER -10.0/2.0/093 IMPLANTABLE COLLAMER LENS (ICL), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT BEGAN TO EXPERIENCE LOW VAULT AND LENS OPACITY. TO DATE, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. CONCLUSION CODE: OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT ((B)(6)), ACD < 3.0MM (2.9MM)]. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE REMAINS IMPLANTED. | -10.0/2.0/093 | OS | Physician | 03/28/2017 | 2017 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873717&pc=MTA | 2023826-2016-01069 | 2015-07-12 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Cataract, Induced; Blurred Vision; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -15.50/+1.50/172 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2004. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (ANTERIOR/POSTERIOR SUBCAPSULAR AND NUCLEAR), WHICH WAS NOTED ON (B)(6) 2015. THE PATIENT EXPERIENCED BLURRED VISION, GLARE AND HALOS. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. AT LAST POST-OP VISIT ON (B)(6) 2016, THE PATIENT'S BEST-CORRECTED VISUAL ACUITY WAS 20/50 +1. Manufacturer Narrative: CONCLUSION: THE PRE-EXISTING "GRAINY PERIPHERAL CORTICAL CLOUDING" WAS LISTED ON THE FORM AND THAT CONDITION MIGHT HAVE CAUSED THE CATARACT DEVELOPMENT SINCE IT IS WELL KNOWN THAT OPACITIES START AS VACUOLES AND CLEFTS BETWEEN THE LENS FIBERS AND PROGRESS TOWARDS ANTERIOR, POSTERIOR OR EQUATORIAL CORTEX. IT SHOULD BE NOTED THAT AT THE TIME OF SURGERY THE PATIENT WAS (B)(6) AND PER DFU INDICATIONS: "VISIAN® IMPLANTABLE COLLAMER® LENSES (VICL) ARE INDICATED FOR USE IN PHAKIC EYE TREATMENTS IN ADULTS 21-45 YEARS OF AGE", AND THEREFORE, THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN PATIENTS. GIVEN ALL THIS INFORMATION, OBSERVED LENS OPACITY AND SUBSEQUENT CATARACT DEVELOPMENT WAS AGE RELATED (PATIENT FACTOR) AND NOT CAUSED BY THE DEVICE. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC BROKEN OFF. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON THE PRODUCT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | -15.50/+1.50/172 | OD | Physician | 07/14/2016 | 2016 | 49
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873711&pc=MTA | 2023826-2016-01070 | 2004-05-19 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract, Induced; Blurred Vision; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -14.50/+3.75/167 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2004. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (POSTERIOR SUBCAPSULAR AND NUCLEAR), WHICH WAS NOTED ON (B)(6) 2015. THE PATIENT EXPERIENCED BLURRED VISION, GLARE AND HALOS. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. AT LAST POST-OP VISIT ON (B)(6) 2016, THE PATIENT'S BEST-CORRECTED VISUAL ACUITY WAS 20/30 -2. Manufacturer Narrative: CORRECTED DATA: (B)(4). ADDITIONAL DATA: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -14.50/+3.75/167 | OS | Physician | 07/14/2016 | 2016 | 49
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6955552&pc=MTA | 2023826-2017-01577 | 2017-04-13 | Injury | STAAR SURGICAL COMPANY | 2017-10-17 | MTA | IMPLANTABLE COLLAMER LENS | Activation Failure | Cataract, Induced; Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT AFTER THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -10.5/+2.0/097 INTO THE PATIENT'S LEFT EYE (OS), THE LENS WOULD NOT PROPERLY UNFOLD. THE SURGEON STATES THAT, "AFTER ALL POSSIBLE MANEUVERS" THE LENS DIDN'T UNFOLD PROPERLY AND THE LENS WAS EXPLANTED. DURING EXPLANT THE LENS WAS SCRATCHED. THIS OCCURRED ON (B)(6) 2017. REPORTEDLY, THE SURGEON THEN PEFORMED PHACO-EMULSIFICATION ON THE EYE AND IMPLANTED AN EXCHANGE IOL ON (B)(6) 2017 "WITHOUT COMPLICATIONS." Manufacturer Narrative: BEST CORRECTED VISUAL ACUITY (BCVA), ON (B)(6) 2017, WAS 20/30 WHICH WAS A DECREASE FROM BASELINE BCVA (20/25). (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4)-NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, PHACO-EMULSIFICATION). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/+2.0/097 | OS | Physician | 06/06/2017 | 2017 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6176008&pc=MTA | 2023826-2016-01763 | 2014-09-17 | Injury | STAAR SURGICAL COMPANY AG | 2016-12-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Cataract, Induced; Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -14.0/+1.50/042 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR), WHICH WAS NOTED ON (B)(6) 2014. THE PATIENT EXPERIENCED BLURRED VISION AND CLOUDY VISUAL ACUITY. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. AT THE LAST POST-OP VISIT ON (B)(6) 2016, THE PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/25 +2. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON THE PRODUCT. THE LENS WAS REMEASURED AND WAS FOUND TO BE IN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4). | -14.0/+1.50/042 | OD | Physician | 07/14/2016 | 2016 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6608173&pc=MTA | 2023826-2017-00900 | 2017-03-20 | Injury | STAAR SURGICAL COMPANY | 2017-06-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract, Induced; Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+1.0/043 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, UNREACTIVE PUPIL (FIXED), ANGLE CLOSURE, AND IRIS ATROPHY. ON (B)(6) 2017 THE PATIENT ALSO EXPERIENCED LENS OPACITY ASCO. THE SURGEON PLANS TO EXCHANGE TO A SHORTER LENS FOR THE PATIENT. AS OF THE DAY OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. Manufacturer Narrative: AN ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS WAS PERFORMED. (B)(4). Manufacturer Narrative: (B)(6). THE LENS WAS EXCHANGED ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). H6_WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.5/+1.0/043 | OS | Physician | 05/04/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6684194&pc=MTA | 2023826-2017-01086 | 2016-10-15 | Injury | STAAR SURGICAL COMPANY | 2017-07-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Cataract, Induced; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/93 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ACCORDING TO THE DR. THAT BECAUSE THE SURGERY OPERATING TIME WAS TOO LONG, THE SUSPENSORY LIGAMENT NEAR THE LENS CAPSULE MEMBRANE DOT LIGHT LENS OPACITY. THE LENS WAS EXPLANTED IN THE SAME SURGERY, AND THE PHACO-EMULSIFICATION (CATARACT SURGERY) WAS PERFORMED, NO IOL LENS WAS IMPLANTED. THE CUSTOMER STATED THAT THE SITUATION IS STABLE, AND OBSERVATION ON REGULAR BASIC. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE : OFF LABEL USE (ACD<3.0MM) (B)(4). | -9.0/+1.0/93 | OD | Physician | 06/09/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5711853&pc=MTA | 2023826-2016-00720 | 2016-05-08 | Injury | STAAR SURGICAL COMPANY | 2016-06-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+0.5/151 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED LENS OPACITY (ANTERIOR SUBCAPSULAR OPACITY - ASCO) ON (B)(6) 2016 AND UNREACTIVE PUPIL (FIXED). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WEIGHT: UNK. OUTCOMES ATTRIBUTED TO ADVERSE EVENTS: UNK. EXPLANT DATE: N/A. (B)(4) LENS REMAINS IMPLANTED. | -12.00/+0.5/151 | OS | Physician | 05/11/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16474967&pc=MTA | 2023826-2023-00725 | 2022-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -4.0/1.0/99 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT FOR A PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THIS WAS DUE TO A EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE PROBLEM WAS RESOLVED. REPORTEDLY, "LENS OPACITY (MYDRIASIS WAS NOT MAXIMAL AT THE TIME OF ICL REMOVAL + ICL SIZE WAS 12.6, WHICH WAS LARGE AND TOOK TIME AND EFFORT TO REMOVE. ALSO, THE LENS CLOUDY DUE TO THE HIGH VISCOSITY OF THE LENS WHEN USING THE SHELL GUN). THE PATIENT EYE IS STILL UNSTABLE POSTOPERATIVELY AND WE ARE STILL DISCUSSING THE FUTURE WITH THE PATIENT." Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.0/+1.0/99 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) AS A REPLACEMENT LENS ON (B)(6) 2022. THE PATIENT EXPERIENCED EARLY LENS OPACITY. REPORTEDLY "LENS OPACITY (MYDRIASIS WAS NOT MAXIMAL AT THE TIME OF ICL REMOVAL + ICL SIZE WAS 13.2, WHICH WAS LARGE AND TOOK TIME AND EFFORT TO REMOVE. ALSO, THE LENS BECAME CLOUDY DUE TO THE HIGH VISCOSITY OF THE LENS WHEN USING THE SHELL GUN). THE PATIENT EYE IS STILL UNSTABLE POSTOPERATIVELY AND WE ARE STILL DISCUSSING THE FUTURE WITH THE PATIENT". THE LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: ADDITONAL INFORMATION: H6- TYPE OF INVESTIGATION CODE; 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | -4.0/1.0/99 | OD | Physician | 01/31/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6721163&pc=MTA | 2023826-2017-01196 | 2017-06-07 | Injury | STAAR SURGICAL COMPANY | 2017-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Cataract; Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/4.5/087 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LENS ROTATION NOT ASSOCIATED WITH LOW VAULT, BLURRED VISION, AND LENS REPOSITIONING. LENS WAS "REMOVED AND CATARACT PROCEDURE HAVE BEEN DONE DUE TO CHANGES ON NATURAL LENS". THE PATIENT'S POST-OP CONDITION IS "WITH APHAKIC LENS AND HIS VISION IS 6/12". Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT, FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE), AND THE LENS MISSING PIECE OF HAPTIC. THE PATIENT'S POST-OP VISION IS 6/12. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S, BUT NOT MARKETED IN THE U.S (B)(4). CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | -16.5/4.5/087 | OS | Physician | 07/01/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6874039&pc=MTA | 2023826-2017-01447 | 2015-08-25 | Injury | STAAR SURGICAL COMPANY | 2017-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANT A 13.2 MM VTICMO13.2, -7.0/+4.5/107 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. ON (B)(6) 2015, THE LENS WAS EXPLANTED DUE TO POST TRAUMATIC DISLOCATION IN A/C AND TRAUMATIC CATARACT DUE TO FIREWORKS MUCH LATER THAN THE IMPLANTATION. IT HAD NO CONNECTION TO THE SURGERY. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND PRESENCE OF FIBERS ON LENS SURFACE. THE PATIENT ALSO HAD BLURRED VISION. (B)(4). Manufacturer Narrative: (B)(4). CONCLUSION CODE:TRAUMA CAUSED BY FIRWORK. NO CONNECTION TO SURGERY. (B)(4). | -7.0/+4.5/107 | OD | Physician | 08/22/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6681259&pc=MTA | 2023826-2017-01089 | 2016-08-08 | Injury | STAAR SURGICAL COMPANY AG | 2017-06-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Cataract; Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -2.0/+1.5/40 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO REFRACTIVE SURPRISE AND THE DEVELOPMENT OF A CATARACT. A REPOSITIONING OF THE LENS, ENLARGEMENT OF THE INCISION (IN ORDER TO PERFORM CATARACT SURGERY), PHACO-EMULSIFICATION, AND IOL IMPLANTATION WAS PERFORMED. THE PATIENT'S POST OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND THE PATIENT'S POST OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/25 WITH STATUS OF THE EYE BEING PSEUDOPHAKIC. Manufacturer Narrative: THE PRODUCT WAS MANUFACTURED IN (B)(4) , NOT IN THE U.S. Manufacturer Narrative: "HOSPITALIZATION-INITIAL OR PROLONGED" WAS CHECKED BY MISTAKE. THERE WAS NO LENS REPOSITIONING. (B)(4). Manufacturer Narrative: THE PRODUCT WAS MANUFACTURED IN (B)(4) , NOT IN THE U.S. | -2.0/+1.5/40 | OD | Physician | 06/06/2017 | 2017 | 57
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5762131&pc=MTA | 2023826-2016-00817 | 2016-02-09 | Injury | STAAR SURGICAL COMPANY | 2016-06-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -11.00/ +6.0/100 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (CORTICAL) BLURRED VISION, THE CATARACT WAS REMOVED, AND A TORIC IOL WAS IMPLANTED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.00/ +6.0/100 | OD | Physician | 05/27/2016 | 2016 | 47
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4513549&pc=MTA | 2023826-2015-00149 | 2014-11-06 | Injury | STAAR SURGICAL COMPANY AG | 2015-02-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. ON (B)(6) 2014 LOW VAULT AND LENS OPACITY WAS OBSERVED AND THE PATIENT EXPERIENCED BLURRED VISON. THE LENS WAS EXCHANGED FOR A LONGER LENGTH WHICH RESOLVED THE VAULTING ISSUE, HOWEVER, THE CATARACT PROGRESED. THE SECOND LENS WAS REMOVED, THE CATARACT WAS EXTRACTED AND AN IOL WAS IMPLANTED. THE PATIENT'S LAST UCVA WAS 20/60. Manufacturer Narrative: .(B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS NOT RETURNED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED DRY WITH A TORN HAPTIC. THERE WAS EVIDENCE OF A CLEAR SURGICAL RESIDUE. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). IN CASES OF POSTOPERATIVE LOW VAULTING STAAR RECOMMENDS THE ICL TO BE LEFT IMPLANTED IF NO CATARACTS OR ROTATION ARE NOTED OR IN CASES WHERE THERE IS A TRACE ANTERIOR SUB-CAPUSLAR CATARACT WITH NO PROGRESSION. STAAR BELIEVES THAT THE ACTION TO REMOVE AND/REPLACE THE ICL INCREASE THE RISK FOR ANTERIOR SUB-CAPSULAR CATARACT. BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. THE MOST PROBABLE CAUSE OF THIS EVENT IS LOW VAULT SECONDARY TO A SHORT LENS. (B)(4). | -4.50/+5.5003 | OD | Other | 07/15/2015 | 2015 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6247963&pc=MTA | 2023826-2017-00071 | 2011-07-02 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -16.00/+1.50/094 DIOPTER, IN THE PATIENT'S LEFT EYE. THE LENS WAS EXPLANTED DUE TO LOW VAULTING, REFRACTIVE SURPRISE, LENS OPACITY, SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNTING OR SIGNIFICANT ENDOTHELIAL CELL LOSS, GLARE/HALOES, AND BLURRED VISION. THE PATIENT UNDERWENT PHACO-EMULSIFICATION AND IOL IMPLANTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -16.00/+1.50/094 | OS | Physician | 11/16/2015 | 2015
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6129538&pc=MTA | 2023826-2016-01682 | 2016-03-03 | Injury | STAAR SURGICAL COMPANY AG | 2016-11-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Cataract; Blurred Vision; Visual Impairment; Visual Disturbances; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, 12.5/+2.00/03 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS ROTATION (NOT ASSOCIATED WITH LOW VAULTING), LENS OPACITY, LENS DISLOCATION/SUBLUXATION, REFRACTIVE CHANGE OVER TIME AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND SURGEON REPOSITIONED THE NEW LENS. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, 12.5/+2.00/03 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS ROTATION (NOT ASSOCIATED WITH LOW VAULTING), LENS OPACITY, LENS DISLOCATION/SUBLUXATION, REFRACTIVE CHANGE OVER TIME AND BLURRED VISION. THE SURGEON EXCHANGED THE LENS FOR A LONGER LENS AND REPOSITIONED THE LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/30. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/2.0/103 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, LENS OPACITY, LENS DISLOCATION/SUBLUXATION, AND BLURRED VISION. REPOSITIONING OF THE LENS WAS PERFORMED AND THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/30 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/30. (B)(4). Manufacturer Narrative: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/2.0/103 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE PATIENT EXPERIENCED LENS ROTATION, DISLOCATION, REFRACTIVE CHANGE, AND BLURRED VISION. THE SURGEON REPOSITIONED THE LENS AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2016, THE SURGEON REPORTED LENS OPACITY (ASCO) AND LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/30 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/30. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE SWITZERLAND BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+2.00/103 | OS | Physician | 04/25/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5694861&pc=MTA | 2023826-2016-00799 | 2016-04-14 | Injury | STAAR SURGICAL COMPANY | 2016-06-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; Blurred Vision; Loss of Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED ASC (ANTERIOR SUBCAPSULAR CATARACT) ON (B)(6) 2016 AND SUBSEQUENTLY LOST BCVA (BEST-CORRECTED VISUAL ACUITY). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS REMAINS IMPLANTED. (B)(4). | -10.5/+1.0/085 | OS | Physician | 05/16/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5572298&pc=MTA | 2023826-2016-00432 | 2015-04-13 | Injury | STAAR SURGICAL COMPANY | 2016-04-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; Blurred Vision; Visual Impairment | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+4.0/X010 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE REPORTER INDICATED THERE WAS LENS OPACITY (CORTICAL), NOTED ON (B)(6) 2015, AND THE PATIENT EXPERIENCED LOSS OF BCVA. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/200. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT EVENT WITHIN ASSOCIATED LOTS WAS FOUND. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). LENS NOT RETURNED. | -18.0/+4.0/010 | OS | Physician | 02/12/2016 | 2016
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5378618&pc=MTA | 2023826-2015-01669 | 2014-11-21 | Injury | STAAR SURGICAL COMPANY AG | 2016-01-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS, -17.0/+3.0/087 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR). THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT'S POST-OP (ON (B)(6) 2015) UNCORRECTED VISUAL ACUITY WAS 20/15. DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4). | -17.0/+3.0/087 | OS | Physician | 11/24/2015 | 2015 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5671142&pc=MTA | 2023826-2016-00741 | 2015-12-21 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -12.0/+0.5./X160 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (CORTICAL), NOTED ON (B)(6) 2015. THE PATIENT HAD CATARACT SURGERY AND AN INTRAOCULAR LENS WAS IMPLANTED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. LENS RETURNED BENT WAS ALSO NOTED. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | -12.0/+0.5./160 | OS | Physician | 04/21/2016 | 2016 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5814335&pc=MTA | 2023826-2016-00916 | 2016-04-18 | Injury | STAAR SURGICAL COMPANY AG | 2016-07-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS, -18.0/+3.5/092 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR)/PERIPHERAL TOUCH. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED AND SPOTS ON LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | -18.0/+3.5/092 | OD | Physician | 06/22/2016 | 2016 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7987375&pc=MTA | 2023826-2018-01533 | 2018-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-10-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Corneal Edema; Blurred Vision; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.50/+1.5/060 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), THE LAST WEEK OF (B)(6) 2018. THE REPORTER INDICATED THE PATIENT HAD A FIXED DILATED PUPIL POST-OP, SEVERE CORNEAL EDEMA, WHICH HAS BEEN PERSISTENT UNTIL NOW AND AN ANTERIOR SUBCAPSULAR CATARACT. THE PATIENT EXPERIENCED PIGMENT DISPERSION, GLARE/HALOS, BLURRED VISION AND PUPIL/IRIS ATROPHY. THE LENS WAS EXPLANTED. THE CATARACT WAS REMOVED ONE MONTH AFTER THE SURGERY AND AN IOL WAS IMPLANTED. THE PUPIL REMAINED DILATED AND THE CORNEA WAS EDEMATOUS AND WHITE. THE PATIENT HAS HAD A CORNEAL TRANSPLANT. THE REPORTER INDICATED THE INTRAOCULAR PRESSURE HAS BEEN STABLE AND NORMAL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING. THE PATIENTS VISION IS STILL HAND MOTION POST OPERATIVELY AND INTRAOCULAR PRESSURE IS 12MMHG. THE FIRST CHECKUP AFTER KERATOPLASTY INDICATED ALL IS GOOD. CORRECTED DATA: (B)(6). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). | -5.50/+1.5/060 | OS |  | 10/04/2018 | 2018 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6220084&pc=MTA | 2023826-2017-00009 | 2016-03-17 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; Glaucoma; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -14.50/+2.5/087 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DESCRIPTION: PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/13 AS OF 04/21/2016. DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. (B)(4). CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4). | -14.50/+2.5/087 | OD | Physician | 05/27/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4937371&pc=MTA | 2023826-2015-00893 | 2015-03-23 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS OPACITY, WHICH WAS FIRST NOTED ON (B)(6) 2015. THE REPORTER INDICATED IT WAS AN ANTERIOR SUBCAPSULAR OPACITY. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE, PIGMENT DISPERSION, GLARE/HALOS AND BLURRED VISION. THE ANTERIOR CHAMBER WAS IRRIGATED/EVACUATED OF REMAINING VISCO/FLUIDS. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). | -13.5/3.5/098 | OS | Physician | 06/29/2015 | 2015 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5926670&pc=MTA | 2023826-2016-00374 | 2015-12-17 | Injury | STAAR SURGICAL COMPANY | 2016-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Intraocular Pressure, Delayed, Uncontrolled; Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+0.5/088 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED AT FOLLOW-UP VISITS, THE LENS HAPTIC WAS FOLDED AND THE PUPIL WAS OVALISED. THE LENS WAS TILTED ON ONE SIDE AND WAS RE-POSITIONED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: FOLLOW-UP #4 FOR MDR # 2023826-2017-00374 WAS SUBMITTED IN ERROR. THE REPORTED INFORMATION WAS MEANT TO BE FOLLOW-UP #4 FOR MDR # 2023826-2016-00374. AS A RESULT, ONLY THE INITIAL MDR # 2023826-2017-00374 PERTAINS TO CLAIM# 710106. FOLLOW-UP #4 FOR MDR # 2023826-2016-00374 WILL BE SUBMITTED. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | -13.0/+0.5/088 | OD | Physician | 02/02/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7794667&pc=MTA | 2023826-2018-01196 | 2018-05-16 | Injury | STAAR SURGICAL COMPANY | 2018-08-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Cataract; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+2.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LENS TEAR/BREAK DURING INJECTION INTO THE EYE. ANOTHER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+2.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), DURING THE SAME SURGERY. THE LENS WAS REMOVED DUE TO LENS TEAR/BREAK DURING INJECTION INTO THE EYE AND THE SURGEON DAMAGED THE CAPSULE. THE SURGEON PERFORMED PHACOEMULSIFICATION, WITH IOL IMPLANTATION. AN INTRAOCULAR LENS (IOL) WAS IMPLANTED IN A SECOND SURGERY AND THE PROBLEM WAS RESOLVED. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: PREVIOUS DATE NOT APPLICABLE, LENS WAS NOT RETURNED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.0/+2.5/084 | OD | Physician | 07/22/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7300934&pc=MTA | 2023826-2018-00292 | 2015-10-28 | Injury | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Cataract; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -11.0/+1.0/150 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. ON (B)(6) 2017 THE LENS WAS EXPLANTED. THE SURGEON REPORTS LENS OPACITY NS. REPORTEDLY, PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED. THE SURGEON STATES "CATARACT SURGERY PERFORMED. ICL REMOVED DUE TO CATARACT. NOW PERFECT VISION." Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -11.0/+1.0/150 | OD | Physician | 02/08/2018 | 2018 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6721184&pc=MTA | 2023826-2017-01190 | 2016-09-12 | Injury | STAAR SURGICAL COMPANY | 2017-07-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Cataract; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.5/+4/87 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED LENS OPACITY (POSTERIOR SUBCAPSULAR CATARACT). ON (B)(6) 2017 THE LENS WAS EXPLANTED FOR PHACO-EMULSIFICATION AND IOL IMPLANTATION. THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. THE PRODUCT HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL AND HAS SOME MOISTURE. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE) AND THE LENS MISSING A PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THIS PRODUCT IS MANUFACTURED IN THE U.S, BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH - NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/+4/87 | OD | Physician | 06/30/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6805946&pc=MTA | 2023826-2017-01283 | 2016-07-25 | Injury | STAAR SURGICAL COMPANY | 2017-08-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.0/+2.0/073 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON THE SAME DAY DUE TO LENS OPACITY (ASCO). PHACO- EMULSIFACTION WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS FOUND TO BE 20/20. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: IN THE PROCESS OF THE IOL IMPLANT, DUE TO THE LACK OF PATIENT'S COOPERATION (TURN THE HEAD) LEAD TO THE OPERATION CAUSED BY NATURAL CRYSTALS BECOME CLOUDY. THE IOL IMPLANTATION SOLVED THE PROBLEM AND THE PATIENT'S CONDITION IS GOOD. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.0/+2.0/073 | OS | Physician | 07/22/2017 | 2017 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6077381&pc=MTA | 2023826-2016-01531 | 2016-04-02 | Injury | STAAR SURGICAL COMPANY | 2016-11-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS - -18.5/+2.0/90 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2008. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT AND CATARACT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/15. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: PMA/510(K) : NA THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.5/+2.0/90 | OS | Physician | 10/11/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6074345&pc=MTA | 2023826-2016-00912 | 2016-04-13 | Injury | STAAR SURGICAL COMPANY | 2016-11-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS, -22.50/+3.0/095 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR OPACITY). Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT UNDER THE AGE OF 21 YEARS OLD AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED. | -22.50/+3.0/095 | OD | Physician | 06/08/2016 | 2016 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6566884&pc=MTA | 2023826-2017-00810 | 2016-08-16 | Injury | STAAR SURGICAL COMPANY | 2017-05-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/82 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE SURGEON NOTICED LOW VAULTING AND LENS OPACITY (ASCO) EARLY POSTOPERATIVELY. THE LENS WAS EXCHANGED FOR A LONGER LENS OF A DIFFERENT MODEL AND DIOPTER AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2016, PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/14 (BETTER THAN PREOPERATIVE). Manufacturer Narrative: THE LENS WAS RETURNED SCOTCHED ON A PAPER. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS AND REST OF SCOTCH). IT WAS NECESSARY TO REHYDRATE BECAUSE THE LENS WAS RETURNED SCOTCHED. (B)(4). Manufacturer Narrative: (B)(4). AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS OVER (B)(6)YEARS OF AGE AND WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM. (B)(4). | -12.5/+2.0/82 | OS | Physician | 10/14/2016 | 2016 | 58
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6590465&pc=MTA | 2023826-2017-00825 | 2016-08-16 | Injury | STAAR SURGICAL COMPANY | 2017-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/88 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED LOW VAULT AND LENS OPACITY (ASCO). THE LENS WAS EXPLANTED ON (B)(6) 2016 AND WAS EXCHANGED WITH A LONGER LENS WITH A DIFFERENT MODEL AND DIOPTER, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED SCOTCHED ON A PAPER. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS AND REST OF SCOTCH). IT WAS NECESSARY TO REHYDRATE BECAUSE OF RETURNED SCOTCHED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE- AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) UNDER 3.0MM AND FOR PATIENTS OVER THE AGE OF (B)(6) AT THE TIME OF IMPLANTATION. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | -12.0/+1.5/88 | OD | Physician | 10/14/2016 | 2016 | 58
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6059292&pc=MTA | 2023826-2016-01530 | 2016-04-02 | Injury | STAAR SURGICAL COMPANY | 2016-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Cataract; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS - -18.5/+1.5/78 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2008. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT, LENS ROTATION AND CATARACT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/13. Manufacturer Narrative: THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/13 AND UNCORRECTED VISUAL ACUITY WAS 20/13. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.5/+1.5/78 | OD | Physician | 10/11/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6850306&pc=MTA | 2023826-2017-01413 | 2017-05-23 | Injury | STAAR SURGICAL COMPANY | 2017-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Cataract; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.5/+1.5/107 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR), WHICH WAS OBSERVED ON (B)(6) 2017. POST CATARACT OPERATION PATIENT IS HAPPY WITH THE VISION. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF LABEL USE (ACD<3.0MM). (B)(4). | -16.5/+1.5/107 | OS | Physician | 08/11/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4704662&pc=MTA | 2023826-2015-00424 | 2015-03-16 | Injury | STAAR SURGICAL COMPANY | 2015-04-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Uveitis; Discomfort; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. DURING EXPLANTATION, ANTERIOR SUBCAPSULAR OPACITIES SUDDENLY DEVELOPED. THE PATIENT EXPERIENCED DISCOMFORT, LIGHT SENSITIVITY AND ANTERIOR UVEITIS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT IS BEING MONITORED AND IS DOING BETTER IN TERMS OF EYE DISCOMFORT BUT HAS COMPLAINED OF BLURRED VISION. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. LENS NOT RETURNED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -4.5/+5.0/101 | OD | Other | 03/19/2015 | 2015 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4704665&pc=MTA | 2023826-2015-00425 | 2015-03-16 | Injury | STAAR SURGICAL COMPANY | 2015-04-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Uveitis; Halo; Discomfort; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. DURING EXPLANTATION, ANTERIOR SUBCAPSULAR OPACITIES SUDDENLY DEVELOPED. THE PATIENT EXPERIENCED PIGMENT DISPERSION, GLARE/HALOS, INFLAMMATION, ANTERIOR UVEITIS, DISCOMFORT, LIGHT SENSITIVITY, PAIN AND TEARING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT IS BEING MONITORED AND IS DOING BETTER IN TERMS OF EYE DISCOMFORT BUT HAS COMPLAINED OF BLURRED VISION. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. LENS NOT RETURNED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -4.5/+3.5/092 | OS | Other | 03/19/2015 | 2015 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7906066&pc=MTA | 2023826-2018-01390 |  | Injury | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Clouding, Central Corneal; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM513.2, -14.00/+3.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. ON (B)(6) 2018 PATIENT WAS OBSERVED WITH SOME CENTRAL CORNEAL EDEMA WITH NO INFLAMMATION OR NO INFECTION. PATIENT IS RECOVERING WELL, AND WILL BE REVIEWED IN NEXT APPOINTMENT. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN SPECIMEN CUP WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -14.00/+3.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. ON (B)(6) 2018 THE PATIENT WAS OBSERVED WITH SOME CENTRAL CORNEAL EDEMA WITH NO INFLAMMATION OR NO INFECTION. PATIENT IS RECOVERING WELL, AND THE CENTRAL CORNEAL EDEMA HAS BEEN RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.00/+3.5/088 | OD | Physician | 08/28/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7791689&pc=MTA | 2023826-2018-01181 | 2018-07-10 | Injury | STAAR SURGICAL COMPANY | 2018-08-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Use of Device Problem; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Corneal Decompensation; Corneal Edema | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -18.00/2.0/129 (SPHERE/CYLINDER/AXIS) IMPLANTBLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. IT WAS REPORTED THAT THE SURGEON REMOVED THE LENS AND IMPLANTED IT AGAIN. ONCE THE LENS WAS INSIDE THE EYE, THE SURGEON REALIZED THAT THE LENS WAS DAMAGED AND REMOVED THE LENS AGAIN. CORNEAL EDEMA AND CORENAL DECOMPENSATION WERE OBSERVED. THE REPORTER STATED THAT THE CAUSE OF THE ADVERSE EVENTS WAS THE SURGEON/SURGICAL PROCEDURE. THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO THE EVENT. PATIENT IS AWAITING ICL REPALCEMENT AND THE EYE IS OKAY. Manufacturer Narrative: (B)(4). Manufacturer Narrative: LENS WAS RETURNED DRY, IN A MICRO-CENTRFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INPSECTION FOUND THE HAPTIC TORN AND DEFORMED; AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: NEW INFORMATION: "SURGEON IMPLANTED THE NEW LENS (SAME PARAMETERS AS THE FIRST ONE SINCE THE PROBLEM WAS THAT HE BROKE THE LENS) VTICMO 13.2 -18/+2.0/108 . PATIENT'S AV IS SATISFYING, VAULT IS NORMAL". CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.00/2.0/129 | OD | Physician | 07/17/2018 | 2018 | 47
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7791893&pc=MTA | 2023826-2018-01244 | 2018-07-12 | Injury | STAAR SURGICAL COMPANY | 2018-08-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Decompensation; Corneal Edema; Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -17.5/+3.0/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO HIGH IOP, MID DILATED PUPIL, CORNEAL EDEMA, CLOSE ANGEL GLAUCOMA AND ANGLE BLOCK. THE SURGEON PERFORMED A PARS PLANA VITRECTOMY, PHACO-EMULSIFICATION (CATARACT SURGERY) WITH IOL IMPLANTATION AND IRIS SUTURE. IT WAS REPORTED THAT AT PATIENTS LAST FOLLOW UP VISIT VISUAL ACUITY WAS 20/40 WITHOUT CORRECTION, IOP WAS NORMAL WITH A REGULAR PUPIL. Manufacturer Narrative: CORRECTED DATA: ADDED 'OTHER SERIOUS' TO INITIAL REPORT. REPORTER ALSO INDICATES CORNEAL DECOMPENSATION. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A TORN HAPTIC AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17.5/+3.0/100 | OD | Physician | 07/17/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6229803&pc=MTA | 2023826-2017-00020 | 2016-12-06 | Injury | STAAR SURGICAL COMPANY | 2017-01-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Decompensation; Corneal Edema; Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/0.96 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT REPORTED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, CORNEAL DECOMPENSATION, PIGMENT DISPERSION, AND ELEVATED IOP. ON (B)(6) 2016 THE LENS WAS EXPLANTED. AS OF THE DATE OF MDR SUBMISSION LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/096 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT REPORTED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, CORNEAL DECOMPENSATION, PIGMENT DISPERSION, AND CORNEAL EDEMA. ON (B)(6) 2016 THE LENS WAS EXPLANTED. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATION PENDING. | -9.5/+2.5/0.96 | OD | Physician | 12/09/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7109978&pc=MTA | 2023826-2017-01877 | 2017-10-30 | Injury | STAAR SURGICAL COMPANY | 2017-12-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Decompensation; Corneal Edema; Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -4.0/+3.5/171 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED. THE SURGEON REPORTS EXCESSIVE VAULT, PUPIL BLOCK, ELEVATED IOP, CORNEAL DECOMPENSATION, UNREACTIVE PUPIL, ANGLE CLOSURE, AND CORNEAL EDEMA. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: REPORTER STATED PATIENT IS OKAY. THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION: 24-OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0 MM (2.92 MM)]. (B)(4). | -4.0/+3.5/171 | OS | Physician | 11/20/2017 | 2017 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7300916&pc=MTA | 2023826-2018-00281 | 2017-12-13 | Injury | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Decompensation; Low Blood Pressure/ Hypotension; Inflammation; Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+2.0/170 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED. THE SURGEON NOTES THE FOLLOWING: EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), CORNEAL DECOMPENSATION, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, GLARE/HALOES, BLURRED VISION, INFLAMMATION, AND "CHOROIDAL DETACHMENT, SEVERE HYPOTENSION." REPORTER STATES THAT HYPOTENSION IS "BEING RESOLVED" AND "IN FACT THE PATIENT IS VERY GOOD." Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND. RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. MEASUREMENT OF THE HYDRATION FOUND MOISTURE CONTENT ON A HIGH LEVEL; BUT NOT EXCEEDING LIMITS. THE OVERALL LENS DIAMETER WAS THEN FOUND OUT OF TOLERANCE. THE LENS SN: (B)(4) WAS A 12.6MM IMPLANTABLE LENS THAT IS 0.22MM LARGER THAN THE THEORETICAL; THE TOLERANCE IS +/-0.20MM. SINCE THE WORK ORDER SEARCH DID NOT FIND ADDITIONAL SIMILAR ERRORS TO THIS COMPLAINT EVENT, IT CAN BE CONCLUDED THAT THE OOS CONDITION IS NOT LIKELY DUE TO MANUFACTURING OR PACKAGING. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: MEASUREMENT OF THE HYDRATION FOUND THE LENS TO BE OVERHYDRATED. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.0/+2.0/170 | OD | Physician | 01/31/2018 | 2018 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6659008&pc=MTA | 2023826-2017-01001 | 2016-09-30 | Injury | STAAR SURGICAL COMPANY | 2017-06-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Corneal Decompensation; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, DIOPTER -3.5/+2/10 IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT BEGAN TO EXPERIENCE CORNEAL DECOMPENSATION, IRIS ATROPHY, IRIS HERNIA, LENS ROTATION, AND IRIS PROLAPSE. SECONDARY INTERVENTIONS INCLUDE ANTERIOR SUBINCISIONAL SINECHIOTOMY, ENLARGEMENT OF INCISION, ANTERIOR CHAMBER IRRIGATION, IRIS SUTURE, AND ADDITIONAL SUTURE. THE LENS REMAINS IMPLANTED AND IS UNLIKELY TO BE REMOVED, ACCORDING TO CLAIM REPORTER. THE INCISION WAS ENLARGED TO PERFORM THE IRIS SUTURE AND THE ANTERIOR CHAMBER IRRIGATION. ONCE PERFORMED, THE INCISION WAS SUTURED AND AN ASTIGMATISM DEVELOPED BECAUSE OF THE STITCHES' TENSION. THE SURGEON HAS REQUESTED MEDICAL CONSULTATION TO DETERMINE IF FURTHER ICL ROTATION (ICL HAD ALREADY ROTATED FROM ORIGINAL POSITION) WOULD COMPENSATE FOR THE ASTIGMATISM. Manufacturer Narrative: PREVIOUS: REPOSITIONING PENDING MEDICAL CONSULTATION. UPDATE: LENS WAS REPOSITIONED (B)(6) 2017. THE PROBLEM HAS BEEN RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED. | -3.5/+2/10 | OS | Physician | 05/26/2017 | 2017 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7213982&pc=MTA | 2023826-2018-00140 | 2017-12-20 | Injury | STAAR SURGICAL COMPANY | 2018-01-23 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Corneal Edema | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.0/+2.0/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES CORNEAL EDEMA THAT WAS "BEYOND STANDARD AFTER OCULAR SURGERY." REPORTEDLY, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION STATES THAT ON THE FOLLOW UP VISIT (B)(6) 2017, THE CORNEA IS "RECOVERED TO TRANSPARENT." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. CONCLUSION CODE: OFF-LABEL USE [LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY] (B)(4). | -14.0/+2.0/083 | OS | Physician | 01/02/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7378746&pc=MTA | 2023826-2018-00466 | 2018-02-26 | Injury | STAAR SURGICAL COMPANY | 2018-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -5.5/+3.5/066 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS NOTICED IN THE PATIENT DUE TO USER ERROR. THE SURGEON USED THE MANIFEST REFRACTION INSTEAD OF THE CYCLOPLEGIC REFRACTION FOR THE LENS CALCULATION. LENS REMAINS IMPLANTED. LENS EXCHANGE IS PLANNED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, IMPLANTABLE COLLAMER LENS OF -5.5/+3.5/066 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6)2018. REFRACTIVE SURPRISE ALONG WITH BLURRED VISION WAS NOTICED IN THE PATIENT DUE TO USER ERROR. THE SURGEON USED THE MANIFEST REFRACTION INSTEAD OF THE CYCLOGENIC REFRACTION FOR LENS CALCULATION. LENS WAS EXCHANGED FOR A SIMILAR LENS OF DIFFERENT DIOPTER ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE IS "LITTLE PERIPHERAL CORNEAL OEDEMA". EXPLANTED DATE: (B)(6) 2018. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN AND BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.5/+3.5/066 | OS | Physician | 02/26/2018 | 2018 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7267530&pc=MTA | 2023826-2018-00214 | 2018-01-03 | Injury | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Corneal Edema; Blurred Vision; Visual Disturbances; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.0/+2.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE SURPRISE AND CORNEAL EDEMA; INCIDENT DATE STATED AS (B)(6) 2018. REPORTEDLY, THE CORNEAL EDEMA DISAPPEARED AT 1-WEEK POST-OP AND NO TREATMENT BEYOND WHAT IS STANDARD FOR OCULAR SURGERY WAS PERFORMED. CLAIMANT STATES THE LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, OF -15.0/+2.0/082 SPHERE/ CYLINDER/ AXIS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE SURPRISE; DOUBLE VISION AND CORNEAL EDEMA ON (B)(6) 2018. REPORTEDLY THE CORNEAL EDEMA DISAPPEARED AT 1 WEEK POST OP AND NO TREATMENT BEYOND WHAT IS STANDARD FOR OCULAR SURGERY WAS PERFORMED. THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MMVTICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.0/+2.0/082 SPHERE/ CYLINDER/ AXIS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE SURPRISE; DOUBLE VISION AND CORNEAL EDEMA ON (B)(6) 2018. REPORTEDLY THE CORNEAL EDEMA DISAPPEARED AT 1 WEEK POST OP AND NO TREATMENT BEYOND WHAT IS STANDARD FOR OCULAR SURGERY WAS PERFORMED. THE LENS WAS EXCHANGED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. IMPLANTED DATE: IT IS CORRECTED FROM "(B)(6) 2017" TO "(B)(6) 2017". EXPLANTED DATE: (B)(6) 2018. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CORRECTED DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION WAS PERFORMED. ALTHOUGH LENS SN: (B)(4) WAS FOUND OUT OF TOLERANCE FOR SPHERIC REFRACTIVE DIOPTER, THERE IS NO REASONABLE PROBABLE CAUSE IN MANUFACTURING THAT WOULD HAVE RESULTED IN A REFRACTIVE CHANGE. THE SPHERIC REFRACTION OF THIS LENS WAS CALCULATED TO BE -13.96D AT THE TIME OF RELEASE. (B)(4) Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: IT IS CORRECTED TO "THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS." (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION WAS PERFORMED. ALTHOUGH LENS SN: (B)(4) WAS FOUND OUT OF TOLERANCE FOR SPHERIC REFRACTIVE DIOPTER, THERE IS NO REASONABLE PROBABLE CAUSE IN MANUFACTURING THAT WOULD HAVE RESULTED IN A REFRACTIVE CHANGE. THE SPHERIC REFRACTION OF THIS LENS WAS CALCULATED TO BE -13.96D AT THE TIME OF RELEASE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.0/+2.0/082 | OS | Physician | 01/18/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6760836&pc=MTA | 2023826-2017-01145 | 2017-07-03 | Injury | STAAR SURGICAL COMPANY | 2017-08-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Corneal Edema; Edema; Pain; Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -4.5/1.5/86 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE NEXT DAY, SIGNS OF TOXIC ANTERIOR SEGMENT SYNDROME, INCLUDE TEARING, PAIN, BLURRY VISION, EYELID EDEMA, HYPEREMIA, CORNEAL EDEMA, FIXED PUPIL AND FIBRIN CLOT IN THE ANTERIOR CHAMBER WAS NOTICED. THERAPY WAS STARTED WITH ORAL PREDNISOLONE (60 MG/DAY) AND ACETAZOLAMIDE 250 MM (1+1+1), TROPICAL PREDNISOLONE 1 DROP PER HOUR, TRAPICAMIDE, OFLOXACINE, AND BROMOFENAC. THE FOLLOWING DAY, THE PATIENT EXPERIENCE LESS PAIN, INTENSE HYPEREMIA, LESS EYELID EDEMA BUT A VERY CLOUDY WHITISH CORNEA. CORNEAL PERIPHERY WAS PARTIALLY PARED AND A BLOOD CLOT WAS SEEN IN THE INFERIOR A.C. AN ANTERIOR CHAMBER WASHING WAS PERFORMED AS WELL AS THE TEMPORAL INCISION WITCH BEING SUTURED WITH 2 10/00 NYLON STICHES AND INTRAVITREOUS INJECTION OF VANCOMICINE 1MG/0.1 ML AND CEFTAZIDIME 1.0 MG/0.1 ML. ACCORDING TO THE SURGEON, THE CORNEA IS CLEARING UP AND WOULD NEED TO MAKE ANOTHER SURGERY TO REMOVE THE MEMBRANE. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -4.5/1.5/86 | OS | Physician | 07/03/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6531883&pc=MTA | 2023826-2017-00727 |  | Injury | STAAR SURGICAL COMPANY | 2017-04-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Corneal Edema; Inflammation; Intraocular Pressure Increased; Pupillary Block; Uveitis; Blurred Vision; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -8.5/+2.5/101 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, LOSS OF BEST-CORRECTED VISUAL ACUITY, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE AND AN UNREACTIVE PUPIL (FIXED). THE PATIENT EXPERIENCED GLARE/HALOS, BLURRED VISION AND INFLAMMATION. ALSO, MILD UVEITIS, PIGMENT DISPERSION AND HAD DEVELOPED URETTS ZAVALIA SYNDROME (PERSISTENT PUPIL DILATION). THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE STATUS OF THE EYE POST-OP SHOWED CORNEAL EDEMA, SHALLOW ANTERIOR CHAMBER AND NORMAL INTRAOCULAR PRESSURE. Manufacturer Narrative: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC DEFORMED AND THE HAPTIC BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.5/+2.5/101 | OD | Physician | 03/30/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6389853&pc=MTA | 2023826-2017-00347 | 2017-02-02 | Injury | STAAR SURGICAL COMPANY | 2017-03-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Corneal Edema; Inflammation; Visual Disturbances; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16/2/111 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED UNREACTIVE PUPIL, GLARE, INFLAMMATION, AND SWOLLEN CORNEA. THE INFERIOR EDGE OF ICL CAUGHT ABOVE IRIS, VAULTED FORWARD. IN ADDITION, THE PATIENT EXPERIENCED IRIS ATROPHY AND A DESCEMET MEMBRANE TEAR. ON (B)(6) 2017, THE LENS WAS EXCHANGED AND REPOSITIONED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE LENS WILL BE RETURNED. BASED ON SURGEON'S COMMENTS DOCTOR IS MONITORING PATIENT'S PROGRESS. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16/2/111 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2017. THE PATIENT EXPERIENCED UNREACTIVE PUPIL, GLARE, INFLAMMATION, AND SWOLLEN CORNEA. THE INFERIOR EDGE OF ICL CAUGHT ABOVE IRIS WHICH CAUSED THE LENS TO VAULT FORWARD SO LENS REPOSITIONING WAS PERFORMED. IN ADDITION, PATIENT EXPERIENCED IRIS ATROPHY AND DESCEMET MEMBRANE TEAR. ON (B)(6)2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. AS OF (B)(6)2017 THE PATIENT WAS STILL EXPERIENCING POSTOPERATIVE SEQUELAE BUT CORNEA WAS CLEAR. THE LENS WILL NOT BE RETURNED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). RESULT: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WAS REPORTED FOR UNITS IN THE SAME LOT. (B)(4). | -16/2/111 | OD | Physician | 02/08/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594294&pc=MTA | 2023826-2018-00898 | 2018-04-16 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/156 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS THE FOLLOWING: ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), ADDITION OF NEW PI (SURGICAL), ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, "SHALLOW AC," DECREASE IN UA, AND CORNEAL EDEMA. STATES THE PROBLEM STARTED TO RESOLVE AFTER "TOW TIME AC WASH" AND "AC REFORMATION." ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ATTEMPTS TO OBTAIN UPDATED INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED FROM REPORTER INDICATED: "THE LAST STATUS WAS AT 3 DAYS POST ICL REMOVAL: CORNEAL EDEMA + ISCHEMIC IRIS; DIFFERENT GRADES. SURELY NO EXCHANGE LENSES HAVE BEEN IMPLANTED." DEVICE EVALUATION: LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT, IN A STERIPACK. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.76MM)]. (B)(4) | -10.0/+2.0/156 | OS | Physician | 05/20/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6850665&pc=MTA | 2023826-2017-01401 | 2017-08-03 | Injury | STAAR SURGICAL COMPANY | 2017-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/X092 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2017. ON (B)(6)2017 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND ANGLE CLOSURE WITH ELEVATED IOP. THE SURGEON REPORTED ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS. SURGEON ALSO REPORTS CORNEAL EDEMA/MIDRIASIS. THE SURGEON STATES THAT THE PATIENT STATUS IS OKAY. Manufacturer Narrative: LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.0/+2.5/092 | OD | Physician | 08/09/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594288&pc=MTA | 2023826-2018-00897 | 2018-04-16 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/014 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS THE FOLLOWING: ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), ADDITION OF NEW PI (SURGICAL), ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. THREE DAYS POSTOP IT IS NOTED THAT THE SITUATION HAS IMPROVED SLIGHTLY, "TOW TIME AC WASH AND AC REFORMATION WITH PI (SURGICAL) AND AIR BUBBLE PLACEMENT." ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A "STERIPACK". SPELLING OF "STERIPACK" WAS INCORRECT PRIOR. (B)(4) Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED FROM REPORTER STATED "THE LAST STATUS WAS AT 3 DAYS POST ICL REMOVAL: CORNEAL EDEMA + ISCHEMIC IRIS; DIFFERENT GRADES. SURELY NO EXCHANGE LENSES HAVE BEEN IMPLANTED." DEVICE EVALUATION - LENS WAS RETURNED DRY IN A STERIPECLE WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND LENS WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE [(SECONDARY SURGERY, LENS EXPLANT, ADDITION OF NEW PI (SURGICAL), SIGNIFICANT REDUCTION OF ICA, ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS] CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.89MM)] (B)(4). | -12.5/+1.0/014 | OS | Physician | 05/20/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594291&pc=MTA | 2023826-2018-00896 | 2018-04-16 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/179 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS THE FOLLOWING: ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), ADDITION OF NEW PI (SURGICAL), ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, "FLAT AC," AND CORNEAL EDEMA. STATES THE PROBLEM STARTED TO RESOLVE AFTER TWO AC WASHES AND SURGICAL PI WITH "AIR BUBBLE IN AC." ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED FROM REPORTER INDICATED: "THE LAST STATUS WAS AT 3 DAYS POST ICL REMOVAL: CORNEAL EDEMA + ISCHEMIC IRIS; DIFFERENT GRADES. SURELY NO EXCHANGE LENSES HAVE BEEN IMPLANTED." LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT, IN A STERIPACK. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE [(SECONDARY SURGERY, LENS EXPLANT, ADDITION OF NEW PI (SURGICAL), SIGNIFICANT REDUCTION OF ICA, ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, "FLAT AC"] CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.93MM)] (B)(4). | -10.0/+1.5/179 | OD | Physician | 05/20/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5221040&pc=MTA | 2023826-2015-01452 | 2015-10-02 | Injury | STAAR SURGICAL COMPANY | 2015-11-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Corneal Edema; Intraocular Pressure Increased; Pain; Pupillary Block; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, UNKNOWN DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE AND UNREACTIVE PUPIL WAS NOTED. THE ICL WAS EXPLANTED. NO OTHER LENS WAS IMPLANTED. SEE MFR. #2023826-2015-01451 FOR RIGHT EYE. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -14.00/0.5/103 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE (IOP), UNREACTIVE PUPIL, ANGLE CLOSURE, AND POST-OP PAIN CAUSED BY ELEVATED IOP AND CORNEAL EDEMA. THE ICL WAS EXPLANTED. NO OTHER LENS WAS IMPLANTED. (B)(4). CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS (B)(4). | -14.00/0.5/103 | OS | Physician | 10/13/2015 | 2015 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7281448&pc=MTA | 2023826-2018-00257 | 2017-07-12 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-19 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Corneal Edema; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+1.00/135 (SPHERE/CYLINDAR/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY ON (B)(6) 2017. ALSO REPORTED IS SIGNIFICANT ENDOTHELIAL CELL LOSS. CLAIM REPORTER STATES "BIG CORNEAL EDEMA AFTER SURGERY. THICKER PAQUIMETRY...REFRACTION MODIFIED." REPORTEDLY THE EYE HEALED PROPERLY AFTER SURGERY. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID INSIDE A LENS VIAL. VISUAL INSPECTION FOUND A TEAR IN THE HAPTIC AND OPTIC. THE LENS WAS ALSO MISSING A PIECE OF ITS HAPTIC. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.00/+1.00/135 | OS | Physician | 01/26/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7152526&pc=MTA | 2023826-2017-01958 | 2017-10-03 | Injury | STAAR SURGICAL COMPANY | 2017-12-28 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Corneal Edema; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.5/+1.0/102 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON REPORTS "PUPIL EXPANDED AND NOT CONSTRICTED (ABNORMAL PUPILLARY MYDRIASIS)." SURGEON ALSO REPORTS INTRAOCULAR INJECTIONS OF "SANPILO (PILOCARPINE HYDROCHLORIDE)" AND STATES THAT CORNEAL EDEMA IS "OBSERVED AT PUPILLARY MYDRIASIS TREATMENT." THE SURGEON REPORTS THAT DURING A FOLLOW-UP VISIT, THE ICL WAS REMOVED AS HE THOUGHT THE ICL POSSIBLY WAS NOT ALLOWING THE PUPIL TO CONSTRICT. HOWEVER, UPON REMOVAL THERE WERE NO IRREGULARITIES FOUND ON THE ICL SO IT WAS "IMPLANTED AGAIN." ON (B)(6) 2017 THE LENS WAS FINALLY EXPLANTED AND REPORTEDLY THE PUPIL DID NOT CONSTRICT AFTER EXPLANT. NO ADDITIONAL INFORMATION IS AVAILABLE AS OF THE TIME OF MDR SUBMISSION. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (RED AND CLEAR RESIDUE ON LENS) AND THE LENS MISSING A PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.5/+1.0/102 | OS | Physician | 11/29/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594282&pc=MTA | 2023826-2018-00895 | 2018-04-16 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/013 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE SURGEON NOTES ADDITION OF NEW PI (SURGICAL). ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A STERIPACK AND THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTIMO13.2 IMPLANTABLE COLLAMER LENS -12.5/+1.0/013 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE SURGEON NOTES ADDITION OF NEW PI (SURGICAL). THE LAST PATIENT STATUS SHOWED CORNEAL EDEMA AND ISCHEMIC IRIS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+1.0/013 | OD | Physician | 05/20/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5291931&pc=MTA | 2023826-2015-01634 | 2015-09-26 | Injury | STAAR SURGICAL COMPANY | 2015-12-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cyclitis; Vitrectomy; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.0/+4.0/X087 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO PATIENT EXPERIENCED CYCLITIS. AFTER THE LENS WAS EXPLANTED, THE PATIENT WAS TRANSFERRED TO TERTIARY HOSPITAL FOR PAS PLANA VITRECTOMY (PPV). Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4). A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -12.0/+4.0/087 | OD | Physician | 11/10/2015 | 2015 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7152455&pc=MTA | 2023826-2017-01962 | 2017-10-09 | Injury | STAAR SURGICAL COMPANY | 2017-12-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Device Dislodged or Dislocated | Cyst(s); Blurred Vision; Halo; Discomfort; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -1.5/+3.5/170 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON REPORTS EXCESSIVE VAULT, LENS ROTATION, REFRACTIVE SURPRISE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA) (ANGLE CLOSURE), HALOES, VISUAL DISCOMFORT, AND ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS. ALSO REPORTED IS SURGICAL ENLARGEMENT/ADDITION OF NEW PI, LOW BCVA AND UCVA, AND LENS REPOSITIONING. THE SURGEON STATES THAT "PROBABLY THE LENS IS TOO BIG FOR THE EYE" BUT BELIEVES THAT THE VISUAL PROBLEMS OF THE PATIENT ARE NOT RELATED TO THE LENS BUT TO THE CORNEA. HE BELIEVES THAT CHANGING THE ICL WILL NOT SOLVE THE VISUAL DISCOMFORT AND THAT THE LENS IS NOT DOING ANY HARM TO THE PATIENT. THE LENS REMAINS IMPLANTED. THE PATIENT IS "NOT HAPPY WITH HIS VISUAL ACUITY." Manufacturer Narrative: THE SURGEON NOTED THAT "THE IRIS CYST WAS NOT SUSPECTED BEFORE THE SURGERY. THE SUSPICION WAS RISEN AFTER THE FIRST BY THE FACT THAT THE VAULT WAS UNEVEN AND THAT WITH DILATION ALWAYS GETS MORE EVEN AND WITH A NORMAL HEIGHT. AND THE CYST SUSPICION WAS EVEN STRONGER AFTER THE SECOND SURGERY SINCE RE-ACCOMMODATION OF THE ICL DID NOT CHANGE THE BEHAVIOR OF THE LENS." WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: PREVIOUS CODE IS NOT APPLICABLE, KERATOCONUS IS NO LONGER CONSIDERED A CONTRAINDICATION. (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -1.5/+3.5/170 | OS | Physician | 11/30/2017 | 2017 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5375453&pc=MTA | 2023826-2016-00050 |  | Injury | STAAR SURGICAL COMPANY | 2016-01-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Cyst(s); Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/091 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD AN EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND LENS ROTATION. THE PATIENT EXPERIENCED GLARES/HALOS. THE REPORTER INDICATED THE LENS ROTATION MAY HAVE BEEN CAUSED BY A CILIARY CYST. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -11.0/+2.5/091 | OD | Physician | 12/01/2015 | 2015 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7047145&pc=MTA | 2023826-2017-01747 |  | Injury | STAAR SURGICAL COMPANY | 2017-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cyst(s); No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.5/+3.0/092 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND MASSIVE CYST TEMPORAL. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+3.0/092 | OD | Physician | 10/23/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7047225&pc=MTA | 2023826-2017-01750 |  | Injury | STAAR SURGICAL COMPANY | 2017-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cyst(s); No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+4.0/083 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND MASSIVE CYST TEMPORAL. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CONCLUSION CODE CHANGED: DEVICE FAILURE RELATED TO PATIENT CONDITION (MASSIVE CYST TEMPORAL). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED. (B)(4). | -9.50/+4.0/083 | OS | Physician | 10/23/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8129216&pc=MTA | 2023826-2018-01760 | 2018-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Discomfort; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -8.00/1.5/107 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. IT WAS REPORTED THAT ON (B)(6) 2018, PATIENT COMPLAINED OF DISCOMFORT. ON (B)(6) 2018 SURGEON TRIED TO REPOSITIONING THE LENS AND NOTICED THE LENS DAMAGED, THE LENS WAS REMOVED AND EXCHANGED WITHIN THE SAME SURGERY. IT WAS REPORTED THAT THERE WAS NO PATIENT INJURY AND THE PATIENTS CONDITION IS GOOD. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.00/1.5/107 | OS |  | 11/06/2018 | 2018 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6629980&pc=MTA | 2023826-2017-00958 | 2016-12-19 | Injury | STAAR SURGICAL COMPANY | 2017-06-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated | Discomfort; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.5/+3.5/X100 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, LENS DISLOCATION/SUBLUXATION, AND UNCOMFORTABLE. THE SURGEON MENTIONED THAT THE LENS WAS ROTATED 100 DEGREE, AND WAS REPOSITIONED. AFTER REPOSITIONING, THE LENS ROTATED A SECOND TIME AND DISPLACED 100 DEGREE. AS OF THE DAY OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED ON (B)(6) 2017. (B)(4). Manufacturer Narrative: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT AND A PIECE OF PAPER STUCK TO THE LENS. VISUAL INSPECTION FOUND HAPTIC TORN, MISSING PIECE OF HAPTIC, AND LENS TORN IN TWO PIECES WITH DEBRIS AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/+3.5/100 | OS | Physician | 01/20/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8133688&pc=MTA | 2023826-2018-01718 | 2018-08-25 | Injury | STAAR SURGICAL COMPANY | 2018-12-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Discomfort; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -9.5/+2.0/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO PATIENT'S RIGHT EYE ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED WITH LOW VAULT WAS OBSERVED THE LENS WAS REPOSITIONED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL BECAUSE THE PATIENT EXPERIENCED DISCOMFORT. CURRENTLY, THE PATIENT'S CONDITION IS GOOD. Manufacturer Narrative: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/+2.0/91 | OD | Physician | 11/19/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7335037&pc=MTA | 2023826-2018-00341 | 2018-01-04 | Injury | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Discomfort; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 VTICM5_12.1, -4.50/+1.0/106 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018, THE LENS WAS EXPLANTED DUE TO PATIENT BEING DISSATISFIED WITH THE VISUAL ACUITY WHEN "LOOKING AT SOMETHING AT NEAR DISTANCE." THE PATIENT ALSO EXPERIENCED SIGNIFICANT DISCOMFORT. THERE IS NO PLAN TO IMPLANT ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON LENS PRODUCT. VISUAL OBSERVATION FOUND THE LENS WITH A BENT HAPTIC AND RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -4.50/+1.0/106 | OD | Physician | 02/13/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6258308&pc=MTA | 2023826-2017-00084 | 2016-10-11 | Injury | STAAR SURGICAL COMPANY | 2017-01-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Discomfort; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/154 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6)2016, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, DISCOMFORT, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS EXCHANGE RESOLVED THE PROBLEM. AS OF THE DATE OF THIS SUBMISSION, THE LENS HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: ON (B)(6) 2016, THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/40 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/40. THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | -12.0/+1.5/154 | OD | Physician | 12/21/2016 | 2016 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6790278&pc=MTA | 2023826-2017-01272 | 2016-12-07 | Injury | STAAR SURGICAL COMPANY | 2017-08-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Discomfort; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17/+6/86 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED DISCOMFORT AND LOSS OF BCVA DUE TO PRESENCE OF A PRE-EXISTING CONDITION,KERATOCONUS EYE. ON (B)(6) 2017 THE LENS WAS EXPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -16.5/5.5/078 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED DISCOMFORT AND LOSS OF BCVA DUE TO PRESENCE OF PRE-EXISTING CONDITION, KERATOCONUS EYE. ON (B)(6) 2017 THE LENS WAS EXPLANTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). DEVICE WAS IMPLANTED ON A PATIENT WITH A PRE-EXISTING CONDITION (KERATOCONUS EYE) WHICH IS CONTRAINDICATED FOR USE IN THE DFU. (B)(4). | -17/+6/86 | OS | Physician | 07/15/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6790320&pc=MTA | 2023826-2017-01269 | 2016-12-07 | Injury | STAAR SURGICAL COMPANY | 2017-08-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Discomfort; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/5.5/078 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED DISCOMFORT AND LOSS BCVA DUE TO PRESENCE OF A PRE-EXISTING CONDITION, KERATOCONUS EYE. ON (B)(6) 2017 THE LENS WAS EXPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -17/+6/86 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED DISCOMFORT AND LOSS OF BCVA DUE TO PRESENCE OF PRE-EXISTING CONDITION KERATOCONUS EYE. ON (B)(6) 2017 THE LENS WAS EXPLANTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). DEVICE WAS IMPLANTED ON A PATIENT WITH A PRE-EXISTING CONDITION (KERATOCONUS EYE) WHICH IS CONTRAINDICATED FOR USE IN THE DFU. (B)(4). | -16.5/5.5/078 | OD | Physician | 07/15/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8042009&pc=MTA | 2023826-2018-01592 | 2018-09-20 | Injury | STAAR SURGICAL COMPANY | 2018-11-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Discomfort; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -7.5/+2.5/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO DISCOMFORT. THIS EXPLANT RESOLVED THE PROBLEM. IT WAS REPORTED THAT IN THE SURGEONS OPINION THE CAUSE OF THIS EVENT WAS DUE TO A PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.5/+2.5/92 | OS | Physician | 10/10/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8042208&pc=MTA | 2023826-2018-01591 | 2018-09-20 | Injury | STAAR SURGICAL COMPANY | 2018-11-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Discomfort; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -8.0/+1.5/86 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO DISCOMFORT. THIS EXPLANT RESOLVED THE PROBLEM. IT WAS REPORTED THAT IN THE SURGEONS OPINION THE CAUSE OF THIS EVENT WAS DUE TO A PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.0/+1.5/86 | OD | Physician | 10/10/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8182707&pc=MTA | 2023826-2018-01762 | 2018-07-12 | Injury | STAAR SURGICAL COMPANY | 2018-12-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Dizziness; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICM5_12.6, -9.00/+2.0/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT EXPERIENCED DIZZINESS AND FEELING SEASICKNESS AFTER INITIAL SURGERY. IT WAS REPORTED THAT IT WAS NOT REFRACTIVE SURPRISE. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A NON TORIC LENS. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.00/+2.0/95 | OD | Physician | 11/21/2018 | 2018 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8182688&pc=MTA | 2023826-2018-01763 | 2018-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Dizziness; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICM5_12.6, -9.00/1.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT EXPERIENCED DIZZINESS AND FEELING SEASICK AFTER INITIAL SURGERY. IT WAS REPORTED THAT IT WAS NOT REFRACTIVE SURPRISE. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A NON TORIC LENS. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.00/1.5/082 | OS |  | 11/21/2018 | 2018 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7555304&pc=MTA | 2023826-2018-00832 | 2018-04-04 | Injury | STAAR SURGICAL COMPANY | 2018-05-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Edema; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.00/+3.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PATIENT RELATED FACTOR OF VERTIGO. THE VERTIGO WAS PRE-EXISTING TO ICL IMPLANTATION. AT ONE DAY POST-OP, THE CRYSTALLINE LENS WAS CLEAR AND THERE WAS 1+ INCISIONAL EDEMA. THE VERTIGO HAS BEEN REDUCED AND PATIENT WILL RETURN TO C/L WEAR. THE PATIENT ELECTED TO HAVE THE ICL EXPLANTED . Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S.BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS RESIDUE ON THE LENS. (B)(4). | -14.00/+3.0/094 | OS | Physician | 04/30/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7555311&pc=MTA | 2023826-2018-00831 | 2018-04-04 | Injury | STAAR SURGICAL COMPANY | 2018-05-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Edema; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.50/+3.5/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PATIENT RELATED FACTOR OF VERTIGO. THE VERTIGO WAS PRE-EXISTING TO ICL IMPLANTATION. AT ONE DAY POST-OP, THE CRYSTALLINE LENS WAS CLEAR AND THERE WAS 1+ INCISIONAL EDEMA. THE VERTIGO HAS BEEN REDUCED AND PATIENT WILL RETURN TO C/L WEAR. THE PATIENT ELECTED TO HAVE THE ICL EXPLANTED . Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S.BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS RESIDUE ON THE LENS. (B)(4). | -12.50/+3.5/097 | OD | Physician | 04/30/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6278117&pc=MTA | 2023826-2017-00073 | 2016-11-16 | Injury | STAAR SURGICAL COMPANY | 2017-01-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Edema; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -13.5/+1.00/090 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. THE PATIENT HAD SOME MILD WOUND EDEMA POST-OP. Manufacturer Narrative: IT SHOULD BE NOTED THAT AT THE TIME OF THE ORIGINAL IMPLANTATION THE PATIENT WAS (B)(6) YEARS OLD AND PER DFU INDICATIONS: "VISIAN® IMPLANTABLE COLLAMER® LENSES (VICL) ARE INDICATED FOR USE IN PHAKIC EYE TREATMENTS IN ADULTS 21-45 YEARS OF AGE", AND THEREFORE, THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR STICKY RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+1.00/090 | OD | Physician | 01/05/2017 | 2017 | 53
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6278120&pc=MTA | 2023826-2017-00074 | 2016-11-16 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Edema; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -13.5/+1.00/078 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. THE PATIENT HAD SOME MILD WOUND EDEMA POST-OP. Manufacturer Narrative: IT SHOULD BE NOTED THAT AT THE TIME OF THE ORIGINAL IMPLANTATION THE PATIENT WAS (B)(6) AND PER DFU INDICATIONS: "VISIAN® IMPLANTABLE COLLAMER® LENSES (VICL) ARE INDICATED FOR USE IN PHAKIC EYE TREATMENTS IN ADULTS 21-45 YEARS OF AGE", AND THEREFORE, THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN PATIENTS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF DEBRIS, CLEAR AND GREY COLORED RESIDUE. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. SECONDARY SURGERY, LENS EXCHANGED FOR A LONGER LENS. (B)(4). | -13.5/+1.00/078 | OS | Physician | 01/05/2017 | 2017 | 53
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5355967&pc=MTA | 2023826-2016-00031 | 2015-12-09 | Injury | STAAR SURGICAL COMPANY | 2016-01-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | Endophthalmitis | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 -5.0/+1.0/X088 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. SUSPECTED ENDOPHTALMITIS WAS NOTED ON (B)(6) 2015. THE LENS REMAINS IMPLANTED. POSSIBLY PLAN TO EXCHANGE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -5.00/+1.0/X088. DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -5.0/+1.0/088 | OS | Physician | 12/10/2015 | 2015 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4618716&pc=MTA | 2023826-2015-00362 | 2015-01-25 | Injury | STAAR SURGICAL COMPANY | 2015-03-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Endophthalmitis; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO SUSPECTED ENDOPHTHALMITIS AND THE PROBLEM WAS RESOLVED. PATIENT WAS ADMINISTERED INTRAOCULAR INJECTION (MEDICATION). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - (INCISION OPENED); (SECONDARY SURGERY); ENDOPHTHALMITIS; (LENS EXPLANTED); , EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): UPON REVIEW OF THE DHR, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION, AND PACKAGING PROCESSES IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON THE INFORMATION PROVIDED BY THE FACILITY, THE LENS WAS NOT INDICATED TO HAVE BEEN RECEIVED DAMAGED. THE RETURNED PRODUCT MET SPECIFICATION AND NO DAMAGE WAS REPORTED. IT IS UNCLEAR IF THE LENS ITSELF CAUSED OR CONTRIBUTED TO THE ONSET OF "SUSPECTED ENDOPHTHALMITIS". THEREFORE, NO ROOT CAUSE CAN BE IDENTIFIED. PROBABLE CAUSES FOR THIS EVENT COULD INCLUDE PRE-EXISTING PATIENT CONDITION OR SURGERY FACILITY OR CONDITION. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -11.0/+1.0/083 | OD | Other | 02/20/2015 | 2015 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7144221&pc=MTA | 2023826-2017-01946 | 2017-11-29 | Injury | STAAR SURGICAL COMPANY | 2017-12-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Endophthalmitis; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.0/+5.5/145 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED AT 8 DAYS POST-OP, THE PATIENT HAD SUSPECTED ENDOPHTHALMITIS. THE PATIENT LIVED FAR AWAY FROM THE SURGERY CENTER, SO ALL CONTROLS AND TREATMENTS (INTRAOCULAR INJECTION (MEDICATION)) WERE ADMINISTERED BY THEIR LOCAL OPHTHALMOLOGIST. THE REPORTER INDICATED THE PATIENT IS GETTING BETTER AND THE LENS REMAINS IMPLANTED. THE PATIENT WILL BE MONITORED. Manufacturer Narrative: PATIENT EYE WAS IMPECCABLE WITH NO SIGNS OF INFLAMMATION AND VISUAL ACUITY WAS BACK TO NORMAL. CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH KERATOCONUS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER INCLUDING THE SUSPECTED DEVICE HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). | -8.0/+5.5/145 | OD | Physician | 12/04/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7676112&pc=MTA | 2023826-2018-01030 | 2016-03-22 | Injury | STAAR SURGICAL COMPANY | 2018-07-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Endophthalmitis; Visual Disturbances; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, SUSPECTED ENDOPHTHALMITIS, AND PIGMENT DISPERSION. THE PROBLEM WAS RESOLVED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS WAS EXCHANGED DUE TO ENDOPHTHALMITIS, EXCESSIVE VAULT, AND PIGMENT DISPERSION. CORRECTED-THE LENS WAS EXCHANGED DUE TO EXCESSIVE VAULT AND PIGMENT DISPERSION ONLY. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE [PRE-EXISTING CONDITION-NARROW ANTERIOR CHAMBER ANGLES (I.E. GRADE II OR LESS), ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.8MM)]. (B)(4) | -15.5/+1.5/090 | OD | Physician | 06/19/2018 | 2018 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7419929&pc=MTA | 2023826-2018-00537 | 2018-03-07 | Injury | STAAR SURGICAL COMPANY | 2018-04-10 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Glaucoma; Intraocular Pressure Increased; Iritis; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.5/+2.5/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS IRITIS. REPORTEDLY, THE PATIENT WAS DIAGNOSED WITH AN ALLERGY TO BRIMONIDINE ON (B)(6) 2018. ON (B)(6) 2018 SHE WAS DIAGNOSED WITH REBOUND IRITIS. AFTER BEGINNING STEROID TREATMENT THE PATIENT IS IMPROVING. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL/CORRECTED INFORMATION: REPORTER STATES THAT, WHILE BEING TREATED FOR IRITIS, THE PATIENT DEVELOPED ELEVATED IOP. THIS WAS TREATED WITH TRAVATAN Z. THE PATIENT WAS ALLERGIC TO THIS MEDICATION, WHICH WAS DISCONTINUED AND WAS GIVEN COSOPT PF. THE PATIENT HAS NOW REPORTEDLY DEVELOPED, "SIGNIFICANT ONGOING PIGMENT DISPERSION, PIGMENTARY GLAUCOMA, WITH IOPS IN THE 30'S DESPITE MAXIMALLY TOLERATED MEDICAL THERAPY." THE PATIENT IS ALSO CURRENTLY PHOTOPHOBIC. THE ICL WAS EXPLANTED ON (B)(6) 2018; STEROID INCREASED TO EVERY 2 HOURS AND CYCLOGEL GIVEN FOR PHOTOPHOBIA. (B)(4). METHOD AND RESULTS CODES CORRECTED. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (STEROID TREATMENT). CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.89MM)] (B)(4). | -11.5/+2.5/093 | OS | Physician | 03/15/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7420014&pc=MTA | 2023826-2018-00535 | 2018-03-07 | Injury | STAAR SURGICAL COMPANY | 2018-04-10 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Glaucoma; Intraocular Pressure Increased; Iritis; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.5/+2.0/109 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS IRITIS. REPORTEDLY, THE PATIENT WAS DIAGNOSED WITH AN ALLERGY TO BRIMONIDINE ON (B)(6) 2018. ON (B)(6) 2018 SHE WAS DIAGNOSED WITH REBOUND IRITIS. AFTER BEGINNING STEROID TREATMENT THE PATIENT IS IMPROVING. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL/CORRECTED INFORMATION: REPORTER STATES THAT, WHILE BEING TREATED FOR IRITIS, THE PATIENT DEVELOPED ELEVATED IOP. THIS WAS TREATED WITH TRAVATAN Z. THE PATIENT WAS ALLERGIC TO THIS MEDICATION, WHICH WAS DISCONTINUED AND WAS GIVEN COSOPT PF. THE PATIENT HAS NOW REPORTEDLY DEVELOPED, "SIGNIFICANT ONGOING PIGMENT DISPERSION, PIGMENTARY GLAUCOMA, WITH IOPS IN THE 30'S DESPITE MAXIMALLY TOLERATED MEDICAL THERAPY." THE PATIENT IS ALSO CURRENTLY PHOTOPHOBIC. THE ICL WAS EXPLANTED ON (B)(6) 2018; STEROID INCREASED TO EVERY 2 HOURS AND CYCLOGEL GIVEN FOR PHOTOPHOBIA. (B)(4). METHOD AND RESULTS CODES CORRECTED. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -10.5/+2.0/109 | OD | Physician | 03/15/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7419890&pc=MTA | 2023826-2018-00531 | 2018-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-04-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Glaucoma; Intraocular Pressure Increased; Iritis; Visual Impairment; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6, -14.5/+2.5/166 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED, REPORTEDLY DUE TO PIGMENT DISPERSION, UNREACTIVE (MID-DILATED FIXED) PUPIL, AND IRITIS. A YAG PI WAS PERFORMED ON (B)(6) 2018. THE SURGEON STATED HE THOUGHT THE PATIENT MAY HAVE HAD AN ASYMPTOMATIC PUPIL BLOCK DUE TO THE VISCOELASTIC USED IN SURGERY, BUT IS UNABLE TO THEN EXPLAIN WHY IT TOOK 2-3 WEEKS FOR IRITIS AND PIGMENT DISPERSION TO DEVELOP. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: AFTER THE LENS EXPLANT, THE PATIENT WAS ADVISED TO TAKE PRED/GATI 3 TIMES DAILY. AFTER DISCONTINUING, THE PATIENT DEVELOPED AN INCREASE IN IOP AND INFLAMMATION IN THE ANTERIOR CHAMBER. CURRENTLY, THE PATIENT IS BEING TREATED WITH PREDNISOLONE AND COMBIGAN, AZOPT, AND TRAVATAN Z FOR PRESSURE SPIKES. IT IS BELIEVED THAT THE SPIKES ARE DUE TO "STEROID USE VS INFLAMMATORY CAUSES." REPORTEDLY, NO INFECTION HAS OCCURRED WITHIN THE EYE. (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE PATIENT REPORTS "PRESSURE ISSUES," "BLOWN IRIS," ONGOING GLAUCOMA, AND LOSS OF VISION. ALSO STATED IS THAT THE IRIS IS PERMANENTLY FIXED AND DILATED. GLAUCOMA SURGERY PERFORMED. G3-REPORT SOURCE ADDED. H6-ADDITIONAL 3191: (BLOWN IRIS, FIXED IRIS, GLAUCOMA SURGERY). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO CODE AVAILABLE (PIGMENT DISPERSION, SECONDARY SURGERY, LENS EXPLANT, FIXED PUPIL) OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT (46YRS), ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.8MM)]. (B)(4). | -14.5/+2.5/166 | OS |  | 03/15/2018 | 2018 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6249746&pc=MTA | 2023826-2017-00051 | 2016-03-20 | Injury | STAAR SURGICAL COMPANY | 2017-01-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Glaucoma; Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/+1.0/098 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ANGLE CLOSURE (WITH ELEVATED IOP), AND DEVELOPMENT OF GLAUCOMA. AS OF THE DATE OF THIS SUBMISSION, THE LENS HAS BEEN RETURNED, BUT THE EVALUATION RESULTS ARE NOT YET AVAILABLE. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION WAS PERFORMED. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.0/+1.0/098 | OD | Physician | 04/20/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7419948&pc=MTA | 2023826-2018-00533 | 2018-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-04-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Glaucoma; Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -14.0/+1.0/007 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS PUPIL BLOCK, ANGLE CLOSURE, ELEVATED IOP, AND PIGMENT DISPERSION. THE SURGEON STATED HE THOUGHT THE HIGH IOP WAS DUE TO THE VISCOELASTIC. REPORTEDLY, HE ASSUMES THE AQUA PORT WAS BLOCKED BY THE VISCOELASTIC. A YAG WAS PERFORMED AND THEN IOP RETURNED TO "NORMAL." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THE PATIENT REPORTS THE LENS WAS EXPLANTED. ALSO REPORTED: "PRESSURE ISSUES," "BLOWN IRIS," ONGOING GLAUCOMA AND VISION LOSS, GLAUCOMA SURGERY PERFORMED, IRIS PERMANENTLY FIXED AND DILATED. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (PIGMENT DISPERSION, SECONDARY SURGERY, ANGLE CLOSURE, YAG) CONCLUSION CODE: OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT ((B)(6)), ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.67MM)] (B)(4). | -14.0/+1.0/007 | OD |  | 03/15/2018 | 2018 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7294520&pc=MTA | 2023826-2018-00267 | 2018-01-23 | Injury | STAAR SURGICAL COMPANY | 2018-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Insufficient Information | Glaucoma; Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | Event Description: MALIGNANT GLACOMA WITH EXPLANT WAS REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 VTICM5_13.2, -13.5/+1.5/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED 5 HOURS LATER DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PUPIL BLOCK WITH ELEVATED IOP (INTRAOCULAR PRESSURE). ADDITIONAL SURGICAL PROCEDURES (ENLARGEMENT OF PI OR ADDITION OF NEW PI AND YAG) WERE REPORTED TO BE PERFORMED. ENTERED STAAR (B)(4)'S ADDRESS DEVICE EVALUATION: LENS WAS RETURNED IN A CONTACT CASE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4). | -13.5/+1.5/95 | OD | Physician | 01/24/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5440040&pc=MTA | 2023826-2016-00152 | 2014-07-02 | Injury | STAAR SURGICAL COMPANY AG | 2016-02-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -16.0/5.5/082 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO GLARES/HALOS AND THE PATIENT WAS NOT SATISFIED WITH THE BEST-CORRECTED VISUAL ACUITY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC TORN AND BENT. FOREIGN MATERIAL IN THE FORM OF FIBERS WERE PRESENT ON THE LENS SURFACE. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | -16.0/5.5/082 | OD | Physician | 12/11/2015 | 2015 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6043822&pc=MTA | 2023826-2016-01111 | 2016-02-20 | Injury | STAAR SURGICAL COMPANY | 2016-10-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -06.50/+1.5/088 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARE/HALOES. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/ VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT OPTIC WAS TORN/ SPLIT, HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). | -06.50/+1.5/088 | OD | Physician | 06/14/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6043825&pc=MTA | 2023826-2016-01112 | 2016-02-20 | Injury | STAAR SURGICAL COMPANY | 2016-10-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.00/+2.5/010 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARE/HALOES. Manufacturer Narrative: (B)(4). PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN, BUT THERE WAS FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS SPOT. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -08.00/+2.5/010 | OS | Physician | 06/14/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5909288&pc=MTA | 2023826-2016-01176 | 2016-04-22 | Injury | STAAR SURGICAL COMPANY | 2016-08-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+2.5/005 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO PIGMENT DISPERSION, GLARE AND HALOS. THE LENS WAS NOT REPLACED WITH ANOTHER ICL LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | -14.0/+2.5/005 | OD | Physician | 07/26/2016 | 2016 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4767214&pc=MTA | 2023826-2015-00519 |  | Injury | STAAR SURGICAL COMPANY | 2015-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | NA | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN HIS RIGHT EYE (OD). THE PATIENT REPORTED HE HAS BAD NIGHT VISION (GLARES/HALOS). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -14.0/+5.0/103 | OD | Patient | 04/13/2015 | 2015
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6404331&pc=MTA | 2023826-2017-00389 | 2016-10-28 | Injury | STAAR SURGICAL COMPANY | 2017-03-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/35 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2016 THE LENS WAS EXPLANTED. AS THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: LENS WAS RETURNED IN A CENTRIFUGE VIAL. VISUAL INSPECTION FOUND CLEAR RESIDUE AND DEBRIS ON LENS SURFACE. LENS IS FOLDED AT THE HAPTICS AND OPTIC IS TWISTED DUE TO POSITION IN VIAL. NO VISIBLE TEARS OR SCRATCHES. UNABLE TO FURTHER EVALUATE DUE TO LENS SHAPE. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4). | -9.5/+1.5/35 | OD | Physician | 02/13/2017 | 2017 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5577529&pc=MTA | 2023826-2016-00440 | 2016-01-23 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -12.5/+2.50/X095 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED GLARES/HALOES. THE ICL WAS EXPLANTED ON (B)(6) 2016 AND WAS NOT REPLACED. NO OTHER LENS WAS IMPLANTED. LAST VISIT, BCVA WAS 20/20. SEE MFR # 2023826-2016-00441 FOR LEFT EYE. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT WITH FOREIGN MATERIAL ON LENS SURFACE (FIBERS) AND A SCRATCH. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -12.5/+2.50/X095. (B)(4). | -12.5/+2.50/095 | OD | Physician | 03/15/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5577591&pc=MTA | 2023826-2016-00441 | 2016-01-23 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -12.00/+3.00/X091 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED GLARES/HALOES. THE ICL WAS EXPLANTED ON (B)(6) 2016 AND WAS NOT REPLACED. NO OTHER LENS WAS IMPLANTED. LAST VISIT, BCVA WAS 20/20. SEE MFR # 2023826-2016-00440 FOR RIGHT EYE. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CHANGE "THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 -12.00/+3.00/X91 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016." TO "THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 -12.00/+3.00/X91 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016." CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -12.00/+3.00/X091. (B)(4). | -12.00/+3.00/091 | OS | Physician | 03/16/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5938281&pc=MTA | 2023826-2016-01247 | 2016-04-22 | Injury | STAAR SURGICAL COMPANY | 2016-09-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+0.5/005 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO PIGMENT DISPERSION, GLARE AND HALOS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. A VISUAL INSPECTION WAS PERFORMED, NO DAMAGE TO THE LENS. UPON DIMENSIONAL INSPECTION, THE LENS WAS FOUND TO BE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -14.0/+0.5/005 | OS | Physician | 08/10/2016 | 2016 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6085625&pc=MTA | 2023826-2016-00882 | 2016-01-26 | Injury | STAAR SURGICAL COMPANY | 2016-11-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00/+1.0/073 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. ON (B)(6) 2016, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING OF THE LENS, GLARE/HALOES, AND A SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE/ DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND HAPTIC BROKEN/ BENT. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: ONE SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -08.00/+1.0/073 | OD | Physician | 05/27/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6701679&pc=MTA | 2023826-2017-01154 | 2017-01-17 | Injury | STAAR SURGICAL COMPANY | 2017-07-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICM5_13.7, DIOPTER -5.5/+1.0/089 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, HALOES, AND WIDE PUPIL. THE SURGEON COMMENTED THAT THE PATIENT'S PUPIL > 8 MM ON (B)(6) 2017, AND PLANS TO EXCHANGE THE LENS FOR THE PATIENT. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (DEBRIS). (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND GLARE/HALOES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP INDICATED NORMAL PUPIL SIZE, BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) TO BE 20/20. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN US BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.5/+1.0/089 | OD | Physician | 06/14/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6836431&pc=MTA | 2023826-2017-01376 | 2017-07-24 | Injury | STAAR SURGICAL COMPANY | 2017-08-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/140 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND THE PATIENT EXPERIENCED GLARE/HALOS. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS. THE PATIENT IS OK. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+2.0/140 | OS | Physician | 08/03/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4724366&pc=MTA | 2023826-2015-00456 | 2015-02-26 | Injury | STAAR SURGICAL COMPANY AG | 2015-04-23 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2007. LOW VAULT WAS NOTED, AS WELL AS REFRACTIVE CHANGE OVER TIME (INCREASED MYOPIA), AND GLARE/HALOS. A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. EXPLANT DATE: N/A. (B)(4). DEVICE EVALUATED BY THE MANUFACTURER? NO - 81(OTHER): LENS REMAINS IMPLANTED. (B)(4). LENS REMAINS IMPLANTED. | -8/+4/93 | OD | Physician | 03/27/2015 | 2015 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5788339&pc=MTA | 2023826-2016-00883 | 2016-01-26 | Injury | STAAR SURGICAL COMPANY | 2016-07-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/073 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND GLARE/HALOES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/13. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -8.0/+1.0/073 | OS | Physician | 05/27/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5462012&pc=MTA | 2023826-2016-00199 |  | Injury | STAAR SURGICAL COMPANY | 2016-02-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND LENS DISLOCATION/SUBLUXATION. THE PATIENT EXPERIENCED GLARE/HALOS. THE LENS WAS EXPLANTED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED. (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -10.00/+2.0/080 | OD | Physician | 01/28/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4600863&pc=MTA | 2023826-2015-00311 | 2015-02-10 | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE REPORTER INDICATED THE LENS HAD A REFRACTIVE SURPRISE AND THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, HEADACHE AND BLURRED VISION. THE LENS REMAINS IMPLANTED. THE PATIENT WAS TREATED FOR CHOROIDAL EFFUSION (SUSPECT) WITH TOPIC CORTICOSTEROIDS, THE ANGLES WERE COMING BACK TO PREVIOUS AND THERE WAS GOOD EVOLUTION OF THE PATIENT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. EXPLANT DATE: NA. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -0.50/+3.5/125 | OD | Other | 02/12/2015 | 2015 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5199928&pc=MTA | 2023826-2015-01357 | 2015-09-17 | Injury | STAAR SURGICAL COMPANY | 2015-11-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Headache; Intraocular Pressure Increased | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+2.0/X093 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ELEVATED INTRAOCULAR PRESSURE (UP TO 25MMHG) STARTED BY THE 4TH DAY. THE PATIENT WAS GIVEN PRESSURE DROPS, THE PRESSURE WENT DOWN TO 22 BUT THEN IT WENT UP TO BETWEEN 32MMHG AND 50MMHG. THE PATIENT'S LAST UCVA WAS 20/20 AND IS SATISFIED WITH THE ICL BUT WAS COMPLAINING OF HEADACHES. THE PHYSICIAN STATED THE PATIENT HAS A HEAVILY PIGMENTED IRIS. RECENT IMAGES SHOWS GOOD VAULT AND NORMAL IOP BUT THE PATIENT IS STILL ON MEDICATION. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+2.0/X093 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS OBSERVED AND ELEVATED INTRAOCULAR PRESSURE (IOP) STARTED BY THE 4TH DAY. THE SURGEON ESTABLISHED MEDICAL MANAGEMENT AND IOP WAS CONTROLLED TO NORMAL LEVELS WITH ONE GLAUCOMA MEDICATION. PATIENT ALSO REPORTED HEADACHES. THE PHYSICIAN STATED THE PATIENT HAS A HEAVILY PIGMENTED IRIS. ON LAST VISIT, PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/30, VAULT WAS NORMAL, AND IOP WAS NORMAL. THE LENS REMAINS IMPLANTED. CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH BELOW 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT IS MARKETED IN THE U.S. DIOPTER: -17.0/+2.0/X093. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -17.0/+2.0/093 | OS | Physician | 10/07/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6444673&pc=MTA | 2023826-2017-00527 | 2017-03-09 | Injury | STAAR SURGICAL COMPANY | 2017-03-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Headache; Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.00/+6.0/085 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS ROTATION, ELEVATED IOP, POOR VISION, AND HEADACHES. REPOSITIONING WAS PERFORMED (B)(6) 2016 AND THE SURGEON NOTICES A GRADUAL INCREASE IN IOP. THE PROBLEM IS NOT RESOLVED AND A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+6.0/85 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS ROTATION, ELEVATED IOP, POOR VISION, AND HEADACHES. REPOSITIONING WAS PERFORMED ON (B)(6) 2016 AND THE SURGEON NOTICED A GRADUAL INCREASE IN IOP. THE LENS WAS EXPLANTED ON (B)(6) 2017 AND EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. (B)(4). CLAIM #(B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+6.0/85 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS ROTATION, ELEVATED IOP, POOR VISION, AND HEADACHES. REPOSITIONING WAS PERFORMED ON (B)(6) 2016 AND THE SURGEON NOTICED A GRADUAL INCREASE IN IOP. THE LENS WAS EXPLANTED ON (B)(6) 2017 AND EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. (B)(4). CLAIM #(B)(4). | -7.00/+6.0/085 | OS | Physician | 03/01/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5112582&pc=MTA | 2023826-2015-01188 | 2015-04-09 | Injury | STAAR SURGICAL COMPANY | 2015-09-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.30/+1.5/075 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO PATIENT SUFFERED FROM HEADACHE. NO OTHER LENS WAS IMPLANTED. SEE MFR #2023826-2015-01187 FOR RIGHT EYE. Manufacturer Narrative: BASED ON THE DHR REVIEW, THE MANUFACTURING AND PACKAGING PROCESSES WAS PERFORMED ACCORDING TO PROCEDURES AND NO FINDINGS OR DEVIATIONS WERE REPORTED. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED, HOWEVER, THE COMPLAINT IS NOT DEVICE RELATED. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THAT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND TWO OF THE FOOTPLATES WERE MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -9.30/+1.5/X075. (B)(4). METHOD CODE: EVALUATION METHOD: WORK ORDER SEARCH. EVALUATION CODE: EVALUATION RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS CODE: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -9.30/+1.5/075 | OS | Physician | 09/01/2015 | 2015 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5146314&pc=MTA | 2023826-2015-01251 | 2015-08-25 | Injury | STAAR SURGICAL COMPANY | 2015-10-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -17.5/+2.0/X087 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE PATIENT EXPERIENCING DISCOMFORT, A HEADACHE. NO OTHER LENS WAS IMPLANTED. THE PATIENT DID NOT GO FOR A FOLLOW-UP. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -17.5/+2.0/X087. LENS NOT RETURNED. (B)(4). A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -17.5/+2.0/087 | OS | Physician | 09/15/2015 | 2015 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5151588&pc=MTA | 2023826-2015-01250 | 2015-08-25 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.50/+2.50/X061 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT EXPERIENCED DISCOMFORT (HEADACHE) AND THE LENS WAS EXPLANTED ON (B)(6) 2015. NO OTHER LENS WAS IMPLANTED. THE PATIENT DID NOT GO FOR FOLLOW UP VISIT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -14.50/+2.50/X061. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | -14.50/+2.50/061 | OD | Physician | 09/15/2015 | 2015 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6587675&pc=MTA | 2023826-2017-00873 | 2016-09-21 | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Headache; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -4.0/+0.5/104 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, IRIS ATROPHY, AND HEADACHE. THE SURGEON COMMENTED THE VAULT ONE MONTH AFTER THE SURGERY IS 600 MICRONS, AND 3 MONTHS AFTER IS 902 MICRONS. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION - PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND CLEAR RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, SIGNIFICANT REDUCTION OF ICA, IRIS ATROPHY). (B)(4). | -4.0/+0.5/104 | OD | Physician | 04/25/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8298844&pc=MTA | 2023826-2019-00191 | 2018-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+2.0/093 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING. THE PATIENT COMPLAINED OF A HEADACHE ON LEFT SIDE AFTER SURGERY. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2018. ANOTHER LENS WAS NOT IMPLANTED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION-THE LENS WAS RETURNED DRY IN LENS CASE/VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.50/+2.0/093 | OS |  | 01/07/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6694745&pc=MTA | 2023826-2017-01119 | 2017-01-04 | Injury | STAAR SURGICAL COMPANY | 2017-07-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Headache; Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -6.5/+1.0/95 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED HEADACHE, GLARE/HALOES. ON (B)(6) 2017 THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. THE CUSTOMER STATED THAT THE PATIENT IS NO LONGER HEADACHE AFTER THE LENS EXPLANT, AND DOES NOT HAVE PLAN TO EXCHANGE THE LENS. AS OF THE DAY OF MDR SUBMISSION THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS WAS RETURNED IN MICRO CENTRIFUGE VIAL WITH LITTLE MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4) WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.5/+1.0/95 | OD | Physician | 06/12/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6694750&pc=MTA | 2023826-2017-01120 | 2017-01-04 | Injury | STAAR SURGICAL COMPANY | 2017-07-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Headache; Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -5.0/+2.0/90 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED HEADACHE, GLARE/HALOES. ON (B)(6) 2017 THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. THE CUSTOMER STATED THAT THE PATIENT IS NO LONGER HEADACHE AFTER THE LENS EXPLANT, AND DOES NOT HAVE PLAN TO EXCHANGE THE LENS. AS OF THE DAY OF MDR SUBMISSION THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS WITH A CURVED SHAPE. THERE WAS EVIDENCE OF CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.0/+2.0/90 | OS | Physician | 06/12/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7186676&pc=MTA | 2023826-2018-00014 | 2014-03-18 | Injury | STAAR SURGICAL COMPANY | 2018-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/077 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2014. AT ONE DAY POST-OP, INFLAMMATION WAS NOTED. THE SURGEON ADMINISTERED ANTIBIOTICS AND THE EYE RETURNED TO NORMAL. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE EVENT WAS NOT LENS RELATED. Manufacturer Narrative: THE REPORTER INDICATED THE PATIENT WAS GIVEN STEROIDS AFTER INFLAMMATION. NOW EVERYTHING IS FINE AND THE LENS IS STILL IN THE EYE. AT THE LAST VISIT ON (B)(6) 2017, THE PATIENT UCVA WAS 20/20. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/077 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2014. IMPLANTED DATE: (B)(6) 2014 CORRECTED TO (B)(6) 2014. (B)(4). Manufacturer Narrative: DATE OF EVENT: CORRECTED FROM (B)(6) 2014 TO (B)(6) 2014. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.0/+2.0/077 | OD | Physician | 12/18/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7186699&pc=MTA | 2023826-2018-00015 | 2014-03-18 | Injury | STAAR SURGICAL COMPANY AG | 2018-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+4.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2014. AT ONE DAY POST-OP, INFLAMMATION WAS NOTED. THE SURGEON ADMINISTERED ANTIBIOTICS AND THE EYE RETURNED TO NORMAL. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE EVENT WAS NOT LENS RELATED. Manufacturer Narrative: THE REPORTER INDICATED THE PATIENT WAS GIVEN STEROIDS AFTER INFLAMMATION. NOW EVERYTHING IS FINE AND THE LENS IS STILL IN THE EYE. AT THE LAST VISIT ON (B)(6) 2017, THE PATIENT UCVA WAS 20/20. (B)(4) Manufacturer Narrative: DATE OF EVENT: CORRECTED FROM (B)(6) 2014. THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+4.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. IMPLANTED DATE: CORRECTED FROM (B)(6) 2014 TO (B)(6) 2014. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND AND NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.0/+4.0/090 | OS | Physician | 12/18/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7987323&pc=MTA | 2023826-2018-01529 | 2018-09-11 | Injury | STAAR SURGICAL COMPANY | 2018-10-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0/+1.0/050 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2016. THE REPORTER INDICATED LOSS OF BEST-CORRECTED VISUAL ACUITY (BCVA), PIGMENT DISPERSION, CELL DEPOSIT AND INFLAMMATION. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.0/+1.0/050 | OS | Physician | 09/21/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13983668&pc=MTA | 2023826-2022-01035 | 2022-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Inflammation; Blurred Vision; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+0.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES (ICA), DISTANCE AND NEAR VISION SLIGHTLY WORSE THAN THE OD EYE, BLEPHARITIS, IRIS BOMBE WITH PREVIOUS APPARENT TOUCH POINT SUPERIOR TO PUPIL, "VH 1/0," NIGHTTIME HALOS WHEN DRIVING AND ANY CEILING LIGHTS, AND IRIS THINNED WITH PIGMENT ON ICL. THE SURGEON STATES, "VERY NARROW ANGLE...ANGLE SHUT OFF." THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "LENS TOO BIG." Manufacturer Narrative: B5- PER UPDATED INFORMATION THE LENS WAS EXPLANTED ON (B)(6) 2022. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED, AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.5/+0.5/074 | OS | Physician | 03/08/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6422330&pc=MTA | 2023826-2017-00444 | 2017-01-31 | Injury | STAAR SURGICAL COMPANY | 2017-03-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Inflammation; Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+3.5/38 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), SIGNIFICANT REDUCTION OF ICA, ACCOMMODATIVE SPASM, AND BLURRED VISION. INTERVENTIONS INCLUDE: ENLARGEMENT OF PI (OR ADDITION OF NEW PI), IRIDECTOMY, AND PILOCARPINE TREATMENT. ON (B)(6) 2017, THE LENS WAS EXPLANTED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR THICK AND REDDISH BROWN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+3.5/38 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), SIGNIFICANT REDUCTION OF ICA, ACCOMMODATIVE SPASM, AND BLURRED VISION. INTERVENTIONS INCLUDED ENLARGEMENT OF PI (OR ADDITION OF NEW PI), IRIDECTOMY, AND PILOCARPINE TREATMENT. AFTER LENS EXPLANT, PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/100 AND THE PATIENT EXPERIENCED A LITTLE INFLAMMATION OF THE ANTERIOR SEGMENT. DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH-NO SIMILAR COMPLAINTS FOUND. CONCLUSION: OFF-LABEL USE-KERATOCONUS EYE. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | -8.5/+3.5/38 | OS | Physician | 02/21/2017 | 2017 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721249&pc=MTA | 2023826-2018-01084 | 2018-06-13 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Inflammation; Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+4.0/082 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT, REFRACTIVE SURPRISE, ELEVATED IOP, ANGLE CLOSURE, INFLAMMATION, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE, THE PATIENT IS "DOING WELL." AZETAZOLAMIDE 250MG GIVEN TO MANAGE SYMPTOMS. THE SURGEON LISTS THE CAUSE OF THE EVENT AS DEVICE. Manufacturer Narrative: DATE OF THIS REPORT: IT IS CORRECTED TO 01-JUL-2018 IN THE INITIAL MDR. DATE RECEIVED BY MFR: -IT IS CORRECTED TO 01-JUL-2018 IN THE INITIAL MDR. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/+4.0/082 | OD | Physician | 07/01/2018 | 2018 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8107278&pc=MTA | 2023826-2018-01723 | 2018-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Contamination /Decontamination Problem; Unintended Movement | Inflammation; Intraocular Pressure Increased; Uveitis; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/114 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE SURGEON REPORTS ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), UVEITIS, SLIGHTLY DISTORTED PUPIL, CELL DEPOSITS ON THE ICL, AND INFLAMMATION. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. PER THE SURGEON, THE LENS "HAS BEEN SLIGHTLY TILTED FROM THE BEGINNING (PATHOLOGY IN THE SULCUS?)." CELL DEPOSITS DISAPPEAR WHEN GIVEN PRED FORTE; REAPPEAR WHEN THE "STERIOD" IS "DISCONSTINUED". REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: EVENT-ADDITIONAL DATA RECEIVED: "THE EXCHANGE RESOLVED THE INFLAMMATION, THE EYE IS CALM. VA SC 0.8, CC 1.0 VAULT 350 MICRON. ICL IN VERTICAL POSITION (AWAY FROM THE NASAL PROBLEM AREA). PUPIL SLIGHTLY NASALLY DISTORTED. THE MD ATTRIBUTES THE INFLAMMATION CLEARLY TO THE OVER VAULTED ICL. EVERYONE IS HAPPY". IMPLANT DATE: (B)(6)2018 DEVICE EVALUATED BY MFR:DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: EVENT-CORRECTED IMPLANT DATE INTO THE RIGHT EYE IS: "(B)(6) 2015" IN THE INITIAL MDR. RELEVANT TESTS/LABORATORY DATA-CORRECTED TO: "NA" IN THE SUPPLEMENTAL MDR#1 CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+1.0/114 | OD |  | 10/31/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6559303&pc=MTA | 2023826-2017-00795 | 2017-04-10 | Injury | STAAR SURGICAL COMPANY | 2017-05-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Insufficient Information | Inflammation; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+2.0/089 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO INFLAMMATION. UPON EXPLANTATION OF THE LENS, THE PROBLEM WAS RESOLVED, THE PATIENTS BCVA WAS 20/20 AND NO INFLAMMATION WAS OBSERVED. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A ALCON VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (RED RESIDUE). DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.5/+2.0/089 | OD | Physician | 04/12/2017 | 2017 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6559302&pc=MTA | 2023826-2017-00796 | 2017-04-11 | Injury | STAAR SURGICAL COMPANY | 2017-05-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Inflammation; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+2.0/093 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO INFLAMMATION. UPON EXPLANTATION OF THE LENS, THE PROBLEM WAS RESOLVED, THE PATIENTS BCVA WAS 20/28 AND NO INFLAMMATION WAS OBSERVED. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17.5/+2.0/093 | OS | Physician | 04/12/2017 | 2017 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6389173&pc=MTA | 2023826-2017-00331 | 2016-11-01 | Injury | STAAR SURGICAL COMPANY | 2017-03-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Inflammation; No Information; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT EXPERIENCED LOW VAULT. MULTIPLE ATTEMPTS HAVE BEEN MADE TO DETERMINE WHETHER THE LENS HAS BEEN EXPLANTED/EXCHANGED, STILL UNABLE TO CLARIFY. ACCORDING TO THE SURGEON, THE EYE IS "QUIET". Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016,THE LENS WAS EXPLANTED DUE TO LOW VAULT AND LENS ROTATION. THE LENS WAS THEN EXCHANGED FOR A LENS, TWO SIZES LARGER, ON (B)(6) 2016. THE SURGEON HAD DIFFICULTY EXPLANTING THE LENS DUE TO PATIENT'S POOR PUPIL DILATION. SURGEON ALSO NOTED THAT IMPLANT OF THE LARGER LENS WAS ESPECIALLY CHALLENGING DUE TO THE PATIENT'S IRIS EDGE WAS SO WIDE IT "APPEARED LIKE A CAT'S EYE." (B)(4). CLAIM # (B)(4) . Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016,THE LENS WAS EXPLANTED DUE TO LOW VAULT AND LENS ROTATION. THE LENS WAS THEN EXCHANGED FOR A LENS, TWO SIZES LARGER, ON (B)(6) 2016. THE SURGEON HAD DIFFICULTY EXPLANTING THE LENS DUE TO PATIENT'S POOR PUPIL DILATION. SURGEON ALSO NOTED THAT IMPLANT OF THE LARGER LENS WAS ESPECIALLY CHALLENGING DUE TO THE PATIENT'S IRIS EDGE WAS SO WIDE IT "APPEARED LIKE A CAT'S EYE." (B)(4). CLAIM # (B)(4) . Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016,THE LENS WAS EXPLANTED DUE TO LOW VAULT AND LENS ROTATION. THE LENS WAS THEN EXCHANGED FOR A LENS, TWO SIZES LARGER, ON (B)(6) 2016. THE SURGEON HAD DIFFICULTY EXPLANTING THE LENS DUE TO PATIENT'S POOR PUPIL DILATION. SURGEON ALSO NOTED THAT IMPLANT OF THE LARGER LENS WAS ESPECIALLY CHALLENGING DUE TO THE PATIENT'S IRIS EDGE WAS SO WIDE IT "APPEARED LIKE A CAT'S EYE." (B)(4). CLAIM # (B)(4) . | -12.0/+2.5/085 | OS | Physician | 12/16/2016 | 2016 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6422283&pc=MTA | 2023826-2017-00442 | 2017-02-20 | Injury | STAAR SURGICAL COMPANY | 2017-03-21 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Inflammation; Uveitis; Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+3.0/091 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT THEN EXPERIENCED BLURRED VISION, INFLAMMATION, CYCLOPLEGIA AND BILATERAL ANTERIOR SEGMENT UVEITIS (TESTS CAME OUT NEGATIVE). PATIENT WAS TREATED WITH PREDNEFRIN FORTE (ALLERGEN). UCVA REMAINS AT 20/20. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT AND THEREFORE, HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: UPDATED INFORMATION-ICL REMAINS IMPLANTED. PATIENT FINISHED WITH PREDNEFRIN FORTE, CURRENTLY IS NOT ON MEDICATION. THE EYE IS WHITE WITH NO ACTIVITY OR INFLAMMATION. BOTH SURGEON AND PATIENT ARE HAPPY WITH THE RESULTS. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED. | -13.00/+3.0/091 | OS | Physician | 02/20/2017 | 2017 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6422303&pc=MTA | 2023826-2017-00443 | 2017-02-20 | Injury | STAAR SURGICAL COMPANY | 2017-03-21 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Inflammation; Uveitis; Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+3.0/109 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT THEN EXPERIENCED BLURRED VISION, INFLAMMATION, CYCLOPLEGIA AND BILATERAL ANTERIOR SEGMENT UVEITIS (TESTS CAME OUT NEGATIVE). PATIENT WAS TREATED WITH PREDNEFRIN COLLYRIUM. UCVA REMAINS AT 20/20. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT AND THEREFORE, HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: ICL REMAINS IMPLANTED. PATIENT FINISHED WITH PREDNEFRIN FORTE, CURRENTLY IS NOT ON MEDICATION. THE EYE IS WHITE WITH NO ACTIVITY OR INFLAMMATION. BOTH SURGEON AND PATIENT ARE HAPPY WITH THE RESULTS. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. PATIENT CODE (B)(4) NO CODE AVAILABLE (CYCLOPEGIA, TREATMENT WITH PREDNEFRIN COLLIRIUM) WORK ORDER SEARCH PERFORMED-NO SIMILAR COMPLAINTS FOUND. (B)(4). DEVICE REMAINS IMPLANTED. | -12.50/+3.0/109 | OD | Physician | 02/20/2017 | 2017 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5355865&pc=MTA | 2023826-2016-00029 | 2015-11-13 | Injury | STAAR SURGICAL COMPANY | 2016-01-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -18.00/+1.50/X084 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. EXCESSIVE VAULTING WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED ON (B)(6) 2015. THE LENS REMAINS IMPLANTED. SEE MFR#2023826-2016-00030 FOR LEFT EYE. Manufacturer Narrative: CONCLUSION: UNABLE TO CONFIRM COMPLAINT - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -18.00/+1.50/X084. DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -18.00/+1.50/084 | OD | Physician | 11/20/2015 | 2015 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5066047&pc=MTA | 2023826-2015-01174 |  | Injury | STAAR SURGICAL COMPANY | 2015-09-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.50/-1.50/089 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD). EXCESSIVE VAULT WAS NOTED AS WELL AS ELEVATED INTRAOCULAR PRESSURE (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION IS AVAILABLE. SEE MFR # 2023826-2015-01175 FOR LEFT EYE. Manufacturer Narrative: THERE WAS NO EXCESSIVE VAULTING. THE INTRAOCULAR PRESSURE RISE WAS UNDER CONTROL AND THE PROBLEM WAS RESOLVED. NO SECONDARY SURGICAL INTERVENTION WAS NEEDED. (B)(4). CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -14.50/+1.50/089. (B)(4). METHOD:EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED. | -14.50/-1.50/089 | OD | Physician | 08/11/2015 | 2015
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4740585&pc=MTA | 2023826-2015-00486 | 2015-02-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-04-30 | MTA | VISIAN ICL (IMPLANTABLE COLLAME LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED, AS WELL AS SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES ASSOCIATED WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST KNOWN UCVA WAS 20/25 AS OF (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT IS NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGERY. NEW INCISION. REDUCTION OF IRIDO-CORNEAL ANGLES. EXPLANTED. VAULTING. DEVICE EVALUATED BY THE MANUFACTURER? NO: LENS WAS NOT RETURNED. (B)(4). Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH AND MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP, ACT.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, PRODUCT EVALUATION AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -6.50/+55/11 | OS | Other | 03/31/2015 | 2015 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6900441&pc=MTA | 2023826-2017-01484 | 2017-08-06 | Injury | STAAR SURGICAL COMPANY | 2017-09-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/082 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT AND THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (CONTROLLED BY MEDICATION). THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS FOR A SMALLER SIZE LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. EXPLANTED DATE IS CORRECTED TO "N/A" FROM "UNK" IN THE INITIAL; AS THE LENS REMAINS IMPLANTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.5/082 | OD | Physician | 09/04/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824489&pc=MTA | 2023826-2018-01252 | 2018-07-28 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -15.0/+1.5/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING. THE LENS REMAINS IMPLANTED. AT ONE WEEK REVIEW, THE PATIENTS INTRAOCULAR PRESSURE WAS 20MMHG AND PUPIL WAS REACTIVE EVEN THOUGH SEEMS SLOWER. THE PATIENT IS BEING MONITORED. Manufacturer Narrative: CORRECTED DATA: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.0/+1.5/097 | OD | Physician | 07/30/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5355968&pc=MTA | 2023826-2016-00030 | 2015-11-13 | Injury | STAAR SURGICAL COMPANY | 2016-01-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.50/+3.00/X083 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. EXCESSIVE VAULTING WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED ON (B)(6) 2015. THE LENS REMAINS IMPLANTED. SEE MFR#2023826-2016-00029 FOR RIGHT EYE. Manufacturer Narrative: CONCLUSION: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH BELOW 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -13.50/+3.00/X083. DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -13.50/+3.00/083 | OS | Physician | 11/20/2015 | 2015 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5066072&pc=MTA | 2023826-2015-01175 |  | Injury | STAAR SURGICAL COMPANY | 2015-09-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.00/2.00/103 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS). EXCESSIVE VAULT WAS NOTED AS WELL AS ELEVATED INTRAOCULAR PRESSURE (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION IS AVAILABLE. SEE MFR # 2023826-2015-01174 FOR RIGHT EYE. Manufacturer Narrative: SURGEON INDICATED THE IOP WAS UNDER CONTROL. NO SECONDARY SURGICAL INTERVENTION WAS NEEDED. LENS IMPLANTED ON (B)(6) 2015. (B)(6). (B)(4). THE PATIENT WAS UNDER THE AGE OF 21 YEARS OLD, AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -14.00/2.00/103. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED | -14.00/2.00/103 | OS | Physician | 08/11/2015 | 2015
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5185751&pc=MTA | 2023826-2015-01447 |  | Injury | STAAR SURGICAL COMPANY | 2015-10-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.50/+1.0/X067 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT AND THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (26MMHG). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -16.50/+1.0/067 | OD | Physician | 09/28/2015 | 2015
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18393718&pc=MTA | 2023826-2023-05723 | 2023-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.5/1.0/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ANGLE CLOSURE WITH ELEVATED IOP 60MMHG(ASSESSED ON (B)(6) 2023). LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - HEALTH EFFECT CLINICAL CODE 4581: ANGLE CLOSURE. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -7.5/1.0/073 | OS | Physician | 12/08/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18525773&pc=MTA | 2023826-2024-00171 | 2023-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.0/3.0/043 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ELEVATED IOP(32MMHG) WITH OUR PUPIL BLOCK AND ANGLE CLOSURE(ASSESSED ON 3RD WEEK POST-OP), PIGMENT DISPERSION WERE OBSERVED. LENS REMAINS IMPLANTED. PATIENT GIVEN STEROIDS AND IOP MEDS. THIS HAS NOT RESOLVED THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -14.0/3.0/043 | OD | Physician | 12/27/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16127557&pc=MTA | 2023826-2023-00115 | 2022-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient Device Interaction Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS THE PATIENT HAD EXCESSIVE VAULTING; ELEVATED INTRAOCULAR PRESSURE AND ANGLE CLOSER; IT WAS LATER ADDITIONALLY NOTED THE PATIENT PRESSURES WERE STABLE ON MEDICATION AND ANGLES NARROW. LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6: INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. B5: "THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.50/1.0/104 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT; ANGLE CLOSURE WITH ELEVATED IOP AND MEDICATION WAS PRESCRIBED. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR THE SAME MODEL/SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "INITIALLY SOME CORNEAL ODEMA AND ENDOTHELIAL CELL LOSS DUE TO TRICKY EXCHANGE (CONSTRICTION OF PUPIL DURING PROCEDURE) WILL COMPLETE ANOTHER COMPLAINT FORM BUT VISION EXCELLENT." THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR." CORRECTED DATA: G2: OTHER: "UNITED KINGDOM" SHOULD BE CORRECTED TO "IRELAND". CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -10.50/1.0/104 | OD | Physician | 12/13/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16128162&pc=MTA | 2023826-2023-00116 | 2022-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient Device Interaction Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS THE PATIENT HAD EXCESSIVE VAULTING; ELEVATED INTRAOCULAR PRESSURE AND ANGLE CLOSER; IT WAS LATER ADDITIONALLY NOTED THE PATIENT PRESSURES WERE STABLE ON MEDICATION AND ANGLES NARROW. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.50/3.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT; ANGLE CLOSURE WITH ELEVATED IOP AND MEDICATION WAS PRESCRIBED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL/SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "SITUATION RESOLVED AND PATIENT HAS EXCELLENT VISION." THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR. CORRECTED DATA: G2: OTHER: "UNITED KINGDOM" SHOULD BE CORRECTED TO "IRELAND". CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -10.50/3.0/092 | OS | Physician | 12/13/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6795682&pc=MTA | 2023826-2017-01320 | 2017-06-28 | Injury | STAAR SURGICAL COMPANY | 2017-08-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+2.0/159 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, REFRACTIVE SURPRISE, ELEVATED IOP. ON (B)(6) 2017, THE SURGEON COMMENTED THAT THE PATIENT IOP IS 20.6MMHG, AND HAS BEEN TREATED BY MEDICATIONS: TRUSOPT, EFFLUMIDEX, FENEFRIN, AND DIAMOX. MULTIPLE ATTEMPS TO GET MORE INFORMATION HAS BEEN UNSUCCESSFUL. Manufacturer Narrative: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES): ON (B)(6) 2017, THE SURGEON COMMENTED THAT THE PATIENT'S IOP IS 20.6MM HG AND HAS BEEN TREATED BY MEDICATIONS: TRUSOPT, EFFLUMIDEX, FENEFRIN AND DIAMOX. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.5/+2.0/159 | OD | Physician | 07/18/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7465307&pc=MTA | 2023826-2018-00607 | 2018-02-18 | Injury | STAAR SURGICAL COMPANY | 2018-04-26 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -8.5/+3.0/104 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS LOSS OF BCVA, ELEVATED IOP, PIGMENT DISPERSION, AND BLURRED VISION. THE LENS WAS EXPLANTED ON (B)(6) 2018. TO DATE, THE PATIENT HAS NOT DECIDED WHETHER TO REPLACE WITH ANOTHER LENS OR NOT. Manufacturer Narrative: ADDITIONAL INFORMATION: REPORTEDLY, THE PATIENT DEVELOPED RAISED IOP AFTER TICL IMPLANT, WHICH WAS "REFRACTORY TO TOPICAL ANTI-GLAUCOMA MEDICATIONS." THE TICL WAS EXPLANTED AND A TRABECULECTOMY PERFORMED. THE CORNEA IS CURRENTLY CLEAR. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND FIBERS AND BLACK SPOTS ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -8.5/+3.0/104 | OD | Physician | 04/05/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6876795&pc=MTA | 2023826-2017-01460 | 2016-05-26 | Injury | STAAR SURGICAL COMPANY | 2017-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/142 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND BLURRED VISION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(6). CLAIM# (B)(4). | -13.5/+1.5/142 | OS | Physician | 08/22/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6876796&pc=MTA | 2023826-2017-01459 | 2016-05-26 | Injury | STAAR SURGICAL COMPANY | 2017-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+2.0/091 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND BLURRED VISION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.0/+2.0/091 | OD | Physician | 08/22/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6641431&pc=MTA | 2023826-2017-00976 | 2017-05-12 | Injury | STAAR SURGICAL COMPANY | 2017-06-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/083 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCING EXCESSIVE VAULT AND ELEVATED IOP. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DIMENSIONAL INSPECTION FOUND LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). NOT RETURNED TO MANUFACTURER. | -9.0/+1.5/083 | OD | Physician | 05/16/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4725289&pc=MTA | 2023826-2015-00428 |  | Injury | STAAR SURGICAL COMPANY | 2015-04-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE REPORTER INDICATED THE LENS HAD A REFRACTIVE SURPRISE AND THE LENS ROTATED (NOT ASSOCIATED TO A LOW VAULT). THE INCISION WAS ENLARGED AND THE LENS WAS REPOSITIONED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED AT THE FIVE MONTH VISIT, THE SURGEON CONFIRMED THE OFF-AXIS ALIGNMENT AND DECIDED TO REPOSITION THE LENS. THE PATIENT EXPERIENCED BLURRED VISION. AFTER THIS INTERVENTION, THE PATIENT EXPERIENCED AN IOP RISE, WHICH WAS SOLVED BY ANTERIOR CHAMBER ASPIRATION AND IRRIGATION. FOR THIS SECONDARY INTERVENTION A NEW INCISION WAS CREATED, LEADING TO INCREASED CORNEAL ASTIGMATISM. THE LENS REMAINS IMPLANTED AND IS STABLE AFTER REPOSITIONING. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW RESULTS: (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON THE ABOVE INVESTIGATION, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - (OTHER): LENS IMPLANTED METHOD - (OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (OTHER): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS IMPLANTED. Manufacturer Narrative: (B)(4). METHOD: MEDICAL REVIEW. RESULTS: MEDICAL REVIEW - SEVERAL FACTORS MAY BE INVOLVED IN A POST-OPERATIVE REFRACTIVE SURPRISE, INCLUDING INAPPROPRIATE REFRACTIVE SURGICAL PLANNING, A WRONG LENS CALCULATION/SELECTION, PREOPERATIVE INACCURATE DETERMINATION OF THE EYE'S MEASUREMENTS TO DETERMINE THE POWER OF THE ICL, INDUCED REFRACTIVE CHANGE BY INCISIONS, UNSTABLE CORNEA AND/OR REFRACTION PRIOR TO IMPLANTATION, CL-INDUCED WARPAGE, UPSIDE-DOWN LENS, WRONG LENS, MISALIGNED LENS, ROTATION, TILTING, ETC. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, THE MOST LIKELY CAUSE OF THE RESIDUAL REFRACTION IS A COMBINATION OF SELECTING A LENS POWER THAT WAS UNDER-CORRECTING THE CYLINDER AND SURGICAL INDUCED ASTIGMATISM SECONDARY TO A NEW INCISION. THE LENS WAS OFF-AXIS AT ONE DAY POSTOPERATIVELY THEREFORE THIS COULD BE DUE TO SURGICAL MISALIGNMENT OR EARLY MOVEMENTS OF THE LENS. (B)(4). | -6.0/+1.0/112 | OS | Other | 03/24/2015 | 2015 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6589772&pc=MTA | 2023826-2017-00878 | 2015-12-01 | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+1.5/080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, PIGMENT DISPERSION AND GLARE/HALOS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE AND CLEAR RESIDUE ON THE LENS. (B)(6). (B)(4). Manufacturer Narrative: WORK ORDER SEARCH PERFORMED: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.0/+1.5/080 | OD | Physician | 04/28/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6412058&pc=MTA | 2023826-2017-00259 | 2017-01-24 | Injury | STAAR SURGICAL COMPANY | 2017-03-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON PARTIALLY INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/102 DIOPTER, BUT DUE TO INTRAOCULAR POSITIVE PRESSURE, THE SURGEON COULD NOT IMPLANT THE LENS COMPLETELY. THE REPORTER INDICATED LOOSE ZONULES WERE SUSPECTED AND THE ANTERIOR CHAMBER COLLAPSED. THE LENS WAS REMOVED WITHOUT ANY PATIENT INJURY OR COMPLICATIONS AND NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS RETURNED IN THREE PIECES. THERE WAS RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.0/+2.0/102 | NA | Physician | 02/16/2017 | 2017 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8077834&pc=MTA | 2023826-2018-01630 | 2018-08-16 | Injury | STAAR SURGICAL COMPANY | 2018-11-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.00/+1.0/X084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP) (13MMHG), WHICH CAUSED THE IRIS TO PROLAPSE (POSTERIOR SYNECHIAE) THROUGH THE WOUND. POST-OP, THE PATIENT STILL HAS ELEVATED IOP AND POSTERIOR SYNECHIAE. IT IS UNKNOWN IF ANOTHER ICL LENS WILL BE IMPLANTED. THE REPORTER INDICATED THE SURGEON FELT THE MEASUREMENTS MAY HAVE BEEN INCORRECT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -8.00/+1.0/084 | OD | Physician | 10/19/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7976037&pc=MTA | 2023826-2018-01513 | 2018-08-20 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM013.2 IMPLANTABLE COLLAMER LENS, -16.50/+1.0/025 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. ON (B)(6) 2018, THE PATIENTS INTRAOCULAR PRESSURE (IOP) WAS ELEVATED AT 21 - 28 MMHG. A YAG WAS PERFORMED FOR ADDITION OF NEW PERIPHERAL IRIDECTOMIES AND IOP WAS 28 MMHG. THE PATIENT WAS PRESCRIBED MEDICATION (SIMBRINZA) AND IOP WAS 23MMHG. THE SURGEON HAS DECIDED TO REDUCE THE MEDICATIONS AND WAIT AND SEE HOW THE IOP DEVELOPS. IF THE IOP INCREASES, HE WILL REPLACE THE LENS. OTHERWISE, THE LENS WILL REMAIN IN THE EYE. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.50/+1.0/025 | OS | Physician | 09/17/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7976051&pc=MTA | 2023826-2018-01512 | 2018-08-20 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM013.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.0/125 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. ON (B)(6) 2018, THE PATIENTS INTRAOCULAR PRESSURE (IOP) WAS ELEVATED AT 21 - 38 MMHG. A YAG WAS PERFORMED FOR ADDITION OF NEW PERIPHERAL IRIDECTOMIES AND IOP WAS 28 MMHG. THE PATIENT WAS PRESCRIBED MEDICATION (SIMBRINZA) AND IOP WAS 23MMHG. THE SURGEON HAS DECIDED TO REDUCE THE MEDICATIONS AND WAIT AND SEE HOW THE IOP DEVELOPS. IF THE IOP INCREASES, HE WILL REPLACE THE LENS. OTHERWISE, THE LENS WILL REMAIN IN THE EYE. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -18.0/+1.0/125 | OD | Physician | 09/17/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8215221&pc=MTA | 2023826-2019-00021 | 2018-11-21 | Injury | STAAR SURGICAL COMPANY | 2019-01-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, - 15.50/+3.5/093 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD). THE REPORTER INDICATED AFTER 1 MONTH POST-OP, THE SURGEON NOTICED RAISED INTRAOCULAR PRESSURE (IOP), WITH OPEN ANGLE BUT SHALLOW CHAMBER. THERE WAS SUSPECTED STEROID RESPONSE, AND THE SURGEON STARTED ANTIGLAUCOMA. AFTER 2 WEEKS, THERE WAS NO IMPROVEMENT TO IOP, SO THE SURGEON PERFORMED LASER PI. THE IOP SHOT UP AFTER 1 WEEK TO 32MMHG AFTER LASER PI. STARTED ANTIGLAUCOMA, BOTH TOPICAL AND SYSTEMIC. AT PRESENT ONLY TOPICAL ANTIGLAUCOMA AND IOP IS 22MMHG. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS IMPLANTED ON (B)(6) 2018 AND EXPLANTED ON (B)(6) 2018 DUE TO ELEVATED INTRAOCULAR PRESSURE AND ANGLE CLOSURE. A YAG WAS PERFORMED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENTS CURRENT CONDITION WAS ANGLE OPENED AND ANTERIOR CHAMBER WELL FORMED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL OBSERVATION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | - 15.50/+3.5/093 | OD | Physician | 12/09/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8359916&pc=MTA | 2023826-2019-00291 | 2019-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICM5_13.2, -9.00/1.5/087 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO ELEVATED IOP (INTRAOCULAR PRESSURE) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. IT WAS REPORTED THAT AFTER EXPLANTING THE LENS THE PROBLEM RESOLVED, IOP AND IRIDO-CORNEAL ANGLES RETURNED TO NORMAL. THE DOCTOR HAS NO PLANS TO IMPLANT ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THAT A HAPTIC WAS TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY:UNK. RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.00/1.5/087 | OD |  | 01/23/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5223685&pc=MTA | 2023826-2015-01409 | 2015-07-01 | Injury | STAAR SURGICAL COMPANY | 2015-11-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -4.50/+0.5/X094 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. UNREACTIVE PUPIL AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION WAS ALSO PERFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -4.50/0.50/X094. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -4.50/+0.5/094 | OS | Physician | 10/13/2015 | 2015 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5756735&pc=MTA | 2023826-2016-00792 | 2016-01-21 | Injury | STAAR SURGICAL COMPANY | 2016-06-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -9/+4/167 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED IOP. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID, IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE BUT PRESENCE OF FOREIGN MATERIAL (FIBERS) ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -9/+4/167 | OD | Physician | 05/25/2016 | 2016
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5891300&pc=MTA | 2023826-2016-00455 | 2016-02-24 | Injury | STAAR SURGICAL COMPANY | 2016-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+0.5/094 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). TWENTY-FOUR REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. PATIENT AGE BELOW 21 AND ANTERIOR ENDOTHELIUM CHAMBER DEPTH IS BELOW 3.0MM FOR VICL/VTICL. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+0.5/094 | OD | Physician | 03/09/2016 | 2016 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5891288&pc=MTA | 2023826-2016-00453 | 2016-03-03 | Injury | STAAR SURGICAL COMPANY | 2016-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/094 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A REFRACTIVE SURPRISE AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+1.0/094 | OD | Physician | 03/09/2016 | 2016 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6077547&pc=MTA | 2023826-2016-01562 | 2016-03-10 | Injury | STAAR SURGICAL COMPANY | 2016-11-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -18.0/./+20 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND PUPIL BLOCK (WITH ELEVATED IOP). Manufacturer Narrative: ADDITIONAL INFORMATION: PREEXISTING MEDICAL CONDITION: KERATOCONUS. CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH KERATOCONUS IS CONTRAINDICATED. CORRECTED DATA: CORRECT LENS DIOPTER IS -18.0/+2.0/X074 (SPHERE/CYLINDER/AXIS). (B)(4). Manufacturer Narrative: PMA/510(K): N/A THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+2.0/074 | OS | Physician | 10/07/2016 | 2016 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6222127&pc=MTA | 2023826-2017-00012 | 2016-10-15 | Injury | STAAR SURGICAL COMPANY | 2017-01-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/096 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS UNDER 21 YEARS OF AGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -11.00/+1.5/096 | OS | Physician | 12/06/2016 | 2016 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6420639&pc=MTA | 2023826-2017-00446 | 2016-12-13 | Injury | STAAR SURGICAL COMPANY | 2017-03-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.5/+2.0/105 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND GONIOSYNECHIA. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -5.5/+2.0/105 | OD | Physician | 02/23/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6482395&pc=MTA | 2023826-2017-00601 | 2016-12-14 | Injury | STAAR SURGICAL COMPANY | 2017-04-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/+1.0/082 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND GONIOSYNECHIAE. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR THE SHORTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DRY, CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. LENS IS TWISTED DUE TO POSITION IN VIAL. (B)(4). | -7.0/+1.0/082 | OS | Physician | 03/13/2017 | 2017 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6260641&pc=MTA | 2023826-2017-00101 | 2016-12-21 | Injury | STAAR SURGICAL COMPANY | 2017-01-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/95 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED TO THE MANUFACTURER BUT NOT YET BEEN EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4). | -12.0/+2.0/95 | OS | Physician | 12/22/2016 | 2016 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6653556&pc=MTA | 2023826-2017-00998 | 2017-03-09 | Injury | STAAR SURGICAL COMPANY | 2017-06-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/072 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, ELEVATED IOP. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A VICMO LENS WITH THE SAME SIZE AND DIFFERENT DIOPTER. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND CLEAR RESIDUE AND DEBRIS ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.0/+1.0/072 | OD | Physician | 05/26/2017 | 2017 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6711929&pc=MTA | 2023826-2017-01185 | 2017-03-11 | Injury | STAAR SURGICAL COMPANY | 2017-07-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/098 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6)2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. UPDATED INFORMATION: THE LENS HAS BEEN RETURNED AND EVALUATED. LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: PREVIOUS ADVERSE EVENTS WERE NOT INCLUDED IN INITIAL MDR: PATIENT HAD ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(6). (B)(4). | -12.0/+1.5/098 | OS | Physician | 06/19/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6554053&pc=MTA | 2023826-2017-00781 | 2017-03-15 | Injury | STAAR SURGICAL COMPANY | 2017-05-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/066 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED IOP. ON (B)(6) 2017 THE LENS WAS EXPLANTED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: PREVIOUS: STATED LENS IS NOT RETURNED. UPDATED: LENS HAS BEEN RETURNED AND EVALUATED. DEVICE EVALUATION: LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION REVEALED LENS WAS WITHIN SPECIFICATIONS. METHOD, RESULT CODES: UPDATED. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.0/+1.5/066 | OS | Physician | 04/13/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6587670&pc=MTA | 2023826-2017-00872 | 2017-03-21 | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+5/X126 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, THE PUPIL CAN NOT CONTRACT, AND ELEVATED IOP. IN THE SAME DAY, THE LENS WAS EXCHANGED FOR SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR AND BROWN RESIDUE ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR AND BROWN RESIDUE ON THE LENS. CLAIM # (B)(4). | -9.0/+5/126 | OD | Physician | 04/24/2017 | 2017 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6637203&pc=MTA | 2023826-2017-00964 | 2017-04-20 | Injury | STAAR SURGICAL COMPANY | 2017-06-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/88 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION. THE PATIENT ALSO UNDERWENT ENLARGEMENT PI OR ADDITION OF NEW PI, AND SECONDARY YAG SURGERY. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS, AND THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RED AND CLEAR RESIDUE ON THE LENS SURFACE. CONCLUSION: PREVIOUS CODE NOT APPLICABLE. AS PER DFU, THIS LENS MODEL IS CONTRAINDICATED FOR IMPLANTATIONS IN PATIENTS WITH ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.5/+1.5/88 | OD | Physician | 05/17/2017 | 2017 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6637164&pc=MTA | 2023826-2017-00961 | 2017-04-25 | Injury | STAAR SURGICAL COMPANY | 2017-06-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0/+1.5/085 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED IOP. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMAGE AND CLEAR RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.0/+1.5/085 | OD | Physician | 05/17/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6637205&pc=MTA | 2023826-2017-00965 | 2017-04-27 | Injury | STAAR SURGICAL COMPANY | 2017-06-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+1.5/88 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION. THE PATIENT ALSO UNDERWENT ENLARGEMENT PI OR ADDITION OF NEW PI, AND SECONDARY YAG SURGERY. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS, AND THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND A MISSING PIECE OF HAPTIC AND RED AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.5/+1.5/88 | OS | Physician | 05/17/2017 | 2017 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6740871&pc=MTA | 2023826-2017-01240 | 2017-06-21 | Injury | STAAR SURGICAL COMPANY | 2017-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -2.5/+3.0/162 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED AS AN EMERGENCY IN THE HOSPITAL. THE LENS WAS EXPLANTED BECAUSE OF EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF PATIENT'S IRIDO-CORNEAL ANGLES (ICA), AND PUPIL BLOCK WITH ELEVATED IOP. THE PROBLEM WAS RESOLVED WITH THE EXPLANT. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -2.5/+3.0/162 | OD | Physician | 07/06/2017 | 2017 | 50
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6850323&pc=MTA | 2023826-2017-01394 | 2017-08-02 | Injury | STAAR SURGICAL COMPANY | 2017-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+1.5/088 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: EVENT DATE: (B)(6) 2017. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CONCLUSION CODE: OFF LABEL USE (PATIENT BELOW 21 YEARS OF AGE AT TIME OF IMPLANT). (B)(4). | -14.5/+1.5/088 | OS | Physician | 08/09/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7114677&pc=MTA | 2023826-2017-01841 | 2017-08-10 | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -15.0/+3.0/079 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT WAS UNDER REVIEW FOR MORE THAN 3 MONTHS DUE TO THE FLUCTUATION OF IOP. THE DOCTOR DECIDED TO DISCHARGE THE PATIENT ON (B)(6) 2017 AS THE IOP WAS 16 MMHG. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS CURVED AND RED AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.0/+3.0/079 | OD | Physician | 11/14/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7114680&pc=MTA | 2023826-2017-01859 | 2017-08-10 | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -14.0/+1.5/090 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT WAS UNDER REVIEW FOR MORE THAN 3 MONTHS DUE TO THE FLUCTUATION OF IOP. THE DOCTOR DECIDED TO DISCHARGE THE PATIENT ON (B)(6) 2017 AS THE IOP WAS 16 MMHG. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF THE HAPTIC AND THE PRESENCE OF CLEAR AND RED RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.0/+1.5/090 | OS | Physician | 11/14/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6955546&pc=MTA | 2023826-2017-01576 | 2017-09-11 | Injury | STAAR SURGICAL COMPANY | 2017-10-17 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. REPORTEDLY, THE LENS WAS EXPLANTED DUE TO IOP SPIKE, EXCESSIVE VAULT, AND NARROW ANGLES. REPORTER STATES LENS IS PLANNED TO BE EXCHANGED WITH A SMALLER LENS; ORIGINAL LENS HAS BEEN LOST BY PHYSICIAN. ADDITIONAL INFORMATION IS NOT KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: REPORTEDLY, ON (B)(6) 2017, A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, DIOPTER -14.5/+2.5/108 WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE. THE LENS WAS EXPLANTED ON (B)(6) 2017. THE LENS WAS DAMAGED DURING EXPLANT. THE DOCTOR IS "WAITING FOR PUPIL TO CONSTRICT" BEFORE DECIDING ON FURTHER ACTION. THE DOCTOR SUSPECTS URRETS-ZAVALIA. PREVIOUS REPORT INCLUDED ELEVATED IOP; CURRENT REPORT FROM COMPLAINANT DOES NOT. CURRENT SYMPTOMS REPORTED: EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), UNREACTIVE (FIXED) PUPIL, IRIS ATROPHY, SUSPECTED URRETS-ZAVALIA. ADDITIONAL PATIENT CODE 3191: NO CODE AVAILABLE (FIXED PUPIL, IRIS ATROPHY, SUSPECTED URRETS-ZAVALIA). (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -14.5/+2.5/108 | OD | Physician | 09/20/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7871042&pc=MTA | 2023826-2018-01367 | 2017-09-30 | Injury | STAAR SURGICAL COMPANY | 2018-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/110 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE PATIENT DID NOT WANT ANOTHER LENS IMPLANTED. Manufacturer Narrative: DEVICE WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION TOLERANCES. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/+1.0/110 | OD | Physician | 08/13/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7871107&pc=MTA | 2023826-2018-01368 | 2017-09-30 | Injury | STAAR SURGICAL COMPANY | 2018-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE PATIENT DID NOT WANT ANOTHER LENS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND DEFORMED. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO BROKEN HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.0/+1.5/092 | OS | Physician | 08/13/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978046&pc=MTA | 2023826-2017-01624 | 2017-10-27 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, -6.0/+1.5/090 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND FIBERS ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.0/+1.5/090 | OD | Physician | 09/29/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7356130&pc=MTA | 2023826-2018-00437 | 2018-01-17 | Injury | STAAR SURGICAL COMPANY | 2018-03-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/071 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND A "PROTUBERANT" IRIS. ON (B)(6) 2018 A SHORTER LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A HAPTIC TEAR AND RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+1.5/071 | OS | Physician | 02/23/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7356135&pc=MTA | 2023826-2018-00435 | 2018-01-17 | Injury | STAAR SURGICAL COMPANY | 2018-03-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND A "PROTUBERANT" IRIS. ON (B)(6) 2018 A SHORTER LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED INA MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.5/095 | OD | Physician | 02/23/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7884850&pc=MTA | 2023826-2018-01266 | 2018-02-12 | Injury | STAAR SURGICAL COMPANY | 2018-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.00/+1.5/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ASSESMENT ON (B)(6) 2018, ANGLE CLOSURE WITH ELEVATED IOP (26MMHG) AND EXCESSIVE VAULT. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL (RESIDUE) ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/+1.5/077 | OS | Physician | 08/21/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7621051&pc=MTA | 2023826-2018-00923 | 2018-03-01 | Injury | STAAR SURGICAL COMPANY | 2018-06-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -12.0/+2.0/089 SPHERE/CYLINDER/AXIS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -12.0/+2.0/089 SPHERE/CYLINDER/AXIS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. SECONDARY SURGERY, LENS EXCHANGE. (B)(4). | -12.0/+2.0/089 | OD | Physician | 05/23/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7621071&pc=MTA | 2023826-2018-00921 | 2018-03-01 | Injury | STAAR SURGICAL COMPANY | 2018-06-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA; DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+2.5/086 | OS | Physician | 05/23/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8275983&pc=MTA | 2023826-2019-00170 | 2018-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -10.00/2.0/85 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT AND ELEVATED IOP (INTRAOCULAR PRESSURE). IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO "FALSE W-TO-W, MEASUREMENT PROVIDED." THE REPORTER STATED THAT THE PATIENTS EYE IS STABLE AND IN GOOD CONDITION. IT WAS REPORTED THAT THE DOCTOR PLANS TO IMPLANT ANOTHER LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/2.0/85 | OS |  | 12/27/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7589676&pc=MTA | 2023826-2018-00879 | 2018-05-05 | Injury | STAAR SURGICAL COMPANY | 2018-06-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL AND THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND SHORTER LENS IMPLANTED PATIENT CODE: (B)(4) SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM AND WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4). | -10.0/+2.0/093 | OS | Physician | 05/14/2018 | 2018 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7772418&pc=MTA | 2023826-2018-01180 | 2018-06-04 | Injury | STAAR SURGICAL COMPANY | 2018-08-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM, VTICM5_13.2, -12.0/1.00/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PREVIOUSLY STATED, "VISUAL INSPECTION FOUND THE HAPTIC TORN." CORRECTED STATEMENT: "VISUAL INSPECTION FOUND THE HAPTICS BENT." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/1.00/93 | OD | Physician | 07/11/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7903849&pc=MTA | 2023826-2018-01373 | 2018-06-12 | Injury | STAAR SURGICAL COMPANY | 2018-09-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.00/6.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 EXCESSIVE VAULT WAS OBSERVED, PATIENT WAS ASSESSED WITH ANGLE CLOSURE WITH ELEVATED IOP (50 MMHG) ON (B)(6) 2018 THE LENS WAS EXPLANTED. SURGEON PLANS TO IMPLANT ANOTHER ICL LENS IN (B)(6) 2018. EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: CORRECTED TO: "ADVERSE EVENT" IN THE INITIAL MDR. CORRECTED TO: "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES)" IN THE INITIAL MDR. ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. DIMENSIONAL INSPECTION FOUND LENS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.00/6.0/089 | OS | Physician | 08/28/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7765108&pc=MTA | 2023826-2018-01169 | 2018-06-19 | Injury | STAAR SURGICAL COMPANY | 2018-08-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+2.0/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK, AND ELEVATED IOP.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+2.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED. THE SURGEON REPORTS VAULT VALUES OF 800UM/1.2CT, ELEVATED IOP, AND REFRACTIVE TREATMENT. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT, PUPIL BLOCK). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). | -14.5/+2.0/086 | OS | Physician | 07/13/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7884841&pc=MTA | 2023826-2018-01392 | 2018-08-21 | Injury | STAAR SURGICAL COMPANY | 2018-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -09.5/6/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: "LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT, SHALLOWING ACD, AND HIGHER IOP" NOT INCLUDED IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -09.5/6/102 | OD | Physician | 08/21/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8090043&pc=MTA | 2023826-2018-01696 | 2018-08-22 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+2.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ANGLE CLOSURE, ELEVATED IOP, AND SHALLOW AC. THE PROBLEM WAS RESOLVED. "SURGEON AND PATIENT ARE BOTH HAPPY." THE CAUSE OF THE EVENT IS REPORTED AS "OVERSIZE LENS." Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.5/+2.5/093 | OD | Physician | 10/25/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8090042&pc=MTA | 2023826-2018-01699 | 2018-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.0/+3.0/072 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ANGLE CLOSURE, ELEVATED IOP, AND SHALLOW AC. THE PROBLEM WAS RESOLVED. "SURGEON AND PATIENT ARE BOTH HAPPY." THE CAUSE OF THE EVENT IS REPORTED AS "OVERSIZE LENS." Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATED BY MFR -DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN, BENT AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+3.0/072 | OS |  | 10/25/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8277800&pc=MTA | 2023826-2019-00177 | 2018-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-25 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -2.50/2.5/178 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT AND ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). ADDITIONAL INFORMATION REPORTED: PATIENT CANCELLED THE SURGERY FOR BOTH EYES AFTER THE EXPLANTED OD LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: IMPLANTED DATE: CORRECTED FROM "(B)(6) 2018" TO (B)(6) 2018" FOR INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -2.50/2.5/178 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT AND ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). PATIENT CANCELLED THE SURGERY AFTER THE EXPLANTED OD LENS AND HAS NOT YET SCHEDULED FOR ANY SURGERIES. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CAS/VIAL. VISUAL INSPECTION FOUND CLEAR SURGICAL RESIDUE ON PRODUCT AND A BENT HAPTIC. DIMENSIONAL INSPECTION WAS WITHIN LENS SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -2.50/2.5/178 | OS |  | 12/26/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7183420&pc=MTA | 2023826-2018-00045 |  | Injury | STAAR SURGICAL COMPANY | 2018-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -15.0/2/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. EXCESSIVE VAULT AND ELEVATED IOP WAS NOTICED. Manufacturer Narrative: DATE OF EVENT: (B)(6) 2017. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2 IMPLANTABLE COLLAMER LENS, OF -15.0/2/110 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXCHANGED ON (B)(6) 2018, DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED AND THE SURGEON STATED THAT THE "VAULT WAS NORMAL". (B)(4) Manufacturer Narrative: IT IS CORRECTED FROM "OTHER SERIOUS" TO "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT / DAMAGE" (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID INSIDE A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF ITS HAPTIC. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.0/2/110 | OD | Physician | 12/15/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7183423&pc=MTA | 2023826-2018-00046 |  | Injury | STAAR SURGICAL COMPANY | 2018-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -14/1/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. EXCESSIVE VAULT AND ELEVATED IOP WAS NOTICED. Manufacturer Narrative: DATE OF EVENT: (B)(6) 2017. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, OF -14/1/101 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXCHANGED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND ELEVATED IOP. THE PROBLEM WAS RESOLVED AND THE SURGEON STATED THAT THE "VAULT WAS NORMAL". (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID INSIDE A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: CORRECTED DATA: CLAIM NUMBER OF THE FILE IN THE PREVIOUS SUPPLEMENTAL IS CORRECTED FROM "CLAIM# (B)(4)" TO " CLAIM# (B)(4)". Manufacturer Narrative: IT IS CORRECTED FROM "OTHER SERIOUS" TO "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT / DAMAGE" (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14/1/101 | OS | Physician | 12/15/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7627647&pc=MTA | 2023826-2018-00932 |  | Injury | STAAR SURGICAL COMPANY | 2018-06-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.0/+3/174 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXCHANGED WITH A SAME MODEL BUT DIFFERENT SIZE LENS DUE TO OBSERVATION OF EXCESSIVE VAULT AND ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). THE REPORTED STATED THAT THE PATIENT'S CONDITION IS VERY GOOD NOW. Manufacturer Narrative: CORRECTED DATA: COMPLAINT "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES)". OUTCOME SHOULD HAVE BEEN MENTIONED IN INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: DATE OF EXCHANGE TOOK PLACE ON (B)(6) 2018. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.0/+3/174 | OS | Physician | 05/23/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5210642&pc=MTA | 2023826-2015-01434 | 2014-12-15 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -12.00/+2.00/X107 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. EXCESSIVE VAULT WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED AND THE ICL WAS EXPLANTED ON (B)(6) 2014. NO OTHER LENS WAS IMPLANTED. SEE MFR. #2023826-2015-01433 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(6) BUT IS NOT MARKETED IN THE U.S. DIOPTER: -12.00/+2.00/X107 (B)(4). | -12.00/+2.00/107 | OD | Physician | 10/13/2015 | 2015 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5210651&pc=MTA | 2023826-2015-01433 | 2014-12-15 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/X064 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. EXCESSIVE VAULT WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED AND THE ICL WAS EXPLANTED ON (B)(6) 2014. NO OTHER LENS WAS IMPLANTED. SEE MFR. #2023826-2015-01434 FOR RIGHT EYE. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT IS NOT MARKETED IN THE U.S. DIOPTER: --11.50/+2.00/X064. (B)(4). CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -11.50/+2.0/064 | OS | Physician | 10/13/2015 | 2015 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5086238&pc=MTA | 2023826-2015-01237 | 2015-01-20 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -23.0/+1.0/075 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | -23.0/+1.0/075 | OS | Physician | 08/20/2015 | 2015 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5086239&pc=MTA | 2023826-2015-01238 | 2015-01-20 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -22.5/+1.5/070 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. PATIENT CODE: (B)(4) ANGLE CLOSURE, LENS EXCHANGE. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). | -22.5/+1.5/070 | OD | Physician | 08/20/2015 | 2015 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4736049&pc=MTA | 2023826-2015-00478 | 2015-01-23 | Injury | STAAR SURGICAL COMPANY | 2015-04-29 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED AS WELL AS ANGLE CLOSURE ASSOCIATED WITH ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST KNOW BCVA WAS 20/20 AS OF (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGERY. NEW INCISION. ANGLE CLOSURE. EXPLANTED. VAULTING. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS TORN INTO TWO PIECES. THE LENS COULD NOT BE RE-MEASURED. THE LENS WAS RETURNED DRY. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, EVALUATION OF THE RETURNED PRODUCT AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -17.5/1.5/109 | OD | Other | 04/02/2015 | 2015 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4759443&pc=MTA | 2023826-2015-00522 | 2015-03-25 | Injury | STAAR SURGICAL COMPANY | 2015-05-08 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED, AS WELL AS SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ANGLE CLOSURE ASSOCIATED WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST KNOWN BCVA WAS 20/100 AS OF (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -4.0/+3//004 | OS | Physician | 04/09/2015 | 2015 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4743860&pc=MTA | 2023826-2015-00499 | 2015-03-27 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-01 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO ELEVATED IOP ASSOCIATED WITH EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. SEE MFR REPORT# 2023826-2015-00498 FOR THE OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4) NEW INCISION. LENS EXPLANTED. VAULTING. METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. (B)(4). Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH & MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A HAPTIC WAS TORN AND THE LENS WAS RETURNED DRY. THE WORK ORDER SEARCH SHOWED NO OTHER COMPLAINT RECORDED RELATED TO THE SAME WORK ORDER. MEDICAL REVIEW - REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -6.0/3.0/095 | OS | Physician | 04/02/2015 | 2015 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5106976&pc=MTA | 2023826-2015-01186 | 2015-04-18 | Injury | STAAR SURGICAL COMPANY | 2015-09-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/X101 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ON(B)(6) 2015 EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE WAS NOTED. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. UNCORRECTED VISUAL ACUITY WAS 20/20 ON (B)(6) 2015. SEE MFR #2023826-2015-01185 FOR RIGHT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -9.5/+2.0/X101. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | -9.5/+2.0/101 | OS | Physician | 09/01/2015 | 2015 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5106991&pc=MTA | 2023826-2015-01185 | 2015-04-18 | Injury | STAAR SURGICAL COMPANY | 2015-09-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+4.0/X091 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ON (B)(6) 2015 EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE WAS NOTED. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. UNCORRECTED VISUAL ACUITY WAS 20/20 ON (B)(6) 2015. SEE MFR #2023826-2015-01186 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -8.0/+4.0/X091. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | -8.0/+4.0/091 | OD | Physician | 09/01/2015 | 2015 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5405797&pc=MTA | 2023826-2016-00045 | 2015-04-25 | Injury | STAAR SURGICAL COMPANY | 2016-02-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -11.5/-4.50/X090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE WAS NOTED AND THE LENS WAS EXPLANTED ON (B)(6) 2015 THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. SEE MFR. #2023826-2016-00046 FOR RIGHT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER:-11.50/+4.50/X090. EVALUATION METHOD: WORK ORDER SEARCH. EVALUATION RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -11.5/-4.50/090 | OS | Physician | 12/11/2015 | 2015 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5178915&pc=MTA | 2023826-2015-01356 | 2015-05-06 | Injury | STAAR SURGICAL COMPANY | 2015-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER -13.00/+1.0/X133 DIOPTER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -13.00/+1.0/X133 (B)(4). METHOD CODES: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -13.00/+1.0/133 | OD | Physician | 09/29/2015 | 2015 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4919323&pc=MTA | 2023826-2015-00835 | 2015-05-26 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-16 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015 AND ELEVATED IOP ASSOCIATED WITH EXCESSIVE VAULT WAS OBSERVED. THE PATIENT'S LEFT PUPIL IS LARGER THAN THE RIGHT. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS OBSERVED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION CODES: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). | -17.0/+0.5/113 | OS | Physician | 06/16/2015 | 2015 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5219006&pc=MTA | 2023826-2015-01410 | 2015-06-25 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+0.5/X094 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED. THE ICL WAS EXPLANTED ON (B)(6) 2015. NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONALS INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -11.50/+0.50/X094. DEVICE EVALUATED BY MANUFACTURER? NO- LENS NOT RETURNED. (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -11.5/+0.5/094 | OD | Physician | 10/13/2015 | 2015 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5208843&pc=MTA | 2023826-2015-01429 | 2015-07-15 | Injury | STAAR SURGICAL COMPANY | 2015-11-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS -16.00/+1.50/X089 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED. THE ICL EXPLANTED AND EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. SEE MFR. #2023826-2015-01428 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S.,BUT IS NOT MARKETED IN THE U.S. DIOPTER: -16.00/1.50/X089. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -16.00/+1.50/089 | OD | Physician | 10/13/2015 | 2015 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5206816&pc=MTA | 2023826-2015-01475 | 2015-07-16 | Injury | STAAR SURGICAL COMPANY | 2015-11-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM13.7 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/X082 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO -LENS NOT RETURNED. (B)(4). | -12.5/+1.0/082 | OS | Physician | 08/28/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5206819&pc=MTA | 2023826-2015-01476 | 2015-07-16 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM13.7 IMPLANTABLE COLLAMER LENS, -13.50/+1.0/X075 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC BROKEN AND THERE WAS FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS FIBERS. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -13.50/+1.0/075 | OD | Physician | 08/28/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5208884&pc=MTA | 2023826-2015-01428 | 2015-07-20 | Injury | STAAR SURGICAL COMPANY | 2015-11-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS -16.50/+2.5/X084 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED. THE ICL EXPLANTED AND EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. SEE MFR. #2023826-2015-01429 FOR RIGHT EYE. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -16.5/+2.5/X084. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -16.50/+2.5/084 | OS | Physician | 10/13/2015 | 2015 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5154950&pc=MTA | 2023826-2015-01354 | 2015-08-08 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+1.5/090 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | -14.0/+1.5/090 | OD | Physician | 09/15/2015 | 2015 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536900&pc=MTA | 2023826-2016-00314 | 2015-09-04 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -12.0/+3.5/91 DIOPTER IMPLANTABLE COLLAMER LENS ON (B)(6) 2015 IN TO THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE AND IOP WAS ELEVATED. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015 AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -12.0/+3.5/91. (B)(4). METHOD CODE (EVALUATION METHOD): WORK ORDER SEARCH RESULTS CODE (EVALUATION RESULT): A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CONCLUSION CODE (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -12.0/+3.5/91 | OD | Physician | 02/29/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536860&pc=MTA | 2023826-2016-00321 | 2015-10-23 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.50/+3.00/X80 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2015. EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT, UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -09.5/+3.00/X80. (B)(4). METHOD CODE(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODE(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -09.50/+3.00/80 | OS | Physician | 02/29/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5308661&pc=MTA | 2023826-2015-01659 | 2015-11-09 | Injury | STAAR SURGICAL COMPANY | 2015-12-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+1.0/X100 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND REFRACTIVE SURPRISE. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER ICL LENS. Manufacturer Narrative: REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, PATIENT AGE BELOW 21 OR OVER 45 YEARS AND ANTERIOR ENDOTHELIUM CHAMBER DEPTH BELOW 3.0MM FOR VICL/VTICL. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO -LENS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -9.50/+1.0/100 | OD | Physician | 11/18/2015 | 2015 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5535797&pc=MTA | 2023826-2016-00320 | 2015-12-31 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.50/+2.50/X091 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2015. EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT, UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. (B)(4). | -09.50/+2.50/091 | OD | Physician | 02/29/2016 | 2016 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5525093&pc=MTA | 2023826-2016-00329 | 2016-01-19 | Injury | STAAR SURGICAL COMPANY | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/+1.5/095 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). CONCLUSION: PATIENT ACD BELOW 3.0 MM. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | -17.50/+1.5/095 | OS | Physician | 02/25/2016 | 2016 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5891295&pc=MTA | 2023826-2016-01131 | 2016-02-25 | Injury | STAAR SURGICAL COMPANY | 2016-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+0.5/083 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). CONCLUSION: REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. PATIENT AGE BELOW 21. ANTERIOR ENDOTHELIUM CHAMBER DEPTH IS BELOW 3.0MM FOR VICL/VTICL. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN LIQUID IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+0.5/083 | OS | Physician | 03/09/2016 | 2016 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5632888&pc=MTA | 2023826-2016-00535 | 2016-03-06 | Injury | STAAR SURGICAL COMPANY | 2016-05-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/097 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SHALLOW ANTERIOR CHAMBER AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. AT ONE WEEK POST-OP THE LENS HAD A HIGH VAULT WITH CONTROLLED INTRAOCULAR PRESSURE ELEVATION. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+2.5/097 | OD | Physician | 03/22/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5670128&pc=MTA | 2023826-2016-00643 | 2016-03-09 | Injury | STAAR SURGICAL COMPANY | 2016-05-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/081 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE (WITH ELEVATED IOP). Manufacturer Narrative: (B)(4). DEVICE EVALUATIONS LENS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (FIBERS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -10.5/+1.5/081 | OS | Physician | 03/17/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5670132&pc=MTA | 2023826-2016-00642 | 2016-03-09 | Injury | STAAR SURGICAL COMPANY | 2016-05-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+3.5/091 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE (WITH ELEVATED IOP). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL (FIBERS) ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). THIS PATIENT WAS (B)(6) AT THE TIME OF IMPLANTATION. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -14.5/+3.5/091 | OD | Physician | 03/17/2016 | 2016 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5955105&pc=MTA | 2023826-2016-01258 | 2016-05-30 | Injury | STAAR SURGICAL COMPANY | 2016-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+3.0/113 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT EVENT WITHIN ASSOCIATED LOTS WAS FOUND. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -18.0/+3.0/113 | OD | Physician | 08/17/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5810832&pc=MTA | 2023826-2016-00942 | 2016-06-20 | Injury | STAAR SURGICAL COMPANY | 2016-07-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.0/+1.0./X180 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. Manufacturer Narrative: CORRECTED INFORMATION: REPORT SOURCE: "USER FACILITY" TO " DISTRIBUTOR". (B)(4) . ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED ON (B)(6) 2016 WITH A SHORTER LENS AND THIS RESOLVED THE PROBLEM. ON (B)(6) 2016, THE PATIENT'S POST-OPERATION BEST-CORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE PRESENCE OF FOREIGN MATERIAL ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -17.0/+1.0./180 | OD | Physician | 06/20/2016 | 2016 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5047723&pc=MTA | 2023826-2015-01113 |  | Injury | STAAR SURGICAL COMPANY | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+1.5/106 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6)2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ELEVATED INTRAOCULAR PRESSURE, IRIS ATROPHY AND UNREACTIVE PUPIL (FIXED). THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -17.5/+1.5/106 | OD | Physician | 08/04/2015 | 2015 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218062&pc=MTA | 2023826-2015-01498 |  | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/079 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EVENT DATE: UNK. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -12.5/+1.5/079 | OS | Physician | 10/13/2015 | 2015 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6064630&pc=MTA | 2023826-2016-01552 | 2016-10-04 | Injury | STAAR SURGICAL COMPANY | 2016-10-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00/+1.0/040 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT EXPERIENCED EXCESSIVE VAULTING OF THE LENS, ELEVATED INTRAOCULAR PRESSURE (IOP), AND LENS DISLOCATION/SUBLUXATION. THE PATIENT HAD TO UNDERGO AN ENLARGEMENT OF PERIPHERAL IRIDOTOMY (PI) OR ADDITION OF NEW PI USING YAG. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -08.00/+1.0/040 | OD | Physician | 10/04/2016 | 2016 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5891293&pc=MTA | 2023826-2016-00454 | 2016-02-18 | Injury | STAAR SURGICAL COMPANY | 2016-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Unintended Movement; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/090 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, LENS ROTATION AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE PATIENT'S POST-OP (ON (B)(6) 2016) UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND BENT. ELEVATED IOP NOT REPORTED FOR PATIENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+2.5/090 | OD | Physician | 03/09/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5308634&pc=MTA | 2023826-2015-01687 | 2015-10-22 | Injury | STAAR SURGICAL COMPANY | 2015-12-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Consequences Or Impact To Patient; No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -10.50/+2.50/X076 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT AND ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE) WAS NOTED. THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENS WITH A SIMILAR DIOPTER SIZE. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. DIMENSION AT INSPECTION FOUND THE LENS WAS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN US, BUT IS NOT MARKETED IN THE U.S. DIOPTER: -10.5/+2.5/X076. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -10.50/+2.50/076 | OD | Physician | 12/04/2015 | 2015 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824382&pc=MTA | 2023826-2018-01253 | 2018-07-28 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/063 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING. THE LENS REMAINS IMPLANTED. AT ONE WEEK REVIEW, THE PATIENTS INTRAOCULAR PRESSURE WAS 20MMHG AND PUPIL WAS REACTIVE EVEN THOUGH SEEMS SLOWER. THE PATIENT IS BEING MONITORED. Manufacturer Narrative: CORRECTED DATA: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -15.5/+1.5/063 | OS | Physician | 07/30/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7021353&pc=MTA | 2023826-2017-01672 | 2017-01-16 | Injury | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7, -15.5/1.0/105 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PATIENT EXPERIENCED RIGHT EYE PAIN BEFORE THE REPLACEMENT SURGERY. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.5/1.0/105 | OD | Physician | 10/18/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6872195&pc=MTA | 2023826-2017-01440 | 2017-04-24 | Injury | STAAR SURGICAL COMPANY | 2017-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -18.00/+3.20/100 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING, EYE PAIN, SHALLOW AC DEPTH, AND HIGH EYE PRESSURE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PATIENT'S POST-OP ICL EXCHANGE IS STABLE. THE STATUS OF THE EYE IS QUITE. THE PATIENT HAS NO COMPLAINT OF PAIN OR OTHER DISCOMFORTS. THE IOP IS ALSO WITHIN NORMAL LIMITS. THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) IS 20/28 AND UNCORRECTED VISUAL ACUITY (UCVA) IS 20/28. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18.00/+3.20/100 | OD | Physician | 08/21/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6962654&pc=MTA | 2023826-2017-01609 | 2017-04-24 | Injury | STAAR SURGICAL COMPANY | 2017-10-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+4.0/073 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND SHALLOW ANTERIOR CHAMBER DEPTH. THE PATIENT EXPERIENCED HIGH EYE PRESSURE AND PAIN. THE LENS WAS DISCARDED BY THE FACILITY. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2017, FOR A SHORTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE PATIENT IS DOING WELL, THE VAULT IS SUFFICIENT AND THE INTRAOCULAR PRESSURE IS WITHIN NORMAL LIMITS. (B)(4). REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+4.0/073 | OS | Physician | 10/02/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7636103&pc=MTA | 2023826-2018-00953 | 2018-06-04 | Injury | STAAR SURGICAL COMPANY | 2018-06-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+0.5/071 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2018 FOR A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED INTRAOCULAR PRESSURE INCREASE (ASSOCIATED WITH PAIN) DUE TO A HYPERVAULT OF THE ICL LEADING TO A NARROW CHAMBER ANGLE AND A FLATTENING OF THE ANTERIOR CHAMBER. PIGMENT DISPERSION AND IRIS ATROPHY DUE TO A HYPERVAULT OF THE ICL. NARROW CHAMBER ANGLE, FLATTENING OF ANTERIOR CHAMBER. PIGMENT DISPERSION, IRIS ATROPHY. DEVICE HISTORY RECORD REVIEW - BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.0/+0.5/071 | OD | Physician | 06/11/2018 | 2018 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4721367&pc=MTA | 2023826-2015-00429 | 2015-02-23 | Injury | STAAR SURGICAL COMPANY | 2015-04-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pain; Vomiting; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE AND AN UNREACTIVE PUPIL (FIXED). THE PATIENT EXPERIENCED SEVERE HEADACHE, VOMITING AND PAIN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. ANGLE CLOSURE/ELEVATED IOP IN THE PRESENCE OF A VICMO/VTICMO MODEL MAY OCCUR DUE TO: REMAINING VISCOELASTIC IN THE POSTERIOR CHAMBER, OVERSIZED ICL WITH ANGLE CLOSURE, TOO NARROW ANGLES/CROWDED AC DUE TO SMALL EYE ANATOMY PREOPERATIVELY, UNEXPECTED ABNORMAL ANATOMY OR TISSUE ABNORMALITIES (PRESENCE OF IRIS/CILIARY BODY CYSTS), BLOCKAGE OF THE PORT BY REMAINING VISCOELASTIC, OR BY AN INFLAMMATORY MEMBRANE, ETC. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -11.0/+2.0/095 | OD | Other | 03/24/2015 | 2015 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5766297&pc=MTA | 2023826-2016-00871 | 2016-05-12 | Injury | STAAR SURGICAL COMPANY | 2016-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Improper or Incorrect Procedure or Method; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.00/+3.5/123 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED A PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE, UNREACTIVE PUPIL (FIXED), GLARE/HALOS AND BLURRED VISION. THE SURGEON ENLARGED/ADDITION THE PERIPHERAL IRIDOTOMIES (YAG) AND ADDED A SUTURE. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 -1 AND THE PATIENT HAS PHOTOPHOBIA. Manufacturer Narrative: (B)(4): IMPROPER OR INCORRECT PROCEDURE OR METHOD - THE SURGEON REALIZED THAT THE IMPLANTED ICL MODEL REQUIRED PREOPERATIVE PERIPHERAL IRIDOTOMIES BUT HE FAILED TO DO IT DUE TO CONFUSION WITH THE ICL MODEL ORDERED. CORRECTION: FROM "13.7MM VTICMO13.7" TO "13.7MM VTICM13.7". (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS IMPLANTED. (B)(4). | -11.00/+3.5/123 | OD | Physician | 06/06/2016 | 2016 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6425385&pc=MTA | 2023826-2017-00441 | 2017-02-07 | Injury | STAAR SURGICAL COMPANY | 2017-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Discomfort; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1/82 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED IOP, AND SIGNIFICANT REDUCTION OF ICA. TO DATE, THE SURGEON DOES NOT PLAN TO EXCHANGE AS THE PUPIL REMAINS FIXED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS HAPTIC TORN. THERE WAS EVIDENCE OF THICK STICKY RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH-NO SIMILAR COMPLAINTS FOUND. OFF-LABEL, CONTRAINDICATED USE. ACD < 3.00MM (2.89MM). (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | -9.5/+1/82 | OS | Physician | 02/20/2017 | 2017 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5686682&pc=MTA | 2023826-2016-00778 | 2016-04-21 | Injury | STAAR SURGICAL COMPANY | 2016-05-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+2.0/064 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER INDICATED THE PATIENT EXPERIENCED PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE, PIGMENT DISPERSION AND GLARE/HALOS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). | -14.5/+2.0/064 | OS | Physician | 04/29/2016 | 2016 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5248647&pc=MTA | 2023826-2015-01460 | 2015-06-23 | Injury | STAAR SURGICAL COMPANY | 2015-11-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.00/-1.50/X155 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE AND UNREACTIVE PUPIL WAS NOTED. THE LENS WAS REMOVED AND NO OTHER LENS WAS IMPLANTED. PATIENT'S LAST VISIT ON (B)(6) 2015, BCVA WAS 20/30 AND UCVA WAS 20/800. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS ALSO CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -14.00/-1.50/X155. (B)(4). METHOD CODE(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -14.00/-1.50/155 | OS | Physician | 08/25/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4923030&pc=MTA | 2023826-2015-00827 |  | Injury | STAAR SURGICAL COMPANY | 2015-07-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PUPIL HAS REMAINED DILATED SINCE SURGERY AND THE INTRAOCULAR PRESSURE SPIKED UP TO 40MM HG (PUPIL BLOCK). THE PATIENT IS PHOTOPHOBIC. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. EVALUATION CODES: RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -8.0/+1.5/080 | OS | Physician | 06/22/2015 | 2015 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6540552&pc=MTA | 2023826-2017-00732 | 2017-03-15 | Injury | STAAR SURGICAL COMPANY | 2017-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+3.0/90 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED PUPIL BLOCK WITH ELEVATED IOP. ON THE SAME DAY, IN A SECOND SURGERY, THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE CUSTOMER STATED THE DR. THOUGHT THE ISSUE WAS CAUSED BY THE LENS SIZE, THOUGH HE DID NOT MEASURE THE VAULT. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A TEAR IN THE HAPTIC. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.5/+3.0/90 | OS | Physician | 04/04/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7219339&pc=MTA | 2023826-2018-00142 | 2017-12-27 | Injury | STAAR SURGICAL COMPANY | 2018-01-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.0/+3.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE, PUPIL BLOCK, PIGMENT DISPERSION AND UNREACTIVE PUPIL (FIXED). THE REPORTER INDICATED THERE IS SUSPECTED URRETS-ZAVALIA DUE TO TEMPORARY AQUAPORT BLOCK AND INTRAOCULAR PRESSURE SPIKE. AT LAST VISIT ON (B)(6) 2017, THE INTRAOCULAR PRESSURE WAS 10 MMHG, DILATED PUPIL NOT VERY REACTIVE AND SOME PIGMENT ON THE LENS. THE LENS REMAINS IMPLANTED. THE PATIENT WAS REPORTED AS BEING WELL AND WITHOUT PRESSURE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.0/+3.0/093 | OD | Physician | 12/27/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7236056&pc=MTA | 2023826-2018-00143 |  | Injury | STAAR SURGICAL COMPANY | 2018-02-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Insufficient Information | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -12.00/+4.5/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE REPORTER INDICATED THE PATIENT EXPERIENCED SIMILAR SYMPTOMS AS THE RIGHT EYE BUT NOT AS STRONG - ELEVATED INTRAOCULAR PRESSURE, PUPIL BLOCK, PIGMENT DISPERSION, UNREACTIVE PUPIL (FIXED) AND IRIS ATROPHY. THE REPORTER INDICATED THERE IS SUSPECTED URRETS-ZAVALIA DUE TO TEMPORARY AQUAPORT BLOCK AND INTRAOCULAR PRESSURE SPIKE. THE LENS REMAINS IMPLANTED. THE PATIENT WAS REPORTED AS BEING WELL AND WITHOUT PRESSURE. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/+4.5/085 | OS | Physician | 01/15/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6031008&pc=MTA | 2023826-2016-01435 | 2016-06-26 | Injury | STAAR SURGICAL COMPANY | 2016-10-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13/1/93 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT, ELEVATED IOP, PUPIL BLOCK (WITH ELEVATED IOP) AND ANGLE CLOSURE (WITH ELEVATED IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BENT/BROKEN/DEFORMED. UNIDENTIFIED FOREIGN MATERIAL WAS PRESENT ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13/1/93 | OD | Physician | 09/19/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6453289&pc=MTA | 2023826-2017-00532 | 2016-07-29 | Injury | STAAR SURGICAL COMPANY | 2017-04-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1/95 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, PUPIL BLOCK, AND ANGLE CLOSURE (WITH ELEVATED IOP). ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION-LENS WAS RETURNED IN LIQUID, IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/+1/95 | OD | Physician | 10/12/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6498926&pc=MTA | 2023826-2017-00646 | 2016-07-29 | Injury | STAAR SURGICAL COMPANY | 2017-04-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/1/81 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, PUPIL BLOCK, AND ANGLE CLOSURE (WITH ELEVATED IOP). ON (B)(6) 2016 THE LENS WAS EXPLANTED, AND THEN EXCHANGED FOR SHORTER LENS. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND FOREIGN MATERIAL PRESENT ON THE LENS SURFACE (FIBERS). DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, ANGLE CLOSURE). RESULT CODE: WORK ORDER SEARCH WAS PERFORMED, NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4). | -12.00/1/81 | OS | Physician | 10/12/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6327728&pc=MTA | 2023826-2017-00122 | 2016-12-21 | Injury | STAAR SURGICAL COMPANY | 2017-02-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -11.0/+3.5/91 DIOPTER, IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT, WHICH LEAD TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PUPIL BLOCK (WITH ELEVATED IOP), AND ANGLE CLOSURE (WITH ELEVATED IOP) IN PATIENT. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2016, AND THE PROBLEM WAS RESOLVED. ON PATIENT'S LAST VISIT ON (B)(6) 2016, THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION-LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. (B)(4). Manufacturer Narrative: (B)(4). | -11.0/+3.5/91 | OD | Physician | 01/17/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6582998&pc=MTA | 2023826-2017-00864 | 2017-04-10 | Injury | STAAR SURGICAL COMPANY | 2017-05-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, DIOPTER -8.50/+1.0/014 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO THE PATIENT EXPERIENCING EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED IOP. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SMALL VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND CLEAR RESIDUE).(B)(4). Manufacturer Narrative: PATIENT WEIGHT: UNK. THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | -8.50/+1.0/014 | OS | Physician | 04/25/2017 | 2017 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6670126&pc=MTA | 2023826-2017-01042 | 2017-05-09 | Injury | STAAR SURGICAL COMPANY | 2017-06-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.50/+2.0/164 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, PUPIL BLOCK WITH ELEVATED IOP. POST-OP FIRST DAY, THE SURGEON NOTED THE LENS PUSHED THE IRIS BACKWARD TO THE CRYSTALLINE AND IRIS HAS TOUCHED THE CRYSTALLINE LENS. HE PERFORMED THE ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS TWICE TO SOLVED THE PROBLEM BUT NOTHING HAS CHANGED. ON (B)(6) 2017 THE LENS WAS EXPLANTED. ON (B)(6) 2017, THE SURGEON COMMENTED THAT THE PUPIL FIXED DILATED POSTERIOR SYNECHIA, AND LESS PUPILLARY MOVEMENT. ON (B)(6) 2017, THE CUSTOMER STATED THE PATIENT'S EYE BACK TO NORMAL AS BEFORE, AND THE DR. IS PLANNING TO IMPLANT ANOTHER LENS FOR THE PATIENT. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.50/+2.0/164 | OD | Physician | 05/31/2017 | 2017 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6673332&pc=MTA | 2023826-2017-01061 | 2017-05-17 | Injury | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1/104 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE. THE SURGEON STATED THAT THE PATIENT ALSO HAS "CORTICAL PERIPHERIC ANTERIOR CATARACT. FREE VISUAL AXIS." THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. THE STATUS OF THE EYE IS PUPIL SPHINCTER PARALYSIS. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. (B)(4). Manufacturer Narrative: VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1/104 | OD | Physician | 06/07/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7076607&pc=MTA | 2023826-2017-01845 | 2017-11-13 | Injury | STAAR SURGICAL COMPANY | 2017-12-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50/+3.0/087 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PATIENT EXPERIENCED PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS REPOSITIONED AND THE PERIPHERAL IRIDECTOMIES WERE ENLARGED/ ADDITIONAL PERIPHERAL IRIDECTOMIES BY YAG. THE SURGEON ALSO DILATED THE PUPIL AND EVACUATED/IRRIGATED THE ANTERIOR CHAMBER OF ANY REMAINING VISCO/FLUIDS. THE LENS REMAINS IMPLANTED BUT THE SURGEON IS CONSIDERING EXPLANT AND EXCHANGE. AT LAST POST-OP VISIT ON (B)(6) 2017, THE PATIENT'S UCVA WAS 20/25++, INTRAOCULAR PRESSURE (IOP) WAS 12 AND THE ICL WAS COVERED WITH IRIS PIGMENT. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2017 AND WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP CONDITION BEFORE THE LENS WAS EXPLANTED WAS HIGH VAULT - 3 CCT AND INTRAOCULAR PRESSURE OF 19. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SMALL PLASTIC BAG. VISUAL INSPECTION FOUND A PIECE OF ONE HAPTIC BROKEN OFF. A DIMENSIONAL INSPECTION OF THE LENS WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.50/+3.0/087 | OD | Physician | 11/13/2017 | 2017 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4786337&pc=MTA | 2023826-2015-00593 | 2014-05-16 | Injury | STAAR SURGICAL COMPANY | 2015-05-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND NOTICED PUPILLARY BLOCK WITH ELEVATED IOP IN THE POST-OPERATIVE PERIOD. THIS WAS ASSOCIATED WITH A SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE ICL VAULTING WAS CONSIDERED TO BE EXCESSIVE. THREE WEEKS LATER THE SURGEON DECIDED TO REMOVE AND REPLACE THE LENS WAS A SHORTER ONE AND THIS RESOLVED THE PROBLEM. AFTER THE EXCHANGE THE PATIENT'S BCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH & MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED DRY WITH NO VISIBLE DAMAGE. THE LENS WAS REHYDRATED AND BOTH THE LENGTH AND WIDTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. MEDICAL REVIEW - PUPIL BLOCK WITH ELEVATED IOP IN THE PRESENCE OF HIGH VAULT EARLY AFTER ICL IMPLANTATION MAY OCCUR DUE TO: NON PATIENT/FUNCTIONING IRIDECTOMIES (TOO SMALL, TOO PERIPHERAL AND/OR NOT FULLY PERMEABLE, BLOCKED BY VISCO), REMAINING VISCOELASTIC IN THE ANTERIOR CHAMBER, OVERSIZED ICL WITH ANGLE CLOSURE, TOO NARROW ANGLES/CROWDED AC DUE TO SMALL EYE ANATOMY PREOPERATIVELY, UNEXPECTED ABNORMAL L ANATOMY OR TISSUE ABNORMALITIES (PRESENCE OF IRIS/CILIARY BODY CYSTS), MAL-POSITIONED FOOTPLATES, ECT. THERE IS NOT ENOUGH CLINICAL DATA TO DETERMINE THE EXACT CAUSE OF THIS EVENT, HOWEVER, WE CANNOT RULE OUT EXCESSIVE VAULTING SECONDARY TO A LONG LENS AS A CONTRIBUTING FACTOR. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP, ECT.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -2.1/1.5/+1.5/010 | OD | Physician | 02/05/2015 | 2015 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4933090&pc=MTA | 2023826-2015-00912 | 2014-11-22 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) /2014. PUPILLARY BLOCK WITH ELEVATED IOP ASSOCIATED WITH EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXPLANTED AND THE SURGEON PLANS ON IMPLANTING A SHORTER LENS. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED (B)(4). | -3.5/4.5/180 | OD | Physician | 06/26/2015 | 2015
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4937377&pc=MTA | 2023826-2015-00902 | 2015-06-16 | Injury | STAAR SURGICAL COMPANY | 2015-07-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, PUPIL BLOCK AND ELEVATED INTRAOCULAR PRESSURE. THE PI WAS EITHER ENLARGED OR THERE WAS AN ADDITION OF A NEW PI. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/80. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS DISCARDED. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS DISCARDED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULT: BASED ON THE RESULTS OF THE DHR REVIEW, THE AVAILABLE INFORMATION GIVEN WITHIN THIS COMPLAINT AND WORK ORDER SEARCH, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. THE MEDICAL REVIEW STATED THE SAME MODEL LENS WAS IMPLANTED IN BOTH EYES OF THIS PATIENT BY THE SAME DOCTOR WITH PREOPERATIVE ACD (ENDO) OF 2.64MM IN OD AND 2.71MM IN OS AND HISTORY OF KERATOCONUS. THE ACD IS OUTSIDE THE RECOMMENDED RANGE OF ICL IMPLANTATION WHICH IS GREATER OR EQUAL TO 3MM. AS REPORTED ON THE MEDICAL REVIEW THE MOST LIKELY CAUSE OF THE EVENT IS TOO SHALLOW EYES (OUTSIDE APPROVED INDICATION). ICL VAULTING MAY HAVE A CONTRIBUTING ROLE IN THIS CASE. CLAIM # (B)(4). | -7.00/+3.0/151 | OD | Physician | 07/07/2015 | 2015 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4937426&pc=MTA | 2023826-2015-00903 | 2015-06-16 | Injury | STAAR SURGICAL COMPANY | 2015-07-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, PUPIL BLOCK AND ELEVATED INTRAOCULAR PRESSURE. THE PI WAS EITHER ENLARGED OR THERE WAS AN ADDITION OF A NEW PI. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/120. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT CODE(S): (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE DHR REVIEW, THE AVAILABLE INFORMATION GIVEN WITHIN THIS COMPLAINT AND WORK ORDER SEARCH, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. THE MEDICAL REVIEW STATED THE SAME MODEL LENS WAS IMPLANTED IN BOTH EYES OF THIS PATIENT BY THE SAME DOCTOR WITH PREOPERATIVE ACD (ENDO) OF 2.64MM IN OD AND 2.71MM IN OS AND HISTORY OF KERATOCONUS. THE ACD IS OUTSIDE THE RECOMMENDED RANGE OF ICL IMPLANTATION WHICH IS GREATER OR EQUAL TO 3MM. AS REPORTED ON THE MEDICAL REVIEW THE MOST LIKELY CAUSE OF THE EVENT IS TOO SHALLOW EYES (OUTSIDE APPROVED INDICATION). ICL VAULTING MAY HAVE A CONTRIBUTING ROLE IN THIS CASE. (B)(4). | -8.00/+4.5/040 | OS | Physician | 07/07/2015 | 2015 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323196&pc=MTA | 2023826-2015-01707 | 2015-07-10 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.0MM TICM130V4 IMPLANTABLE COLLAMER LENS, -16.0/+4.5/079 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, PUPIL BLOCK AND ELEVATED INTRAOCULAR PRESSURE. THE PATIENT EXPERIENCED IRIS ATROPHY. THE PERIPHERAL IRIDOTOMIES WERE SURGICAL ENLARGED AND THE LENS WAS REPOSITIONED. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | -16.0/+4.5/079 | OS | Physician | 11/16/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5223686&pc=MTA | 2023826-2015-01408 | 2015-07-22 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -4.00/+2.00/X096 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED ON (B)(6) 2015. A YAG AND ENLARGEMENT OF PI WAS PERFORMED. THE LENS WAS EXPLANTED AND WAS EXCHANGED FOR A SHORTER LENS. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015; UCVA WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4), BUT IS NOT MARKETED IN THE U.S. DIOPTER: -4.00/+2.00/X096. EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -4.00/+2.00/096 | OS | Physician | 10/13/2015 | 2015 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5632893&pc=MTA | 2023826-2016-00536 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SHALLOW ANTERIOR CHAMBER, ANGLE CLOSURE, PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. AT 4 DAYS POST-OP THE LENS HAD A HIGH VAULT WITH NON-CONTROLLED INTRAOCULAR PRESSURE ELEVATION. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION - THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE (UNSPECIFIED). PATIENT CODE - (B)(4) ANGLE CLOSURE, SHALLOW ANTERIOR CHAMBER. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+2.0/085 | OS | Physician | 03/22/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5685192&pc=MTA | 2023826-2016-00669 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+1.50/180 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND REMAINS IMPLANTED. THE LENS IS PLANNED TO EXCHANGE WITH ANOTHER LENS. THE PATIENT EXPERIENCED PUPIL BLOCK (WITH ELEVATED IOP). Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND OPTIC TORN/SPLIT AND HAPTIC TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED AT AN UNKNOWN DATE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT AGE AND WEIGHT: UNK. COMMON DEVICE NAME: UNK. EVENT DATE: UNK. EXPLANT DATE: UNK. EVENT PROBLEM AND EVALUATION CODES: (B)(4). | -7.0/+1.50/180 | OS | Physician | 05/02/2016 | 2016
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6487402&pc=MTA | 2023826-2017-00630 | 2017-02-27 | Injury | STAAR SURGICAL COMPANY | 2017-04-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+4/101 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS ROTATION, PUPIL BLOCK WITH ELEVATED IOP, AND AN IRIS THAT WILL NOT DECREASE IN SIZE. THE LENS WAS REPOSITIONED. THE SURGEON NOTES THAT AN EXCHANGE IS PLANNED, REPOSITIONING DID NOT WORK. TO DATE, UNABLE TO OBTAIN INFORMATION REGARDING IF/WHEN AN EXCHANGE OCCURRED. THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS REMAINS IMPLANTED. | -12.5/+4/101 | OD | Physician | 03/14/2017 | 2017 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5750966&pc=MTA | 2023826-2016-00862 | 2016-05-03 | Injury | STAAR SURGICAL COMPANY | 2016-06-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, DIOPTER UNKNOWN, IN THE PATIENT'S EYE AND THE SURGERY WAS UNEVENTFUL. SOON AFTER THE INITIAL SURGERY, THE PATIENT EXPERIENCED SEVERE PUPIL BLOCK DESPITE TWO ADEQUATE IRIDOTOMIES. INITIALLY, RETAINED VISCOELASTIC WAS THE CAUSE BECAUSE DILATING THE PUPIL DID NOT COMPLETELY RESOLVE THE PUPIL BLOCK. AFTER A FEW DAYS, THE DILATING DID RESOLVE IT, BUT EACH TIME THE PUPIL CAME DOWN, THE PATIENT HAD A PUPIL BLOCK. THE EVENT WAS RESOLVED BY GOING BACK TO THE OR AND PERFORMING A SURGICAL IRIDECTOMY. THE LENS REMAINS IMPLANTED BUT THE PATIENT HAS REQUESTED THE LENS BE REMOVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -15.5/+3.0/100 DIOPTER, WAS IMPLANTED IN THE LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DATE OF BIRTH - UNK. EXPIRATION DATE - UNK. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4). | -15.5/+3.0/100 | OS | Physician | 05/25/2016 | 2016 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6577252&pc=MTA | 2023826-2017-00848 |  | Injury | STAAR SURGICAL COMPANY | 2017-05-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pupillary Block; No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED PROBLEMS WITH A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18/+2.0/X082 DIOPTER. HOWEVER, WE RECEIVED CONFLICTING INFORMATION REGARDING THIS LENS. REPORTER LISTS BOTH EYES AS BEING AFFECTED, STATES THAT THE LENS WAS NEITHER IMPLANTED NOR EXPLANTED DUE TO HIGH VAULT, PUPIL BLOCK, AND ELEVATED IOP. SURGERY DATE IS LISTED AS (B)(6) 2016, HOWEVER, LENS WAS NOT SHIP TO THE CUSTOMER UNTIL (B)(6) 2016. ATTEMPTS TO CLARIFY HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: THE REPLACEMENT LENS WAS SUBMITTED IN THE INITIAL MDR, INCORRECTLY. THERE ARE NO ADVERSE EVENTS ASSOCIATED WITH THE REPLACEMENT LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18/+2.0/082 | NA | Physician | 03/03/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6037544&pc=MTA | 2023826-2016-01480 | 2016-06-09 | Injury | STAAR SURGICAL COMPANY | 2016-10-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS 16.50/25/74 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS REMAINS IMPLANTED. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOS. THE ELEVATED IOP WAS CONTROLLED WITH TWO GLAUCOMA MEDICATIONS. DIPLOPIA EXPERIENCED IN LOW LIGHT. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.5/2.5/74 DIOPER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND GLARE/HALOES. THE ELEVATED IOP WAS CONTROLLED WITH TWO GLAUCOMA MEDICATIONS. PATIENT ALSO EXPERIENCED DIPLOPIA IN LOW LIGHT. THE LENS WAS EXCHANGED ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. AS OF (B)(6) 2016, PATIENT'S POST-OP BEST-CORRECTED AND UNCORRECTED VISUAL ACUITIES WERE 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: DATE OF EVENT LISTED ON THE INITIAL REPORT IS INCORRECT; CORRECT DATE IS (B)(6) 2016. DATE RETURNED TO MANUFACTURER ON THE INITIAL REPORT IS INCORRECT; CORRECT DATE IS (B)(6) 2017. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.5/2.5/74 | OD | Physician | 09/23/2016 | 2016 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6037545&pc=MTA | 2023826-2016-01481 | 2016-06-11 | Injury | STAAR SURGICAL COMPANY | 2016-10-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS - -16.00/25/133 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS REMAINS IMPLANTED. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOS. THE ELEVATED IOP WAS CONTROLLED WITH TWO GLAUCOMA MEDICATIONS. DIPLOPIA EXPERIENCED IN LOW LIGHT. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). METHOD: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.00/2.5/133 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND GLARES/HALOES. DIPLOPLIA WAS ALSO EXPERIENCED IN LOW LIGHT. THE ELEVATED IOP WAS CONTROLLED WITH TWO GLAUCOMA MEDICATIONS. THE LENS WAS EXCHANGED ON (B)(6) 2016 FOR A SHORTER LENS OF A DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. AS OF (B)(6) 2016, PATIENT'S POST-OP BEST-CORRECTED AND UNCORRECTED VISUAL ACUITIES WERE 20/20. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.00/25/133 | OS | Physician | 09/23/2016 | 2016 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8254868&pc=MTA | 2023826-2019-00062 | 2018-05-18 | Injury | STAAR SURGICAL COMPANY | 2019-01-16 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+4.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, GLARE/HALOS, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS "OVERSIZE OF LENS BECAUSE OF RIGHT EYE ANATOMY". Manufacturer Narrative: THE LENS WAS RETURNED MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.; (B)(4). | -15.5/+4.0/094 | OD | Physician | 12/17/2018 | 2018 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5985749&pc=MTA | 2023826-2016-01347 | 2016-05-13 | Injury | STAAR SURGICAL COMPANY | 2016-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure, Delayed, Uncontrolled; Intraocular Pressure Increased; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -4.5/+2.0/090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. A SHORTER LENS WAS IMPLANTED ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -4.5/+2.0/090 | OS | Physician | 08/30/2016 | 2016 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7409133&pc=MTA | 2023826-2018-00523 | 2018-02-12 | Injury | STAAR SURGICAL COMPANY | 2018-04-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure, Delayed, Uncontrolled; Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOS AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+1.5/093 | OS | Physician | 03/10/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6015356&pc=MTA | 2023826-2016-01434 | 2016-08-30 | Injury | STAAR SURGICAL COMPANY | 2016-10-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure, Delayed, Uncontrolled; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11/+6/10 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. Manufacturer Narrative: LENS DESCRIPTION IS -11.0/+6.0/140 (SPHERE/CYLINDER/AXIS). VAULT VALUE PRIOR TO EXPLANT WAS REPORTED AS 1500 UM. THE LENS WAS EXCHANGED FOR A SHORTER SAME MODEL LENS WITH POST-OP VAULT VALUE OF 1 MM AND NO REPORTED PATIENT ISSUES. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). CORRECTED DATA: REPORT SOURCE IS DISTRIBUTOR, NOT USER FACILITY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+6.0/140 | OD | Physician | 09/19/2016 | 2016 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7629830&pc=MTA | 2023826-2018-00945 |  | Injury | STAAR SURGICAL COMPANY | 2018-06-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure, Delayed, Uncontrolled; Pain; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+2.0/123 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PATIENT EXPERIENCED OCULAR PAIN IN DARKNESS, WITH IOP SPIKES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.5/+2.0/123 | OD | Physician | 05/28/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7403687&pc=MTA | 2023826-2018-00512 | 2018-03-02 | Injury | STAAR SURGICAL COMPANY | 2018-04-06 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Iritis; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -9.5/+0.5/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED, REPORTEDLY DUE TO PIGMENT DISPERSION (MORE INFERIOR THAN SUPERIOR), UNREACTIVE (MID-DILATED FIXED) PUPIL, AND IRITIS. THE SURGEON STATED HE THOUGHT THE PATIENT MAY HAVE HAD AN ASYMPTOMATIC PUPIL BLOCK DUE TO THE VISCOELASTIC, BUT IS UNABLE TO THEN EXPLAIN WHY IT TOOK 2-3 WEEKS FOR IRITIS AND PIGMENT DISPERSION TO DEVELOP. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4): NO CODE AVAILABLE (PIGMENT DISPERSION, SECONDARY SURGERY, LENS EXPLANT, FIXED PUPIL). CONCLUSION CODE: 24-OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.86MM)] (B)(4). | -9.5/+0.5/094 | OD | Physician | 03/09/2018 | 2018 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7419858&pc=MTA | 2023826-2018-00530 | 2018-03-02 | Injury | STAAR SURGICAL COMPANY | 2018-04-10 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Iritis; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.5/+2.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED, REPORTEDLY DUE TO PIGMENT DISPERSION (MORE INFERIOR THAN SUPERIOR), UNREACTIVE (MID-DILATED FIXED) PUPIL, AND IRITIS. ON (B)(6) 2018 THE SURGEON HAD PERFORMED A YAG PROCEDURE. THE SURGEON STATED HE THOUGHT THE PATIENT MAY HAVE HAD AN ASYMPTOMATIC PUPIL BLOCK DUE TO THE VISCOELASTIC, BUT IS UNABLE TO THEN EXPLAIN WHY IT TOOK 2-3 WEEKS FOR IRITIS AND PIGMENT DISPERSION TO DEVELOP. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -11.5/+2.5/095 | OS | Physician | 03/15/2018 | 2018 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7419871&pc=MTA | 2023826-2018-00536 | 2018-03-02 | Injury | STAAR SURGICAL COMPANY | 2018-04-10 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Iritis; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO 12.6, -16.5/+2.0/075 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS IRITIS AND STATES THE PATIENT IS BEING TREATED WITH STEROIDS. THE LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: ADDITIONAL INFORMATION: REPORTEDLY, NO ANTIBIOTICS WERE ADMINISTERED BEYOND WHAT IS STANDARD FOR POST-OP TREATMENT AFTER OCULAR SURGERY. THE REPORTER STATES THE PATIENT CONTINUES TO HAVE A VERY MILD IRITIS AND IS BEING CLOSELY MONITORED. (B)(4) Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -16.5/+2.0/075 | OS | Physician | 03/15/2018 | 2018 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8070859&pc=MTA | 2023826-2018-01681 | 2018-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Iritis; Patient Problem/Medical Problem; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.00/3.5/164 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). THE LENS WAS REMOVED ON THE SAME DAY, DIFFERENT SESSION (1 HOUR LATER). ACUTE IRITIS, ANTERIOR CHAMBER IRRIGATION, AND INTRAOCULAR INJECTION WERE REPORTED. PATIENT RELATED FACTOR WAS REPORTED AS CAUSE OF THIS EVENT. THE REPORTER STATED "I BELIEVE THE CAUSE IS PATIENT IDIOSYNCRASY, DEVELOPED ACUTE IRITIS (MECHANICAL FRICTION) NEVER SEEN BEFORE. THE PATIENT COULDN'T TOLERATE THE PAIN (NO INCREASE IOP) AND ASKED FOR IMMEDIATE REMOVAL, ALMOST 30 MIN AFTER IMPLANTATION. "KERATOCONUS WAS REPORTED AS A PRE-EXISTING CONDITION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED: TOBRADEX DROPS 1X5, NEVANAC 1MG TODAY, AND ACETAZOLAMIDE 250MG 1/2 TAB EVERY 6 HOURS, WERE PRESCRIBED FOR ONE WEEK TO ADDRESS THE INFLAMMATION. EVACUATION OF VISCOELASTIC IS ALWAYS PERFORMED AS PART OF THE NORMAL PROCEDURE. AFTER EXPLANTING THE LENS THE PROBLEM RESOLVED, PATIENT IS DOING FINE. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. NO INTRAOCULAR INJECTION WAS USED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS, CLAIM#: (B)(4). | -7.00/3.5/164 | OD |  | 10/18/2018 | 2018 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119131&pc=MTA | 2023826-2015-01310 | 2015-08-31 | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Loss of Vision | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+0.5/X057 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THERE WAS LOSS OF BEST-CORRECTED VISUAL ACUITY POST-OP. THE LENS REMAINS IMPLANTED. THERE WERE STRONG CENTRAL DEPOSITS ON THE CRYSTALLINE LENS (VISCO). Manufacturer Narrative: METHOD (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION - INVESTIGATION SUMMARY: THE MANUFACTURER'S DEVICE ANALYSIS RESULTS - 1 THE EVENTS REPORTED BY PROF. (B)(6) ARE LIMITED TO A SINGLE SITE BY A SINGLE SURGEON IMPLANTING IMPLANTABLE COLLAMER LENSES (LCLS) BETWEEN THE DATES OF (B)(6) 2015. TEN PATIENTS WERE REPORTED TO HAVE MILKY WHITE PRECIPITATES FORM ON THE SURFACE OF THE ICL AND ON THE ANTERIOR LENS CAPSULE OF THE PHYSIOLOGIC CRYSTALLINE LENS. NO ADDITIONAL REPORTS OF PRECIPITATES HAVE BEEN RECEIVED FROM ANY OTHER SURGEON OR SITE WITH THE SAME LOT NUMBERS AS THE REPORTED EVENTS FROM PROF. (B)(6). ADDITIONALLY, IT IS IMPORTANT TO NOTE THAT IN THE DIRECTIONS FOR USE (DFU) OF THE LCL, THE COMPLICATIONS AND ADVERSE REACTIONS SECTION IDENTIFIES THE POTENTIAL RISK FOR "SECONDARY SURGICAL INTERVENTION TO REMOVE/REPLACE THE LENS." THE PRECAUTIONS SECTION ALSO SUGGESTS "THE LENS SHOULD NOT BE EXPOSED TO ANY SOLUTIONS OTHER THAN THE NORMALLY USED INTRAOCULAR IRRIGATING SOLUTIONS (E.G. ISOTONIC SALINE, BSS, VISCOELASTIC. ETC.)." DURING THE IN-DEPTH INVESTIGATION WITH PROF. (B)(6) TO IDENTIFY ROOT CAUSE, THE INTRAOPERATIVE REGIMEN REVEALED THAT INTRACAMERAL 0.1% CEFUROXIME ANTIBIOTIC FOR PROPHYLAXIS WAS USED IN THE 10 REPORTED CASES, ALONG WITH RINGER'S SOLUTION. LN AN ARTICLE PUBLISHED BY VASAVADA ET AL. (COMPARISON BETWEEN RINGER'S LACTATE AND BALANCED SALT SOLUTION ON POSTOPERATIVE OUTCOMES AFTER PHACOEMULSIFICATION: A RANDOMIZED CLINICAL TRIAL, LNDIAN J OPHTHALMOL. 2009 MAY-JUN; 57(3): 191-195), THE AUTHORS CONCLUDE THAT RINGER'S LACTATE IS HYPOTONIC AND SLIGHTLY ACIDIC (OSMOLALITY 280MMOL, PH 6.0) AS COMPARED TO BSS (OSMOLALITY 302 MMOL, PH 7.4) AND AQUEOUS (OSMOLALITY 302 MMOL, PH 7.4). THEY FURTHER CONCLUDED THAT A STATISTICALLY SIGNIFICANT LARGER NUMBER OF PATIENTS WITH FLARE AND CELLS MAY BE DUE TO THE SLIGHTLY ACIDIC FLUCTUATING PH AND LACK OF BUFFER SYSTEM IN RINGER'S LACTATE, WHICH COULD AFFECT THE BLOOD AQUEOUS BARRIER STABILITY. ON THE OTHER HAND, 855 HAS AN ACETATE CITRATE BUFFER SYSTEM AND THE PH, THOUGH SLIGHTLY ALKALINE, IS STABLE. SUBSEQUENT TO THE SERIES OF REPORTED EVENTS, ADDITIONAL PATIENTS HAVE BEEN TREATED BY PROF. (B)(6) WITH AN ICL USING ONLY COMMERCIALLY AVAILABLE BALANCED SALINE SOLUTION (BSS, ALCON) AND OCUCOAT (BAUSCH & LOMB) INTRAOPERATIVE, WITHOUT THE USE OF 0.1% INTRACAMERAL CEFUROXIME ANTIBIOTIC. THERE WERE NO OBSERVED INCIDENTS OF DEPOSITS FOLLOWING THIS PROTOCOL REGIMEN. 2 UPON REVIEW OF THE DEVICE HISTORY RECORDS (DHR), ALL PRODUCTS REPORTED IN THE INCIDENCES, BASED UPON LOT AND SERIAL NUMBER REVIEW, INDICATED SPECIFICATIONS WERE MET DURING MANUFACTURING PROCESS. ALL PRODUCTS PASSED QUALITY INSPECTIONS AND WERE RELEASED PER ACCEPTANCE CRITERIA. FINAL COMMENTS FROM THE MANUFACTURER: BASED UPON A THOROUGH REVIEW OF ALL AVAILABLE DATA AND ON CLINICAL PERSPECTIVE AS A PRACTICING OPHTHALMIC SURGEON WITH EXPERTISE IN THE LCL, DR. (B)(6) (MEDICAL CONSULTANT) AND STAAR CONCLUDED THAT THE OBSERVED PRECIPITATES WERE MOST LIKELY THE RESULT OF INTERACTION BETWEEN THE INTRAOCULAR MEDICATIONS USED AT THE TIME OF SURGERY. (B)(4). Manufacturer Narrative: THE CUSTOMER HAS REPORTED SINCE USING THE NEW SURGICAL PROTOCOL (USING BSS AND OCUCOAT ONLY) NO NEW COMPLAINT CASES OF LOSS OF BCVA, WITH PRECIPITATE ON THE CRYSTALLINE LENS, HAVE BEEN RECEIVED SINCE THE BEGINNING OF 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS IMPLANTED. | -10.0/+0.5/057 | OS | Physician | 09/03/2015 | 2015 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5124740&pc=MTA | 2023826-2015-01307 | 2015-09-01 | Injury | STAAR SURGICAL COMPANY | 2015-10-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Loss of Vision | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+0.5/X094 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE REPORTER INDICATED THERE WAS LOSS OF BEST-CORRECTED VISUAL ACUITY POST-OP. THE LENS REMAINS IMPLANTED. THERE WERE DEPOSITS ON THE SURFACE OF THE CRYSTALLINE LENS (VISCO). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: INVESTIGATION SUMMARY: THE MANUFACTURER'S DEVICE ANALYSIS RESULTS - THE EVENTS REPORTED BY PROF. (B)(6) ARE LIMITED TO A SINGLE SITE BY A SINGLE SURGEON IMPLANTING IMPLANTABLE COLLAMER LENSES (LCLS) BETWEEN THE DATES OF (B)(6) 2015. TEN PATIENTS WERE REPORTED TO HAVE MILKY WHITE PRECIPITATES FORM ON THE SURFACE OF THE ICL AND ON THE ANTERIOR LENS CAPSULE OF THE PHYSIOLOGIC CRYSTALLINE LENS. NO ADDITIONAL REPORTS OF PRECIPITATES HAVE BEEN RECEIVED FROM ANY OTHER SURGEON OR SITE WITH THE SAME LOT NUMBERS AS THE REPORTED EVENTS FROM PROF. (B)(6). ADDITIONALLY, IT IS IMPORTANT TO NOTE THAT IN THE DIRECTIONS FOR USE (DFU) OF THE LCL, THE COMPLICATIONS AND ADVERSE REACTIONS SECTION IDENTIFIES THE POTENTIAL RISK FOR "SECONDARY SURGICAL INTERVENTION TO REMOVE/REPLACE THE LENS." THE PRECAUTIONS SECTION ALSO SUGGESTS "THE LENS SHOULD NOT BE EXPOSED TO ANY SOLUTIONS OTHER THAN THE NORMALLY USED INTRAOCULAR IRRIGATING SOLUTIONS (E.G. ISOTONIC SALINE, BSS, VISCOELASTIC. ETC.)." DURING THE IN-DEPTH INVESTIGATION WITH PROF. (B)(6) TO IDENTIFY ROOT CAUSE, THE INTRAOPERATIVE REGIMEN REVEALED THAT INTRACAMERAL 0.1% CEFUROXIME ANTIBIOTIC FOR PROPHYLAXIS WAS USED IN THE (B)(4) REPORTED CASES, ALONG WITH RINGER'S SOLUTION. IN AN ARTICLE PUBLISHED BY VASAVADA ET AL. (COMPARISON BETWEEN RINGER'S LACTATE AND BALANCED SALT SOLUTION ON POSTOPERATIVE OUTCOMES AFTER PHACOEMULSIFICATION: A RANDOMIZED CLINICAL TRIAL, LNDIAN J OPHTHALMOL. 2009 MAY-JUN; 57(3): 191-195), THE AUTHORS CONCLUDE THAT RINGER'S LACTATE IS HYPOTONIC AND SLIGHTLY ACIDIC (OSMOLALITY 280MMOL, PH 6.0) AS COMPARED TO BSS (OSMOLALITY 302 MMOL, PH 7.4) AND AQUEOUS (OSMOLALITY 302 MMOL, PH 7.4). THEY FURTHER CONCLUDED THAT A STATISTICALLY SIGNIFICANT LARGER NUMBER OF PATIENTS WITH FLARE AND CELLS MAY BE DUE TO THE SLIGHTLY ACIDIC FLUCTUATING PH AND LACK OF BUFFER SYSTEM IN RINGER'S LACTATE, WHICH COULD AFFECT THE BLOOD AQUEOUS BARRIER STABILITY. ON THE OTHER HAND, 855 HAS AN ACETATE CITRATE BUFFER SYSTEM AND THE PH, THOUGH SLIGHTLY ALKALINE, IS STABLE. SUBSEQUENT TO THE SERIES OF REPORTED EVENTS, ADDITIONAL PATIENTS HAVE BEEN TREATED BY PROF. (B)(6) WITH AN ICL USING ONLY COMMERCIALLY AVAILABLE BALANCED SALINE SOLUTION (BSS, ALCON) AND OCUCOAT (BAUSCH & LOMB) INTRAOPERATIVE, WITHOUT THE USE OF 0.1% INTRACAMERAL CEFUROXIME ANTIBIOTIC. THERE WERE NO OBSERVED INCIDENTS OF DEPOSITS FOLLOWING THIS PROTOCOL REGIMEN. UPON REVIEW OF THE DEVICE HISTORY RECORDS (DHR), ALL PRODUCTS REPORTED IN THE INCIDENCES, BASED UPON LOT AND SERIAL NUMBER REVIEW, INDICATED SPECIFICATIONS WERE MET DURING MANUFACTURING PROCESS. ALL PRODUCTS PASSED QUALITY INSPECTIONS AND WERE RELEASED PER ACCEPTANCE CRITERIA. FINAL COMMENTS FROM THE MANUFACTURER: BASED UPON A THOROUGH REVIEW OF ALL AVAILABLE DATA AND ON CLINICAL PERSPECTIVE AS A PRACTICING OPHTHALMIC SURGEON WITH EXPERTISE IN THE LCL, DR. (B)(6) (MEDICAL CONSULTANT) AND STAAR CONCLUDED THAT THE OBSERVED PRECIPITATES WERE MOST LIKELY THE RESULT OF INTERACTION BETWEEN THE INTRAOCULAR MEDICATIONS USED AT THE TIME OF SURGERY. (B)(4). Manufacturer Narrative: THE CUSTOMER HAS REPORTED SINCE USING THE NEW SURGICAL PROTOCOL (USING BSS AND OCUCOAT ONLY) NO NEW COMPLAINT CASES OF LOSS OF BCVA, WITH PRECIPITATE ON THE CRYSTALLINE LENS, HAVE BEEN RECEIVED SINCE THE BEGINNING OF 2016. CLAIM # (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS IMPLANTED. WORK ORDER SEARCH. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -11.5/+0.5/094 | OD | Physician | 09/03/2015 | 2015 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5151863&pc=MTA | 2023826-2015-01256 | 2015-09-15 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Loss of Vision | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+0.5/X071 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. POST-OP. THERE WERE DEPOSITS ON THE CRYSTALLINE LENS. SEE MFR #2023826-2015-01257 LEFT EYE. Manufacturer Narrative: ADDITIONAL INFORMATION - INVESTIGATION SUMMARY: THE MANUFACTURER'S DEVICE ANALYSIS RESULTS - THE EVENTS REPORTED BY (B)(6) ARE LIMITED TO A SINGLE SITE BY A SINGLE SURGEON IMPLANTING IMPLANTABLE COLLAMER LENSES (LCLS) BETWEEN THE DATES OF MARCH 2015 AND OCTOBER 2015. TEN PATIENTS WERE REPORTED TO HAVE MILKY WHITE PRECIPITATES FORM ON THE SURFACE OF THE ICL AND ON THE ANTERIOR LENS CAPSULE OF THE PHYSIOLOGIC CRYSTALLINE LENS. NO ADDITIONAL REPORTS OF PRECIPITATES HAVE BEEN RECEIVED FROM ANY OTHER SURGEON OR SITE WITH THE SAME LOT NUMBERS AS THE REPORTED EVENTS FROM (B)(6). ADDITIONALLY, IT IS IMPORTANT TO NOTE THAT IN THE DIRECTIONS FOR USE (DFU) OF THE LCL, THE COMPLICATIONS AND ADVERSE REACTIONS SECTION IDENTIFIES THE POTENTIAL RISK FOR "SECONDARY SURGICAL INTERVENTION TO REMOVE/REPLACE THE LENS". THE PRECAUTIONS SECTION ALSO SUGGESTS "THE LENS SHOULD NOT BE EXPOSED TO ANY SOLUTIONS OTHER THAN THE NORMALLY USED INTRAOCULAR IRRIGATING SOLUTIONS (E.G. ISOTONIC SALINE, BSS, VISCOELASTIC. ETC.)." DURING THE IN-DEPTH INVESTIGATION WITH (B)(6) TO IDENTIFY ROOT CAUSE, THE INTRAOPERATIVE REGIMEN REVEALED THAT INTRACAMERAL 0.1% CEFUROXIME ANTIBIOTIC FOR PROPHYLAXIS WAS USED IN THE 10 REPORTED CASES, ALONG WITH RINGER'S SOLUTION. LN AN ARTICLE PUBLISHED BY VASAVADA ET AL. (COMPARISON BETWEEN RINGER'S LACTATE AND BALANCED SALT SOLUTION ON POSTOPERATIVE OUTCOMES AFTER PHACOEMULSFICATION: A RANDOMIZED CLINICAL TRIAL, LNDIAN J OPHTHALMOL. 2009 MAY-JUN; 57(3): 191-195), THE AUTHORS CONCLUDE THAT RINGER'S LACTATE IS HYPOTONIC AND SLIGHTLY ACIDIC (OSMOLALITY 280MMOL, PH 6.0) AS COMPARED TO BSS (OSMOLALITY 302 MMOL, PH 7.4) AND AQUEOUS (OSMOLALITY 302 MMOL, PH 7.4). THEY FURTHER CONCLUDED THAT A STATISTICALLY SIGNIFÌCANT LARGER NUMBER OF PATIENTS WITH FLARE AND CELLS MAY BE DUE TO THE SLIGHTLY ACIDIC FLUCTUATING PH AND LACK OF BUFFER SYSTEM IN RINGER'S LACTATE, WHICH COULD AFFECT THE BLOOD AQUEOUS BARRIER STABILITY. ON THE OTHER HAND, 855 HAS AN ACETATE CITRATE BUFFER SYSTEM AND THE PH, THOUGH SLIGHTLY ALKALINE, IS STABLE. SUBSEQUENT TO THE SERIES OF REPORTED EVENTS, ADDITIONAL PATIENTS HAVE BEEN TREATED BY (B)(6) WITH AN ICL USING ONLY COMMERCIALLY AVAILABLE BALANCED SALINE SOLUTION (BSS, ALCON) AND OCUCOAT (BAUSCH & LOMB) INTRAOPERATIVE, WITHOUT THE USE OF 0.1% INTRACAMERAL CEFUROXIME ANTIBIOTIC. THERE WERE NO OBSERVED INCIDENTS OF DEPOSITS FOLLOWING THIS PROTOCOL REGIMEN. UPON REVIEW OF THE DEVICE HISTORY RECORDS (DHR), ALL PRODUCTS REPORTED IN THE INCIDENCES, BASED UPON LOT AND SERIAL NUMBER REVIEW, INDICATED SPECIFICATIONS WERE MET DURING MANUFACTURING PROCESS. ALL PRODUCTS PASSED QUALITY INSPECTIONS AND WERE RELEASED PER ACCEPTANCE CRITERIA. FINAL COMMENTS FROM THE MANUFACTURER: BASED UPON A THOROUGH REVIEW OF ALL AVAILABLE DATA AND ON CLINICAL PERSPECTIVE AS A PRACTICING OPHTHALMIC SURGEON WITH EXPERTISE IN THE LCL, DR. (B)(6) (MEDICAL CONSULTANT) AND STAAR CONCLUDED THAT THE OBSERVED PRECIPITATES WERE MOST LIKELY THE RESULT OF INTERACTION BETWEEN THE INTRAOCULAR MEDICATIONS USED AT THE TIME OF SURGERY. (B)(4). Manufacturer Narrative: THE CUSTOMER HAS REPORTED SINCE USING THE NEW SURGICAL PROTOCOL (USING BSS AND OCUCOAT ONLY) NO NEW COMPLAINT CASES OF LOSS OF BCVA, WITH PRECIPITATE ON THE CRYSTALLINE LENS, HAVE BEEN RECEIVED SINCE THE BEGINNING OF 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -6.5/+0.5/X071. (B)(4). METHOD: EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -6.5/+0.5/071 | OD | Physician | 09/18/2015 | 2015 | 51
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4672087&pc=MTA | 2023826-2015-00420 |  | Injury | STAAR SURGICAL COMPANY | 2015-04-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Loss of Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) AND A REFRACTIVE SURPRISE WAS NOTED. PATIENT ALSO EXPERIENCED A LOSS OF BCVA. THE REPORTER INDICATED THE EVENT WAS THE RESULT OF A SURGEON'S ERROR WHEN ENTERING THE PRE-OP DATA IN OCOS. THE LENS REMAINS IMPLANTED, HOWEVER, THE SURGEON PLANS TO EXCHANGE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - (REFRACTIVE SURPRISE03); VISION, LOSS OF; NO KNOWN PROBLEM WITH THE DEVICE. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH & MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW- WHEN A WRONG LENS POWER IS IMPLANTED AND EXCHANGED WITH THE CORRECT LENS POWER RESOLVES THE PROBLEM. THE SECONDARY INTERVENTION IS USUALLY DONE BY MEANS OF SAME INCISION WHICH WOULD NOT REQUIRE INCISION ENLARGEMENT AND/OR SUTURING DUE TO THE FOLDABLE PROPERTIES OF THE COLLAMER MATERIAL (ICL). (B)(4). | -07/-3.5/15 | OS | Physician | 03/10/2015 | 2015
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4786463&pc=MTA | 2023826-2015-00594 | 2014-04-11 | Injury | STAAR SURGICAL COMPANY | 2015-05-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Loss of Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND THE PATIENT EXPERIENCED A LOSS OF VISION (PRE-OP 20/20,POST-OP 20/50. THE EVENT WAS THE RESULT OF A SURGEONS ERROR WHEN ENTERING CYLINDER SIGN IN OCOS. THE LENS WAS EXCHANGED FOR A DIFFERENT POWER, SAME SIZE ON (B)(6) 2014 AND THIS RESOLVED THE PROBLEM. UCVA ON (B)(6) 2014 WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH, WIDTH AND DIOPTER POWER WERE MEASURED AND THE RESULTS OF THE MEASUREMENTS WERE COMPARED AGAINST THE ORIGINAL VALUES AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - REVIEW OF THE COMPLAINT FILE INDICATES THAT THE SURGEON IMPLANTED A WRONG LENS POWER BY MISTAKE. THE SURGEON CONFIRMS THE ERROR IN THE COMPLAINT FORM OR ADDITIONAL COMMUNICATIONS TO STAAR. THE POWER OF THE IMPLANTED LENS IS DIFFERENT FROM THE POWER LENS USED FOR EXCHANGE. THIS RESOLVED THE PROBLEM. SYMPTOMS REPORTED COULD INCLUDE: BLURRED VISION, DECREASED UCVA. WHEN A WRONG LENS POWER IS IMPLANTED AN EXCHANGE WITH THE CORRECT LENS POWER RESOLVES THE PROBLEM. THE SECONDARY INTERVENTION IS USUALLY DONE BY MEANS OF SAME INCISION WHICH WOULD NOT REQUIRE INCISION ENLARGEMENT AND/OR SUTURING DUE TO THE FOLDABLE PROPERTIES OF THE COLLAMER MATERIAL (ICL). (B)(4). | -15.0/+2.5/072 | OD | Other | 02/05/2015 | 2015 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5352440&pc=MTA | 2023826-2016-00013 | 2015-12-04 | Injury | STAAR SURGICAL COMPANY | 2016-01-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Loss of Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -13.00/+1.5/X069 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. THE IENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THERE WERE DEPOSITS ON THE CRYSTALLINE LENS. Manufacturer Narrative: CORRECT SN IS (B)(4). METHOD: LENS WORK ORDER SEARCH/DEVICE HISTORY REPORT REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. DEVICE HISTORY REPORT REVIEW: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). Manufacturer Narrative: THE CUSTOMER HAS REPORTED SINCE USING THE NEW SURGICAL PROTOCOL (USING BSS AND OCUCOAT ONLY) NO NEW COMPLAINT CASES OF LOSS OF BCVA, WITH PRECIPITATE ON THE CRYSTALLINE LENS, HAVE BEEN RECEIVED SINCE THE BEGINNING OF 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -13.00/+1.50/X069. DEVICE EVALUATED BY MANUFACTURER: DEVICE NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -13.00/+1.5/069 | OS | Physician | 12/08/2015 | 2015 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119125&pc=MTA | 2023826-2015-01308 |  | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Loss of Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/X084 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THERE WERE DEPOSITS ON THE CRYSTALLINE LENS AFTER ICL EXPLANTATION. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION - INVESTIGATION SUMMARY: THE MANUFACTURER'S DEVICE ANALYSIS RESULTS - THE EVENTS REPORTED BY PROF. (B)(6) ARE LIMITED TO A SINGLE SITE BY A SINGLE SURGEON IMPLANTING IMPLANTABLE COLLAMER LENSES (LCLS) BETWEEN THE DATES OF (B)(6) 2015. TEN PATIENTS WERE REPORTED TO HAVE MILKY WHITE PRECIPITATES FORM ON THE SURFACE OF THE ICL AND ON THE ANTERIOR LENS CAPSULE OF THE PHYSIOLOGIC CRYSTALLINE LENS. NO ADDITIONAL REPORTS OF PRECIPITATES HAVE BEEN RECEIVED FROM ANY OTHER SURGEON OR SITE WITH THE SAME LOT NUMBERS AS THE REPORTED EVENTS FROM PROF. (B)(6). ADDITIONALLY, IT IS IMPORTANT TO NOTE THAT IN THE DIRECTIONS FOR USE (DFU) OF THE LCL, THE COMPLICATIONS AND ADVERSE REACTIONS SECTION IDENTIFIES THE POTENTIAL RISK FOR "SECONDARY SURGICAL INTERVENTION TO REMOVE/REPLACE THE LENS". THE PRECAUTIONS SECTION ALSO SUGGESTS "THE LENS SHOULD NOT BE EXPOSED TO ANY SOLUTIONS OTHER THAN THE NORMALLY USED INTRAOCULAR IRRIGATING SOLUTIONS (E.G. ISOTONIC SALINE, BSS, VISCOELASTIC. ETC.)." DURING THE IN-DEPTH INVESTIGATION WITH PROF. (B)(6) TO IDENTIFY ROOT CAUSE, THE INTRAOPERATIVE REGIMEN REVEALED THAT INTRACAMERAL 0.1% CEFUROXIME ANTIBIOTIC FOR PROPHYLAXIS WAS USED IN THE 10 REPORTED CASES, ALONG WITH RINGER'S SOLUTION. IN AN ARTICLE PUBLISHED BY VASAVADA ET AL. (COMPARISON BETWEEN RINGER'S LACTATE AND BALANCED SALT SOLUTION ON POSTOPERATIVE OUTCOMES AFTER PHACOEMULSIFICATION: A RANDOMIZED CLINICAL TRIAL, INDIAN J OPHTHALMOL. 2009 MAY-JUN; 57(3): 191-195), THE AUTHORS CONCLUDE THAT RINGER'S LACTATE IS HYPOTONIC AND SLIGHTLY ACIDIC (OSMOLALITY 280MMOL, PH 6.0) AS COMPARED TO BSS (OSMOLALITY 302 MMOL, PH 7.4) AND AQUEOUS (OSMOLALITY 302 MMOL, PH 7.4). THEY FURTHER CONCLUDED THAT A STATISTICALLY SIGNIFICANT LARGER NUMBER OF PATIENTS WITH FLARE AND CELLS MAY BE DUE TO THE SLIGHTLY ACIDIC FLUCTUATING PH AND LACK OF BUFFER SYSTEM IN RINGER'S LACTATE, WHICH COULD AFFECT THE BLOOD AQUEOUS BARRIER STABILITY. ON THE OTHER HAND, 855 HAS AN ACETATE CITRATE BUFFER SYSTEM AND THE PH, THOUGH SLIGHTLY ALKALINE, IS STABLE. SUBSEQUENT TO THE SERIES OF REPORTED EVENTS, ADDITIONAL PATIENTS HAVE BEEN TREATED BY PROF. (B)(6) WITH AN ICL USING ONLY COMMERCIALLY AVAILABLE BALANCED SALINE SOLUTION (BSS, ALCON) AND OCUCOAT (BAUSCH & LOMB) INTRAOPERATIVE, WITHOUT THE USE OF 0.1% INTRACAMERAL CEFUROXIME ANTIBIOTIC. THERE WERE NO OBSERVED INCIDENTS OF DEPOSITS FOLLOWING THIS PROTOCOL REGIMEN. UPON REVIEW OF THE DEVICE HISTORY RECORDS (DHR), ALL PRODUCTS REPORTED IN THE INCIDENCES, BASED UPON LOT AND SERIAL NUMBER REVIEW, INDICATED SPECIFICATIONS WERE MET DURING MANUFACTURING PROCESS. ALL PRODUCTS PASSED QUALITY INSPECTIONS AND WERE RELEASED PER ACCEPTANCE CRITERIA. FINAL COMMENTS FROM THE MANUFACTURER: BASED UPON A THOROUGH REVIEW OF ALL AVAILABLE DATA AND ON CLINICAL PERSPECTIVE AS A PRACTICING OPHTHALMIC SURGEON WITH EXPERTISE IN THE LCL, DR. (B)(6) (MEDICAL CONSULTANT) AND STAAR CONCLUDED THAT THE OBSERVED PRECIPITATES WERE MOST LIKELY THE RESULT OF INTERACTION BETWEEN THE INTRAOCULAR MEDICATIONS USED AT THE TIME OF SURGERY. (B)(4). Manufacturer Narrative: THE CUSTOMER HAS REPORTED SINCE USING THE NEW SURGICAL PROTOCOL (USING BSS AND OCUCOAT ONLY) NO NEW COMPLAINT CASES OF LOSS OF BCVA, WITH PRECIPITATE ON THE CRYSTALLINE LENS, HAVE BEEN RECEIVED SINCE THE (B)(6) 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -15.5/+1.5/084 | OS | Physician | 09/03/2015 | 2015 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119128&pc=MTA | 2023826-2015-01305 |  | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Loss of Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+1.5/179 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THERE WERE POST-OP DEPOSITS ON CRYSTALLINE LENS (CENTRAL) WHICH DECREASED VISUAL ACUITY AFTER ICL IMPLANTATION. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: METHOD (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION - INVESTIGATION SUMMARY: THE MANUFACTURER'S DEVICE ANALYSIS RESULTS - THE EVENTS REPORTED BY PROF. (B)(6) ARE LIMITED TO A SINGLE SITE BY A SINGLE SURGEON IMPLANTING IMPLANTABLE COLLAMER LENSES (LCLS) BETWEEN THE DATES OF (B)(6) 2015. TEN PATIENTS WERE REPORTED TO HAVE MILKY WHITE PRECIPITATES FORM ON THE SURFACE OF THE ICL AND ON THE ANTERIOR LENS CAPSULE OF THE PHYSIOLOGIC CRYSTALLINE LENS. NO ADDITIONAL REPORTS OF PRECIPITATES HAVE BEEN RECEIVED FROM ANY OTHER SURGEON OR SITE WITH THE SAME LOT NUMBERS AS THE REPORTED EVENTS FROM PROF. (B)(6). ADDITIONALLY, IT IS IMPORTANT TO NOTE THAT IN THE DIRECTIONS FOR USE (DFU) OF THE LCL, THE COMPLICATIONS AND ADVERSE REACTIONS SECTION IDENTIFIES THE POTENTIAL RISK FOR "SECONDARY SURGICAL INTERVENTION TO REMOVE/REPLACE THE LENS". THE PRECAUTIONS SECTION ALSO SUGGESTS "THE LENS SHOULD NOT BE EXPOSED TO ANY SOLUTIONS OTHER THAN THE NORMALLY USED INTRAOCULAR IRRIGATING SOLUTIONS (E.G. ISOTONIC SALINE, BSS, VISCOELASTIC. ETC.)." DURING THE IN-DEPTH INVESTIGATION WITH PROF. (B)(6) TO IDENTIFY ROOT CAUSE, THE INTRAOPERATIVE REGIMEN REVEALED THAT INTRACAMERAL 0.1% CEFUROXIME ANTIBIOTIC FOR PROPHYLAXIS WAS USED IN THE (B)(4) REPORTED CASES, ALONG WITH RINGER'S SOLUTION. LN AN ARTICLE PUBLISHED BY VASAVADA ET AL. (COMPARISON BETWEEN RINGER'S LACTATE AND BALANCED SALT SOLUTION ON POSTOPERATIVE OUTCOMES AFTER PHACOEMULSIFICATION: A RANDOMIZED CLINICAL TRIAL, LNDIAN J OPHTHALMOL. 2009 MAY-JUN; 57(3): 191-195), THE AUTHORS CONCLUDE THAT RINGER'S LACTATE IS HYPOTONIC AND SLIGHTLY ACIDIC (OSMOLALITY 280MMOL, PH 6.0) AS COMPARED TO BSS (OSMOLALITY 302 MMOL, PH 7.4) AND AQUEOUS (OSMOLALITY 302 MMOL, PH 7.4). THEY FURTHER CONCLUDED THAT A STATISTICALLY SIGNIFICANT LARGER NUMBER OF PATIENTS WITH FLARE AND CELLS MAY BE DUE TO THE SLIGHTLY ACIDIC FLUCTUATING PH AND LACK OF BUFFER SYSTEM IN RINGER'S LACTATE, WHICH COULD AFFECT THE BLOOD AQUEOUS BARRIER STABILITY. ON THE OTHER HAND, (B)(4) HAS AN ACETATE CITRATE BUFFER SYSTEM AND THE PH, THOUGH SLIGHTLY ALKALINE, IS STABLE. SUBSEQUENT TO THE SERIES OF REPORTED EVENTS, ADDITIONAL PATIENTS HAVE BEEN TREATED BY PROF. (B)(6) WITH AN ICL USING ONLY COMMERCIALLY AVAILABLE BALANCED SALINE SOLUTION (BSS, ALCON) AND OCUCOAT (BAUSCH & LOMB) INTRAOPERATIVE, WITHOUT THE USE OF 0.1% INTRACAMERAL CEFUROXIME ANTIBIOTIC. THERE WERE NO OBSERVED INCIDENTS OF DEPOSITS FOLLOWING THIS PROTOCOL REGIMEN. UPON REVIEW OF THE DEVICE HISTORY RECORDS (DHR), ALL PRODUCTS REPORTED IN THE INCIDENCES, BASED UPON LOT AND SERIAL NUMBER REVIEW, INDICATED SPECIFICATIONS WERE MET DURING MANUFACTURING PROCESS. ALL PRODUCTS PASSED QUALITY INSPECTIONS AND WERE RELEASED PER ACCEPTANCE CRITERIA. FINAL COMMENTS FROM THE MANUFACTURER: BASED UPON A THOROUGH REVIEW OF ALL AVAILABLE DATA AND ON CLINICAL PERSPECTIVE AS A PRACTICING OPHTHALMIC SURGEON WITH EXPERTISE IN THE LCL, DR. (B)(6) (MEDICAL CONSULTANT) AND STAAR CONCLUDED THAT THE OBSERVED PRECIPITATES WERE MOST LIKELY THE RESULT OF INTERACTION BETWEEN THE INTRAOCULAR MEDICATIONS USED AT THE TIME OF SURGERY. CLAIM # (B)(4). Manufacturer Narrative: THE CUSTOMER HAS REPORTED SINCE USING THE NEW SURGICAL PROTOCOL (USING BSS AND OCUCOAT ONLY) NO NEW COMPLAINT CASES OF LOSS OF BCVA, WITH PRECIPITATE ON THE CRYSTALLINE LENS, HAVE BEEN RECEIVED SINCE THE BEGINNING OF 2016. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -14.5/+1.5/179 | OD | Physician | 09/03/2015 | 2015 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5124750&pc=MTA | 2023826-2015-01306 |  | Injury | STAAR SURGICAL COMPANY | 2015-10-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Loss of Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+4.0/074 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED DUE TO LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THERE WERE DEPOSITS ON CRYSTALLINE LENS (VISCO) WHICH DECREASED VISUAL ACUITY AFTER ICL IMPLANTATION. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION - INVESTIGATION SUMMARY: THE MANUFACTURER'S DEVICE ANALYSIS RESULTS - 1 THE EVENTS REPORTED BY PROF. (B)(6) ARE LIMITED TO A SINGLE SITE BY A SINGLE SURGEON IMPLANTING IMPLANTABLE COLLAMER LENSES (LCLS) BETWEEN THE DATES OF (B)(6) 2015. TEN PATIENTS WERE REPORTED TO HAVE MILKY WHITE PRECIPITATES FORM ON THE SURFACE OF THE ICL AND ON THE ANTERIOR LENS CAPSULE OF THE PHYSIOLOGIC CRYSTALLINE LENS. NO ADDITIONAL REPORTS OF PRECIPITATES HAVE BEEN RECEIVED FROM ANY OTHER SURGEON OR SITE WITH THE SAME LOT NUMBERS AS THE REPORTED EVENTS FROM PROF. (B)(6). ADDITIONALLY, IT IS IMPORTANT TO NOTE THAT IN THE DIRECTIONS FOR USE (DFU) OF THE LCL, THE COMPLICATIONS AND ADVERSE REACTIONS SECTION IDENTIFIES THE POTENTIAL RISK FOR "SECONDARY SURGICAL INTERVENTION TO REMOVE/REPLACE THE LENS." THE PRECAUTIONS SECTION ALSO SUGGESTS "THE LENS SHOULD NOT BE EXPOSED TO ANY SOLUTIONS OTHER THAN THE NORMALLY USED INTRAOCULAR IRRIGATING SOLUTIONS (E.G. ISOTONIC SALINE, BSS, VISCOELASTIC. ETC.)." DURING THE IN-DEPTH INVESTIGATION WITH PROF. (B)(6) TO IDENTIFY ROOT CAUSE, THE INTRAOPERATIVE REGIMEN REVEALED THAT INTRACAMERAL 0.1% CEFUROXIME ANTIBIOTIC FOR PROPHYLAXIS WAS USED IN THE 10 REPORTED CASES, ALONG WITH RINGER'S SOLUTION. LN AN ARTICLE PUBLISHED BY VASAVADA ET AL. (COMPARISON BETWEEN RINGER'S LACTATE AND BALANCED SALT SOLUTION ON POSTOPERATIVE OUTCOMES AFTER PHACOEMULSIFICATION: A RANDOMIZED CLINICAL TRIAL, LNDIAN J OPHTHALMOL. 2009 MAY-JUN; 57(3): 191-195), THE AUTHORS CONCLUDE THAT RINGER'S LACTATE IS HYPOTONIC AND SLIGHTLY ACIDIC (OSMOLALITY 280MMOL, PH 6.0) AS COMPARED TO BSS (OSMOLALITY 302 MMOL, PH 7.4) AND AQUEOUS (OSMOLALITY 302 MMOL, PH 7.4). THEY FURTHER CONCLUDED THAT A STATISTICALLY SIGNIFICANT LARGER NUMBER OF PATIENTS WITH FLARE AND CELLS MAY BE DUE TO THE SLIGHTLY ACIDIC FLUCTUATING PH AND LACK OF BUFFER SYSTEM IN RINGER'S LACTATE, WHICH COULD AFFECT THE BLOOD AQUEOUS BARRIER STABILITY. ON THE OTHER HAND, 855 HAS AN ACETATE CITRATE BUFFER SYSTEM AND THE PH, THOUGH SLIGHTLY ALKALINE, IS STABLE. SUBSEQUENT TO THE SERIES OF REPORTED EVENTS, ADDITIONAL PATIENTS HAVE BEEN TREATED BY PROF. (B)(6) WITH AN ICL USING ONLY COMMERCIALLY AVAILABLE BALANCED SALINE SOLUTION (BSS, ALCON) AND OCUCOAT (BAUSCH & LOMB) INTRAOPERATIVE, WITHOUT THE USE OF 0.1% INTRACAMERAL CEFUROXIME ANTIBIOTIC. THERE WERE NO OBSERVED INCIDENTS OF DEPOSITS FOLLOWING THIS PROTOCOL REGIMEN. UPON REVIEW OF THE DEVICE HISTORY RECORDS (DHR), ALL PRODUCTS REPORTED IN THE INCIDENCES, BASED UPON LOT AND SERIAL NUMBER REVIEW, INDICATED SPECIFICATIONS WERE MET DURING MANUFACTURING PROCESS. ALL PRODUCTS PASSED QUALITY INSPECTIONS AND WERE RELEASED PER ACCEPTANCE CRITERIA. FINAL COMMENTS FROM THE MANUFACTURER: BASED UPON A THOROUGH REVIEW OF ALL AVAILABLE DATA AND ON CLINICAL PERSPECTIVE AS A PRACTICING OPHTHALMIC SURGEON WITH EXPERTISE IN THE LCL, DR. (B)(6) (MEDICAL CONSULTANT) AND STAAR CONCLUDED THAT THE OBSERVED PRECIPITATES WERE MOST LIKELY THE RESULT OF INTERACTION BETWEEN THE INTRAOCULAR MEDICATIONS USED AT THE TIME OF SURGERY. (B)(4). Manufacturer Narrative: THE CUSTOMER HAS REPORTED SINCE USING THE NEW SURGICAL PROTOCOL (USING BSS AND OCUCOAT ONLY) NO NEW COMPLAINT CASES OF LOSS OF BCVA, WITH PRECIPITATE ON THE CRYSTALLINE LENS, HAVE BEEN RECEIVED SINCE THE BEGINNING OF 2016. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -11.5/+4.0/074 | OS | Physician | 09/03/2015 | 2015 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5408629&pc=MTA | 2023826-2016-00114 |  | Injury | STAAR SURGICAL COMPANY AG | 2016-02-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Loss of Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+3.0/036 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND LOSS OF BEST CORRECTED VISUAL ACUITY. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | -9.0/+3.0/036 | OD | Physician | 10/22/2015 | 2015
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4704671&pc=MTA | 2023826-2015-00426 | 2014-12-15 | Injury | STAAR SURGICAL COMPANY | 2015-04-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | Loss of Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND LOSS OF BEST-CORRECTED VISUAL ACUITY. ALSO REFRACTIVE CHANGE OVERTIME AND AN AXIS GAP OF ABOUT 90 DEGREES. THE LENS WAS RE-POSITIONED. THE LENS WAS EXCHANGED FOR A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) Manufacturer Narrative: METHOD: - MEDICAL REVIEW. RESULTS.: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH, WIDTH AND DIOPTER POWER WERE MEASURED AND THE RESULTS OF THE MEASUREMENTS WERE COMPARED AGAINST THE ORIGINAL VALUES AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CLINICAL STUDIES HAVE SHOWN THAT TORIC ICL ROTATION OCCURS IN 0.5% OF PATIENTS WHERE A TICL HAS BEEN IMPLANTED. SEVERAL FACTORS MAY CONTRIBUTE TO THIS PHENOMENON (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -10.0/2.0/083 | OS | Other | 03/19/2015 | 2015 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7381601&pc=MTA | 2023826-2018-00491 | 2018-02-26 | Injury | STAAR SURGICAL COMPANY | 2018-03-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Macular Edema; Inflammation; Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -16.5/+0.5/169 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) IN (B)(6) OF 2017. THE SURGEON REPORTS MACULAR EDEMA WITHOUT OTHER SIGNS. IT WAS ALSO REPORTED THAT THIS WAS A MYOPIC CASE. THE SURGEON IS CONSIDERING REMOVAL OF ICL BUT IT HAS NOT BEEN CONFIRMED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS IMPLANTED ON (B)(6) 2017. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, LOSS OF BCVA, INFLAMMATION, AND MACULAR EDEMA. ON (B)(6) 2018 THE REPORTER STATED "DISAPPEARANCE OF MACULAR EDEMA OR POSTERIOR SYNECHIA WITH PUPILLARY DEFORMITY RIGHT EYE." SECONDARY SURGICAL INTERVENTION, LENS EXPLANT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | -16.5/+0.5/169 | OD | Physician | 02/28/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7890073&pc=MTA | 2023826-2018-01388 | 2018-07-18 | Injury | STAAR SURGICAL COMPANY | 2018-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Macular Edema; Inflammation; Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOSS OF BCVA (POST-OP BCVA 20/50) AND CYSTOID MACULAR EDEMA. ANTIBIOTICS AND ANTI-INFLAMMATORIES WERE USED TO ADDRESS INFLAMMATION. NO INFECTION OCCURED. IT WAS REPORTED THAT THE EXPLANT RESOLVED THE PROBLEM. PATIENT OK AFTER EDEMA ADSORPTION. NO PLANS TO IMPLANT AN ICL LENS. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT - IS ADDED TO INITIAL MDR. (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -15/+0.5/038 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOSS OF BCVA (POST-OP BCVA 20/50) AND CYSTOID MACULAR EDEMA. ANTIBIOTICS AND ANTI-INFLAMMATORIES WERE USED TO ADDRESS INFLAMMATION. NO INFECTION OCCURRED. IT WAS REPORTED THAT THE EXPLANT RESOLVED THE PROBLEM. PATIENT OK AFTER EDEMA ADSORPTION. NO PLANS TO IMPLANT AN ICL LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15/+0.5/038 | OS | Physician | 08/28/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6867141&pc=MTA | 2023826-2017-01464 | 2017-08-11 | Injury | STAAR SURGICAL COMPANY | 2017-09-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Nausea; Blurred Vision; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -6.0/+5.5/158 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS, VTICMO13.7, -6.0/+5.5/166 DIOPTER, AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE SURGEON FELT THE LENS WAS MISLABELED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. INVESTIGATION CONFIRMED THE LABELLING ISSUE. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -6.0/+5.5/158 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER STATED THAT THE PATIENT COULD SEE "NOTHING WITH OD AND CANNOT MOVE DUE TO NAUSEA WITH HIS ASTIGMATISM." THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS, VTICMO13.7, -6.0/+5.5/166 DIOPTER, AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE SURGEON FELT THE LENS WAS MISLABELED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.0/+5.5/158 | OD | Physician | 08/16/2017 | 2017 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14492494&pc=MTA | 2023826-2022-01542 |  | Injury | STAAR SURGICAL COMPANY | 2022-05-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED AN ICL REMOVAL AND REPLACEMENT ON (B)(6) 2022 FOR THE PATIENT'S RIGHT EYE (OD). REPORTEDLY POST-OP HAS SEEN IMPROVEMENT IN THE RIGHT EYE AND FLUCTUATION OF VISION. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 -13.5/2.5/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). IT WAS EXPLANTED ON (B)(6) 2022, DUE TO AN UNDESIRABLE REFRACTIVE OUTCOME. IT WAS LATER REPLACED FOR A SAME LENGTH LENS OF DIFFERENT DIOPTER. REPLACEMENT HAS IMPROVED THE PATIENT'S VISION. H6- INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -13.5/2.5/102 | OD | Physician | 05/26/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17962120&pc=MTA | 2023826-2023-04569 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | UNK | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). EXCESSIVE VAULT WAS OBSERVED. Manufacturer Narrative: A4-A6:UNK. B3: UNK. D4 (EXPERIATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4). | -5.0/4.5/046 | OS | Physician | 10/10/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11065590&pc=MTA | 2023826-2020-03100 | 2020-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.00/3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REMOVED FROM THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020 DUE TO EXCESSIVE VAULT. A SHORTER LENGTH LENS WAS IMPLANTED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.00/3.0/093 | OD |  | 12/03/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11598748&pc=MTA | 2023826-2021-00795 | 2020-11-11 | Injury | STAAR SURGICAL COMPANY | 2021-03-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -9.00/1.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH A SHORTER LENGTH LENS INTRAOPERATIVELY ON (B)(6) 2020 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/1.0/109 | OD | Physician | 03/19/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11778998&pc=MTA | 2023826-2021-01332 | 2021-04-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). LOW VAULT WAS OBSERVED. THEN LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - PER NEW INFORMATION "PATIENT CURRENTLY DOSEN'T WISH A FURTHER SURGERY." CLAIM # (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 -6.50/+2.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT WAS OBSERVED. REPORTEDLY THE PATIENT DOES NOT WANT FURTHER SURGERY. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -6.50/+2.0/093 | OS |  | 04/16/2021 | 2021 | 1 MO
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11779035&pc=MTA | 2023826-2021-01331 | 2021-04-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). LOW VAULT WAS OBSERVED. THEN LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 -8.00/+3.0/135 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LOW VAULT WAS OBSERVED. REPORTEDLY THE PATIENT DOES NOT WANT FURTHER SURGERY. THE LENS REMAINS IMPLANTED. CLAIM # (B)(4). Manufacturer Narrative: B5 - PER NEW INFORMATION "PATIENT CURRENTLY DOESN'T WISH A FURTHER SURGERY." CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.00/+3.0/135 | OD |  | 04/16/2021 | 2021 | 1 MO
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15042300&pc=MTA | 2023826-2022-02169 | 2022-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6, -11.00/3.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). | -11.00/3.5/087 | OS | Physician | 06/30/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15674295&pc=MTA | 2023826-2022-03626 | 2022-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO A PATIENT'S EYE. LOW VAULT AND LENS ROTATION WAS OBSERVED. LENS REMOVAL AND REPLACEMENT IS PLANNED. Manufacturer Narrative: A2: AGE AT TIME OF EVENT, DATE OF BIRTH ON (B)(6) 1998 SHOULD BE ADDED TO THE INITIAL MDR. B3: DATE OF EVENT 30-JUN-2022 SHOULD BE ADDED TO THE INITIAL MDR. B5: "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.50/3.0/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS LATER EXCHANGED FOR A LONGER LENGTH/ SAME MODEL LENS ON (B)(6) 2022 DUE TO LOW VAULT. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE." SHOULD BE ADDED TO THE INITIAL MDR. D4: MODEL#: VTICM5 12.1 SHOULD BE ADDED TO THE INITIAL MDR. D4: SERIAL#: (B)(6) SHOULD BE ADDED TO THE INITIAL MDR. D4: EXPIRATION DATE 31-MAR-2025 SHOULD BE ADDED TO THE INITIAL MDR. D4: UNIQUE IDENTIFIER (UDI) NUMBER: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. D6A: IMPLANT DATE: ON (B)(6) 2022 SHOULD BE ADDED TO THE INITIAL MDR. D6B: EXPLANT DATE: ON (B)(6) 2022 SHOULD BE ADDED TO THE INITIAL MDR. H1: CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H4: DEVICE MANUFACTURE DATE 02-MAY-2022 SHOULD BE ADDED TO THE INITIAL MDR. H6: CORRECTED TO HEALTH EFFECT - IMPACT CODE: 4629 SHOULD BE ADDED TO THE INITIAL MDR. H6:- TYPE OF INVESTIGATION: 4110 WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -10.50/3.0/085 | OS | Physician | 09/27/2022 | 2022 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035427&pc=MTA | 2023826-2022-04553 | 2022-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -15.0/1.0/126 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -15.0/1.0/126 | OD | Physician | 11/28/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17122144&pc=MTA | 2023826-2023-02348 | 2023-02-02 05:00 | Injury | STAAR SURGICAL CO. | 2023-06-13 04:00 | MTA | Unknown Brand Name | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/0.5/006 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH/DIFFERENT MODEL LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "ASTIGMATIC CRYSTALS WERE IMPLANTED FOR THE FIRST TIME. THE PATIENT'S SUPPLY WAS TOO HIGH, AND THE DOCTOR FINALLY ADJUSTED 12.6 IMPLANTATION OF NON-ASTIGMATIC CRYSTALS BECAUSE SUITABLE ASTIGMATIC CRYSTALS COULD NOT BE MATCHED." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -6.5/0.5/006 | OD | Physician | 06/09/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549371&pc=MTA | 2023826-2023-03458 | 2023-07-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE(OD). HIGH VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2, -12.5/1.0/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2 - UNK. A3 - UNK. A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). D4 - UNK. D6A - UNK. H4 - UNK. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: CLAIM# (B)(4). | -12.5/1.0/072 | OD | Physician | 07/28/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17368426&pc=MTA | 2023826-2023-02973 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF A -10.50/1.0/068 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 DUE TO LOW VAULT. THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "PATIENT HAS STILL A LOW VAULT," AND "MAY BE THE DOCTOR WILL EXCHANGE AGAINST A 13.7MM LENS.". Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | -10.50/1.0/068 | OS | Physician | 07/04/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17633331&pc=MTA | 2023826-2023-03543 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). HIGH VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). CLAIM# (B)(4). | -11.0/4.0/085 | OS | Physician | 08/11/2023 | 2023 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16155877&pc=MTA | 2023826-2023-00146 | 2022-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.5/1.0/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022 AND THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL/SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE WAS "ALL FINE". ADDITIONAL INFORMATION PROVIDED WAS "PATIENT IS HAPPY". CLAIM# (B)(4). | -12.5/1.0/107 | OD | Physician | 12/14/2022 | 2022 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17322211&pc=MTA | 2023826-2023-02773 | 2023-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | UNK | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REPORTEDLY, "PATIENT WITH A MARKEDLY ELEVATED VAULT IN ONE EYE." Manufacturer Narrative: ADDITIONAL DATA: H6-HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND CODE "4629" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "2993" SHOULD BE ADDED. B5-THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.00/3.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE( OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL, SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "DOING WELL. EXCELLENT VISION." THE REPORTER INDICATED THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. B3: UNK. D4 (EXPERIATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4). | -16.00/3.0/092 | OD | Physician | 06/27/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17777896&pc=MTA | 2023826-2023-04087 | 2022-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | UNK | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). EXCESSIVE VAULT WAS OBSERVED. Manufacturer Narrative: B5: LENS WAS EXCHANGED. CLAIM# (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED EXCESSIVE VAULT. REPORTEDLY " EXCHANGING LENS BECAUSE OF OVERVAULTING". ON (B)(6) 2023 THE LENS WAS EXPLANTED AND REPLACED. H6-HEALTH EFFECT- IMPACT CODE: "4629" SHOULD BE REMOVED AND CODE "4627" SHOULD BE ADDED. H6- MEDICAL DEVICE CODE: "3189" SHOULD BE REMOVED AND "2993" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: ADDITONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.50/1.0/088(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT. ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THE SAME DAY SEPARATE SURGERY THE LENS WAS REPLACED WITH THE SAME MODEL, SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. H6- TYPE OF INVESTIGATION CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. B3: UNK. D4 (EXPIRIATION DATE); UNK NO SERIAL NUMBER REPORTED.. D6A: UNK. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4). | -8.50/1.0/088 | OS | Physician | 09/05/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17777897&pc=MTA | 2023826-2023-04086 | 2022-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | UNK | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). EXCESSIVE VAULT WAS OBSERVED. Manufacturer Narrative: H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4). Manufacturer Narrative: LENS WAS EXCHANGED. CLAIM# (B)(4). Manufacturer Narrative: ADDITONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.50/2.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT. ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THE SAME DAY SEPARATE SURGERY THE LENS WAS REPLACED WITH THE SAME MODEL, SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. H6- TYPE OF INVESTIGATION CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED EXCESSIVE VAULT. REPORTEDLY " EXCHANGING LENS BECAUSE OF OVERVAULTING". ON (B)(6) 2023 THE LENS WAS EXPLANTED AND REPLACED. H6-HEALTH EFFECT- IMPACT CODE: "4629" SHOULD BE REMOVED AND CODE "4627" SHOULD BE ADDED. H6- MEDICAL DEVICE CODE: "3189" SHOULD BE REMOVED AND "2993" SHOULD BE ADDED. CLAIM# (B)(4). | -9.50/2.0/089 | OD | Physician | 09/05/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15299315&pc=MTA | 2023826-2022-02905 | 2022-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -18.0/4.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE IS REPORTED AS UNKNOWN. REPORTEDLY. "WITH A BIGGER LENS EXCHANGE, PATIENT IS DOING WELL AND HAPPY." Manufacturer Narrative: CORRECTED DATA: D4: EXPIRATION DATE: ADD (B)(6) 2024. H6: MEDICAL DEVICE PROBLEM CODE: ADD CODE 2923. CLAIM# (B)(4). Manufacturer Narrative: H3: PRODUCT EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND THE OPTIC DEFORMED AND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -18.0/4.5/092 | OD | Physician | 07/29/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11119736&pc=MTA | 2023826-2021-00010 | 2020-05-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS LEFT EYE (OS). THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/1.5/078 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS WAS EXCHANGED ON (B)(6) 2021 FOR A LOWER POWER/LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: H6: INVESTIGATION TYPE 4110 - A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION : LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/1.5/078 | OS | Physician | 12/09/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14538892&pc=MTA | 2023826-2022-01664 | 2021-05-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS RIGHT EYE (OD) . THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- TWO SIMILAR COMPLAINTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2 VTICM5_13.2 -2.0/2.0/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LENS WAS REPLACED WITH A SHORTER LENGTH LENS ON (B)(6) 2022 DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. DEVICE MANUFACTURE DATE: UNK. CLAIM #: (B)(4). | -2.0/2.0/088 | OD | Physician | 05/03/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14598790&pc=MTA | 2023826-2022-01736 | 2022-01-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -4.5/5.0/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS AN EXCHANGE LENS IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. A LITTLE ASTIGMATISM REMAINS. REPORTER STATES LENS SIZE IS NOT THE PROBLEM. LENS REMAINS IMPLANTED. SEE MFR# 2023826-2022-00139 FOR INITIAL LENS. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. INVESTIGATION TYPE (B)(4): LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -4.5/5.0/086 | OD | Physician | 05/08/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14961423&pc=MTA | 2023826-2022-01730 | 2022-05-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -13.0/2.5/75 (SPHERE/CYLINDER/AXIS) LENS TEAR/BREAK DURING LOADING. THE LENS WAS NOT IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: A SHORTER LENS WAS IMPLANTED IN A SUBSEQUENT SURGERY AND PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -13.0/2.5/75 | OS | Physician | 05/03/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16075120&pc=MTA | 2023826-2022-04774 | 2022-09-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6; -10.5/2.00/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS HAD TORE DURING LOADING AND NOTED "LENS RUPTURE DUE TO SNAGGING WHILE LOADING INJECTOR". THIS OCCURRED ON (B)(6) 2022. THE PROBLEM WAS RESOLVED. IT WAS NOTED "NEXT OPERATION DATE TO BE DETERMINED". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: "H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE." CLAIM# (B)(4). | -10.5/2.00/95 | NA | Physician | 12/22/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16065225&pc=MTA | 2023826-2022-04760 | 2022-12-01 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -9.0/3.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THERE WAS NO PATIENT INJURY. Manufacturer Narrative: (B)(4). | -9.0/3.0/090 | OD | Physician | 12/01/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16508351&pc=MTA | 2023826-2023-00787 | 2023-02-06 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF A -17.5/2.0/95 (SPHERE/CYLINDER/AXIS) TORE/BROKE DURING LOADING. THE REPLACEMENT LENS WAS IMPLANTED. CAUSE OF THE EVENT WAS USER ERROR. REPORTEDLY, "EYE QUIET, "DOCTOR AND PATIENT HAPPY." Manufacturer Narrative: CLAIM# (B)(4). | -17.5/2.0/95 | NA | Physician | 02/09/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17007084&pc=MTA | 2023826-2023-02041 | 2023-04-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A IMPLANTABLE COLLAMER LENS HAD BECAME DAMAGED ON (B)(6) 2023. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -16.0/-2.50/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AND REMOVED INTRAOPERATIVELY ON (B)(6) 2023 INTO THE PATIENT'S LEFT EYE (OS), DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. PATIENT CONTACT BUT NO PATIENT INJURY. LENS WAS REPLACED ON (B)(6) 2023 WITH SAME MODEL LENS AND PROBLEM WAS RESOLVED. EYE QUIET, SURGEON AND PATIENT HAPPY. CAUSE DETAILS WERE "PLUNDGER GOT STUCK IN THE CARTRIDGE." CLAIM# (B)(4). Manufacturer Narrative: H6 - INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: CLAIM (B)(4). | -16.0/-2.50/90 | OS | Physician | 05/01/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17099048&pc=MTA | 2023826-2023-02238 | 2023-05-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.50/2.5/109 TORE DURING LOADING. A REPLACEMENT LENS WAS IMPLANTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THE LENS WAS TORE DURING LOADING AND INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT, BUT NO PATIENT INJURY. CAUSE OF THE EVENT WAS OTHER. REPORTEDLY, "LENS TORE WHILE IMPLANTING. REMOVED & RELOADED NEW LENS WITHOUT INCIDENT." THE PROBLEM WAS RESOLVED. STATUS OF THE EYE IS SAID, "CORTICAL SPOKING ON LENS. DOING WELL OTHERISE." CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS. CLAIM# (B)(4). | -10.50/2.5/109 | NA | Physician | 05/25/2023 | 2023 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17401155&pc=MTA | 2023826-2023-03123 | 2023-06-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF A -7.0/1.5/087 (SPHERE/CYLINDER/AXIS) TORE/BROKE DURING LOADING WHILE ADVANCING ON (B)(6) 2023. THE LENS WAS NOT IMPLANTED. ON (B)(6) 2023, A REPLACEMENT LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). CAUSE OF THE EVENT WAS THE DEVICE. REPORTEDLY, "CARTRIDGE/INJECTOR DEFECTIVE," AND "SEE (B)(4) - BOTH PRIMARY AND B/U LENSES FOR OS WERE DESTROYED BY DEFECTIVE INJECTOR/CARTRIDGE. A 3RD LENS WAS ORDERED & IMPLANTED (B)(6) 2023." Manufacturer Narrative: D10: LIOLI LOT# FCK1702. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS FOUND WITHIN ASSOCIATED LOTS. CLAIM# (B)(4). CLAIM# (B)(4). | -7.0/1.5/087 | OS | Physician | 07/06/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17401156&pc=MTA | 2023826-2023-03124 | 2023-06-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF A -7.0/1.5/087 (SPHERE/CYLINDER/AXIS) TORE/BROKE DURING LOADING WHILE ADVANCING ON (B)(6) 2023. THE LENS WAS INTENDED TO BE A REPLACEMENT LENS AND WAS NOT IMPLANTED. ON (B)(6) 2023, A REPLACEMENT LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). CAUSE OF THE EVENT WAS THE DEVICE. REPORTEDLY, "CARTRIDGE/INJECTOR DEFECTIVE." Manufacturer Narrative: D10: LIOLI LOT# FCK1702. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS FOUND WITHIN ASSOCIATED LOTS. CLAIM# (B)(4). CLAIM# (B)(4). | -7.0/1.5/087 | OS | Physician | 07/06/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410535&pc=MTA | 2023826-2023-03007 | 2023-07-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/1.5/069 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK, CLAIM #: (B)(4). | -12.0/1.5/069 | OD | Physician | 07/17/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17928151&pc=MTA | 2023826-2023-04365 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-10-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -5.50/1.0/108 (SPHERE/CYLINDER/AXIS) HAPTIC RIP DURING LOADING, NO PATIENT HARM REPORTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -5.50/1.0/108 | NA | Other Health Care Professional | 09/12/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147473&pc=MTA | 2023826-2021-02253 | 2021-02-06 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -4.5/1.0/147 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -4.5/1.0/147 | OS | Physician | 06/14/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12169908&pc=MTA | 2023826-2021-02577 | 2021-02-18 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.0/3.0/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -11.0/3.0/92 | OD | Physician | 06/14/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12790687&pc=MTA | 2023826-2021-04191 | 2021-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/1.50/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.50/1.50/088 | OD | Physician | 10/22/2021 | 2021 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15459739&pc=MTA | 2023826-2022-03236 | 2022-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS THE PATIENT HAS EXCESSIVE VAULT (1200 MICRON). LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 -5.5/5.5/173 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED, ON (B)(6) 2022 THE LENS WAS REPLACED FOR A SHORTER LENGTH LENS. PROBLEM RESOLVED. CAUSE REPORTED AS DEVICE. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: 1494 OFF-LABEL SHOULD BE REMOVED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE: 1494-PRE-EXISTING KERATOCONUS. CLAIM#: (B)(4). | -5.5/5.5/173 | OD | Physician | 10/07/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17084047&pc=MTA | 2023826-2023-02279 | 2023-03-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.50/+1.5/104 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -11.50/+1.5/104 | OD | Physician | 05/10/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18097779&pc=MTA | 2023826-2023-05036 | 2023-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/2.5/096 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL/LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "THE SECOND TICL IS IN A VERTICAL POSITION." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -15.0/2.5/096 | OD | Physician | 11/07/2023 | 2023 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12965129&pc=MTA | 2023826-2021-04692 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-12-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS IMPLANTED INTO A PATIENT'S EYE THAT IS PLANNED FOR AN EXCHANGE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/3.0/092 INTO THE PATIENT'S LEFT EYE (OS) ON AN UNKNOWN DATE. IT WAS REPORTED TO HAVE BEEN ASSOCIATED WITH AN UNSPECIFIED LOW VAULT. TO THE BEST OF OUR KNOWLEDGE THE LENS REMAINS IMPLANTED." SHOULD BE ADDED TO THE INITIAL MDR. H6 - DEVICE CODE: 1494, OFF-LABEL USE, ACD < 3.0MM. CLAIM #: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/3.0/092 | OS | Physician | 11/15/2021 | 2021 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15741078&pc=MTA | 2023826-2022-03870 | 2022-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -15.0/2.0/039 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS IMPLANTED UPSIDE DOWN. ON (B)(6)2022 THE LENS WAS EXCHANGED FOR A SAME MODEL; AND SIZE BUT DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. REPORTER STATES, "THE FIRST PIECE IMPLANTED TICL AT 30 °, TURNED COUNTERCLOCKWISE 11 ° AND ALIGNED 41 °. THE SECOND PIECE IMPLANTED TICL AT 120 °, TURNED COUNTERCLOCKWISE 15 ° AND ALIGNED 135 °. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM #(B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D4- "VTICM5_12.1" SHOULD BE CORRECTED TO "VTICMO 12.1". CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D4: "VTICM5_12.1" SHOULD BE CORRECTED TO "VTICMO 12.6". CLAIM#: (B)(4). | -15.0/2.0/039 | OD | Physician | 10/07/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5102470&pc=MTA | 2023826-2015-01181 | 2015-07-17 | Injury | STAAR SURGICAL COMPANY | 2015-09-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50/+2.0/X113 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. REFRACTIVE SURPRISE WAS NOTED ON (B)(6) 2015. THE LENS REMAINS IMPLANTED. ERROR MADE ON CALCULATION. A SECONDARY SURGERY IS PENDING. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -13.50/+2.0/X113. DEVICE EVALUATED BY MANUFACTURER?: NO, THE LENS REMAINS IMPLANTED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. (B)(4). LENS REMAINS IMPLANTED. | -13.50/+2.0/113 | OD | Physician | 08/27/2015 | 2015 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5085969&pc=MTA | 2023826-2015-01184 | 2015-08-18 | Injury | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -16.0/+4.0/X112 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THERE WAS REFRACTIVE SURPRISE DUE TO DATA ERROR ENTERED ON AXIS AND SPHERICAL POWER. A SUPPLEMENTAL WIL BE SUBMITTED WHEN ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -16.0/+4.0/X112. DEVICE EVALUATED BY MANUFACTURER? NO: LENS REMAINS IMPLANTED. (B)(4) REFRACTIVE SURPRISE. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). | -16.0/+4.0/112 | OS | Physician | 08/19/2015 | 2015 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6152000&pc=MTA | 2023826-2016-01703 | 2016-09-06 | Injury | STAAR SURGICAL COMPANY | 2016-12-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM512.6 IMPLANTABLE COLLAMER LENS, -11.0/+0.5/97 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016 AND IT WAS NOTED THE PATIENT HAD A MACULAR HOLE, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -11.0/+0.5/97 | OD | Physician | 11/11/2016 | 2016 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6665472&pc=MTA | 2023826-2017-01027 | 2016-12-15 | Injury | STAAR SURGICAL COMPANY | 2017-06-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -14.5/+3.0/093 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS REMOVED BECAUSE THE DOCTOR HAD ORDERED THE WRONG REFRACTION FOR THE LENS. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SAME SIZE, DIFFERENT DIOPTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). NOT RETURNED TO MANUFACTURER. | -14.5/+3.0/093 | OS | Physician | 12/15/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8254844&pc=MTA | 2023826-2019-00064 | 2017-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.5/+3.0/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS DUE TO USER ERROR. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS "REFRACTION MEASUREMENT ERROR BY USER." Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. METHOD CODE ADDED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.5/+3.0/087 | OS |  | 12/17/2018 | 2018 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6653520&pc=MTA | 2023826-2017-01004 | 2017-03-31 | Injury | STAAR SURGICAL COMPANY | 2017-06-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10/+6/120 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ACCORDING TO THE CUSTOMER, BECAUSE THE PATIENT BECAME UNCOOPERATIVE DURING IMPLANTING, THE LENS DID COMING CONTACT TO THE PATIENT, BUT IT WAS NOT FULLY IMPLANTED. THEREFORE, THE SURGERY WAS ABORTED. ON (B)(6) 2017 THE LENS WAS EXCHANGED, AND THE PATIENT IS HAPPY WITH THE VISION. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: CONCLUSION CODE: PREVIOUS CODE NOT APPLICABLE; EVENT CAUSED DUE TO USER ERROR. (B)(4). Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMAGE AND CLEAR RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10/+6/120 | OS | Physician | 05/22/2017 | 2017 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8323888&pc=MTA | 2023826-2019-00206 | 2018-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+1.5/003 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS WAS EXPLANTED DUE TO THE PATIENT'S REQUEST. THE SURGEON DOES NOT PLAN TO IMPLANT ANOTHER ICL LENS. THE PATIENT COULD NOT ADAPT TO POST-OPERATIVE CONDITIONS. NO ADDITIONAL INFORMATION WAS PROVIDED. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.5/+1.5/003 | OD |  | 01/10/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7886285&pc=MTA | 2023826-2018-01408 | 2018-08-06 | Injury | STAAR SURGICAL COMPANY | 2018-09-18 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+1.5/160 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE. THE REPORTER STATES THE HAPTIC BROKE DURING THE INJECTION INTO THE EYE. THIS OCCURRED ON (B)(6) 2018. THE LENS WAS EXCHANGED ON (B)(6) 2018 WITH AN ALTERNATE LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -16.5/+1.5/160 | OD | Physician | 08/24/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8176856&pc=MTA | 2023826-2018-01859 | 2018-10-15 | Injury | STAAR SURGICAL COMPANY | 2018-12-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -3.50/+3.5/033 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE SURGEON WANTED THE LENS POWER TO MATCH THE LENS IN THE LEFT EYE (OS). THE LENS WAS EXCHANGED FOR A SAME MODEL BUT DIFFERENT POWER LENS (-4.0/+3.5/136 SPHERE/CYLINDER/AXIS) AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. "DUE TO THE SURGEON WANTED THE LENS POWER TO MATCH THE LENS IN THE LEFT EYE (OS)" SHOULD BE CORRECTED TO "DUE TO THE SURGEON IMPLANTING THE LENS MEANT FOR THE LEFT EYE (OS)". (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -3.50/+3.5/033 | OD | Physician | 11/19/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8227434&pc=MTA | 2023826-2019-00058 | 2018-11-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/052 (SPHERE/"CYLINDAR"/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, THE LENS TORE DURING INJECTION INTO THE EYE. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH AN ALTERNATE LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC TORN, HAPTIC TORN AND CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.0/+1.0/052 | OS |  | 12/14/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5670137&pc=MTA | 2023826-2016-00572 | 2015-10-09 | Injury | STAAR SURGICAL COMPANY | 2016-05-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+2.0/127 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO PATIENT PROBLEM. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -15.0/+2.0/127 | OD | Physician | 02/04/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5670139&pc=MTA | 2023826-2016-00573 | 2015-10-09 | Injury | STAAR SURGICAL COMPANY | 2016-05-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/84 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO PATIENT PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT; VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -15.5/+1.5/84 | OS | Physician | 02/04/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7740043&pc=MTA | 2023826-2018-01152 | 2018-04-03 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/085 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. INTRAOPERATIVELY, THE LENS WAS EXPLANTED. THE SURGEON REPORTS A CILIARY RING BLOCK, "AFTER VISCOELASTIC REMOVED, BECAUSE OF THE POSTERIOR WAS OVER PRESSURE, THE ANTERIOR CHAMBER CAN'T BE FORMED WELL SO THE LENS WAS REMOVED". THE PROBLEM WAS RESOLVED BY THE LENS EXPLANT. Manufacturer Narrative: CORRECTED DATA: (B)(4).. ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.5/+1.5/085 | OS | Physician | 07/10/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5082911&pc=MTA | 2023826-2015-01229 | 2014-01-08 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4, -22.5/+1.0/X166 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO INADEQUATE VAULTING. NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: DATE OF THIS REPORT - 08/20/2015. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) , BUT NOT MARKETED IN THE U.S. DIOPTER: -22.5/+1.0/X166. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -22.5/+1.0/166 | OD | Physician | 08/20/2015 | 2015 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5481990&pc=MTA | 2023826-2016-00166 | 2014-06-20 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS, -21.0/+2.0/113 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2007. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNTING OR SIGNIFICANT ENDOTHELIAL CELL LOSS. PHACO-EMULSIFICATION WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP STATUS WAS UCVA .09 2 MONTHS AFTER LENS REMOVAL. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | -21.0/+2.0/113 | OD | Physician | 12/23/2015 | 2015 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5308737&pc=MTA | 2023826-2015-01674 | 2014-07-24 | Injury | STAAR SURGICAL COMPANY | 2015-12-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/150 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS HAD AN EXCESSIVE VAULT AND THERE WAS A REFRACTIVE SURPRISE. THE SURGEON PLANS TO EXPLANT THE LENS BUT TO DATE THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A SHORTER LENS WITH A DIFFERENT DIOPTER ON THE SAME DAY. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/40. DEVICE EVALUATION?: NO - LENS WAS EXPLANTED BUT NOT RETURNED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -14.00/+2.0/150 | OS | Physician | 12/01/2015 | 2015 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5047754&pc=MTA | 2023826-2015-01118 | 2014-08-12 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -4.5/+1.0/24 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -4.5/+1.0/24 | OD | Physician | 08/04/2015 | 2015 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5048952&pc=MTA | 2023826-2015-01114 | 2014-08-29 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/84 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS BUT DIFFERENT DIOPTER. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS WAS WITHIN SPECIFICATION. FUNCTIONAL INSPECTION WAS PERFORMED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4). LENS NOT RETURNED. | -10.0/+1.5/84 | OS | Physician | 08/04/2015 | 2015 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5287689&pc=MTA | 2023826-2015-01606 | 2014-12-12 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.5/+3.0/089 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4). | -14.5/+3.0/089 | OS | Physician | 10/27/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5089548&pc=MTA | 2023826-2015-01233 | 2014-12-26 | Injury | STAAR SURGICAL COMPANY | 2015-09-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -22.5/1.5/055 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO- LENS NOT RETURNED. (B)(4). METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -22.5/1.5/055 | OS | Physician | 08/20/2015 | 2015 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5054881&pc=MTA | 2023826-2015-01100 | 2015-01-07 | Injury | STAAR SURGICAL COMPANY | 2015-09-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.50/+3.50/083 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ONE DAY POST-OP THERE WAS EXCESSIVE VAULTING, LENS ROTATION AND REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. A SUPPLEMENTAL WILL BE SUBMITTED WHEN ADDITIONAL IS AVAILABLE. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS CODE(S): (EVALUATION RESULT): BASED ON THE ABOVE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. EVALUATION CONCLUSION CODE(S): BASED ON THE COMPLAINT HISTORY, DEVICE HISTORY RECORD REVIEW AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S BUT NOT MARKETED IN THE U.S. DIOPTER: -14.50/3.5/083 LENS REMAINS IMPLANTED. (B)(4). EVALUATION: METHOD: LENS WORK ORDER SEARCH. EVALUATION: RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WAS NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED. | -14.50/+3.50/083 | OD | Physician | 06/19/2015 | 2015 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4743750&pc=MTA | 2023826-2015-00496 | 2015-01-12 | Injury | STAAR SURGICAL COMPANY | 2015-05-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND GLARE/HALOS WERE NOTED ON (B)(6) 2015. THE LENS REMAINS IMPLANTED BUT THE SURGEON IS PLANNING TO EXCHANGE FOR A V4B . SEE MFR REPORT #2023826-2015-00497 FOR THE OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) GLARES/HALOS. LENS IMPLANTED. DEVICE EVALUATED BY MANUFACTURER? NO : LENS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE-EVALUATION IN PROGRESS. (B)(4). LENS IMPLANTED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: GLARE AND HALOS MAY BE NOTED BY PATIENTS EVEN IF THEY NEVER HAD ANY OCULAR SURGERY. THIS IS COMMONLY DUE TO ABERRATIONS OR IRREGULARITIES IN THE CORNEA. IT MAY BE NOTED MORE AT LOW LIGHT CONDITIONS WHEN THE PUPIL IS DILATED. IN ICL SURGERY, THE CENTRAL AND MID PERIPHERAL CORNEA IS NOT TOUCHED THEREFORE, PATIENTS WHO HAVE GLARE/HALOS BEFORE SURGERY, WILL COMMONLY RETRAIN THESE SYMPTOMS BUT LITTLE OR NO INCREASE IN SEVERITY SHOULD BE EXPERIENCED. SINCE THE CENTRAL AND MID PERIPHERAL REGION REMAINS UNABLATED AND ONLY A CORNEA INCISION WITH A MAXIMUM LENGTH OF 3.2MM AT THE TEMPORAL PERIPHERAL REGION IS MADE. GLARE AND HALOS MAY ALSO BE PERCEIVED MORE DUE TO INCREASED CLARITY OF VISION AFTER ICL SURGERY OR IF THE PATIENT HAS RETAINED ASTIGMATISM. ANOTHER FACTOR COULD ALSO BE THE EFFECTIVE OPTICAL CORNEAL ZONES (EOCZ) OF THE ICLS, WHICH HAVE MAXIMUM DIAMETER OF 6.17 TO 7.30 MM DEPENDING ON ICL POWERS. THIS IS DESIGN LIMITATION WHICH MAY CRATE SIMILAR SYMPTOMS OF GLARE AND HALOS DUE TO THE INTERFACE OF THE OPTIC ZONE AND THE PATIENT'S OPTICAL FIELD OF VISION, WHICH MAY BE NOTED WHEN THE PUPIL SIZE IS GREATER THAN THE EOCZ. THE 0.36 HOLE AT THE CENTER OF THE ICL HAS BEEN COMPARED TO THE ICL WITHOUT A HOLE WHICH SHOWED THAT MODULATION-TRANSFER FUNCTIONS (MFT) ARE SMALL AND CLINICALLY NEGLIGIBLE. GLARE AND HALOS ARE COMMONLY TEMPORARY, LASTING 1-6 MONTHS, BUT MAY ON VERY RARE OCCASIONS BECOME PERMANENT. THERE IS NOT ENOUGH CLINICAL INFORMATION TO DETERMINE A ROOT CAUSE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, LENS WORK ORDER SEARCH AND MEDICAL REVIEW, WE WERE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4). | -9.50/+1.5/91 | OD | Physician | 04/02/2015 | 2015 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4743798&pc=MTA | 2023826-2015-00497 | 2015-01-12 | Injury | STAAR SURGICAL COMPANY | 2015-05-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND GLARE/HALOS WERE NOTED ON (B)(6) 2015. THE LENS REMAINS IMPLANTED BUT THE SURGEON IS PLANNING TO EXCHANGE FOR A V4B . SEE MFR REPORT #2023826-2015-00496 FOR THE OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EXPIRATION DATE - NOT AVAILABLE. (B)(4) GLARES/HALOS. LENS IMPLANTED. DEVICE EVALUATED BY MANUFACTURER? NO: LENS IMPLANTED. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE-EVALUATION IN PROGRESS. (B)(4). | -9.50/+1.5/88 | OD | Physician | 04/02/2015 | 2015 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4695350&pc=MTA | 2023826-2015-00421 | 2015-02-16 | Injury | STAAR SURGICAL COMPANY | 2015-04-15 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. REFRACTIVE SURPRISE WAS NOTED, THE LENS REMAINS IMPLANTED. PATIENT'S LAST KNOWN BCVA WAS 20/20 AS OF (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT:UNK. EXPLANT DATE: N/A. (B)(4). DEVICE EVALUATED BY THE MANUFACTURER? NO. LENS REMAINS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW RESULTS: (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -3.00/+6.0/002 | OD | Other | 03/16/2015 | 2015 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4760617&pc=MTA | 2023826-2015-00515 | 2015-04-04 | Injury | STAAR SURGICAL COMPANY | 2015-05-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) INCISION OPENED. SECONDARY SURGERY. LENS EXPLANTED. VAULTING. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -11.0/2.5/80 | OD | Other | 04/13/2015 | 2015 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4760619&pc=MTA | 2023826-2015-00495 | 2015-04-04 | Injury | STAAR SURGICAL COMPANY | 2015-05-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -10.5/3.0/98 | OS | Other | 04/13/2015 | 2015 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5072524&pc=MTA | 2023826-2015-01130 | 2015-04-21 | Injury | STAAR SURGICAL COMPANY | 2015-09-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2, -15.00/+2.00/X103 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT. NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -15.00/+2.00/X103. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -15.00/+2.00/103 | OD | Physician | 08/20/2015 | 2015 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5072732&pc=MTA | 2023826-2015-01132 | 2015-04-30 | Injury | STAAR SURGICAL COMPANY | 2015-09-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/+0.5/X099 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT. NO OTHER LENS WAS IMPLANTED. THE PATIENT DOES NOT WANT SURGERY ONCE AGAIN. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -12.5/+0.50/X099. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -12.50/+0.5/099 | OD | Physician | 08/20/2015 | 2015 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5112652&pc=MTA | 2023826-2015-01190 | 2015-05-29 | Injury | STAAR SURGICAL COMPANY | 2015-09-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/X169 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED DUE TO A REFRACTIVE SURPRISE ON (B)(6) 2015. LENS WAS EXCHANGE FOR DIFFERENT POWER AND SAME SIZE. THE PROBLEM WAS RESOLVED. PATIENT'S LAST VISIT ON (B)(6) 2015, UCVA WAS 20/24. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -12.0/+2.5/X169. (B)(4). METHOD CODE: EVALUATION METHOD: WORK ORDER SEARCH. EVALUATION CODE: EVALUATION RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS CODE: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -12.0/+2.5/169 | OD | Physician | 09/01/2015 | 2015 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4954855&pc=MTA | 2023826-2015-00982 | 2015-06-12 | Injury | STAAR SURGICAL COMPANY | 2015-07-30 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND THERE WAS AN INCIDENT OF REFRACTIVE SURPRISE. SHALLOW ANTERIOR CHAMBER ASSOCIATED WITH EXCESSIVE VAULT WAS NOTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS DEVICE IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND WITHIN THE WORK ORDER. LENS IMPLANTED. Manufacturer Narrative: BASED ON THE ABOVE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -18.0/+2.5/56 | OD | Other | 07/08/2015 | 2015 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5155159&pc=MTA | 2023826-2015-01350 | 2015-07-02 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/089 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND A SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | -12.5/+2.5/089 | OS | Physician | 09/15/2015 | 2015 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5325950&pc=MTA | 2023826-2015-01709 | 2015-07-22 | Injury | STAAR SURGICAL COMPANY | 2015-12-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/X057 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -9.5/+1.5/057 | OS | Physician | 11/16/2015 | 2015 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5443094&pc=MTA | 2023826-2016-00169 | 2015-07-29 | Injury | STAAR SURGICAL COMPANY | 2016-02-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+4.5/057 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS HAD ROTATED (NOT ASSOCIATED TO A LOW VAULT) AND WAS RE-POSITIONED. THERE WAS LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, ACD BELOW 3.0MM FOR VICL/VTICL. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -12.0/+4.5/057 | OS | Physician | 01/18/2016 | 2016 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5248746&pc=MTA | 2023826-2015-01549 | 2015-08-16 | Injury | STAAR SURGICAL COMPANY | 2015-11-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -1.00/+6.0/100 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ELEVATED INTRAOCULAR PRESSURE, BLURRED VISION AND INFLAMMATION. THE SURGEON IS PLANNING TO IMPLANT A SHORTER LENS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO -LENS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -1.00/+6.0/100 | OD | Physician | 11/02/2015 | 2015 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5325936&pc=MTA | 2023826-2015-01714 | 2015-09-09 | Injury | STAAR SURGICAL COMPANY | 2015-12-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/X086 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -12.5/+2.5/086 | OS | Physician | 11/16/2015 | 2015 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5225080&pc=MTA | 2023826-2015-01510 | 2015-10-14 | Injury | STAAR SURGICAL COMPANY | 2015-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/X092 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD A LOW VAULT AND LENS ROTATION. THE SURGEON IS PLANNING TO EXPLANT THE LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: LENS EXCHANGED, PROBLEM RESOLVED. EXPLANTED DATE (B)(6) 2015. (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS REMAINS IMPLANTED. (B)(4) METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -11.0/+1.0/092 | OD | Physician | 10/15/2015 | 2015 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5304424&pc=MTA | 2023826-2015-01685 | 2015-11-19 | Injury | STAAR SURGICAL COMPANY | 2015-12-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -08.50/+1.50/X008 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT AND MOVING PUPIL WAS NOTED ON (B)(6) 2015. THE LENS REMAINS IMPLANTED AND THE PATIENT IS BEING MONITORED. PATIENT'S LAST VISIT ON (B)(6) 2015, UCVA WAS 20/20. SEE MFR# 2023826-2015-01686 FOR LEFT EYE. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SMALLER LENS. THIS RESOLVED THE PROBLEM. (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT AND THE HAPTIC TORN/BROKEN. DIMENSIONAL INSPECTION FOUND LENS MEASUREMENTS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -08.50/+1.50/X008. DEVICE EVALUATED BY MANUFACTURER?: NO. LENS REMAINS IMPLANTED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. EVALUATION CONCLUSION CODE(S): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED. | -08.50/+1.50/008 | OD | Physician | 12/01/2015 | 2015 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5305157&pc=MTA | 2023826-2015-01686 | 2015-11-19 | Injury | STAAR SURGICAL COMPANY | 2015-12-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -08.50/+1.00/X176 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES AND ANGLE CLOSURE WITH ELEVATED IOP WAS NOTED ON (B)(6) 2015. THE ICL WAS EXPLANTED AND WAITING FOR IMPLANTATION OF SHORTER LENGTH LENS IN THE FUTURE. SEE MFR#2023826-2015-01685 FOR RIGHT EYE. Manufacturer Narrative: THE LENS WAS EXCHANGED FOR A SHORTER LENS AND SIMILAR DIOPTER ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2015, THE PATIENT'S BEST CORRECTED VISUAL ACCUITY (BCVA) WAS 20/20. (B)(4). CORRECTED DATA: THE LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. (B)(4). LENS RETURNED OVER A YEAR AGO, NO EVAL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -08.50/+1.0/X0176 DEVICE EVALUATED BY MANUFACTURER?: NO - PRODUCT NOT RETURNED. (B)(4). METHOD CODE(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODE(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). DEVICE NOT RETURNED. | -08.50/+1.00/176 | OS | Physician | 12/01/2015 | 2015 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5518879&pc=MTA | 2023826-2016-00297 | 2015-12-02 | Injury | STAAR SURGICAL COMPANY | 2016-03-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.50/+1.0/100 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD A REFRACTIVE SURPRISE AND REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). DHR EVALUATION FOUND THAT ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS, INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -5.50/+1.0/100 | OS | Physician | 01/15/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5596429&pc=MTA | 2023826-2016-00493 | 2016-02-10 | Injury | STAAR SURGICAL COMPANY | 2016-04-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+6.0/072 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD A REFRACTIVE SURPRISE AND REMAINS IMPLANTED. Manufacturer Narrative: DHR EVALUATION: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4). | -8.5/+6.0/072 | OS | Physician | 02/11/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6077471&pc=MTA | 2023826-2016-01564 | 2016-02-18 | Injury | STAAR SURGICAL COMPANY | 2016-11-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+3/75 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON AN UNKNOWN DATE. THE LENS WAS EXPLANTED ON (B)(6) 2016, NO INFORMATION WAS PROVIDED AS TO WHY THE LENS WAS EXPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED FOR A LONGER LENS. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: PMA/510(K): N/A THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+3/75 | OS | Physician | 10/07/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6091602&pc=MTA | 2023826-2016-01606 | 2016-06-09 | Injury | STAAR SURGICAL COMPANY | 2016-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/170 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A REFRACTIVE CHANGE OVERTIME AND LENS DISLOCATION/SUBLUXATION. LENS WAS EXCHANGED FOR A LONGER LENS, AND PROBLEM WAS RESOLVED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.0/170 | OS | Physician | 10/13/2016 | 2016 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4655007&pc=MTA | 2023826-2015-01121 |  | Injury | STAAR SURGICAL COMPANY | 2015-04-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -7.0/+0.5/X105 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THE LENS WAS EXPLANTED DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP, ETC.) TO PREVENT ANY OF THESE COMPLICATIONS FORM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -7.0/+0.5/105 | OD | Physician | 08/11/2015 | 2015
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218821&pc=MTA | 2023826-2015-01500 |  | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.50/+0.5/155 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -7.50/+0.5/155 | OS | Physician | 10/13/2015 | 2015 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4937376&pc=MTA | 2023826-2015-00909 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-07-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. AT ONE DAY POST-OP, THERE WAS CORNEAL EROSION, HYPOTONY AND THE ANTERIOR CHAMBER WAS NOT FILLED (THERE MAY HAVE BEEN A WOUND LEAK). THE LENS VAULT WAS OK AND THE EYE WAS SLIGHTLY MYDRIATIC. THE PATIENT WAS BROUGHT BACK TO SURGERY AND THE SURGEON IRRIGATED/FILLED UP THE ANTERIOR CHAMBER. MIOSTAT WAS ADMINISTERED. A COUPLE OF WEEKS LATER, THE LENS WAS RE-POSITIONED. AT ONE DAY POST-OP, THE PUPIL IS SEMI-MYDRIATIC. THE LENS REMAINS IMPLANTED AND THE PATIENT IS HAPPY. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO, LENS IMPLANTED. (B)(4). | -3.5/+5.0/064 | OD | Physician | 06/03/2015 | 2015
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5174781&pc=MTA | 2023826-2015-01401 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-10-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+4.0/103 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS ROTATED ONE YEAR AFTER IMPLANTATION AND WAS RE-POSITIONED. THERE WAS REFRACTIVE CHANGE OVERTIME. THE LENS IS NOT COMPLETELY STABLE INSIDE THE EYE AND ROTATES. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/90. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE VAULTING WAS 350 MICRONS. A NEW LENS HAS BEEN CALCULATED IN OCOS WITH SAME POWER BUT LONGER LENS. MEDICAL REVIEW CONCLUDED THAT THE LENS WAS EXPLANTED AND EXCHANGED WITH A LONGER ONE. REPORT SOURCE IS FROM (B)(6), NOT (B)(6). (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). | -14.5/+4.0/103 | OD | Physician | 09/25/2015 | 2015 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5086231&pc=MTA | 2023826-2015-01133 | 2015-05-28 | Malfunction | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -12.0/+1.5/X101 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS TORE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED DURING THE SAME PROCEDURE AND NO ADVERSE EVENT WAS REPORTED. NO OTHER LENS WAS IMPLANTED. CAUSE OF LENS TEAR IS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -12.0/+1.5/X101. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). METHOD CODE: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODE: (EVALUATION RESULT) : A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS CODE: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -12.0/+1.5/101 | OD | Physician | 08/20/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5018310&pc=MTA | 2023826-2015-01064 | 2015-07-18 | Malfunction | STAAR SURGICAL COMPANY | 2015-08-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6, -15.00/+1.0/094 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. IT WAS NOTED WHILE IMPLANTING THE LENS, THE LENS BROKE. THE LENS WAS REMOVED. NO FURTHER INFORMATION. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO LENS NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS NOT RETURNED. | -15.00/+1.0/094 | OS | Physician | 07/21/2015 | 2015 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6042652&pc=MTA | 2023826-2016-01427 | 2016-05-02 | Injury | STAAR SURGICAL COMPANY | 2016-10-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+3.5/079 DIOPTER, AND THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED ON (B)(6) 2016 AND ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS WAS IMPLANTED ON (B)(6) 2016. THE EXCHANGED LENS RESOLVED THE PROBLEM AND THERE WAS NO INDICATION OF ANY APPARENT INJURY. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION - THE LENS WAS RETURNED IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND FIBERS ON THE LENS SURFACE. (B)(4). WORK ORDER SEARCH - NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -14.0/+3.5/079 | NA | Physician | 09/27/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5167031&pc=MTA | 2023826-2015-01369 | 2015-02-13 | Injury | STAAR SURGICAL COMPANY | 2015-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -18.5/+2.5/086 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND THE LENS TORE/BROKE. THE LENS WAS EXPLANTED ON (B)(6) 2015, WITH NO REPORTED PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL, DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -18.5/+2.5/086 | OD | Physician | 09/23/2015 | 2015 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4789551&pc=MTA | 2023826-2015-00574 | 2015-03-19 | Injury | STAAR SURGICAL COMPANY | 2015-05-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material; Unintended Movement | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6)2015. THE LENS WAS EXPLANTED ON (B)(6)2015 DUE TO THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THE LENS DISLOCATED/SUBLUXED AND THERE WAS REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED FOR ANOTHER SAME SIZE/POWER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: (B)(4). EVALUATION CODES: METHOD - (OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (OTHER): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF. THE LENS WAS RETURNED IN LIQUID. DUE TO THE RETURNED CONDITION OF THE LENS, A REMEASUREMENT OF THE LENS COULD NOT BE PERFORMED. A REVIEW OF THE DEVICE HISTORY RECORD FOUND NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. NO DEFINITE ROOT CAUSE OF THE COMPLAINT IS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -14.00/+3.0/095 | OS | Other | 04/27/2015 | 2015 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4773384&pc=MTA | 2023826-2015-00553 | 2015-04-20 | Malfunction | STAAR SURGICAL COMPANY | 2015-05-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Unintended Movement | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) FOUR TIMES AND EACH TIME THE LENS WENT INTO THE PATIENT'S EYE UPSIDE DOWN. THE SURGEON DECIDED NOT TO IMPLANT THE LENS DUE TO THE EYE HAD SUFFERED AND THE SURGERY WAS NOT FINISHED. THE LENS WAS BROKEN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). EVALUATION CODES: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. PER MEDICAL REVIEW - WHILE THE LENS IS BEING EJECTED INTO THE ANTERIOR CHAMBER THE SURGEON MUST CHECK FOR LENS LANDMARKS TO ENSURE PROPER LENS ORIENTATION. IF THIS IS NOT DONE, THE LENS MAY END-UP BEING IMPLANTED UPSIDE-DOWN. ONCE THE SURGEON REALIZES OF THE IMPROPER ORIENTATION, THE LENS MUST BE REMOVED TO AVOID COMPLICATIONS AND A NEW LENS SHOULD BE IMPLANTED. THIS IS NORMALLY DONE USING THE SAME INCISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, THE MOST PROBABLE CAUSE OF THIS EVENT WAS USER'S ERROR WHILE LOADING/INJECTING THE LENS IN THE ANTERIOR CHAMBER. (B)(4). | -8.5/+1.5/072 | OS | Other | 04/23/2015 | 2015 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6600868&pc=MTA | 2023826-2017-00907 | 2015-03-11 | Injury | STAAR SURGICAL COMPANY | 2017-05-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0/+4.5/90 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 AND NOTICED LENS ROTATION EARLY POSTOPERATIVELY. THE LENS WAS REPOSITIONED ON (B)(6) 2015 BUT THE LENS ROTATED AGAIN. THE LENS WAS EXCHANGED ON (B)(6) 2016 WITH A SPHERICAL MODEL LENS WITH SAME DIAMETER. THIS RESOLVED THE PROBLEM. PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 AND UNCORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: THIS DEVICE IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). | -16.0/+4.5/90 | OS | Physician | 03/08/2016 | 2016 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6591027&pc=MTA | 2023826-2017-00891 | 2015-10-20 | Injury | STAAR SURGICAL COMPANY | 2017-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.0/+2.5/80 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE SURGEON NOTICED LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND DECIDED TO REPOSITION THE LENS. DESPITE THREE ATTEMPTS OF LENS REPOSITIONING, THE PROBLEM WAS NOT RESOLVED. THE SURGEON PLANS TO EXCHANGE THE LENS. Manufacturer Narrative: (B)(4). CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS WITH PRE-EXISTING KERATOCONUS. (B)(4). | -8.0/+2.5/80 | OS | Physician | 08/08/2016 | 2016 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6604411&pc=MTA | 2023826-2017-00916 | 2015-11-18 | Injury | STAAR SURGICAL COMPANY | 2017-06-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO IMPLANTABLE COLLAMER LENS, -15.0/+3.0/61 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2016 BUT THE REPOSITIONING DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXCHANGED ON (B)(6) 2016 FOR A LONGER LENS OF A SIMILAR DIOPTER AND THE PROBLEM WAS RESOLVED. PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY ON (B)(6) 2016 WAS 20/20. Manufacturer Narrative: THIS DEVICE IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS BUT FOREIGN MATERIAL ON LENS SURFACE (DRIED AND DISCOLORED MATERIAL). (B)(4). NO CODE AVAILABLE: SECONDARY SURGICAL INTERVENTION, REPOSITIONING, LENS EXCHANGE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.0/+3.0/61 | OD | Physician | 07/26/2016 | 2016 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6616474&pc=MTA | 2023826-2017-00932 | 2015-11-23 | Injury | STAAR SURGICAL COMPANY | 2017-06-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -3.5/+6/064 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT BEGAN EXPERIENCING LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND ELEVATED IOP. ON (B)(6) 2015 AND (B)(6) 2016 THE LENS WAS REPOSITIONED. THE LENS REMAINS IMPLANTED AND THE SURGEON PLANS TO EXCHANGE FOR A LARGER LENS. THE SURGEON STATES THAT THE LENS "ROTATES ALL THE TIME." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXCHANGED ON (B)(6) -2017 FOR A LONGER LENS. THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED OR MARKETED IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED. | -3.5/+6/064 | OS | Physician | 05/10/2017 | 2017 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6055916&pc=MTA | 2023826-2016-01445 | 2015-12-04 | Injury | STAAR SURGICAL COMPANY | 2016-10-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+3.0/074 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON 07/04/2016 DUE TO LENS DISLOCATION/SUBLUXATION AND REPOSITIONING. THE LENS WAS EXCHANGED FOR A LENS WITH A DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT. FOREIGN MATERIAL (DRIED, DISCOLORED MATERIAL) WAS PRESENT ON THE LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). LENS EXPLANTED AND EXCHANGED FOR A LENS THE SAME LENGTH WITH A DIFFERENT DIOPTER. (B)(4). | -15.00/+3.0/074 | OS | Physician | 09/26/2016 | 2016 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6608039&pc=MTA | 2023826-2017-00905 | 2016-01-15 | Injury | STAAR SURGICAL COMPANY AG | 2017-06-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0 MM TICM120V4 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/97 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. LENS ROTATION NOT ASSOCIATED WITH LOW VAULT WAS OBSERVED. ON (B)(6) 2015, THE LENS WAS REPOSITIONED. THE LENS WAS THEN EXCHANGED FOR ANOTHER LENS ON (B)(6) 2016. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). | -13.0/+2.5/97 | OD | Physician | 05/17/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6585641&pc=MTA | 2023826-2017-00868 | 2016-05-18 | Injury | STAAR SURGICAL COMPANY | 2017-05-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -8.50/+2.5/083 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO THE PATIENT EXPERIENCING LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. PRIOR TO LENS EXCHANGE, ON (B)(6) 2016 AND (B)(6) 2016 THE LENS WAS REPOSITIONED. THE LENS EXCHANGE RESOLVED THE PROBLEM. AS OF THE DATE OF MDR REPORTING THE LENS HAS NOT YET BEEN RETURNED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: PATIENT WEIGHT: UNK. THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4): DEVICE NOT RETURNED TO MANUFACTURER. | -8.50/+2.5/083 | OS | Physician | 04/28/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6242846&pc=MTA | 2023826-2017-00033 | 2016-06-01 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -15.5/+3.0/92 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS THEN REPOSITIONED IN (B)(6) 2013, (B)(6) 2014, AND (B)(6) 2016 (EXACT DATES NOT KNOWN). ON (B)(6) 2016, WAS EXCHANGED WITH ANOTHER LENS DUE TO LENS ROTATION (NOT ASSOCIATED TO A LOW VAULT). THE PROBLEM WAS RESOLVED. THE LENS HAS BEEN RETURNED, HOWEVER THE EVALUATION IS PENDING AS OF THE DATE OF THIS REPORT. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND NOT MARKETED IN THE U.S. (B)(4). | -15.5/+3.0/92 | OS | Physician | 11/08/2016 | 2016 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6591833&pc=MTA | 2023826-2017-00894 | 2016-06-02 | Injury | STAAR SURGICAL COMPANY AG | 2017-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -14.5/+4.5/70 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2008. THE SURGEON NOTICED LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND DECIDED TO REPOSITION THE LENS. THE LENS WAS EXCHANGED ON (B)(6) 2016 FOR ANOTHER LENS OF A DIFFERENT MODEL/DIOPTER AND THE PROBLEM WAS RESOLVED. ON PATIENT'S LAST VISIT ON (B)(6) 2016, THE SURGEON NOTED THAT THE PATIENT IS DOING FINE WITH PRETTY GOOD VAULT. Manufacturer Narrative: ADDITIONAL INFORMATION: THIS DEVICE IS NOT MANUFACTURED OR MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). | -14.5/+4.5/70 | OS | Physician | 08/12/2016 | 2016 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6606585&pc=MTA | 2023826-2017-00924 | 2016-06-02 | Injury | STAAR SURGICAL COMPANY AG | 2017-06-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -16.5/+3.5/120 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2008. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED AND THE LENS WAS REPOSITIONED. THE LENS WAS EXCHANGED ON (B)(6) 2016 FOR LONGER LENS OF A DIFFERENT MODEL AND DIOPTER. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS DEVICE IS NOT MARKETED IN THE U.S. THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. PATIENT CODE: (B)(4)- NO CODE AVAILABLE: SECONDARY SURGICAL INTERVENTION, REPOSITIONING, LENS EXCHANGE. METHOD CODE: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.5/+3.5/120 | OD | Physician | 08/12/2016 | 2016 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6227650&pc=MTA | 2023826-2017-00028 | 2016-06-17 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -14.0/+4.0/82 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2009. THE LENS WAS REPOSITIONED, AND ON (B)(6) 2016, WAS EXCHANGED WITH ANOTHER LENS DUE TO LENS ROTATION (NOT ASSOCIATED TO A LOW VAULT). THE PROBLEM WAS RESOLVED. THE LENS HAS BEEN RETURNED, HOWEVER THE EVALUATION IS PENDING AS OF THE DATE OF THIS REPORT. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DRIED, DISCOLORED MATERIAL ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND AND NOT MARKETED IN THE U.S. (B)(4): (SECONDARY SURGERY, LENS REPOSITIONING, LENS EXCHANGE). (B)(4). | -14.0/+4.0/82 | OS | Physician | 11/08/2016 | 2016 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6606756&pc=MTA | 2023826-2017-00923 | 2016-07-11 | Injury | STAAR SURGICAL COMPANY | 2017-06-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.0/3.0/109 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED AND THE LENS WAS REPOSITIONED. THE LENS WAS EXCHANGED ON (B)(6) 2016 FOR LONGER LENS OF A DIFFERENT DIOPTER. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS DEVICE IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. PATIENT CODE: (B)(4)- NO CODE AVAILABLE: SECONDARY SURGICAL INTERVENTION, REPOSITIONING, LENS EXCHANGE. METHOD CODE: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.0/3.0/109 | OD | Physician | 08/12/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6608047&pc=MTA | 2023826-2017-00906 | 2016-07-11 | Injury | STAAR SURGICAL COMPANY | 2017-06-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.0/+3.5/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LENS ROTATION NOT ASSOCIATED WITH LOW VAULT WAS OBSERVED. ON (B)(6) 2016, THE LENS WAS REPOSITIONED AND THEN EXCHANGED FOR ANOTHER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION- PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). | -18.0/+3.5/085 | OS | Physician | 08/12/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6598741&pc=MTA | 2023826-2017-00904 | 2016-07-18 | Injury | STAAR SURGICAL COMPANY | 2017-05-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -4.0/+1.0/110 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. THE LENS WAS ROTATED THREE TIMES. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED FOR PATIENTS WITH KERATOCONUS EYE. (B)(4). LENS NOT RETURNED. | -4.0/+1.0/110 | OD | Physician | 08/08/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6577102&pc=MTA | 2023826-2017-00797 | 2016-07-23 | Injury | STAAR SURGICAL COMPANY | 2017-05-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+0.5/102 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2016. THE LENS WAS REPOSITIONED BUT WAS UNSUCCESSFUL. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A SAME LENGTH (12.6MM) BUT DIFFERENT MODEL AND DIOPTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: NOTE: THE CUSTOMER DID NOT PROPERLY LABEL/IDENTIFY RETURNED LENSES. THREE LENS WERE RETURNED ASSOCIATED WITH THREE DIFFERENT COMPLAINTS. PRODUCT EVALUATIONS FOR THE THREE LENSES ARE IN THE EACH FILE. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) - SECONDARY SURGICAL INTERVENTION, REPOSITIONED, EXPLANT AND EXCHANGE FOR SAME LENGTH BUT DIFFERENT MODEL AND DIOPTER LENS (B)(4). | -15.5/+0.5/102 | OD | Physician | 04/21/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6577103&pc=MTA | 2023826-2017-00798 | 2016-07-23 | Injury | STAAR SURGICAL COMPANY | 2017-05-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+0.5/096 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2016. THE LENS WAS REPOSITIONED BUT WAS UNSUCCESSFUL. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A SAME LENGTH (12.6MM) BUT DIFFERENT MODEL AND DIOPTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: NOTE: THE CUSTOMER DID NOT PROPERLY LABEL/IDENTIFY RETURNED LENSES. THREE LENS WERE RETURNED ASSOCIATED WITH THREE DIFFERENT COMPLAINTS. PRODUCT EVALUATION FOR THE THREE LENSES ARE IN THE EACH FILE. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. SECONDARY SURGICAL INTERVENTION, REPOSITIONED, EXPLANT AND EXCHANGE FOR SAME LENGTH BUT DIFFERENT MODEL AND DIOPTER LENS (B)(4). | -17.0/+0.5/096 | OS | Physician | 04/21/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6620566&pc=MTA | 2023826-2017-00935 | 2016-10-13 | Injury | STAAR SURGICAL COMPANY | 2017-06-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -10.0/3.5/98 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2016. ON (B)(6) 2016 AND (B)(6) 2016 THE LENS WAS REPOSITIONED BECAUSE THE PATIENT BEGAN EXPERIENCING LENS ROTATION NOT ASSOCIATED TO A LOW VAULT ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A LARGER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE INITIAL LENS WAS IMPLANTED ON (B)(6) 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION-LENS WAS RETURNED DRY IN A SMALL VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS.. (B)(4). | -10.0/3.5/98 | OD | Physician | 05/11/2017 | 2017 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6619009&pc=MTA | 2023826-2017-00941 | 2016-11-02 | Injury | STAAR SURGICAL COMPANY | 2017-06-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -3.0/+5/109 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. ON (B)(6) 2017 THE LENS WAS EXPLANTED AFTER THE REPOSITIONING ON (B)(6) 2016. THE SURGEON COMMENTED THAT THE LENS WAS EXPLANTED AFTER "CONSICUTIVE LENS TURNS", AND HE OPTS FOR PRK SURGERY. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: LENS WAS RETURNED IN SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS, AND CLEAR RESIDUE ON THE LENS SURFACE. (B)(4). | -3.0/+5/109 | OS | Physician | 05/12/2017 | 2017 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6539263&pc=MTA | 2023826-2017-00746 | 2016-11-24 | Injury | STAAR SURGICAL COMPANY | 2017-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+3.0/091 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS ROTATED, NOT ASSOCIATED TO A LOW VAULT, ON (B)(6) 2016 AND WAS REPOSITIONED ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 AND EXCHANGED FOR A LONGER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND MISSING A PIECE WITH WHITE/CLEAR RESIDUE ON LENS SURFACE. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+3.0/091 | OD | Physician | 04/04/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6247129&pc=MTA | 2023826-2017-00063 | 2016-12-01 | Injury | STAAR SURGICAL COMPANY | 2017-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11/1/162 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS THEN REPOSITIONED (EXACT DATE NOT KNOWN). ON (B)(6) 2016, THE LENS WAS EXPLANTED DUE TO LENS ROTATION (NOT ASSOCIATED TO A LOW VAULT). THE LENS HAS BEEN RETURNED, HOWEVER THE EVALUATION IS PENDING AS OF THE DATE OF THIS REPORT. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11/1/162 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS THEN REPOSITIONED (EXACT DATE NOT KNOWN) DUE TO LENS ROTATION (NOT ASSOCIATED TO A LOW VAULT). ON (B)(6) 2016, THE LENS WAS EXPLANTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11/1/162 | OD | Physician | 12/14/2016 | 2016 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7509196&pc=MTA | 2023826-2018-00685 | 2017-01-25 | Injury | STAAR SURGICAL COMPANY | 2018-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.5/+2.5/167 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPOSITIONED ON (B)(6) 2017 DUE TO LENS ROTATION NO ASSOCIATED TO A LOW VAULT. THE LENS REPOSITIONING DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR LONGER LENGTH LENS. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND PRESENCE OF DEBRIS ON LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE:(B)(4)- SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGED. WORK ORDER SEARCH- NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.5/+2.5/167 | OS | Physician | 04/11/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6701722&pc=MTA | 2023826-2017-01157 | 2017-02-03 | Injury | STAAR SURGICAL COMPANY | 2017-07-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+4.5/91 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED WITH LOW VAULT, AND UNDERWENT LENS REPOSITIONING TWICE ON (B)(6) 2017, AND (B)(6) 2017 BUT UNSUCCESSFUL. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A LONGER LENS, AND THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC SLIGHTLY BENT. THERE WAS EVIDENCE OF CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/+4.5/91 | OD | Physician | 06/19/2017 | 2017 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050901&pc=MTA | 2023826-2017-01717 | 2017-02-09 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1, -16.0/+1.5/108 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. IT DID NOT RESOLVE THE PROBLEM. THE LENS WAS THEN EXCHANGED WITH A LONGER LENS ON (B)(6) 2017. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN AND CLEAR AND WHITE RESIDUE ON LENS. A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US.(B)(4). | -16.0/+1.5/108 | OD | Physician | 10/25/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6440672&pc=MTA | 2023826-2017-00501 | 2017-02-14 | Injury | STAAR SURGICAL COMPANY | 2017-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -17.00/2.50/093 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. ON (B)(6) 2016, THE LENS WAS REPOSITIONED. THE PROBLEM WAS NOT RESOLVED. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR THE SAME SIZE, SIMILAR DIOPTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. Manufacturer Narrative: PREVIOUSLY STATED THE LENS HAD NOT BEEN RETURNED. UPDATED: THE LENS HAS BEEN RETURNED AND EVALUATED. DEVICE EVALUATION-LENS WAS RETURNED DRY A SMALL VIAL, WITH CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND TWO PIECES MISSING FROM THE HAPTIC AND THE HAPTIC TORN.(B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | -17.00/2.50/093 | OD | Physician | 02/27/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6491470&pc=MTA | 2023826-2017-00642 | 2017-02-14 | Injury | STAAR SURGICAL COMPANY | 2017-04-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -14.0/+2.0/166 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS REPOSITIONED. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SAME SIZE LENS, SIMILAR DIOPTER, DUE TO THE PATIENT EXPERIENCING LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR REPORTING THE LENS HAS NOT YET BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4) . Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4) . | -14.0/+2.0/166 | OS | Physician | 03/15/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7006961&pc=MTA | 2023826-2017-01688 | 2017-02-25 | Injury | STAAR SURGICAL COMPANY | 2017-11-07 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 DIOPTER -12.0/+3.5/097 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS RE-POSITIONED. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A LARGER LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. REPORTEDLY, THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. PATIENT CODE (B)(4) (NO CODE AVAILABLE): SECONDARY SURGERY, LENS EXCHANGE, LENS REPOSITIONING. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.0/+3.5/097 | OD | Physician | 10/13/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6800883&pc=MTA | 2023826-2017-01340 | 2017-02-27 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -17.0/+1.5/82 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED WITH LOW VAULT. ON (B)(6) 2017 THE LENS WAS REPOSITIONED AND EXCHANGED FOR A LONGER LENS IN THE SAME SURGERY. THE PROBLEM WAS RESOLVED. THE REPORTER MENTIONED THAT THE SURGEON TRIED TO REPOSITION THE LENS FIRST, BUT IT WAS NOT SATISFIED TO RESOLVED THE ISSUE. SO HE EXCHANGED THE LENS IN ONE SIZE LARGER. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17.0/+1.5/82 | OS | Physician | 07/25/2017 | 2017 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6899547&pc=MTA | 2023826-2017-01511 | 2017-03-13 | Injury | STAAR SURGICAL COMPANY | 2017-09-27 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 DIOPTER -12.0/+1.5/092 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 AND (B)(6) 2017 THE LENS WAS RE-POSITIONED. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SAME SIZE, BUT DIFFERENT MODEL LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND WHITE AND CLEAR RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.5/092 | OS | Physician | 09/01/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6648849&pc=MTA | 2023826-2017-00992 | 2017-04-26 | Injury | STAAR SURGICAL COMPANY | 2017-06-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, DIOPTER -10.0/+5.5/X126 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. WHEN RESEARCHING WHY THE LENS HAD ROTATED, IT WAS DISCOVERED THAT THE SURGEON HAD USED AN INCORRECT WTW VALUE TO CALCULATE THE SIZE OF THE LENS. THE LENS ORDERED WAS SMALLER THAN WHAT IT SHOULD HAVE BEEN. THE VALUE USED WAS 11.0 BUT THE ACTUAL MEASUREMENT IS 11.7. ON (B)(6) 2017 THE LENS WAS EXPLANTED. A NEW LENS IS PLANNED TO BE IMPLANTED. THE LENS HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN THE LENS CASE/VIAL, NO VISIBLE DAMAGE TO THE LENS. METHOD CODE: CHANGED TO REFLECT PRODUCT EVALUATION. (B)(4). Manufacturer Narrative: A NEW LENS IS PLANNED TO BE IMPLANTED. THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. UPDATED: THE LENS WAS EXCHANGED ON (B)(6) 2017 FOR A LONGER, SIMILAR DIOPTER LENS. THE PROBLEM IS RESOLVED. LENS HAS BEEN RETURNED FOR EVALUATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). NOT RETURNED TO MANUFACTURER. | -10.0/+5.5/126 | OD | Physician | 05/22/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6645721&pc=MTA | 2023826-2017-00990 | 2017-05-02 | Injury | STAAR SURGICAL COMPANY | 2017-06-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -9.5/+2.0/076 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS REPOSITIONED BECAUSE THE PATIENT BEGAN EXPERIENCING LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A LARGER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION THE LENS HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: LENS WAS RETURNED DRY IN A SMALL VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH RED AND CLEAR RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). NOT RETURNED TO MANUFACTURER. | -9.5/+2.0/076 | OS | Physician | 05/18/2017 | 2017 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6899537&pc=MTA | 2023826-2017-01510 | 2017-05-12 | Injury | STAAR SURGICAL COMPANY | 2017-09-27 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 DIOPTER -11.0/+1.0/088 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS RE-POSITIONED. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SAME SIZE, BUT DIFFERENT MODEL LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. PATIENT CODE (B)(4) (NO CODE AVAILABLE): SECONDARY SURGERY, LENS EXCHANGELM LENS REPOSITIONING. CONCLUSION CODE: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))]. (B)(4). | -11.0/+1.0/088 | OD | Physician | 09/01/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6873950&pc=MTA | 2023826-2017-01444 | 2017-05-15 | Injury | STAAR SURGICAL COMPANY | 2017-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -10.0/+2.0/080 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. LENS ROTATION NOT ASSOCIATED WITH LOW VAULT WAS NOTICED. THE LENS WAS THEN REPOSITIONED. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR ANOTHER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT UNDER THE AGE OF 21 YEARS OLD. (B)(4). | -10.0/+2.0/080 | OS | Physician | 08/22/2017 | 2017 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6667702&pc=MTA | 2023826-2017-01036 | 2017-05-22 | Injury | STAAR SURGICAL COMPANY | 2017-06-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0 MM TICM120V4, DIOPTER -12.0/+6.0/91 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2010. ON (B)(6) 2016 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2017 THE LENS WAS EXPLANTED. THERE IS NO PLAN TO IMPLANT A NEW LENS. THE LENS HAS BEEN RETURNED FOR EVALUATION BUT HAS NOT YET BEEN EVALUATED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTIONS FOUND LENS IS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED NOR MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.0/+6.0/91 | OD | Physician | 05/31/2017 | 2017 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6899531&pc=MTA | 2023826-2017-01505 | 2017-07-11 | Injury | STAAR SURGICAL COMPANY | 2017-09-27 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A VTICM5_13.2 DIOPTER -11.0/+1.5/083 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS RE-POSITIONED. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SAME SIZE, BUT DIFFERENT MODEL LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE LENS HAS A CURVED SHAPE. (B)(4). Manufacturer Narrative: PATIENT POST-OP BCVA WAS 20/20 AND UCVA WAS 20/25 ON (B)(6) 2017. A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS FOUND. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -11.0/+1.5/083 | OD | Physician | 08/31/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6970588&pc=MTA | 2023826-2017-01596 | 2017-07-30 | Injury | STAAR SURGICAL COMPANY | 2017-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1, -11/+4.5/007 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. AFTER A WEEK, THE LENS ROTATED. ON (B)(6) 2017, THE LENS WAS REPOSITIONED BUT THEN ROTATED AGAIN ONE DAY LATER. AS OF (B)(6) 2017, THE VAULT IS 409 MICRONS AND THE LENS IS NEARLY VERTICAL. Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS FOUND. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11/+4.5/007 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. AFTER A WEEK, THE LENS ROTATED. ON (B)(6) 2017, THE LENS WAS REPOSITIONED BUT THEN ROTATED AGAIN ONE DAY LATER. THE LENS WAS EXPLANTED ON (B)(6) 2018 AND EXCHANGED WITH A SPHERICAL LENS. THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11/+4.5/007 | OD | Physician | 09/26/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7164791&pc=MTA | 2023826-2018-00023 | 2017-07-30 | Injury | STAAR SURGICAL COMPANY | 2018-01-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VITCMO12.1 IMPLANTABLE COLLAMER LENS, -18.0/+3.0/108 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. LENS ROTATION NOT ASSOCIATED TO LOW VAULT WAS NOTICED. THE LENS WAS THEN REPOSITIONED. IT WAS STATED THAT THAT THE LENS REMAINS INPLANTED, AND THE SURGEON PLANS TO EXCHANGE THE LENS. KERATOCONUS EYE WAS REPORTED AS A PRE-EXISTING CONDITION. Manufacturer Narrative: THE LENS WAS REPOSITIONED ON (B)(6) 2017 AND RESOLVED THE PROBLEM FOR ONLY A MONTH. THE LENS THEN GOT ROTATED AGAIN. (B)(4) Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS) AND THE LENS HAS A CURVED SHAPE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.0/+3.0/108 | OD | Physician | 12/06/2017 | 2017 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7165291&pc=MTA | 2023826-2018-00024 | 2017-07-30 | Injury | STAAR SURGICAL COMPANY | 2018-01-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VITCMO12.1 IMPLANTABLE COLLAMER LENS, -18.0/+3.5/73 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. LENS ROTATION NOT ASSOCIATED TO LOW VAULT WAS NOTICED. THE LENS WAS THEN REPOSITIONED. IT WAS STATED THAT THE LENS REMAINS INPLANTED, AND THE SURGEON PLANS TO EXCHANGE THE LENS. KERATOCONUS EYE WAS REPORTED AS A PRE-EXISTING CONDITION. Manufacturer Narrative: THE LENS WAS REPOSITIONED ON (B)(6) 2017 AND RESOLVED THE PROBLEM FOR ONLY A MONTH. THE LENS THEN GOT ROTATED AGAIN. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -18.0/+3.5/73 | OS | Physician | 12/06/2017 | 2017 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8092554&pc=MTA | 2023826-2018-01711 | 2017-08-26 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+2.0/109 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2017 AND AGAIN ON (B)(6) 2018 BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.(B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.5/+2.0/109 | OS | Physician | 10/25/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7089429&pc=MTA | 2023826-2017-01843 | 2017-08-30 | Injury | STAAR SURGICAL COMPANY | 2017-12-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -11.5/2/93 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS ROTATED AGAIN AND THE LENS WAS THEN EXPLANTED ON (B)(6) 2017. THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND CAUGHT IN THE VIAL CAP AND FOLDED IN AND WHITE RESIDUE WAS ON LENS. (B)(4) Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/2/93 | OS | Physician | 11/08/2017 | 2017 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8129228&pc=MTA | 2023826-2018-01768 | 2017-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -5.50/5.0/019 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)() 2017. LENS ROTATION WAS OBSERVED, REPOSITIONING SURGERY WAS ON (B)(6) 2017 AND THEN AGAIN ON (B)(6) 2017. ON (B)(6) 2018 THE SURGEON OBSERVED THE LENS HAD ROTATED AGAIN. IT IS REPORTED THAT THE SURGEON WILL REMOVE THE LENS AND EXCHANGE WITH A LONGER LENGTH LENS (B)(6) 2019. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THE LENS WAS RECEIVED IN A GLASS VIAL, STUCK ON THE BOTTOM; FILLED VIAL WITH BSS TO BE ABLE TO REMOVE IT. VISUAL INSPECTION FOUND LENS OPTIC SPLIT AND HAPTIC TORN AND BROKEN. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO LENS SPLIT. CLAIM# (B)(4). Manufacturer Narrative: LENS REMOVED AND EXCHANGED ON (B)(6) 2019. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.50/5.0/019 | OS |  | 11/06/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7947531&pc=MTA | 2023826-2018-01483 | 2017-10-19 | Injury | STAAR SURGICAL COMPANY | 2018-10-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -9.00/1.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED WITH NO RESOLVE. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED, THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. CORRECTED DATA: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -9.00/1.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 30-MAY-2017. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON 13-JULY-2018 WITH NO RESOLVE. ON 02-AUG-2018 THE LENS WAS REMOVED AND EXCHANGED, THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.00/1.5/085 | OD | Physician | 10/02/2018 | 2018 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7947537&pc=MTA | 2023826-2018-01484 | 2017-10-19 | Injury | STAAR SURGICAL COMPANY | 2018-10-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -8.00/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2018. ON (B)(6) 2018 THE LENS EXCHANGED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: DEVICE WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH FOREIGN MATERIAL (RESIDUE) ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: CORRECTED DATA: INITIAL MDR STATES THAT LENS WAS IMPLANTED ON "(B)(6) 2018". CORRECT IMPLANT DATE IS (B)(6) 2017. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(6). WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.00/1.0/093 | OS | Physician | 10/02/2018 | 2018 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7322086&pc=MTA | 2023826-2018-00313 | 2017-10-21 | Injury | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2017 BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED WITH A DIFFERENT TYPE OF LENS (DIFFERENT MANUFACTURER). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF ITS HAPTIC AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.0/+2.5/090 | OD | Physician | 02/05/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7156009&pc=MTA | 2023826-2017-01972 | 2017-11-20 | Injury | STAAR SURGICAL COMPANY | 2017-12-29 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/081 (SPHERE/ CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS REPOSITIONED. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS REPORTED. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SAME SIZE, DIFFERENT TYPE OF LENS. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. (B)(4) Manufacturer Narrative: VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND CLEAR AND RED RESIDUE ON LENS. (B)(4) Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.5/+2.0/081 | OS | Physician | 12/07/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8273955&pc=MTA | 2023826-2019-00088 | 2017-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -18.00/3.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2017. ON (B)(6) 2018 THE LENS EXCHANGED AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND RESIDUE AND DEBRIS ON LENS AND THE LENS WAS RETURNED IN A PIECE SMASHED TOGETHER AND WAS UNABLE TO BE EVALUATED. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.00/3.5/082 | OS |  | 01/09/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8182232&pc=MTA | 2023826-2018-01867 | 2017-12-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.5/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER INDICATED LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED BUT IT WAS UNSUCCESSFUL. THE LENS WAS EXPLANTED ON (B)(6) 2018 AND WAS EXCHANGED FOR A DIFFERENT MODEL LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: EVENT: THE REPORTER INDICATED THE LENS ROTATED A 2ND TIME AND THE LENS WAS EXPLANTED. DEVICE EVALUATED BY MFR : DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+2.5/080 | OS |  | 11/22/2018 | 2018 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7740965&pc=MTA | 2023826-2018-01136 | 2018-01-08 | Injury | STAAR SURGICAL COMPANY | 2018-08-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -15/2/88 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. AT ONE MONTH POST-OP LENS ROTATION WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2018 BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A 12.6MM VICMO12.6, -14.0 DIOPTER AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS SURFACE. CORRECTED DATA: (B)(4). CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4) - SECONDARY SURGICAL INTERVENTION, EXCHANGE. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15/2/88 | OS | Physician | 07/04/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7557877&pc=MTA | 2023826-2018-00823 | 2018-01-30 | Injury | STAAR SURGICAL COMPANY | 2018-05-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.5/+3.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON NOTED LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REPOSITIONED THE LENS ON (B)(6) 2018 AND (B)(6) 2018. REPORTEDLY, THE LENS ROTATED IN THE SAME DIRECTION AND THE SAME AMOUNT REPEATEDLY. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT DIOPTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATIONS: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, LENS REPOSITIONING). CONCLUSION CODE: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))] (B)(4). | -12.5/+3.5/091 | OD | Physician | 05/09/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8182114&pc=MTA | 2023826-2018-01866 | 2018-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+3.0/099 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED BUT IT WAS UNSUCCESSFUL. THE LENS WAS EXPLANTED ON (B)(6) 2018 AND WAS EXCHANGED FOR A DIFFERENT MODEL LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: EVENT: THE REPORTER INDICATED THE LENS ROTATED A 2ND TIME AND THE LENS WAS EXPLANTED. DEVICE EVALUATED BY MFR: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC DEFORMED. CORRECTED DATA: DATE OF EVENT - (B)(6)2018. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/+3.0/099 | OD |  | 11/22/2018 | 2018 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7571620&pc=MTA | 2023826-2018-00844 | 2018-02-19 | Injury | STAAR SURGICAL COMPANY | 2018-06-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.0/+4.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTED LENS ROTATION. REPORTEDLY, THE LENS WAS REPOSITIONED ON (B)(6) 2018 BUT THAT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.0/+4.0/091 | OS | Physician | 05/09/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7518808&pc=MTA | 2023826-2018-00733 | 2018-02-26 | Injury | STAAR SURGICAL COMPANY | 2018-05-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.0/+2.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD ROTATED (NOT ASSOCIATED TO A LOW VAULT) AND WAS REPOSITIONED ON (B)(6) 2018 BUT IT WAS UNSUCCESSFUL. THE LENS WAS EXPLANTED ON (B)(6) 2018 AND WAS EXCHANGED FOR A LONGER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A PIECE OF ONE HAPTIC MISSING. THERE WAS EVIDENCE OF RESIDUE ON THE LENS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.0/+2.5/090 | OD | Physician | 04/17/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7639920&pc=MTA | 2023826-2018-00963 | 2018-03-20 | Injury | STAAR SURGICAL COMPANY | 2018-06-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LENS ROTATION NOT ASSOCIATED WITH A LOW VAULT, THE SURGEON ALSO NOTING SURGICAL ENLARGEMENT OF PI. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, SURGICAL ENLARGEMENT OF PI) LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/+2.0/095 | OD | Physician | 06/01/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7947546&pc=MTA | 2023826-2018-01486 | 2018-03-26 | Injury | STAAR SURGICAL COMPANY | 2018-10-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: HE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.00/3.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2018. ON (B)(6) 2018 THE LENS EXCHANGED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/3.0/089 | OD | Physician | 10/02/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7571596&pc=MTA | 2023826-2018-00846 | 2018-04-20 | Injury | STAAR SURGICAL COMPANY | 2018-06-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+3.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 AND (B)(6) 2018 THE LENS WAS OBSERVED TO HAVE ROTATED 90 DEGREES. AT EACH TIME OF OBSERVATION, THE LENS WAS REPOSITIONED. ON (B)(6) 2018, THE LENS HAD AGAIN ROTATED AND THEREFORE WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, LENS REPOSITIONING). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17.0/+3.0/090 | OD | Physician | 05/11/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7704671&pc=MTA | 2023826-2018-01052 | 2018-04-29 | Injury | STAAR SURGICAL COMPANY | 2018-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.5/+6.0/169 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS ROTATED (NOT ASSOCIATED TO A LOW VAULT). THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED. THE REPORTER INDICATED THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS BUT TO DATE THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO WRONG MEASUREMENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.5/+6.0/169 | OS | Physician | 06/19/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8062286&pc=MTA | 2023826-2018-01620 | 2018-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -14.00/3.5/071, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE REPOSITIONING DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2018 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATED BY MFR-DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.00/3.5/071 | OS |  | 10/25/2018 | 2018 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7740027&pc=MTA | 2023826-2018-01145 | 2018-05-18 | Injury | STAAR SURGICAL COMPANY | 2018-08-01 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1, -16.5/+1.5/096 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS RE-POSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS. REPORTEDLY, THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.5/+1.5/096 | OD | Physician | 07/06/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7765094&pc=MTA | 2023826-2018-01165 | 2018-05-18 | Injury | STAAR SURGICAL COMPANY | 2018-08-08 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM12.6, -16.5/+3.5/010 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS RE-POSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: "THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL, LONGER LENGTH LENS ON (B)(6) 2019. THE EXCHANGE RESOLVED THE PROBLEM." EXPLANT DATE: " (B)(4) 2019". DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.5/+3.5/010 | OD | Physician | 07/13/2018 | 2018 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8337581&pc=MTA | 2023826-2019-00252 | 2018-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -3.0/+3.5/154 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A SAME MODEL BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2018. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -3.0/+3.5/154 | OD |  | 01/14/2019 | 2019 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721253&pc=MTA | 2023826-2018-01072 | 2018-05-31 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -16.5/+2.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS, THE SURGEON NOTING LENS ROTATION. ON (B)(6) 2018 THE LENS HAD BEEN REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED. AFTER THE LENS EXCHANGE, THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS DEVICE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, LENS REPOSITIONING). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.5/+2.5/095 | OS | Physician | 06/28/2018 | 2018 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8229450&pc=MTA | 2023826-2019-00026 | 2018-06-04 | Injury | STAAR SURGICAL COMPANY | 2019-01-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/071 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS, WAS ROTATED 90 DEGREES AXIALLY AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.5/+2.0/071 | OS | Physician | 12/09/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824394&pc=MTA | 2023826-2018-01254 | 2018-07-11 | Injury | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.50/+1.50/113 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2018 BUT WAS NOT SUCCESSFUL. THE LENS WAS EXCHANGED FOR ANOTHER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DATE OF EVENT: (B)(6) 2018. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -8.50/+1.50/113 | OS | Physician | 07/29/2018 | 2018 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8229339&pc=MTA | 2023826-2019-00024 | 2018-07-13 | Injury | STAAR SURGICAL COMPANY | 2019-01-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.5/+2.0/084 | OS | Physician | 12/09/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8092574&pc=MTA | 2023826-2018-01712 | 2018-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.0/+4.0/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WASRESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATED BY MFR -DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.0/+4.0/084 | OS |  | 10/25/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8352267&pc=MTA | 2023826-2019-00290 | 2018-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICMO12.6, -13.50/3.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SAME LENGTH, NON-TORIC LENS DUE TO OBSERVATION OF LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. IT WAS REPORTED THAT THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.50/3.0/083 | OS |  | 01/23/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8129207&pc=MTA | 2023826-2018-01728 | 2018-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.50/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2018 AND IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND. CORRECTED DATA: DATE REC'D BY MFR - DATE IN PREVIOUS MDR SHOULD BE CORRECTED TO "07-JAN-2019." (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATED BY MFR - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/2.0/093 | OD |  | 11/06/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7916607&pc=MTA | 2023826-2018-01452 | 2018-08-28 | Injury | STAAR SURGICAL COMPANY | 2018-09-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICH12.6, -5.50/5.5/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND LENS DISLOCATION/SUBLUXATION WAS OBSERVED. ON (B)(6) 2018 THE SURGEON REPOSTION THE LENS, THE REPOSITIONING DID NOT RESOLVE THE PROBLEM. IT WAS REPORTED THAT THE SURGEON PLANS TO EXCHANGE TO A LONGER LENS. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED, NO EXCHANGE TILL NOW. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: INITIAL MDR STATES DIOPTER IS "-5.50/5.5/109 (SPHERE/CYLINDER/AXIS)"; CORRECT LENS DIOPTER IS "+5.5/5.5/109 (SPHERE/CYLINDER/AXIS)". (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.50/5.5/109 | OS | Physician | 08/29/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8172143&pc=MTA | 2023826-2018-01848 | 2018-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.5/+3.0/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+3.0/080 | OS |  | 11/23/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8129043&pc=MTA | 2023826-2018-01782 | 2018-09-28 | Injury | STAAR SURGICAL COMPANY | 2018-12-04 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.5/+2.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS RE-POSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SAME SIZE LENS, DIFFERENT AXIS LENS. REPORTEDLY, THE LENS EXCHANGE RESOLVED THE PROBLEM. THE SURGEON STATES "VERTICAL IMPLANTATION OF THE SAME SIZE ICL, PROBLEM SOLVED." THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS; VISUAL INSPECTION FOR NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+2.5/087 | OD | Physician | 11/09/2018 | 2018 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372527&pc=MTA | 2023826-2019-00317 | 2018-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.5/+3.0/114 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO LENS ROTATION. ON (B)(6) 2018 A SAME SIZE, DIFFERENT MODEL LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.5/+3.0/114 | OS |  | 01/31/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8092604&pc=MTA | 2023826-2018-01713 | 2018-10-12 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.5/+2.5/089 | OS | Physician | 10/25/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8274639&pc=MTA | 2023826-2019-00175 | 2018-10-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM,VTICM5_12.6, -12.50/4.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.50/4.0/093 | OD |  | 12/26/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8307832&pc=MTA | 2023826-2019-00221 | 2018-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+5.5/109 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED; THE LENS IS "STABLE," THE PATIENT IS "HAPPY." Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A TEAR IN THE LENS HAPTIC AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: THE LENS RETURNED TO THE MANUFACTURER HAS THE WRONG SERIAL NUMBER ON THE RETURNED PACKAGE. THE RETURNED LENS DOES NOT BELONG TO CLAIM (B)(4). DEVICE AVAILABLE FOR EVALUATION: CHANGED FROM YES TO NO FOR SUPPLEMENTAL MDR#1. DEVICE EVALUATION REPORT NOT APPLICABLE FOR THIS CLAIM. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/+5.5/109 | OD |  | 01/10/2019 | 2019 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8227442&pc=MTA | 2023826-2019-00055 | 2018-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.0/+3.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | -10.0/+3.5/091 | OD |  | 12/14/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8266784&pc=MTA | 2023826-2019-00015 | 2018-11-25 | Injury | STAAR SURGICAL COMPANY | 2019-01-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5 12.6, -5.50/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2018. ON (B)(6) 2018 THE LENS REMOVED AND EXCHANGED WITH A LONGER LENS. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.50/1.0/090 | OS | Physician | 12/25/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372524&pc=MTA | 2023826-2019-00319 | 2018-12-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.0/+1.0/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -13.0/+1.0/080 | OD |  | 02/01/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8352212&pc=MTA | 2023826-2019-00289 | 2018-12-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICMO12.6, -13.50/3.0/094 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SAME LENGTH, NON-TORIC LENS DUE TO OBSERVATION OF LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. IT WAS REPORTED THAT THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: LENS WAS RETURNED IN DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.50/3.0/094 | OD |  | 01/23/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372518&pc=MTA | 2023826-2019-00320 | 2018-12-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO_12.1 IMPLANTABLE COLLAMER LENS, -15.5/+5.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC BENT AND RESIDUE ON LENS SURFACE. CORRECTED DATA: EXPLANTED DATE CORRECTED FROM "(B)(6) 2019" TO "(B)(6) 2019" FOR INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.5/+5.5/090 | OS |  | 02/01/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8595421&pc=MTA | 2023826-2019-00768 | 2019-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.5/+3.0/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -12.5/+3.0/084 | OD |  | 04/12/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8092497&pc=MTA | 2023826-2018-01708 |  | Injury | STAAR SURGICAL COMPANY | 2018-11-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00/+2.5/092 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2018. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY; IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND ONE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: EVENT-ADDITIONAL INFORMATION: THE EXCHANGED RESOLVED THE ISSUE. THE DOCTOR DECIDED TO CORRECT ASTIGMATISM WITH PRK TREATMENT LASER. PATIENT CODE: (B)(4). CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.00/+2.5/092 | OD |  | 10/22/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8092515&pc=MTA | 2023826-2018-01709 |  | Injury | STAAR SURGICAL COMPANY | 2018-11-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00/+2.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2018. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY; IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND ONE HAPTIC BROKEN. CORRECTED DATA: CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: EVENT-REPORTER STATED DIFFERENT IMPLANT AND REPLACEMENT DATES: "THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMR LENS -11.00/+2.5/093 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS) ON (B)(6)2018. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6)2018 DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT". THE REPORTER STATED THAT "THE LENS WAS EXCHANGED FOR A SPHERIC LENS AND THE DOCTOR DECIDED TO DO PRK TREATMENT LASER". THE PROBLEM WAS RESOLVED. PATIENT CODE: (B)(4). CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.00/+2.5/093 | OS |  | 10/22/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5506331&pc=MTA | 2023826-2016-00217 | 2015-12-16 | Injury | STAAR SURGICAL COMPANY | 2016-03-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -16.00/+6.00/X022 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE SURPRISE WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016. LENS WAS NOT REPLACED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT'S POST-OP (ON (B)(6) 2015) BEST CORRECTED VISUAL ACUITY WAS 20/40 AND UNCORRECTED VISUAL ACUITY WAS 20/200. DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS WITH A KERATOCONUS EYE. CORRECTED DATA: CHANGE "VTICMO 13.2 -16.00/+6.00/X022" TO "VTICMO 13.2 -15.00/+6.00/X022". (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -16.00/+6.00/X022. DEVICE EVALUATED BY MANUFACTURER? NO. EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WAS NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -16.00/+6.00/022 | OS | Physician | 02/15/2016 | 2016 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5583325&pc=MTA | 2023826-2016-00451 |  | Injury | STAAR SURGICAL COMPANY AG | 2016-04-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated; Positioning Problem; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 -11.00/+06.00/X058 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2008. THERE WAS A LENS ROTATION AND LENS REPOSITIONING. THE PATIENT EXPERIENCED A SECOND LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. LENS REMAINS IMPLANTED. PLANNING TO EXCHANGE LENS. LAST VISIT ON (B)(6) 2016, UCVA WAS 20/300. SEE MFR. #2023826-2016-00450 FOR RIGHT EYE. Manufacturer Narrative: PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. CONCLUSION: PATIENT HAD CATARACT IN THE OPERATIVE EYE OR NONTRAUMATIC CATARACT IN THE FELLOW EYE, THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE (B)(4). BUT NOT MARKETED IN THE U.S. DIOPTER: -11.00/+06.00/X058. DEVICE EVALUATED BY MANUFACTURER: NO, DEVICE REMAINS IMPLANTED. (B)(4). CONCLUSIONS: BASED ON THE COMPLAINT INFORMATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -11.00/+06.00/058 | OS | Physician | 03/17/2016 | 2016 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6519202&pc=MTA | 2023826-2017-00600 | 2015-09-25 | Injury | STAAR SURGICAL COMPANY | 2017-04-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+4.0/89 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNTING OR SIGNIFICANT ENDOTHELIAL CELL LOSS. THE LENS WAS NOT EXCHANGED. THE PATIENT'S POST-OP (ON (B)(6) 2015) BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN, BROKEN, AND BENT. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.5/+4.0/89 | OS | Physician | 10/20/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6473526&pc=MTA | 2023826-2017-00609 | 2015-09-25 | Injury | STAAR SURGICAL COMPANY AG | 2017-04-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.0/+3.5/91 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNTING OR SIGNIFICANT ENDOTHELIAL CELL LOSS. THE LENS WAS NOT EXCHANGED. THE PATIENT'S POST-OP (ON (B)(6) 2015) BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE CORRECT DATE FOR THE INITIAL IS 10/20/2015. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THIS PRODUCT WAS NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.0/+3.5/91 | OD | Physician | 10/20/2015 | 2015 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6606817&pc=MTA | 2023826-2017-00919 | 2015-11-10 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-01 | MTA | INTRAOCULAR COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Code Available |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+1.5/93 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS TORE/BROKE DURING LOADING. THE LENS WAS NOT IMPLANTED. ON (B)(6) 2015, THE LENS WAS EXCHANGED FOR ANOTHER SIMILAR LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FIBERS ON LENS SURFACE. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17.0/+1.5/93 | OD | Physician | 08/02/2016 | 2016 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6365381&pc=MTA | 2023826-2017-00288 | 2015-12-23 | Injury | STAAR SURGICAL COMPANY | 2017-02-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -14.5/1.5/169 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. AS OF (B)(6) 2017, THE LENS IS STILL NOT EXCHANGED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A SMALL PLASTIC VIAL. VISUAL INSPECTION FOUND LENS COVERED WITH CLEAR, STICKY RESIDUE AND THE HAPTIC WITH THREE MISSING PIECES. (B)(4). Manufacturer Narrative: CONCLUSION CODE: PER DFU CONTRAINDICATIONS VICL'S IS CONTRAINDICATED IN THE PRESENCE OF ANY CATARACT IN THE OPERATIVE EYE OR NON-TRAUMATIC CATARACT IN THE FELLOW EYE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.5/1.5/169 | OD | Physician | 02/01/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6592262&pc=MTA | 2023826-2017-00827 | 2016-01-22 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON TRIED TO IMPLANT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/+2.5/73 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) AND THE LENS TORE/BROKE BEFORE INJECTION INTO THE EYE. THE LENS WAS NOT IMPLANTED. PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY ON (B)(6) 2016 WAS 20/20. Manufacturer Narrative: THIS DEVICE IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT, HAPTIC TORN/BROKEN/BENT/DEFORMED, FOREIGN MATERIAL ON LENS SURFACE (LENS BROKEN IN TWO PIECES). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -17.50/+2.5/73 | OD | Physician | 04/19/2016 | 2016 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6370399&pc=MTA | 2023826-2017-00294 | 2016-04-21 | Injury | STAAR SURGICAL COMPANY | 2017-03-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12/1/046 DIOPTER, WAS NOT IMPLANTED DUE TO A TEAR/BREAK BEFORE INJECTION/DELIVERY OF THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR A NEW LENS ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN THE LENS VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS OPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12/1/046 | OS | Physician | 02/01/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6598690&pc=MTA | 2023826-2017-00828 | 2016-05-04 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/92 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) AND THE LENS TORE/BROKE BEFORE INJECTION INTO THE EYE. THE LENS WAS NOT IMPLANTED. Manufacturer Narrative: THIS DEVICE IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS). (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4). | -18.0/+2.0/92 | OD | Physician | 07/11/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6597936&pc=MTA | 2023826-2017-00814 | 2016-05-15 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/1.0/082 DIOPTER, IN THE PATIENT'S LEFT EYE (OS). LENS TEAR/BREAK BEFORE INJECTION INTO THE EYE WAS NOTED. THE LENS WAS NOT IMPLANTED. THE LENS WAS EXCHANGED FOR SIMILAR LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/1.0/082 | OS | Physician | 06/22/2016 | 2016
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6370433&pc=MTA | 2023826-2017-00297 | 2016-07-08 | Injury | STAAR SURGICAL COMPANY | 2017-03-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/2.5/112 DIOPTER, WAS NOT IMPLANTED DUE TO A TEAR/BRREAK BEFORE INJECTION INTO THE EYE. THE SURGEON IMPLANTED A NEW LENS OF THE SAME MODEL AND SIMILAR DIOPTER ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF THE LENS OPTIC AND ONE HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4). | -16.5/2.5/112 | NA | Physician | 02/01/2017 | 2017 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6271081&pc=MTA | 2023826-2017-00131 | 2016-09-01 | Malfunction | STAAR SURGICAL COMPANY | 2017-01-23 | MTA | INTRAOCULAR LENS | Device Operates Differently Than Expected | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -2.0/1.0/120 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS). THIS WAS PERFORMED ON (B)(6) 2016. DURING INJECTION/DELIVERY INTO THE EYE, THE PHYSICIAN NOTICED THE LENS WAS TORN/BROKE. THE LENS WAS THEN EXPLANTED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED AND EVALUATED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY, IN THE LENS CONTAINER, WITH FIBER MATERIAL ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC AND OPTIC TORN. PATIENT WEIGHT-UNK. PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -2.0/1.0/120 | OS | Physician | 09/01/2016 | 2016 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6420620&pc=MTA | 2023826-2017-00445 | 2016-09-30 | Injury | STAAR SURGICAL COMPANY | 2017-03-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+4.0/54 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED WITH LOW VAULT. ON (B)(6) 2016 THE LENS WAS REPOSITIONED AND THEN EXCHANGED WITH A LONGER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DR IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC PIECE MISSING, WHITE AND CLEAR RESIDUE ON LENS, AND THE LENS CURVED UP. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.5/+4.0/54 | OS | Physician | 02/21/2017 | 2017 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6339264&pc=MTA | 2023826-2017-00188 | 2016-12-21 | Injury | STAAR SURGICAL COMPANY | 2017-02-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.00/ 25/ 091 DIOPTER, TORE/BROKE DURING THE INJECTION/DELIVERY ON (B)(6) 2016. THE LENS WAS REMOVED ON (B)(6) 2016 AND IMPLANTED WITH ANOTHER LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM WHICH WAS CONFIRMED DURING THE PATIENT 'S LAST VISIT ON (B)(6) 2016 WITH UCVA 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE US BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.00/ 25/ 091 | NA | Physician | 01/25/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6637202&pc=MTA | 2023826-2017-00970 | 2017-01-07 | Injury | STAAR SURGICAL COMPANY AG | 2017-06-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -9.0/+4.5/178 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNTING OR SIGNIFICANT ENDOTHELIAL CELL LOSS. ON (B)(6) 2017 THE LENS WAS EXPLANTED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FIBERS ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.0/+4.5/178 | OD | Physician | 05/15/2017 | 2017 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6637213&pc=MTA | 2023826-2017-00971 | 2017-01-07 | Injury | STAAR SURGICAL COMPANY AG | 2017-06-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/174 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNTING OR SIGNIFICANT ENDOTHELIAL CELL LOSS. ON (B)(6) 2017 THE LENS WAS EXPLANTED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. LENS WAS RETURNED DRY IN A SMALL VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.5/+1.5/174 | OS | Physician | 05/15/2017 | 2017 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6447049&pc=MTA | 2023826-2017-00526 | 2017-01-26 | Injury | STAAR SURGICAL COMPANY | 2017-03-30 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.5/093 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD). THIS WAS PERFORMED ON (B)(6) 2017. DURING INJECTION/DELIVERY, THE LENS TORE/BROKE. ON THE SAME DAY THE LENS WAS EXCHANGED FOR A LENS OF THE SAME TYPE, SIZE, AND DIOPTER. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+2.5/093 | OD | Physician | 02/28/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6361176&pc=MTA | 2023826-2017-00263 | 2017-01-30 | Injury | STAAR SURGICAL COMPANY AG | 2017-02-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -3/-4/007 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE REPORTER STATED REFRACTIVE CHANGE OVERTIME WAS OBSERVED ON (B)(64) 2013. KERARING / INTACS WAS PERFORMED AND THE LENS REMAINS IMPLANTED AS OF (B)(6) 2017. Manufacturer Narrative: SECONDARY SURGERY, LENS EXCHANGE FOR A DIFFERENT MODEL AND DIFFERENT DIOPTER BUT SAME SIZE, AND THE PROBLEM WAS RESOLVED. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -3/-4/7 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE REPORTER STATED A REFRACTIVE CHANGE OVER TIME WAS OBSERVED ON (B)(6) 2017. LENS WAS EXCHANGED FOR A DIFFERENT LENS MODEL AND DIOPTER ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: (B)(4). THIS PRODUCT IS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). LENS NOT RETURNED. | -3/-4/007 | OD | Physician | 01/31/2017 | 2017 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6437977&pc=MTA | 2023826-2017-00483 | 2017-02-06 | Injury | STAAR SURGICAL COMPANY | 2017-03-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/085 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR THE SAME SIZE AND SIMILAR DIOPTER LENS DUE TO TEAR/BREAK DURING INJECTION INTO THE EYE. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND BLACK DEBRIS ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). RESULT: WORK ORDER SEARCH PERFORMED: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4). | -13.5/+2.0/085 | OD | Physician | 02/26/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6503624&pc=MTA | 2023826-2017-00670 | 2017-02-24 | Injury | STAAR SURGICAL COMPANY | 2017-04-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.5/1.5/038 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. IN THE SAME SURGERY, THE LENS WAS EXPLANTED DUE TO LENS TEAR/BREAK DURING INJECTION INTO THE EYE. THE LENS WAS THEN EXCHANGED FOR THE SAME SIZE AND DIOPTER LENS ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THE CUSTOMER STATED THAT THE LENS TORE BECAUSE AT THE TIME OF INJECTION, IT GOT CAUGHT IN PLUNGER AND DAMAGED. THE SURGEON COULD SEE THE BREAK AFTER IMPLANTING INTO THE EYE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: ON FOLLOW UP #2, THE DATE RECEIVED BY MFR IS INCORRECTLY LISTED AS 07-JUL-2017. THE CORRECT DATE SHOULD BE 29-JUN-2017. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION: OFF LABEL USE (ACD<3.0MM)). (B)(4). | -17.5/1.5/038 | OS | Physician | 03/21/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6661531&pc=MTA | 2023826-2017-01002 | 2017-03-03 | Injury | STAAR SURGICAL COMPANY | 2017-06-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+3.0/089 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD), IT TORE/BROKE DURING INJECTION. THIS OCCURRED ON (B)(6) 2017 AND THE LENS WAS IMMEDIATELY EXPLANTED. THE CAUSE FOR THE TEAR/BREAK IS UNKNOWN AT THIS TIME. ON (B)(6) 2017 A REPLACEMENT LENS OF THE SAME SIZE AND SIMILAR DIOPTER WAS IMPLANTED SUCCESSFULLY. Manufacturer Narrative: PREVIOUS: THIS OCCURRED ON (B)(6) 2017 AND THE LENS WAS IMMEDIATELY EXPLANTED. THE CAUSE FOR THE TEAR/BREAK IS UNKNOWN AT THIS TIME. ON (B)(6) 2017 A REPLACEMENT LENS OF THE SAME SIZE AND SIMILAR DIOPTER WAS IMPLANTED SUCCESSFULLY. CORRECTED: THIS OCCURRED ON (B)(6) 2017 AND THE LENS WAS EXPLANTED IN THE SAME DAY. THE CAUSE FOR THE TEAR/BREAK IS UNKNOWN AT THIS TIME. ON (B)(6) 2017 A REPLACEMENT LENS OF THE SAME SIZE AND SIMILAR DIOPTER WAS IMPLANTED SUCCESSFULLY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION-LENS WAS RETURNED DRY IN A SMALL VIAL, WITH CLEAR RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS WAS TORN INTO TWO PIECES. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+3.0/089 | OD | Physician | 05/26/2017 | 2017 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6616433&pc=MTA | 2023826-2017-00930 | 2017-04-20 | Injury | STAAR SURGICAL COMPANY | 2017-06-06 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON ATTEMPTED TO IMPLANT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+2.0/092 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE SURGEON STATED THAT AS HE WAS INJECTING THE LENS INTO THE EYE, THE LENS GOT STUCK ("ENTANGLED") ON THE PLUNGER. THIS OCCURRED ON (B)(6) 2017 AND ALTHOUGH THE LENS WAS NOT COMPLETELY IMPLANTED, THERE WAS PATIENT CONTACT WITH THE LENS. ON (B)(6) 2017, AN ALTERNATE LENS WAS IMPLANTED-OF THE SAME TYPE AND SIZE, AND SIMILAR DIOPTER. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION - LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS WITH HAPTIC TEARS, THE LENS TORN INTO TWO PIECES AND CLEAR RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE NOT RETURNED TO MANUFACTURER. | -12.50/+2.0/092 | OS | Physician | 05/09/2017 | 2017 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6584362&pc=MTA | 2023826-2017-00856 | 2017-04-23 | Injury | STAAR SURGICAL COMPANY | 2017-05-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/40 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. IN THE SAME SURGERY, THE LENS WAS EXPLANTED DUE TO LENS TEAR/BREAK DURING INJECTION INTO THE EYE. THE CUSTOMER STATED THAT THE SURGEON FOUND THE LENS BROKE SINCE IT WAS IMPLANTED FULLY UNFOLD IN THE PATIENT EYE. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.0/+1.5/40 | OD | Physician | 04/28/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6680914&pc=MTA | 2023826-2017-01078 | 2017-04-27 | Injury | STAAR SURGICAL COMPANY | 2017-06-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/2.00/86 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR THE SAME SIZE AND DIOPTER LENS DUE TO LENS TEAR/BEAK DURING INJECTION INTO THE EYE. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: LENS WAS RETURNED DRY, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND PIECE OF HAPTIC MISSING, CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: DATE OF EVENT: IT SHOULD BE CORRECTED TO (B)(6) 2017. PREVIOUS STATED "FOLLOW UP" #2 BUT IT WAS "FOLLOW UP" #1. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR THE UNIT WITHIN THE SAME LOT. CLAIM#(B)(4). | -11.5/2.00/86 | OS | Physician | 06/04/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6637773&pc=MTA | 2023826-2017-00973 | 2017-05-11 | Injury | STAAR SURGICAL COMPANY | 2017-06-13 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.00/3.5/87 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE SURGEON STATED THAT THE LENS "GOT CRACKED" DURING THE LOADING PROCESS AND THE "POST-OP OUTCOME WAS NOT GOOD." THE LENS WAS IMPLANTED ON (B)(6) 2017. ON (B)(6) 2017, AN ALTERNATE LENS WAS IMPLANTED-OF THE SAME TYPE AND SIZE, AND SIMILAR DIOPTER. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC SPLIT AND HAPTIC BROKEN/BENT. (B)(4). Manufacturer Narrative: PATIENT WEIGHT-UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | -16.00/3.5/87 | OS | Physician | 05/16/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6673571&pc=MTA | 2023826-2017-01055 | 2017-05-23 | Injury | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -9.0/+2.5/93 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. DURING IMPLANTATION, THE LENS WAS CAUGHT IN THE FOAM TIP PLUNGER DURING THE PROCESS OF LOADING AND PUSHING IT INTO THE EYE AND IT TORE. AFTER APPROXIMATELY AN HOUR AND A HALF, THE LENS WAS EXCHANGED WITH A SIMILAR LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: PREVIOUS REPORT STATEMENT "SINGLE-USE DEVICE WAS REPROCESSED AND REUSED ON A PATIENT" IS INCORRECT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND THE LENS AND HAPTIC TORN AND A PIECE OF THE LENS MISSING. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.0/+2.5/93 | OD | Physician | 06/05/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6711499&pc=MTA | 2023826-2017-01174 | 2017-05-23 | Injury | STAAR SURGICAL COMPANY | 2017-07-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, DIOPTER -14.5/2.0/098 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. IN THE SAME DAY, THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL LENS DUE TO LENS TEAR/BREAK DURING INJECTION INTO THE EYE. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND CLEAR RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN US BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.5/2.0/098 | OD | Physician | 06/19/2017 | 2017 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6729084&pc=MTA | 2023826-2017-01230 | 2017-06-23 | Injury | STAAR SURGICAL COMPANY | 2017-07-20 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/137 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE "AFTER THE LENS WAS INJECTED INTO THE EYE, THE DOCTOR FOUND THAT THE OPTICAL ZONE WAS BROKEN, SO HE EXPLANT THE LENS AND RESERVED A NEW ONE. DOCTOR LI DIDN'T KNOW WHAT CAUSED THE DAMAGE." Manufacturer Narrative: A PRODUCT EVALUATION WAS COMPLETED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIMILAR DIOPTER LENS ON (B)(6) 2017. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.0/+1.5/137 | OD | Physician | 06/28/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6850694&pc=MTA | 2023826-2017-01417 | 2017-07-09 | Injury | STAAR SURGICAL COMPANY | 2017-09-07 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -13.0/+1.0/090 INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE SURGEON IMPLANTED AND EXPLANTED THE LENS ON (B)(6) 2017. ON (B)(6) 2017 AN ALTERNATE LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT. THE LENS WAS RETURNED IN TWO PIECES. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+1.0/090 | OS | Physician | 08/09/2017 | 2017 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7346995&pc=MTA | 2023826-2018-00428 | 2017-10-03 | Injury | STAAR SURGICAL COMPANY | 2018-03-16 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.5/+1.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. PER THE REPORTER THE PUPIL CONSTRICTED RIGHT AFTER ICL IMPLANTATION. AS PUPIL EXTENSION WAS ABSORBED AT 12 DAYS POST OP, THE SURGEON APPLIED SANPILO (PILOCARPINE HYDROCHLORIDE) TO CONSTRICT THE PUPIL. AT A MEDICAL CHECK ON (B)(6) 2017, THE PUPIL WAS NOT CONSTRICTED SO THE SURGEON ADDITIONALLY APPLIED SANPILO. THE RIGHT EYE (OD) WAS OBSERVED TO "REACT BY THE STRONG DIAGNOSTIC EQUIPMENT AND PUPIL CONSTRICTED". PER THE REPORTER THE PUPIL ISSUE WAS RESOLVED AND THE LENS REMAINS IMPLANTED WITHOUT ISSUES. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.5/+1.0/086 | OD | Physician | 11/29/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7180671&pc=MTA | 2023826-2018-00031 | 2017-10-11 | Injury | STAAR SURGICAL COMPANY | 2018-01-10 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -17.00/+1.50/150 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTES ANISOCORIA, A REACTIVE BUT "LARGER" PUPIL, SHALLOW ANTERIOR CHAMBER, "IRIS PIGMENTS ON ICL." REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: "THE LATEST UPDATES THAT THERE IS NO PROBLEM WITH THE PATIENT (NORMAL VAULT, NORMAL IOP) BUT THE ONLY COMPLAIN IS ABOUT THE IRIS SHAPE JUST." (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))]. (B)(4). | -17.00/+1.50/150 | OD | Physician | 12/12/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7073566&pc=MTA | 2023826-2017-01821 | 2017-10-23 | Injury | STAAR SURGICAL COMPANY | 2017-11-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+3.0/082 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. INTRAOPERATIVELY, THE LENS WAS EXCHANGED WITH "ANOTHER TYPE" OF IOL LENS. THE SURGEON REPORTS A CAPSULAR TEAR AND PERFORMANCE OF PHACOEMULSIFICATION ALONG WITH THE LENS EXCHANGE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.94MM)] (B)(4). | -9.0/+3.0/082 | OS | Physician | 11/03/2017 | 2017 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7245231&pc=MTA | 2023826-2018-00202 | 2017-10-30 | Injury | STAAR SURGICAL COMPANY | 2018-02-06 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/081 (SPHERE/CYLINDAR/AXIS) INTO THE PATIENT'S LEFT EYE (OS). ON THE SAME DATE, THE LENS WAS EXPLANTED, THE SURGEON NOTING LENS TEAR/BREAK DURING INJECTION INTO THE EYE. IT IS UNKNOWN WHETHER THE LENS WAS EXPLANTED INTRA-OPERATIVELY OR IN A 2ND SURGERY. ON (B)(6) 2017, AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF HAPTIC AND FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND RESIDUE ON LENS SURFACE). (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.0/+1.5/081 | OS | Physician | 01/11/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7286001&pc=MTA | 2023826-2018-00245 | 2018-01-16 | Injury | STAAR SURGICAL COMPANY | 2018-02-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -14/1.5/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. IT WAS NOTICED THAT THE LENS TORE DURING INJECTION. THE LENS WAS EXPLANTED ON THE SAME DAY. ATTEMPTS AT OBTAINING ADDITION INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: PATIENT STATUS IS "IN GOOD CONDITION". RETURNED PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON LENS. VISUAL INSPECTION FOUND TEARS IN THE HAPTIC AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT UNDER THE AGE OF (B)(6) YEARS OLD. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | -14/1.5/93 | OD | Physician | 01/23/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7275472&pc=MTA | 2023826-2018-00239 | 2018-01-17 | Injury | STAAR SURGICAL COMPANY | 2018-02-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -6.5/+1.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL AND THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED. CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). | -6.5/+1.0/087 | OS | Physician | 01/21/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7301460&pc=MTA | 2023826-2018-00274 | 2018-01-18 | Injury | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.0/+1.5/095 | OS | Physician | 01/29/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7294641&pc=MTA | 2023826-2018-00271 | 2018-01-23 | Injury | STAAR SURGICAL COMPANY | 2018-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+3.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THERE WAS NO REPORT OF ANY APPARENT INJURY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF HAPTIC, THE HAPTIC BENT, AND RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.0/+3.0/093 | OS | Physician | 01/27/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7301478&pc=MTA | 2023826-2018-00294 | 2018-01-27 | Injury | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0/+3.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. AT ONE DAY POST-OP THE INTRAOCULAR PRESSURE WAS 14MMHG. Manufacturer Narrative: THE PATIENT IS IN GOOD CONDITION. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN, HAPTIC TORN AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4). | -16.0/+3.0/094 | OS | Physician | 01/29/2018 | 2018 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7580667&pc=MTA | 2023826-2018-00866 | 2018-03-06 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -3.5/1.0/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. OBSERVATION OF EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS REPORTED. THE REPORTER STATED THAT THE LENS REMAINS IMPLANTED AND THAT THE PATIENT IS FINE. Manufacturer Narrative: CORRECTED DATA: DIOPTER CORRECTED TO -9.5/1.0/92 (SPHERE/CYLINDER/AXIS) IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -3.5/1.0/92 | OD | Other | 05/09/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7607794&pc=MTA | 2023826-2018-00918 | 2018-03-08 | Injury | STAAR SURGICAL COMPANY | 2018-06-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+2.0/035 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THERE WAS NO REPORT OF ANY PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THEY FELT THE CAUSE OF THE EVENT WAS UNKNOWN. AT POST-OP VISIT, THE PATIENTS EYE WAS PERFECTLY QUIET AND REFRACTION AS EXPECTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4) | -15.5/+2.0/035 | OS | Physician | 05/28/2018 | 2018 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7571708&pc=MTA | 2023826-2018-00801 | 2018-04-13 | Injury | STAAR SURGICAL COMPANY | 2018-06-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/115 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LENS TEAR DURING INJECTION INTO THE EYE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. | -11.0/+1.5/115 | OD | Physician | 05/06/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6370389&pc=MTA | 2023826-2017-00293 |  | Injury | STAAR SURGICAL COMPANY | 2017-03-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12/3.5/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO A LENS TEAR DURING THE SURGERY. THE LENS WAS EXCHANGED FOR A NEW LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12/3.5/085 | OS | Physician | 02/01/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6805970&pc=MTA | 2023826-2017-01354 |  | Injury | STAAR SURGICAL COMPANY | 2017-08-18 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/78 DIOPTER, INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+1.0/78 | NA | Physician | 07/24/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4695373&pc=MTA | 2023826-2015-00422 | 2015-02-16 | Injury | STAAR SURGICAL COMPANY | 2015-04-15 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. REFRACTIVE SURPRISE WAS NOTED, THE LENS REMAINS IMPLANTED. PATIENT'S LAST KNOWN BCVA WAS 20/20 AS OF (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. EXPLANT DATE: N/A. (B)(4). DEVICE EVALUATED BY THE MANUFACTURER? NO. LENS REMAINS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW RESULTS: (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). THE PATIENT HAS PRE-EXISTING CORNEAL DISEASE (PELLUCID MARGINAL DEGENERATION) WHICH MAY BE ASSOCIATED WITH ECTATIC AND IRREGULAR CORNEAS. (B)(4). | -4.50/+6.0/001 | OS | Other | 03/16/2015 | 2015 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4889319&pc=MTA | 2023826-2015-00768 | 2015-06-02 | Injury | STAAR SURGICAL COMPANY | 2015-07-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REFRACTIVE SURPRISE WAS NOTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND THERE WERE NO SIMILAR COMPLAINTS FOUND. CONCLUSION: BASED ON THE COMPLAINT INFORMATION AND LENS WORK ORDER SEARCH, WE WERE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4). LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.5/5.5/012 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. REFRACTIVE SURPRISE WAS NOTED AND THE LENS WAS EXPLANTED ON (B)(6) 2015. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 6/24 AND BEST CORRECTED VISUAL ACUITY WAS 6/6. (B)(4). CORRECTED DATA: THE CORRECT REPORT COUNTRY IS (B)(6). THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.5/5.5/012 | OS | Physician | 06/10/2015 | 2015 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5040395&pc=MTA | 2023826-2015-01111 | 2015-07-31 | Injury | STAAR SURGICAL COMPANY | 2015-08-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0/+2.5/X001 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THE LENS HAD AN UNEXPECTED REFRACTIVE OUTCOME. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: UPON REVIEW OF THE DEVICE HISTORY RECORD, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON THE ABOVE INVESTIGATION, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2015 WITH SAME LENS MODEL, DIFFERENT DIOPTER, WHICH RESOLVED THE PROBLEM. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS IS IN SPECIFICATIONS. (B)(4). DATE OF REPORT WAS '12/11/2015'. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. (B)(4). METHOD -(PROCESS EVALUATION): - WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION -(UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS IMPLANTED. | -8.0/+2.5/001 | OS | Physician | 12/11/2015 | 2015 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5248744&pc=MTA | 2023826-2015-01548 | 2015-08-25 | Injury | STAAR SURGICAL COMPANY | 2015-11-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.00/+1.5/171 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON 08/25/2015. THE LENS HAD AN EXCESSIVE VAULT AND THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE SURGEON IS PLANNING TO EXPLANT AND EXCHANGE FOR A SHORTER LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(6). (B)(4). CONCLUSION: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED FOR PATIENTS WITH KERATOCONUS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS IMPLANTED. (B)(4). METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | -6.00/+1.5/171 | OS | Physician | 11/02/2015 | 2015 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5174741&pc=MTA | 2023826-2015-01400 | 2015-09-07 | Injury | STAAR SURGICAL COMPANY | 2015-10-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE. THE LENS WAS EXPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -2.0/+2.0/068 DIOPTER IN THE PATIENT'S LEFT EYE (OS). ON (B)(6) 2015, THE LENS WAS EXCHANGED DUE TO REFRACTIVE SURPRISE, WHICH RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/25. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS IS IN SPECIFICATIONS. CORRECTIONAL DATA: THE CORRECT SERIAL # (B)(4). THE CORRECT REPORT SOURCE IS 'DISTRIBUTOR' AND NOT USER FACILITY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - (OTHER): LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). LENS NOT RETURNED. | -2.0/+2.0/068 | OS | Physician | 09/25/2015 | 2015 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5613777&pc=MTA | 2023826-2016-00530 | 2015-09-22 | Injury | STAAR SURGICAL COMPANY | 2016-04-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+5.5/091 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED.(B)(4). | -10.5/+5.5/091 | OD | Physician | 02/16/2016 | 2016 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5613013&pc=MTA | 2023826-2016-00529 | 2015-09-25 | Injury | STAAR SURGICAL COMPANY | 2016-04-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+6.0/093 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. (B)(4). | -10.5/+6.0/093 | OS | Physician | 02/16/2016 | 2016 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5481729&pc=MTA | 2023826-2016-00168 | 2015-09-30 | Injury | STAAR SURGICAL COMPANY | 2016-03-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+0.50/008 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD A REFRACTIVE SURPRISE AND THE IRIS WAS SUTURED. THE LENS IS IMPLANTED BUT WILL BE EXPLANTED AT A LATER DATE. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+0.50/008 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO REFRACTIVE SURPRISE AND AN ENLARGEMENT OF INCISION WAS PERFORMED. THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD: WORK ORDER SEARCH. EVALUATION RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -11.0/+0.50/008 | OS | Physician | 11/19/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5381880&pc=MTA | 2023826-2016-00068 | 2015-10-17 | Injury | STAAR SURGICAL COMPANY | 2016-01-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -1.0/+3.0/086 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD A REFRACTIVE SURPRISE AND WAS EXPLANTED ON (B)(6) 2015. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. (B)(4). CONCLUSION: (USER ERROR CAUSED OR CONTRIBUTED TO EVENT) - REFRACTIVE SURPRISE DUE TO SURGEON'S ERROR WHEN ENTERING THE DATA IN OCOS, LENS EXCHANGED FOR SAME MODEL AND DIFFERENT POWER. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO- LENS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -1.0/+3.0/086 | OS | Physician | 12/22/2015 | 2015 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5806430&pc=MTA | 2023826-2016-00904 | 2015-11-14 | Injury | STAAR SURGICAL COMPANY AG | 2016-07-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -4.50/+1.5/052 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/28. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION WAS PERFORMED. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE (B)(4) BUT IS NOT MARKETED IN THE U.S. (B)(4). | -4.50/+1.5/052 | OS | Physician | 05/27/2016 | 2016 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5442765&pc=MTA | 2023826-2016-00173 | 2015-11-25 | Injury | STAAR SURGICAL COMPANY AG | 2016-02-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -1.50/+2.5/073 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE WRONG DIOPTER LENS WAS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXCHANGED FOR A SIMILAR LENS WITH A DIFFERENT DIOPTER. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | -1.50/+2.5/073 | OD | Physician | 12/14/2015 | 2015 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5743552&pc=MTA | 2023826-2016-00850 | 2016-02-10 | Injury | STAAR SURGICAL COMPANY | 2016-06-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.50/+6.0/087 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER INDICATED THE LENS HAD A REFRACTIVE SURPRISE AND REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: CONCLUSION CODE: PATIENT IS UNDER THE AGE OF 21 AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS IMPLANTED. (B)(4) REFRACTIVE SURPRISE. IMPLANTED. (B)(4). | -6.50/+6.0/087 | OS | Physician | 06/02/2016 | 2016 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5572540&pc=MTA | 2023826-2016-00398 | 2016-03-07 | Injury | STAAR SURGICAL COMPANY | 2016-04-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -1.50/+5.5/107 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS HAD A REFRACTIVE SURPRISE, WITH INCREASED CYLINDER. THE LENS WAS REPOSITIONED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/80. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) BELOW 3.0MM AND WHOSE AGE IS ABOVE 45 YEARS IS CONTRAINDICATED. THIS PATIENT WAS 66 YEARS OLD AND HAD AN ACD OF 2.90MM AT THE TIME OF IMPLANTATION. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). | -1.50/+5.5/107 | OD | Physician | 03/15/2016 | 2016 | 66
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5937758&pc=MTA | 2023826-2016-01233 | 2016-06-04 | Injury | STAAR SURGICAL COMPANY | 2016-09-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/127 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT DEVELOPED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED FOR A SIMILAR LENS WITH A DIFFERENT DIOPTER. THIS RESOLVED THE PROBLEM. THE EXPLANT DATE IS UNKNOWN. DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED. | -10.5/+1.0/127 | OD | Physician | 08/09/2016 | 2016 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6109407&pc=MTA | 2023826-2016-01593 | 2016-09-09 | Injury | STAAR SURGICAL COMPANY | 2016-11-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/085 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016, DUE TO LENS UPSIDE DOWN. THE LENS WAS EXCHANGED WITH A LENS OF THE SAME MODEL, SIZE AND DIOPTER. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.5/085 | OD | Physician | 10/26/2016 | 2016 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5193796&pc=MTA | 2023826-2015-01442 |  | Injury | STAAR SURGICAL COMPANY | 2015-11-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -3.0/+1.0/087 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS BUT DIFFERENT DIOPTER POWER. Manufacturer Narrative: PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20 AS OF (B)(6) 2015. DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. FOREIGN MATERIAL (FIBERS) WAS PRESENT ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS MEASUREMENTS WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS DIOPTER MEASUREMENTS WITHIN SPECIFICATIONS. (B)(4). CONCLUSION: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.96MM AT THE TIME OF IMPLANTATION. PATIENT WAS ALSO REPORTED TO HAVE PREVIOUS CORNEAL REFRACTIVE SURGERY PRIOR TO LENS IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -3.0/+1.0/087 | OS | Physician | 10/02/2015 | 2015 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5442795&pc=MTA | 2023826-2016-00170 |  | Injury | STAAR SURGICAL COMPANY | 2016-02-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -18.0/+5.5/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS). THE REPORTER INDICATED THE LENS HAD A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 VTIMO13.7 IMPLANTABLE COLLAMER LENS, -18.0/+5.5/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) AND NOTICED REFRACTIVE SURPRISE. THE SURGEON REPOSITIONED THE LENS TO ADJUST THE ERROR AND THIS RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -18.0/+5.5/085 | OS | Physician | 01/19/2016 | 2016
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5664146&pc=MTA | 2023826-2016-00694 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+0.5/047 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL/RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). CONCLUSION: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS WITH ANTERIOR CHAMBER DEPTHS (ACD) BELOW 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WEIGHT: UNK. DATE OF EVENT: UNK. (B)(4). LENS NOT RETURNED. | -11.0/+0.5/047 | OS | Physician | 05/09/2016 | 2016 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5697777&pc=MTA | 2023826-2016-00371 |  | Injury | STAAR SURGICAL COMPANY | 2016-06-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+3.5/9 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED AND IS PLANNED TO EXCHANGE FOR ANOTHER LENS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS REMAINS IMPLANTED. (B)(4). | -15.5/+3.5/9 | OS | Physician | 03/24/2016 | 2016
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5937841&pc=MTA | 2023826-2016-01234 |  | Injury | STAAR SURGICAL COMPANY | 2016-09-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/93 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT DEVELOPED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED AND EXCHANGED FOR ANOTHER LENS OF THE SAME MODEL AND SIZE, BUT DIFFERENT DIOPTER. THIS RESOLVED THE PROBLEM. DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION - DIMENSIONAL AND FUNCTIONAL INSPECTIONS WERE PERFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4). DEVICE NOT RETURNED. | -10.5/+1.5/93 | OS | Physician | 08/09/2016 | 2016 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5998770&pc=MTA | 2023826-2016-01378 |  | Injury | STAAR SURGICAL COMPANY | 2016-10-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -3.00/+6.00X23 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. IT WAS INDICATED THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADD'L INFO: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE LENS WAS IMPLANTED ON (B)(6) 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -3.00/+6.0023 | OS | Physician | 09/05/2016 | 2016
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5678933&pc=MTA | 2023826-2016-00763 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Device Dislodged or Dislocated; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+3.0/097 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS ROTATION AND A REFRACTIVE SURPRISE. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN AND FOREIGN MATERIAL (FIBERS) ON LENS SURFACE. (B)(4). CLAIM #(B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE LENS WAS EXCHANGED. AFTER THE EXCHANGE, PATIENT'S VISION IMPROVED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -18.00/+3.0/097 | OS | Physician | 11/03/2015 | 2015 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5304423&pc=MTA | 2023826-2015-01630 | 2015-09-22 | Injury | STAAR SURGICAL COMPANY | 2015-12-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Insufficient Information | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -3.0/+2.00/X085 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. A REFRACTIVE SURPRISE WAS REPORTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -3.0/+2.00/085 | OS | Physician | 11/19/2015 | 2015 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6024467&pc=MTA | 2023826-2016-01411 | 2015-11-24 | Injury | STAAR SURGICAL COMPANY | 2016-10-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS - -21.5/+4.0/095 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS TEAR DURING DELIVERY INTO THE EYE. ANOTHER SAME MODEL/DIFFERENT DIOPTER WAS IMPLANTED ON (B)(6) 2015. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: RESULT: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF THE LENS OPTIC AND ONE HAPTIC TORN OFF. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). | -21.5/+4.0/095 | OS | Physician | 09/14/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6057748&pc=MTA | 2023826-2016-01476 | 2016-05-12 | Injury | STAAR SURGICAL COMPANY | 2016-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON WAS INSERTING AN 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+0.5/99 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) AND THE LENS TORE DURING INJECTION. THE LENS WAS NOT IMPLANTED. ANOTHER LENS WAS IMPLANTED ON (B)(6) 2016 WITH THE SAME LENGTH, DIFFERENT MODEL AND DIOPTER AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC WAS SPLIT AND THE HAPTIC WAS BROKEN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -7.0/+0.5/99 | OD | Physician | 09/27/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6094563&pc=MTA | 2023826-2016-00880 |  | Injury | STAAR SURGICAL COMPANY | 2016-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50/+1.5/090 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS EXCHANGED FOR A LENS OF THE SAME MODEL AND SIZE, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -15.50/+1.5/090 | OD | Physician | 05/25/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5822522&pc=MTA | 2023826-2016-00982 | 2016-05-24 | Injury | STAAR SURGICAL COMPANY | 2016-07-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Unintended Movement | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.00/+5.50/080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING ROTATION AND A REFRACTIVE SURPRISE AND REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/50. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION: AS PER THE DFU OF THIS LENS MODEL, IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRADICTED. THIS PATIENT HAD AN ACD OF 2.9 MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: LENS REMAINS IMPLANTED. (B)(4). | -5.00/+5.50/080 | OD | Physician | 07/05/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7268656&pc=MTA | 2023826-2018-00177 |  | Injury | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Difficult to Fold, Unfold or Collapse; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -10.50/+2.0/X093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS). EXCESSIVE VAULT AND NARROWING OF THE ANGLES WERE NOTICED. THE LENS WAS THEN REPOSITIONED. AFTER THE REPOSITION, THE SURGEON BELIEVED THAT THE FOOTPLATESS MAY HAVE NOT BEEN POSITIONED PROPERLY. THE PATIENT'S IOP IS FINE AND VISION IS GOOD. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.50/+2.0/093 | OS | Physician | 01/15/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6657575&pc=MTA | 2023826-2017-01019 | 2016-08-31 | Injury | STAAR SURGICAL COMPANY | 2017-06-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Improper or Incorrect Procedure or Method | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/1.5/74 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. IN THE SAME SURGERY, THE LENS WAS EXPLANTED DUE TO LENS TEAR/BEAK DURING INJECTION INTO THE EYE. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR THE SAME SIZE AND DIOPTER LENS. THE PROBLEM WAS RESOLVED. THE CUSTOMER STATED AFTER LOADING THE LENS, IT WAS NOT PUT IN BSS. THE CUSTOMER THOUGH THAT IS THE REASON CAUSING THE LENS TORE, AND THE SURGEON HAS BEEN TRAINED AGAIN. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE LENS CUT IN TWO PIECES, MISSING A PIECE OF HAPTIC AND RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.0/1.5/74 | OD | Physician | 05/26/2017 | 2017 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7484280&pc=MTA | 2023826-2018-00676 | 2018-03-15 | Injury | STAAR SURGICAL COMPANY | 2018-05-03 | MTA | IMPLANTABLE COLLAMER LENS | Improper or Incorrect Procedure or Method | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.50/+4.50/094 (SPHERE/CYLINDAR/AXIS) INTO THE PATIENT'S LEFT EYE (OS). ONE DAY POST-OP IT WAS DISCOVERED THAT THE LENS TORE/BROKE BEFORE OR DURING LOADING. THE REPORTED CAUSE IS USER/TECHNIQUE ERROR. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH AN ALTERNATE LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN, HAPTIC TORN/BROKEN/BENT/DEFORMED, AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) (B)(4). | -5.50/+4.50/094 | OS | Physician | 04/10/2018 | 2018 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8354784&pc=MTA | 2023826-2019-00276 | 2019-01-05 | Injury | STAAR SURGICAL COMPANY | 2019-02-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -4.00/2.5/176 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. SIGNIFICANT ENDOTHELIAL CELL LOSS WAS OBSERVED. THE SURGEON STATES "THE ROOT CAUSE IS CLEAR." REPORTEDLY, THE TIP OF THE IA CONTACTED THE ANTERIOR CHAMBER DURING IRRIGATION. SURGEON "RECOGNIZES IIT IS ROOT CAUSE" OF THE SIGNIFICANT ENDOTHELIAL CELL LOSS. THE LENS REMAINS IMPLANTED. THIS INCIDENT OCCURRED DURING DOCTOR TRAINING/QUALIFICATION SURGERY. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -4.00/2.5/176 | OS | Physician | 01/29/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6850675&pc=MTA | 2023826-2017-01404 | 2017-07-25 | Injury | STAAR SURGICAL COMPANY | 2017-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, DIOPTER -8.50/2.5/90 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REPORTEDLY, THE LENS WAS IMPLANTED "INSIDE OUT." COMPLAINT FORM STATES THAT AS THE LENS WAS REMOVED, IT WAS DISCOVERED TO HAVE A "CRACK" ON THE LENS. IT IS UNKNOWN AS TO WHEN OR HOW THE CRACK OCCURRED ON THE LENS. EXPLANT DATE IS NOT REPORTED. ACCORDING TO CLAIM REPORTER, THE SURGERY HAS BEEN POSTPONED AND ALTERNATE LENS IS ORDERED. Manufacturer Narrative: THE EXPLANT DATE IS SAME AS THE IMPLANT DATE I.E.,(B)(6) 2017. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.97MM)]. (B)(4). | -8.50/2.5/90 | OS | Physician | 08/09/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4724434&pc=MTA | 2023826-2015-00455 | 2014-05-20 | Injury | STAAR SURGICAL COMPANY AG | 2015-04-23 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM13.7 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED, AS WELL AS ANGLE CLOSURE ASSOCIATED WITH ELEVATED IOP AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXPLANTED ON (B)(6) 2014 AND ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS WAS PERFORMED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DOB UNK. (B)(4) -(SECONDARY SURGERY); (NEW INCISION); INTRAOCULAR PRESSURE RISE; (REDUCTION OF IRIDO-CORNEAL ANGLES); SIZE INCORRECT FOR PATIENT; (EXPLANTED); (VAULTING). DEVICE EVALUATED BY THE MANUFACTURER? NO. LENS WAS NOT RETURNED. (B)(4). | -16.5/+4/50 | OS | Physician | 03/26/2015 | 2015 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4614296&pc=MTA | 2023826-2015-00330 | 2014-06-20 | Injury | STAAR SURGICAL COMPANY | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE, RESULTING IN AN EXPLANT OF THE LENS ON (B)(6) 2014 AND EXCHANGE FOR A LENS WITH A SMALLER VAULT. THE ISSUE WAS RESOLVED WITH THE IMPLANTATION OF THE NEW LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4) - (SECONDARY SURGERY); (NEW INCISION); SIZE INCORRECT FOR PATIENT; (EXPLANTED); (VAULTING). DEVICE EVALUATED BY THE MANUFACTURER? NO. LENS WAS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -10.5/+3.5/090 | OS | Physician | 02/17/2015 | 2015 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4784566&pc=MTA | 2023826-2015-00590 | 2014-10-23 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND LOW VAULT WAS NOTED IN THE POST-OPERATIVE PERIOD. THE LENS WAS EXCHANGED FOR A LONGER LENGTH ON (B)(6) 2014 AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS DRY WITH NO VISIBLE DAMAGE. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND WITHIN THE WORKORDER. MEDICAL REVIEW: ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -7.5/+2.5/079 | OD | Physician | 02/05/2015 | 2015 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4599021&pc=MTA | 2023826-2015-00307 | 2014-11-05 | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(4) 2014. THE LENS WAS EXPLANTED ON (B)(4) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/50. Manufacturer Narrative: (B)(4). EVALUATION CODES: METHOD -(OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. MEDICAL REVIEW -ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP, ECT.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -18.00/+4.0/090 | OD | Physician | 02/10/2015 | 2015 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4789535&pc=MTA | 2023826-2015-00595 | 2014-11-16 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-21 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND LOW VAULT WAS NOTED. THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS ON EXCHANGING THE LENS FOR A LONGER LENGTH. THE PATIENT'S LAST UCVA WAS 20/50 ONE DAY POSTOP. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). STAAR RECOMMENDS THE ICL TO BE LEFT IMPLANTED IF NO CATARACTS OR WHERE THERE IS A TRACE ANTERIOR SUB-CAPSULAR CATARACT WITH NO PROGRESSION. STAAR BELIEVES THAT THE ACTION TO REMOVE AND/REPLACE THE ICL INCREASE THE RISK FOR ANTERIOR SUB-CAPSULAR CATARACT. BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. THE LENS REMAINS IMPLANTED. (B)(4). | -9.00/1.50/99 | OD | Other | 04/22/2015 | 2015 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6112657&pc=MTA | 2023826-2016-01669 | 2015-03-11 | Injury | STAAR SURGICAL COMPANY | 2016-11-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50/+4.5/085 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO NO VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID AND IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. CONCLUSION: AS PER THE DFU OF THIS LENS MODEL, IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRADICTED. THIS PATIENT HAD AN ACD OF 2.82 MM AT THE TIME OF IMPLANTATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -08.50/+4.5/085 | OS | Physician | 02/24/2016 | 2016
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6460519&pc=MTA | 2023826-2017-00558 | 2015-03-11 | Injury | STAAR SURGICAL COMPANY | 2017-04-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/2.5/170 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.5/2.5/170 | OS | Physician | 03/07/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6855503&pc=MTA | 2023826-2017-01425 | 2015-03-14 | Injury | STAAR SURGICAL COMPANY | 2017-09-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -9.5/+2.00/73 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.5/+2.00/73 | OS | Physician | 08/14/2017 | 2017 | 16
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4789555&pc=MTA | 2023826-2015-00577 | 2015-04-19 | Injury | STAAR SURGICAL COMPANY | 2015-05-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) NEW INCISION. SECONDARY SURGERY. LENS EXPLANTED. VAULTING. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. PER MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -8.5/4.5/061 | OS | Other | 04/28/2015 | 2015 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4789586&pc=MTA | 2023826-2015-00576 | 2015-04-19 | Injury | STAAR SURGICAL COMPANY | 2015-05-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - 81 (OTHER): LENS NOT RETURNED. EVALUATION METHOD : WORK ORDER SEARCH. RESULTS - 100 (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) . Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. PER MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -7.5/2.0/109 | OD | Other | 04/28/2015 | 2015 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4919446&pc=MTA | 2023826-2015-00836 | 2015-04-28 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-16 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND EXCESSIVE VAULT WAS OBSERVED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2015 AND THIS RESOLVED THE PROBLEM. PATIENT'S LAST BCVA WAS 20/20. SEE MFR REPORT# 2023826-2015-00837 FOR THE OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. BASED ON THE COMPLAINT INFORMATION AND LENS WORK ORDER SEARCH, WE WERE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4). | -9/+3/91 | OD | Physician | 06/16/2015 | 2015 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4902967&pc=MTA | 2023826-2015-00806 | 2015-05-07 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-09 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015 AND LOW VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE PATIENT'S LAST UCVA WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT INFORMATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -10.0/+3/96 | OS | Physician | 06/09/2015 | 2015
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6151959&pc=MTA | 2023826-2016-01705 | 2015-07-31 | Injury | STAAR SURGICAL COMPANY AG | 2016-12-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0 MM TICM120V4 IMPLANTABLE COLLAMER LENS, -22.0/+5.0/97 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2010. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(6) THIS PRODUCT IS MANUFACTURED IN THE U.S. IS INCORRECT, CORRECT DATA IS THIS PRODUCT IS NOT MANUFACTURED IN THE U.S. PRODUCT EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -22.0/+5.0/97 | OD | Physician | 11/08/2016 | 2016 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6094543&pc=MTA | 2023826-2016-01423 | 2015-08-25 | Injury | STAAR SURGICAL COMPANY | 2016-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00/+4.0/095 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS, BUT HAVE FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS DRIED, DISCOLORED MATERIAL. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -09.00/+4.0/095 | OD | Physician | 04/28/2016 | 2016 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6855476&pc=MTA | 2023826-2017-01426 | 2015-08-25 | Injury | STAAR SURGICAL COMPANY AG | 2017-09-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0 MM TICM120V4, -16.00/1.50/90 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS). THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT UNDER THE AGE OF 21 YEARS OLD AND IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | -16.00/1.50/90 | OD | Physician | 08/14/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6094551&pc=MTA | 2023826-2016-01416 | 2015-10-19 | Injury | STAAR SURGICAL COMPANY | 2016-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+2.5/078 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND A SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAPTIC TORN (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.00/+2.5/078 | OD | Physician | 12/09/2015 | 2015 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6043895&pc=MTA | 2023826-2016-01087 | 2015-12-17 | Injury | STAAR SURGICAL COMPANY | 2016-10-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+3.0/091 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND A SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.00/+3.0/091 | OS | Physician | 06/14/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6059227&pc=MTA | 2023826-2016-01414 | 2016-01-12 | Injury | STAAR SURGICAL COMPANY AG | 2016-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -6.50/+1.5/086 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/16. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | -6.50/+1.5/086 | OD | Physician | 09/26/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6059229&pc=MTA | 2023826-2016-01415 | 2016-01-12 | Injury | STAAR SURGICAL COMPANY AG | 2016-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -5.50/+1.5/093 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/26. Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2016 FOR A LONGER LENS AND THIS RESOLVED THE PROBLEM. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4). | -5.50/+1.5/093 | OS | Physician | 09/26/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6043835&pc=MTA | 2023826-2016-01113 | 2016-01-23 | Injury | STAAR SURGICAL COMPANY | 2016-10-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.50/+1.5/105 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND A SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER AND THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.50/+1.5/105 | OD | Physician | 06/14/2016 | 2016 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6043845&pc=MTA | 2023826-2016-01114 | 2016-01-23 | Injury | STAAR SURGICAL COMPANY | 2016-10-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.50/+2.0/078 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND A SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BENT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.50/+2.0/078 | OS | Physician | 06/14/2016 | 2016 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6036797&pc=MTA | 2023826-2016-01467 | 2016-01-27 | Injury | STAAR SURGICAL COMPANY AG | 2016-10-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.5 MM TICM125V4 IMPLANTABLE COLLAMER LENS, -20.0/+2.5/080 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY AND BENT IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. (B)(4). | -20.0/+2.5/080 | OS | Physician | 04/19/2016 | 2016 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6048722&pc=MTA | 2023826-2016-01081 | 2016-01-30 | Injury | STAAR SURGICAL COMPANY | 2016-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+1.5/091 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND A SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL EVENT DESCRIPTION: PATIENT'S LAST VISIT WAS ON (B)(6) 2016, WITH BEST-CORRECTED VISUAL ACUITY (BCVA) 20/20. DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). | -18.00/+1.5/091 | OD | Physician | 06/14/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6124427&pc=MTA | 2023826-2016-01672 | 2016-02-15 | Injury | STAAR SURGICAL COMPANY | 2016-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+2.5/087 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL; SURGICAL RESIDUE WAS CLEAR; VISUAL INSPECTION FOUND NO VISIBLE DAMAGE (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.50/+2.5/087 | OS | Physician | 10/25/2016 | 2016 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6158833&pc=MTA | 2023826-2016-01742 | 2016-02-15 | Injury | STAAR SURGICAL COMPANY | 2016-12-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16/2.5/047 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) WAS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE PER INTERNATIONAL DECISION TREE. (B)(4). | -16/2.5/047 | OD | Physician | 04/20/2016 | 2016 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6031444&pc=MTA | 2023826-2016-01446 | 2016-02-17 | Injury | STAAR SURGICAL COMPANY AG | 2016-10-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/085 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 . THE LENS WAS EXPLANTED ON AN UNKNOWN DATE DUE TO LOW VAULT AND EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Event Description: THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT AND EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. EVENT PROBLEM AND EVALUATION CODES: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. (B)(4). | -10.5/+1.5/085 | OS | Physician | 09/26/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7700569&pc=MTA | 2023826-2018-01063 | 2016-03-17 | Injury | STAAR SURGICAL COMPANY | 2018-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICMO13.2, -13.50/+2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ON (B)(6) 2016 THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SUSPECTED ENDOPHTALMITIS AND PIGMENT DISPERSION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PREVIOUSLY REPORTED THAT THE LENS WAS EXCHANGED DUE TO EXCESSIVE VAULT, ENDOPHTHALMITIS, AND PIGMENT DISPERSION. SURGEON CONFIRMS THAT THERE WAS NO ENDOPHTHALMITIS. THE LENS WAS EXCHANGED DUE TO EXCESSIVE VAULT AND PIGMENT DISPERSION ONLY. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICMO13.2 WITH A DIOPTER OF -13.50/1.5/097 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, LENS REMOVED AND EXCHANGED FOR A SHORTER LENS. SUSPECTED ENDOPHTALMITIS AND PIGMENT DISPERSION. (B)(4). | -13.50/+2.0/090 | OD | Physician | 06/19/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6048665&pc=MTA | 2023826-2016-01080 | 2016-03-19 | Injury | STAAR SURGICAL COMPANY | 2016-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+2.5/079 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND A SHALLOW CHAMBER. Manufacturer Narrative: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BENT/BROKEN. FOREIGN MATERIAL (FIBERS) WAS PRESENT ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.50/+2.5/079 | OS | Physician | 06/14/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6052746&pc=MTA | 2023826-2016-01079 | 2016-03-19 | Injury | STAAR SURGICAL COMPANY | 2016-10-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/092 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND A SHALLOW CHAMBER. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.00/+2.0/092 | OD | Physician | 06/14/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6351569&pc=MTA | 2023826-2017-00205 | 2016-03-31 | Injury | STAAR SURGICAL COMPANY AG | 2017-02-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -9.50/+3/83 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2008. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE REPORTER STATED ENLARGEMENT OF PI OR ADDITION TO NEW PI (SURGICAL). THE LENS WAS EXCHANGED FOR A 13.2MM VTICM5_13.2, -8.00/1.50/80 DIOPTER, ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. ON PATIENT'S LAST VISIT WHICH WAS ON (B)(6) 2016, THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 (THE SAME AS PREOPERATIVE DATA ON (B)(6) 2008) AND UNCORRECTED VISUAL ACUITY (UCVA) WAS ALSO 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: LENS WAS EXPLANTED DUE TO LOW VAULTING AND LENS ROTATION. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS NOT MANUFACTURED IN THE U.S. AND NOT MARKETED IN THE U.S. (B)(4). | -9.50/+3/83 | OS | Physician | 01/31/2017 | 2017 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6486089&pc=MTA | 2023826-2017-00595 | 2016-04-03 | Injury | STAAR SURGICAL COMPANY | 2017-04-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/083 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXCHANGED FOR SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS, CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4). | -12.5/+2.5/083 | OD | Physician | 06/24/2016 | 2016 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6595544&pc=MTA | 2023826-2017-00816 | 2016-04-15 | Injury | STAAR SURGICAL COMPANY AG | 2017-05-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5 MM TICM125V4 IMPLANTABLE COLLAMER LENS, -15/+2.0/88 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2006. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL ON LENS SURFACE. (B)(4). | -15/+2.0/88 | OS | Physician | 10/14/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6042647&pc=MTA | 2023826-2016-01424 | 2016-04-22 | Injury | STAAR SURGICAL COMPANY | 2016-10-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/082 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND PIGMENT DISPERSION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION - THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH - NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.5/082 | OS | Physician | 09/27/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6077578&pc=MTA | 2023826-2016-01567 | 2016-05-10 | Injury | STAAR SURGICAL COMPANY | 2016-11-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+1.0/145 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.00/+1.0/145 | OS | Physician | 10/19/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6879896&pc=MTA | 2023826-2017-01481 | 2016-05-14 | Injury | STAAR SURGICAL COMPANY | 2017-09-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -10.50/+4.0/090 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF DEBRIS AND CLEAR RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | -10.50/+4.0/090 | OS | Physician | 08/25/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6879905&pc=MTA | 2023826-2017-01479 | 2016-05-14 | Injury | STAAR SURGICAL COMPANY | 2017-09-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -10.0/+1.5/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PATIENT POST-OP BCVA WAS 20/20 ON (B)(6) 2016. PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN, HAPTIC PIECE MISSING, AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.0/+1.5/093 | OD | Physician | 08/25/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5987401&pc=MTA | 2023826-2016-01342 | 2016-05-16 | Injury | STAAR SURGICAL COMPANY | 2016-09-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5 MM TICM 12.5 IMPLANTABLE COLLAMER LENS - -18.5/+1.0/024 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5 MM TICM125V4 IMPLANTABLE COLLAMER LENS, 18.5/+1.0/024 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: TICM125V4 IS THE CORRECT MODEL NUMBER, REPLACING TICM 12.5. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM #(B)(4). | -18.5/+1.0/024 | OD | Physician | 09/03/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5987409&pc=MTA | 2023826-2016-01343 | 2016-05-16 | Injury | STAAR SURGICAL COMPANY | 2016-09-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM 12.5 IMPLANTABLE COLLAMER LENS - -14.5/+1.0/150 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE FOREIGN MATERIAL ON LENS SURFACE (FIBERS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.5/+1.0/150 | OS | Physician | 09/03/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5985710&pc=MTA | 2023826-2016-01345 | 2016-05-19 | Injury | STAAR SURGICAL COMPANY | 2016-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/096 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN CASE/VIAL; VISUAL INSPECTION FOUND HAPTIC TORN AND HAIR ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.5/096 | OD | Physician | 08/30/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5985711&pc=MTA | 2023826-2016-01346 | 2016-05-19 | Injury | STAAR SURGICAL COMPANY | 2016-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/22. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.5/085 | OS | Physician | 08/30/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6094544&pc=MTA | 2023826-2016-01170 | 2016-05-20 | Injury | STAAR SURGICAL COMPANY | 2016-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+1.5/090 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/90 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE REPORTER STATED THAT THE PATIENT "DOES NOT ACCEPT TO BE LOADED AGAIN." WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -12.50/+1.5/090 | OS | Physician | 06/16/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6055958&pc=MTA | 2023826-2016-01477 | 2016-05-23 | Injury | STAAR SURGICAL COMPANY | 2016-10-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5/+2.5/92 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/10. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND HAIR ON LENS SURFACE, BUT NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -7.5/+2.5/92 | OS | Physician | 09/27/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6107624&pc=MTA | 2023826-2016-01619 | 2016-05-24 | Injury | STAAR SURGICAL COMPANY | 2016-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/2.5/086 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND SCRATCHES ON THE LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -10.0/2.5/086 | OD | Physician | 10/18/2016 | 2016 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6107654&pc=MTA | 2023826-2016-01589 | 2016-05-24 | Injury | STAAR SURGICAL COMPANY | 2016-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/2.0/097 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: THA PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20 AND UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/2.0/097 | OS | Physician | 10/18/2016 | 2016 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5988652&pc=MTA | 2023826-2016-01229 | 2016-05-27 | Injury | STAAR SURGICAL COMPANY | 2016-09-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -/5.5/+2/086 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FIBERS ON LENS SURFACE. CONCLUSION: THE PATIENT'S ANTERIOR CHAMBER DEPTH IS LESS THAN 3.0 MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -5.5/+2/086 | OS | Physician | 08/29/2016 | 2016 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6117680&pc=MTA | 2023826-2016-01639 | 2016-06-02 | Injury | STAAR SURGICAL COMPANY | 2016-11-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/+1.0/093 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL; SURGICAL RESIDUE WAS CLEAR; VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/+1.0/093 | OS | Physician | 10/25/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6270897&pc=MTA | 2023826-2017-00134 | 2016-06-02 | Injury | STAAR SURGICAL COMPANY | 2017-01-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+2.5/089 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON IT. VISUAL INSPECTION FOUND HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -13.00/+2.5/089 | OD | Physician | 12/27/2016 | 2016 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6271795&pc=MTA | 2023826-2017-00135 | 2016-06-03 | Injury | STAAR SURGICAL COMPANY | 2017-01-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/092 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED BENT AND DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE OPTIC WAS TORN AND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -12.00/+2.0/092 | OS | Physician | 12/27/2016 | 2016 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6800858&pc=MTA | 2023826-2017-01333 | 2016-06-08 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/095 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL AND HAD CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO BE CURVED AND CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.0/095 | OD | Physician | 07/22/2017 | 2017 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6800862&pc=MTA | 2023826-2017-01334 | 2016-06-08 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/085 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL AND HAD CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND DEFORMED. SERIAL #: IT IS CORRECTED FROM "(B)(4)" TO "(B)(4)" IN THE INITIAL. EXPIRATION DATE: IT IS CORRECTED FROM "31-JUL-2018" TO "30-APR-2018" IN THE INITIAL. DEVICE MANUFACTURE DATE: IT IS CORRECTED FORM "25-AUG-2015" TO "04-JUN-2015" IN THE INITIAL. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.0/085 | OS | Physician | 07/22/2017 | 2017 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6679730&pc=MTA | 2023826-2017-01069 | 2016-06-09 | Injury | STAAR SURGICAL COMPANY | 2017-06-29 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6MM IMPLANTABLE COLLAMER LENS WITH A -11.5/+1/066 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXCHANGED ON (B)(6) 2016 WITH A SHORTER LENS, 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS WITH A -11.50/+1.0/069 DIOPTER AND THE EXCHANGE RESOLVED THE PROBLEM. ON PATIENT'S LAST VISIT (B)(6) 2016 THE PATIENT'S BEST CORRECTED VISUAL ACUITY IS 20/20 WITH A CORNEAL THICKNESS OF 0.5. PATIENT'S ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS 2.90MM. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE), A PIECE OF HAPTIC MISSING, AND THE LENS IS CURVED UP. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL DATA: PRODUCT NOT MANUFACTURED IN THE U.S. (B)(4). CONCLUSION CODE: REVIEW OF THE FILE INDICATES THAT THE SURGEON DID NOT MEET DFU REQUIREMENTS, PATIENT'S ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM. (B)(4). | -11.5/+1/066 | OD | Physician | 06/09/2017 | 2017 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6487750&pc=MTA | 2023826-2017-00628 | 2016-06-16 | Injury | STAAR SURGICAL COMPANY | 2017-04-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.5/+1.5/093 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS DRY, CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. LENS HAD A SLIGHT CURVE DUE TO THE POSITION IN THE VIAL. PATIENT CODE: (B)(4) NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CONCLUSION CODE: OFF-LABEL USE [OVER 45YEARS OF AGE AT TIME OF IMPLANT ((B)(6))]. (B)(4). | -5.5/+1.5/093 | OD | Physician | 03/13/2017 | 2017 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6487780&pc=MTA | 2023826-2017-00627 | 2016-06-16 | Injury | STAAR SURGICAL COMPANY | 2017-04-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.5/+2.0/077 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS NOT RESOLVED. REFERENCE MDR# 2023826-2017-00624 FOR CASE RESOLUTION. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A SMALL VIAL. VISUAL INSPECTION FOUND A TEAR ON THE HAPTIC AND DRY, CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. LENS HAD A SLIGHT CURVE DUE TO THE POSITION IN THE VIAL. (B)(4). CONCLUSION: OFF-LABEL USE [OVER 45YEARS OF AGE AT TIME OF IMPLANT ((B)(6))]. (B)(4). | -5.5/+2.0/077 | OS | Physician | 03/13/2017 | 2017 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6415428&pc=MTA | 2023826-2017-00437 | 2016-06-23 | Injury | STAAR SURGICAL COMPANY | 2017-03-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MMVTICMO13.2 IMPLANTABLE COLLAMER LENS, -13/+4/99 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2016 ,THE LENS WAS EXCHANGED FOR A SHORTER LENS. AT THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN THE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -13/+4/99 | OD | Physician | 02/21/2017 | 2017 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5987411&pc=MTA | 2023826-2016-01344 | 2016-06-29 | Injury | STAAR SURGICAL COMPANY | 2016-09-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS - -7.50/+1.0/090 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.50/+1.0/090 | OD | Physician | 09/03/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5992358&pc=MTA | 2023826-2016-01364 | 2016-06-29 | Injury | STAAR SURGICAL COMPANY | 2016-09-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS - -7.00/+1.5/075 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND A SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.00/+1.5/075 | OS | Physician | 09/03/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6133891&pc=MTA | 2023826-2016-01691 | 2016-06-30 | Injury | STAAR SURGICAL COMPANY | 2016-11-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/3.5/102 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. CONCLUSION: AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.99MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -18.00/3.5/102 | OS | Physician | 11/01/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6111672&pc=MTA | 2023826-2016-01658 | 2016-07-02 | Injury | STAAR SURGICAL COMPANY | 2016-11-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.0/+4.0/100 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT AND THE HAPTIC TORN/BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS MEASUREMENTS WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+4.0/100 | OD | Physician | 10/18/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6487966&pc=MTA | 2023826-2017-00634 | 2016-07-02 | Injury | STAAR SURGICAL COMPANY | 2017-04-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.50/2.5/120 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, LENS VIAL. VISUAL INSPECTION FOUND HAPTIC TORN, MISSING PIECE, AND CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4). | -9.50/2.5/120 | OD | Physician | 03/14/2017 | 2017 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6487865&pc=MTA | 2023826-2017-00539 | 2016-07-03 | Injury | STAAR SURGICAL COMPANY | 2017-04-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -14.50/3.0/094 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO THE PATIENT EXPERIENCING LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/3.0/94 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2016. THE LENS WAS EXPLANTED ON (B)(6)2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. CORRECTED DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY, IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC PARTIALLY TORN. (B)(4). | -14.50/3.0/094 | OD | Physician | 11/28/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6065005&pc=MTA | 2023826-2016-01541 | 2016-07-05 | Injury | STAAR SURGICAL COMPANY | 2016-10-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/2.5/87 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; VISUAL INSPECTION FOUND OPTIC AND HAPTIC TORN AND FIBERS ON LENS SURFACE (B)(4). Manufacturer Narrative: PMA/510(K): N/A THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/2.5/87 | OS | Physician | 09/30/2016 | 2016 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6131789&pc=MTA | 2023826-2016-01684 | 2016-07-06 | Injury | STAAR SURGICAL COMPANY | 2016-11-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/90 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED. THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+1.0/90 | OD | Physician | 11/02/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6365396&pc=MTA | 2023826-2017-00292 | 2016-07-08 | Injury | STAAR SURGICAL COMPANY | 2017-02-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/2.5/112 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND MYDRIASIS. THE LENS WAS EXCHANGED ON (B)(6) 2016 WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: ON (B)(6) 2016, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A PLASTIC TUBE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED AND LENS HAD CLEAR DRY RESIDUE ALL OVER. CONCLUSION CODE: PATIENT HAD PRE-EXISTING OCULAR CONDITIONS AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS THAT HAD CATARACT IN THE OPERATIVE EYE OR NONTRAUMATIC CATARACT IN THE FELLOW EYE AND ANY PREVIOUS CORNEAL OR REFRACTIVE SURGERY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: 1 SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17.5/2.5/112 | OD | Physician | 02/01/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6577095&pc=MTA | 2023826-2017-00843 | 2016-07-11 | Injury | STAAR SURGICAL COMPANY | 2017-05-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+0.5/X149 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND CLEAR RESIDUE). (B)(4). Manufacturer Narrative: CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+0.5/149 | OD | Physician | 10/28/2016 | 2016 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6482814&pc=MTA | 2023826-2017-00624 | 2016-07-13 | Injury | STAAR SURGICAL COMPANY | 2017-04-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.5/+1/088 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: PATIENT WEIGHT-UNK. DATE OF EVENT-UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TEAR ON THE OPTIC. (B)(4). | -5.5/+1/088 | OS | Physician | 03/13/2017 | 2017 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6015335&pc=MTA | 2023826-2016-01390 | 2016-07-18 | Injury | STAAR SURGICAL COMPANY | 2016-10-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00/0.0/0 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -14.00/0.0/0 | OS | Physician | 09/12/2016 | 2016 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6462376&pc=MTA | 2023826-2017-00585 | 2016-07-18 | Injury | STAAR SURGICAL COMPANY | 2017-04-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+4.5/109 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND EXCHANGED FOR A SHORTER LENS. THE EXCHANGE RESOLVED THE PROBLEM. ON (B)(6) 2016, PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS). (B)(4). | -13.0/+4.5/109 | OD | Physician | 10/14/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6105840&pc=MTA | 2023826-2016-01615 | 2016-07-19 | Injury | STAAR SURGICAL COMPANY | 2016-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/83 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. PATIENT CODE: (B)(4) SECONDARY SURGERY, LENS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM # (B)(4). | -11.0/+1.5/83 | OS | Physician | 10/25/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6105846&pc=MTA | 2023826-2016-01614 | 2016-07-19 | Injury | STAAR SURGICAL COMPANY | 2016-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.0/+3.5/91 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/22. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION WAS ALSO PERFORMED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+3.5/91 | OD | Physician | 10/25/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6001942&pc=MTA | 2023826-2016-01401 | 2016-07-20 | Injury | STAAR SURGICAL COMPANY | 2016-10-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+4.5/097 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. AS PER DFU FOR THIS LENS MODEL, VICLS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM. CLAIM #:(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -18.0/+4.5/097 | OS | Physician | 09/21/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6222369&pc=MTA | 2023826-2017-00018 | 2016-07-29 | Injury | STAAR SURGICAL COMPANY | 2017-01-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.00/+1.0/035 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND A SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.00/+1.0/035 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. AS OF (B)(6) 2016, PATIENT'S POST-OP BEST-CORRECTED AND UNCORRECTED VISUAL ACUITIES WERE 20/20. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL AND HAD CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND A PIECE OF LENS HAPTIC MISSING AND PRESENCE OF CLEAR AND WHITE RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -16.00/+1.0/035 | OS | Physician | 12/07/2016 | 2016 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6015344&pc=MTA | 2023826-2016-01412 | 2016-08-01 | Injury | STAAR SURGICAL COMPANY | 2016-10-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.50/0.0/0 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. (B)(4) (B)(6) . | -9.50/0.0/0 | OD | Physician | 09/12/2016 | 2016 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6486094&pc=MTA | 2023826-2017-00599 | 2016-08-01 | Injury | STAAR SURGICAL COMPANY | 2017-04-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+2.5/105 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAUTING ON THE SAME DAY. ON (B)(6) 2016, THE LENS WAS EXCHANGED FOR THE SHORTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND LENS HAPTIC BROKEN, AND BENT. (B)(4). | -15.0/+2.5/105 | OD | Physician | 11/16/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6111683&pc=MTA | 2023826-2016-01659 | 2016-08-04 | Injury | STAAR SURGICAL COMPANY | 2016-11-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+1.0/106 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016. A LONGER LENS WAS IMPLANTED ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THE PATIENT EXPERIENCED LOW VAULT. THE PATIENT'S POST-OP (ON (B)(6) 2016) UNCORRECTED VISUAL ACUITY WAS 20/25. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+1.0/106 | OD | Physician | 10/20/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5985709&pc=MTA | 2023826-2016-01354 | 2016-08-09 | Injury | STAAR SURGICAL COMPANY | 2016-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+1.0/X085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ON (B)(6) 2016, THE PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -13.0/+1.0/085 | OS | Physician | 08/30/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6015327&pc=MTA | 2023826-2016-01391 | 2016-08-10 | Injury | STAAR SURGICAL COMPANY | 2016-10-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00/2.5/084 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -14.00/2.5/084 | OS | Physician | 09/12/2016 | 2016 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6264663&pc=MTA | 2023826-2017-00105 | 2016-08-11 | Injury | STAAR SURGICAL COMPANY | 2017-01-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+3.0/096 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2016 THE LENS WAS REPLACED WITH A SHORTER LENS. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. Manufacturer Narrative: ADDITIONAL EVENT DESCRIPTION: PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20 AS OF (B)(6) 2016. DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | -13.5/+3.0/096 | OD | Physician | 12/23/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6133896&pc=MTA | 2023826-2016-01692 | 2016-08-12 | Injury | STAAR SURGICAL COMPANY | 2016-11-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.50/2.5/078 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC TORN/BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -17.50/2.5/078 | OD | Physician | 11/01/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6447618&pc=MTA | 2023826-2017-00536 | 2016-08-12 | Injury | STAAR SURGICAL COMPANY | 2017-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18/+2.0/X082 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXPLANTED DUE TO PATIENT EXCESSIVE VAULTING AND ELEVATED IOP. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN A GLASS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN A GLASS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS SURFACE. CLAIM# (B)(4). | -18/+2.0/082 | OS | Physician | 03/03/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6124351&pc=MTA | 2023826-2016-01670 | 2016-08-13 | Injury | STAAR SURGICAL COMPANY | 2016-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+1.0/088 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED. FOREIGN MATERIAL (DEPOSITS) WAS FOUND ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.50/+1.0/088 | OD | Physician | 10/25/2016 | 2016 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5992374&pc=MTA | 2023826-2016-01367 | 2016-08-23 | Injury | STAAR SURGICAL COMPANY | 2016-09-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS - -15.0/2.0/090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: CONCLUSION: REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. THIS LENS MODEL IS INDICATED FOR USE IN ADULTS 21 - 45 YEARS OF AGE. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.0/2.0/090 | OS | Physician | 09/06/2016 | 2016 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5996469&pc=MTA | 2023826-2016-01371 | 2016-08-24 | Injury | STAAR SURGICAL COMPANY | 2016-10-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS - -6.5/2.5/104 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ DEFORMED AND FOREIGN MATERIAL ON THE LENS SURFACE - HAIR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.5/2.5/104 | OD | Physician | 09/06/2016 | 2016 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6015351&pc=MTA | 2023826-2016-01413 | 2016-08-25 | Injury | STAAR SURGICAL COMPANY | 2016-10-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -5.00/0.0/0 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -5.00/0.0/0 | OS | Physician | 09/12/2016 | 2016 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6327713&pc=MTA | 2023826-2017-00120 | 2016-08-25 | Injury | STAAR SURGICAL COMPANY | 2017-02-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, - 10.0/+3.0/91 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE PATIENT'S LAST VISIT ON (B)(6) 2016 HAD A BEST CORRECTED VISUAL ACUITY (BCVA) OF 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND LENS HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL BELOW THE AGE OF 21 YEARS IS CONTRAINDICATED. THIS PATIENT'S AGE IS (B)(6). (B)(4). | - 10.0/+3.0/91 | OS | Physician | 01/17/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6156833&pc=MTA | 2023826-2016-01738 | 2016-08-30 | Injury | STAAR SURGICAL COMPANY | 2016-12-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.5/1.5/64 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID AND IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S SECONDARY SURGERY, LENS EXCHANGED FOR A LONGER LENS. (B)(4). | -5.5/1.5/64 | OS | Physician | 11/11/2016 | 2016 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6012314&pc=MTA | 2023826-2016-01405 | 2016-09-01 | Injury | STAAR SURGICAL COMPANY | 2016-10-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS - -10.50/2.0/082 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL; SURGICAL RESIDUE WAS CLEAR; VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.50/2.0/082 | OD | Physician | 09/12/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6094474&pc=MTA | 2023826-2016-01585 | 2016-09-01 | Injury | STAAR SURGICAL COMPANY AG | 2016-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -14.50/+2.50/048 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND PIGMENT DISPERSION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT OPTIC WAS TORN AND HAPTIC WAS TORN AND DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(6) BUT NOT MARKETED IN THE U.S. (B)(4). | -14.50/+2.50/048 | OD | Physician | 10/13/2016 | 2016 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6214738&pc=MTA | 2023826-2016-01842 | 2016-09-02 | Injury | STAAR SURGICAL COMPANY | 2016-12-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/108 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THERE WAS NO REPORT OF ANY PATIENT INJURY. Manufacturer Narrative: (B)(6) DATE OF EVENT: (B)(6) 2016. LENS WAS EXCHANGED ON (B)(6) 2016. AS OF THE DATE OF SUPPLEMENTAL MDR SUBMISSION, LENS HAS NOT YET BEEN EVALUATED. DEVICE MFR DATE: 11/04/2015; (CORRECTED)-11/09/2015. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT. CONCLUSION: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL OVER THE AGE OF 45 YEARS AND WITH ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM IS CONTRAINDICATED. THIS PATIENT WAS 47 YEARS OLD AT THE TIME OF IMPLANTATION AND HAD AN ACD OF 2.88MM. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.5/108 | OD | Physician | 11/29/2016 | 2016 | 47
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5991826&pc=MTA | 2023826-2016-01360 | 2016-09-05 | Injury | STAAR SURGICAL COMPANY | 2016-09-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/+1.5/161 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN/ DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.50/+1.5/161 | OS | Physician | 09/06/2016 | 2016 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6125983&pc=MTA | 2023826-2016-01679 | 2016-09-05 | Injury | STAAR SURGICAL COMPANY | 2016-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+1.00/130 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THE HAPTIC TO BE BROKEN AND THE OPTIC TORN. CORRECTED DATA: CONCLUSION CODE: PER DFU, THIS LENS MODEL IS CONTRAINDICATED IN THE PATIENT'S ANTERIOR CHAMBER DEPTH (ACD) BEING UNDER 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -11.50/+1.00/130 | OD | Physician | 10/28/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6410791&pc=MTA | 2023826-2017-00406 | 2016-09-05 | Injury | STAAR SURGICAL COMPANY | 2017-03-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/1.0/088 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED IN A SMALL VIAL, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS, BUT THERE IS DRY, CLEAR RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.0/088 | OS | Physician | 02/15/2017 | 2017 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6105836&pc=MTA | 2023826-2016-01613 | 2016-09-09 | Injury | STAAR SURGICAL COMPANY | 2016-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+1.5/120 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/12. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. PATIENT CODE: (B)(4) SECONDARY SURGERY, LENS EXCHANGED FOR A SHORTER LENS CLAIM # (B)(4). | -14.5/+1.5/120 | OS | Physician | 10/25/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6463861&pc=MTA | 2023826-2017-00578 | 2016-09-10 | Injury | STAAR SURGICAL COMPANY | 2017-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+2.0/083 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXCHANGED FOR SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THE DEVICE WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.0/+2.0/083 | OD | Physician | 03/09/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6048904&pc=MTA | 2023826-2016-01513 | 2016-09-17 | Injury | STAAR SURGICAL COMPANY | 2016-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS - -2.0/+1.5/169 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY AND BENT IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND LENS MEASUREMENTS WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S (B)(4). | -2.0/+1.5/169 | OS | Physician | 09/29/2016 | 2016 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6222057&pc=MTA | 2023826-2017-00010 | 2016-09-19 | Injury | STAAR SURGICAL COMPANY | 2017-01-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+1.5/046 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: ADDITIONAL DATA: WORK ORDER SEARCH: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -18.00/+1.5/046 | OD | Physician | 09/19/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6449798&pc=MTA | 2023826-2017-00531 | 2016-09-20 | Injury | STAAR SURGICAL COMPANY | 2017-03-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/093 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/+1.0/093 | OS | Physician | 11/08/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6453280&pc=MTA | 2023826-2017-00545 | 2016-09-21 | Injury | STAAR SURGICAL COMPANY | 2017-04-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+1.0/087 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -14.0/+1.0/087 | OS | Physician | 03/06/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6453282&pc=MTA | 2023826-2017-00544 | 2016-09-21 | Injury | STAAR SURGICAL COMPANY | 2017-04-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+5/092 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A SMALL ZIPLOC BAG. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS BUT DEBRIS ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION: OFF LABEL USE (ACD<3.00MM). (B)(4). | -13.5/+5/092 | OD | Physician | 03/06/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6060852&pc=MTA | 2023826-2016-01524 | 2016-09-22 | Injury | STAAR SURGICAL COMPANY | 2016-10-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.00/+2.0/087 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. BRAND NAME: THE CORRECT DIOPTER IS -17.00/+2.0/087. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. PATIENT'S ACD IS 2.81 MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -07.00/+2.0/087 | OD | Physician | 10/08/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6399841&pc=MTA | 2023826-2017-00383 | 2016-09-25 | Injury | STAAR SURGICAL COMPANY | 2017-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/X123 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. BASED ON SURGEON'S COMMENTS DOCTOR IS MONITORING THE PATIENT'S STATUS. AT THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR RESIDUE. (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/X123 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. THE DOCTOR IS MONITORING THE PATIENT'S STATUS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/+1.0/123 | OS | Physician | 02/11/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6094714&pc=MTA | 2023826-2016-01587 | 2016-09-28 | Injury | STAAR SURGICAL COMPANY | 2016-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+1.0/100 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PATIENT'S LAST VISIT WAS ON (B)(6) 2016 WITH UNCORRECTED VISUAL ACUITY (UCVA) 20/20. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -14.5/+1.0/100 | OD | Physician | 10/17/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6064956&pc=MTA | 2023826-2016-01542 | 2016-09-29 | Injury | STAAR SURGICAL COMPANY | 2016-10-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -7.5/+1.0/108 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A LENS THE SAME SIZE, DIFFERENT MODEL AND DIOPTER AND PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. CONCLUSION: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS UNDER 21 YEARS OLD. (B)(4). Manufacturer Narrative: PMA/510(K): N/A THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.5/+1.0/108 | OD | Physician | 09/30/2016 | 2016 | 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6463849&pc=MTA | 2023826-2017-00582 | 2016-09-29 | Injury | STAAR SURGICAL COMPANY | 2017-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+4.5/80 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXCHANGED FOR SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF HAPTIC BROKEN AND BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/+4.5/80 | OD | Physician | 03/07/2017 | 2017 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6463851&pc=MTA | 2023826-2017-00581 | 2016-09-29 | Injury | STAAR SURGICAL COMPANY | 2017-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+3.5/80 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXCHANGED FOR SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF HAPTIC BROKEN AND BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/+3.5/80 | OS | Physician | 03/07/2017 | 2017 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6094777&pc=MTA | 2023826-2016-01588 | 2016-09-30 | Injury | STAAR SURGICAL COMPANY | 2016-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+3.5/100 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PATIENT'S LAST VISIT WAS ON (B)(6) 2016 WITH BEST CORRECTED VISUAL ACUITY (BCVA) 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) 20/20. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.50/+3.5/100 | OS | Physician | 10/17/2016 | 2016 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6125955&pc=MTA | 2023826-2016-01638 | 2016-10-02 | Injury | STAAR SURGICAL COMPANY | 2016-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -18.00/+0.50/114 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. - WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED. (B)(4). LENS NOT RETURNED. | -18.00/+0.50/114 | OS | Physician | 10/28/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6449791&pc=MTA | 2023826-2017-00529 | 2016-10-02 | Injury | STAAR SURGICAL COMPANY | 2017-03-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/090 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/+2.5/090 | OD | Physician | 10/28/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6478568&pc=MTA | 2023826-2017-00547 | 2016-10-02 | Injury | STAAR SURGICAL COMPANY | 2017-04-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+2.5/087 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXCHANGED WITH A SHORTER LENS AND SIMILAR DIOPTER DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND SCRATCHES ON THE LENS AND DRY, CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4). | -11.00/+2.5/087 | OS | Physician | 10/28/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6140168&pc=MTA | 2023826-2016-01697 | 2016-10-10 | Injury | STAAR SURGICAL COMPANY | 2016-12-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, - 5.50/+5.5/005 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THERE WAS NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | - 5.50/+5.5/005 | OS | Physician | 11/04/2016 | 2016 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6631965&pc=MTA | 2023826-2017-00966 | 2016-10-11 | Injury | STAAR SURGICAL COMPANY | 2017-06-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.0/080 | OD | Physician | 05/14/2017 | 2017 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6060687&pc=MTA | 2023826-2016-01532 | 2016-10-12 | Injury | STAAR SURGICAL COMPANY | 2016-10-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/102 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT EXPERIENCED EXCESSIVE VAULTING OF THE LENS AND A SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. Event Description: THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENSE ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAPTIC BENT AND FIBER ON THE SURFACE. (B)(4). | -09.00/+1.5/102 | OS | Physician | 10/12/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6077588&pc=MTA | 2023826-2016-01597 | 2016-10-18 | Injury | STAAR SURGICAL COMPANY | 2016-11-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+2.50/091 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: PATIENT'S LAST VISIT WAS ON (B)(6) 2016 WITH UCVA 20/20. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CONCLUSION: AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.96MM AT THE TIME OF IMPLANTATION. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -12.00/+2.50/091 | OS | Physician | 10/19/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6077590&pc=MTA | 2023826-2016-01596 | 2016-10-18 | Injury | STAAR SURGICAL COMPANY | 2016-11-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.50/093 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CONCLUSION: AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH(ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.93MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -10.00/+1.50/093 | OD | Physician | 10/19/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6108112&pc=MTA | 2023826-2016-01618 | 2016-10-18 | Injury | STAAR SURGICAL COMPANY | 2016-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.00/+3.00/103 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -14.00/+3.00/103 | OD | Physician | 10/19/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6679291&pc=MTA | 2023826-2017-01092 | 2016-10-18 | Injury | STAAR SURGICAL COMPANY | 2017-06-29 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+5/91 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND EXCHANGED WITH SHORTER LENS OF SIMILAR MODEL. THIS RESOLVED THE PROBLEM AND THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 AND UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+5.0/91 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH SOME MOISTURE ON THE LENS. VISUAL INSPECTION FOUND A TEAR ON THE HAPTIC. (B)(4). Manufacturer Narrative: (B)(4). | -11.0/+5/91 | OS | Physician | 06/08/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6126258&pc=MTA | 2023826-2016-01675 | 2016-10-20 | Injury | STAAR SURGICAL COMPANY | 2016-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/80 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODES: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED FOR A SHORTER LENS. (B)(4). LENS NOT RETURNED. | -10.50/+1.5/80 | OD | Physician | 10/25/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6214492&pc=MTA | 2023826-2016-01847 | 2016-10-20 | Injury | STAAR SURGICAL COMPANY | 2016-12-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/90 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY AND BENT IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -12.0/+2.0/90 | OD | Physician | 11/30/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6684197&pc=MTA | 2023826-2017-01085 | 2016-10-20 | Injury | STAAR SURGICAL COMPANY | 2017-07-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/137 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CUSTOMER STATED THE PATIENT IS WEARING GLASSES, AND DOES NOT WANT TO IMPLANT ANOTHER LENS. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN, OPTIC BROKEN, AND FIBERS ON THE LENS. DIMENSIONAL INSPECTION WAS UNABLE TO BE PERFORMED DUE TO THE HAPTIC BEING BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.0/+1.0/137 | OS | Physician | 06/09/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6247110&pc=MTA | 2023826-2017-00069 | 2016-10-21 | Injury | STAAR SURGICAL COMPANY | 2017-01-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.50/+1.0/080 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: PRODUCT EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -8.50/+1.0/080 | OS | Physician | 12/23/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6247111&pc=MTA | 2023826-2017-00068 | 2016-10-21 | Injury | STAAR SURGICAL COMPANY | 2017-01-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.00/+1.0/090 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -9.00/+1.0/090 | OD | Physician | 12/23/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6573680&pc=MTA | 2023826-2017-00829 | 2016-10-26 | Injury | STAAR SURGICAL COMPANY | 2017-05-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, DIOPTER -12.0/1.5/118 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON THE SAME DATE THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO THE PATIENT EXPERIENCING LOW VAULT. IT IS UNKNOWN WHETHER THIS OCCURRED DURING THE SAME SURGERY OR NOT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL ON LENS SURFACE. (B)(4). Manufacturer Narrative: PREVIOUS: ON THE SAME DATE THE LENS WAS EXCHANGED FOR A LONGER LENS. CORRECTED: THE LENS WAS EXPLANTED ON THE SAME DATE OF THE IMPLANTATION. ON (B)(6) 2016, THE LENS WAS EXCHANGED FOR LONGER LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/1.5/118 | OD | Physician | 11/16/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6140107&pc=MTA | 2023826-2016-01695 | 2016-10-27 | Injury | STAAR SURGICAL COMPANY | 2016-12-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/2.00/94 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -10.0/2.00/94 | OD | Physician | 11/02/2016 | 2016 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6258420&pc=MTA | 2023826-2017-00086 | 2016-10-27 | Injury | STAAR SURGICAL COMPANY | 2017-01-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/091 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT EXPERIENCED EXCESSIVE VAULT. ON (B)(6) 2016, THE LENS WAS EXCHANGED WITH A SHORTER LENS AND AN ADDITIONAL SUTURE WAS PERFORMED. AS OF THE MDR SUBMISSION DATE, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | -12.5/+1.5/091 | OD | Physician | 12/22/2016 | 2016 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6172567&pc=MTA | 2023826-2016-01755 | 2016-10-28 | Injury | STAAR SURGICAL COMPANY | 2016-12-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.50/0.0/0 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -9.50/0.0/0 | OS | Physician | 11/14/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6179434&pc=MTA | 2023826-2016-01777 | 2016-10-28 | Injury | STAAR SURGICAL COMPANY | 2016-12-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/93 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. Manufacturer Narrative: ADDITIONAL INFORMATION: THE SUSPECT LENS WAS EXCHANGED ON (B)(6) 2016 FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) AS OF (B)(6) 2016 WAS 20/40. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THERE WAS NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -9.5/+2.0/93 | OD | Physician | 11/16/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6239382&pc=MTA | 2023826-2017-00044 | 2016-10-28 | Injury | STAAR SURGICAL COMPANY | 2017-01-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -7.50/+1.5/082 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT, AND THE LENS WERE RETURNED BENT. VISUAL INSPECTION FOUND HAPTIC BENT. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.50/+1.5/082 | OD | Physician | 12/20/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6258363&pc=MTA | 2023826-2017-00089 | 2016-10-29 | Injury | STAAR SURGICAL COMPANY | 2017-01-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.50/+1.5/082 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY AND BENT IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CONCLUSION: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -15.50/+1.5/082 | OS | Physician | 01/09/2017 | 2017 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6414316&pc=MTA | 2023826-2017-00432 | 2016-10-30 | Injury | STAAR SURGICAL COMPANY | 2017-03-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/107 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE LENS WAS THEN EXCHANGED FOR A SHORTER LENS. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS BUT CLEAR AND WHITE RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18.0/+1.5/107 | OS | Physician | 02/20/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6172541&pc=MTA | 2023826-2016-01751 | 2016-10-31 | Injury | STAAR SURGICAL COMPANY | 2016-12-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/86 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: CONCLUSION: REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH IS BELOW 3.0MM FOR VICL/VTICL. PRODUCT EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -11.50/+1.5/86 | OS | Physician | 11/14/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6140081&pc=MTA | 2023826-2016-01698 | 2016-11-02 | Injury | STAAR SURGICAL COMPANY | 2016-12-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50/1.0/069 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -08.50/1.0/069 | OS | Physician | 11/04/2016 | 2016 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6216085&pc=MTA | 2023826-2016-01849 | 2016-11-04 | Injury | STAAR SURGICAL COMPANY | 2016-12-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.0/+6.0/092 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS EXPLANTED DUE TO LOW VAULTING AND ROTATION. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -08.0/+6.0/092 | OD | Physician | 12/01/2016 | 2016 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6239044&pc=MTA | 2023826-2017-00043 | 2016-11-04 | Injury | STAAR SURGICAL COMPANY | 2017-01-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.0/+3.0/085 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(6). PRODUCT EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+3.0/085 | OS | Physician | 12/20/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6521074&pc=MTA | 2023826-2017-00708 | 2016-11-07 | Injury | STAAR SURGICAL COMPANY | 2017-04-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/5.5/111 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR RESIDUE ON LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.50/5.5/111 | OD | Physician | 03/28/2017 | 2017 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6218478&pc=MTA | 2023826-2017-00003 | 2016-11-08 | Injury | STAAR SURGICAL COMPANY | 2017-01-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/086 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -09.50/+1.5/086 | OD | Physician | 12/06/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6521075&pc=MTA | 2023826-2017-00702 | 2016-11-10 | Injury | STAAR SURGICAL COMPANY | 2017-04-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/4.5/091 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR AND BROWNISH/RED RESIDUE ON THE LENS. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.50/4.5/091 | OS | Physician | 03/28/2017 | 2017 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6491531&pc=MTA | 2023826-2017-00643 | 2016-11-11 | Injury | STAAR SURGICAL COMPANY | 2017-04-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -11.5/+2.0/086 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO THE PATIENT EXPERIENCING LOW VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR REPORTING THE LENS HAS NOT YET BEEN RETURNED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER | -11.5/+2.0/086 | OS | Physician | 03/15/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6329730&pc=MTA | 2023826-2017-00123 | 2016-11-16 | Injury | STAAR SURGICAL COMPANY | 2017-02-14 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+2.5/91 DIOPTER, WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON DATE (B)(6) 2016. ON DATE (B)(6) 2016, THE REPORTER INDICATED EXCESSIVE VAULTING AND WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2016. THIS EXCHANGE RESOLVED THE PROBLEM. ON THE PATIENT'S LAST VISIT DATE ON DATE (B)(6) 2016, BCVA WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.0/+2.5/91 | OD | Physician | 01/17/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6786758&pc=MTA | 2023826-2017-01194 | 2016-11-17 | Injury | STAAR SURGICAL COMPANY | 2017-08-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTED INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14/+1.5/6 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF ICA. THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S BCVA WAS 20/20 AND UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC CURVED AND RED AND CLEAR RESIDUE ON LENS SURFACE. (B)(4). | -14/+1.5/6 | OD | Physician | 07/17/2017 | 2017 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6447612&pc=MTA | 2023826-2017-00535 | 2016-11-20 | Injury | STAAR SURGICAL COMPANY | 2017-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17/+1.0/031 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED IN LIQUID IN A GLASS VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND PIECE OF HAPTIC MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17/+1.0/031 | OS | Physician | 03/03/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6218480&pc=MTA | 2023826-2017-00002 | 2016-11-24 | Injury | STAAR SURGICAL COMPANY | 2017-01-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.00/+1.5/094 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE WAS RETURNED AND EVALUATED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.00/+1.5/094 | OD | Physician | 12/06/2016 | 2016 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6312844&pc=MTA | 2023826-2017-00139 | 2016-11-25 | Injury | STAAR SURGICAL COMPANY | 2017-02-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICMO12.6MM IMPLANTABLE COLLAMER LENS, -13.0/+3.0/81 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2016, PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS ALSO CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -13.0/+3.0/81 | OD | Physician | 01/17/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6429994&pc=MTA | 2023826-2017-00448 | 2016-11-25 | Injury | STAAR SURGICAL COMPANY | 2017-03-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15/+2.0/079 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, NARROW ANGLE, AND SHALLOW AC. ON (B)(6) 2016 THE YAG WAS PERFORMED TO SOLVE THE VAULT ISSUE BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND THE HAPTIC MISSING A PIECE AND DEBRIS AND CLEAR RESIDUE ON LENS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND THE HAPTIC MISSING A PIECE AND DEBRIS AND CLEAR RESIDUE ON LENS. CLAIM # (B)(4). | -15/+2.0/079 | OS | Physician | 02/23/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6303742&pc=MTA | 2023826-2017-00109 | 2016-12-01 | Injury | STAAR SURGICAL COMPANY | 2017-02-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0/ +4.0/ *089 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING (VAULT VALUE OF 90 MICROMETERS). THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S, BUT NOT MARKETED IN THE U.S. (B)(4). NOT RETURNED TO MANUFACTURER. | -8.0/ +4.0/ 089 | OD | Physician | 01/15/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6361161&pc=MTA | 2023826-2017-00267 | 2016-12-05 | Injury | STAAR SURGICAL COMPANY | 2017-02-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/3/84 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SMALLER LENS AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2016, THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/25. Manufacturer Narrative: CORRECTED DATA: PATIENT DIDN'T HAVE AN ACD BELOW 3.0MM AT THE TIME OF IMPLANTATION. ADDITIONAL DATA: DEVICE EVALUATION-LENS RETURNED DRY IN SMALL CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. FOREIGN MATERIAL (CLEAR RESIDUE AND DEBRIS) WAS PRESENT ON LENS SURFACE. LENS IS CURVED DUE TO POSITION IN VIAL. PER DFU OF THIS LENS MODEL, VTICMOS ARE CONTRAINDICATED IN PATIENTS BELOW 21 YEARS OLD. THIS PATIENT WAS BELOW 21 AT THE TIME OF IMPLANTATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IT IS NOT MARKETED IN THE U.S. (B)(4). THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4). | -9.5/3/84 | OS | Physician | 01/31/2017 | 2017 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6413976&pc=MTA | 2023826-2017-00421 | 2016-12-05 | Injury | STAAR SURGICAL COMPANY | 2017-03-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+4.0/095 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A SMALL VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN/BENT/DEFORMED, DEBRIS/CLEAR RESIDUE ON LENS SURFACE, AND A PIECE OF HAPTIC MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/+4.0/095 | OD | Physician | 02/19/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6482396&pc=MTA | 2023826-2017-00602 | 2016-12-05 | Injury | STAAR SURGICAL COMPANY | 2017-04-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9/+2/102 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN SMALL VIAL. VISUAL INSPECTION FOUND HAPTIC TORN, SMALL PIECE OF HAPTIC WAS MISSING. (B)(4). | -9/+2/102 | OD | Physician | 03/13/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6373589&pc=MTA | 2023826-2017-00312 | 2016-12-06 | Injury | STAAR SURGICAL COMPANY | 2017-03-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16/2.5/105 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS BUT DID NOT RESOLVE THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: PATIENT IS UNDER THE AGE OF 21 AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS. (B)(4). | -16/2.5/105 | OS | Physician | 02/02/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6522307&pc=MTA | 2023826-2017-00714 | 2016-12-06 | Injury | STAAR SURGICAL COMPANY | 2017-04-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/93 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THE DEVICE WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND DEBRIS AND CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/+1.5/93 | OD | Physician | 03/30/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6325636&pc=MTA | 2023826-2017-00172 | 2016-12-07 | Injury | STAAR SURGICAL COMPANY | 2017-02-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/3.5/82 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING (VAULT VALUE OF 1107 MICROMETERS) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. ON PATIENT'S LAST VISIT ON (B)(6) 2017, THE PATIENT'S BEST CORRECTED VISUAL ACUITY WAS 20/22, AND THE UNCORRECTED VISUAL ACUITY WAS 20/25, WITH A VAULT OF 580 MICROMETERS. THE PRE-OPERATIVE DATA FOR THE PATIENT'S BEST CORRECTED VISUAL ACUITY IS 20/22, BUT THE UNCORRECTED VISUAL ACUITY IS 20/200 ON (B)(6) 2016. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED FOR A SHORTER LENS. (B)(4). LENS NOT RETURNED. | -11.5/3.5/82 | OS | Physician | 01/15/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6658289&pc=MTA | 2023826-2017-01008 | 2016-12-07 | Injury | STAAR SURGICAL COMPANY | 2017-06-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+2.5/090 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND MISSING PIECE OF HAPTIC, CLEAR RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+2.5/090 | OD | Physician | 05/24/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6373620&pc=MTA | 2023826-2017-00308 | 2016-12-08 | Injury | STAAR SURGICAL COMPANY | 2017-03-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16/2.5/102 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION OF THE LENS FOUND THE LENS OPTIC TORN. PIECES OF BOTH HAPTICS WERE TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR AND GREY RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGICAL INTERVENTION, LENS WAS EXPLANTED. THE PATIENT IS UNDER THE AGE OF 21 AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS. CLAIM #(B)(4). LENS NOT RETURNED. | -16/2.5/102 | OS | Physician | 02/02/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6373621&pc=MTA | 2023826-2017-00311 | 2016-12-08 | Injury | STAAR SURGICAL COMPANY | 2017-03-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -14/2.5/075 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) /2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGICAL INTERVENTION, LENS WAS EXPLANTED. PATIENT IS UNDER THE AGE OF 21 AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS. CLAIM #(B)(4). | -14/2.5/075 | OD | Physician | 02/02/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6658298&pc=MTA | 2023826-2017-01009 | 2016-12-08 | Injury | STAAR SURGICAL COMPANY | 2017-06-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+3.5/096 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE, AND CLEAR RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+3.5/096 | OS | Physician | 05/24/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6668791&pc=MTA | 2023826-2017-01046 | 2016-12-08 | Injury | STAAR SURGICAL COMPANY | 2017-06-26 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/012 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. Manufacturer Narrative: LENS WAS RETURNED DRY, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TEARS, CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGICAL INTERVENTION, LENS WAS EXCHANGED FOR A SHORTER LENS. CONCLUSION CODE: THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) WAS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4). LENS NOT RETURNED. | -13.0/+2.0/012 | OD | Physician | 06/08/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050140&pc=MTA | 2023826-2017-01772 | 2016-12-09 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -6.5/+2.0/101 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND CLEAR AND WHITE RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -6.5/+2.0/101 | OD | Physician | 10/27/2017 | 2017 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7051393&pc=MTA | 2023826-2017-01773 | 2016-12-09 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.5/+3.0/93 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED FOR SHORTER LENS. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | -8.5/+3.0/93 | OS | Physician | 10/27/2017 | 2017 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6442673&pc=MTA | 2023826-2017-00524 | 2016-12-12 | Injury | STAAR SURGICAL COMPANY | 2017-03-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/85 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXCHANGED FOR SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.5/+2.5/85 | OD | Physician | 03/01/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6327382&pc=MTA | 2023826-2017-00184 | 2016-12-13 | Injury | STAAR SURGICAL COMPANY | 2017-02-13 | MTA | ICL(IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -1.5/+2/+4 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING (VAULT VALUE OF 127 MICROMETERS). THE PATIENT EXPERIENCED A SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: EVALUATION FOUND LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN AND BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -1.5/+2/+4 | OD | Physician | 01/24/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6465203&pc=MTA | 2023826-2017-00577 | 2016-12-13 | Injury | STAAR SURGICAL COMPANY | 2017-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/1.5/108 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A SMALL VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC MISSING A PIECE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | -11.0/1.5/108 | OD | Physician | 03/08/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6271869&pc=MTA | 2023826-2017-00136 | 2016-12-16 | Injury | STAAR SURGICAL COMPANY | 2017-01-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+1.5/067 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -16.50/+1.5/067 | OS | Physician | 12/27/2016 | 2016 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6365367&pc=MTA | 2023826-2017-00282 | 2016-12-21 | Injury | STAAR SURGICAL COMPANY | 2017-02-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/2.0/090 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE SURGEON NOTED "VAULT REDUCED VERY WELL WITH EXCHANGING OF 12.6 MM LENS. PATIENT IS EXTREMELY HAPPY" ON (B)(6) 2017 WHICH WAS THE PATIENT'S LAST VISIT. Manufacturer Narrative: CONCLUSION CODE: PER DFU OF THIS LENS MODEL, VTICMOS ARE CONTRAINDICATED TO BE IMPLANTED INTO A KERATOCONUS EYE. DEVICE EVALUATION-LENS WAS RETURNED DRY IN A SMALL CENTRIFUGE VIAL. LENS HAD CLEAR SURGICAL RESIDUE/DEBRIS ON IT. VISUAL INSPECTION FOUND HAPTIC DEFORMED. APPEARS LENS WAS DAMAGED DURING HANDLING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/2.0/090 | OD | Physician | 02/01/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6478542&pc=MTA | 2023826-2017-00608 | 2016-12-22 | Injury | STAAR SURGICAL COMPANY | 2017-04-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+2.5/88 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR RESIDUE ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR RESIDUE ON THE LENS. CLAIM # (B)(4). | -14.5/+2.5/88 | OS | Physician | 03/12/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6399360&pc=MTA | 2023826-2017-00381 | 2016-12-23 | Injury | STAAR SURGICAL COMPANY | 2017-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50/3.5/078 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION-LENS RETURNED IN SMALL CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN, DRY RESIDUE ON LENS AND MISSING PIECE OF HAPTIC. LENS HAPTICS ARE BENT DUE TO POSITION IN VIAL. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.50/3.5/078 | OS | Physician | 02/11/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6399432&pc=MTA | 2023826-2017-00382 | 2016-12-23 | Injury | STAAR SURGICAL COMPANY | 2017-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50/3.5/101 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS RETURNED IN A SMALL VIAL. THERE WAS NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.50/3.5/101 | OD | Physician | 02/11/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6258514&pc=MTA | 2023826-2017-00090 | 2016-12-24 | Injury | STAAR SURGICAL COMPANY | 2017-01-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/093 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC SPLIT AND HAPTIC BROKEN. (B)(4). Manufacturer Narrative: FOLLOWUP #1 WAS INADVERTENTLY SKIPPED AND NOT SUBMITTED. FOLLOWUP #2 WAS SUBMITTED INSTEAD. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -12.50/+1.0/093 | OS | Physician | 01/09/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6404600&pc=MTA | 2023826-2017-00395 | 2016-12-24 | Injury | STAAR SURGICAL COMPANY | 2017-03-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+2.0/X110 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT, A PIECE OF THE HAPTIC MISSING, THE LENS IS CURVED, AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.0/+2.0/110 | OD | Physician | 02/16/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6800939&pc=MTA | 2023826-2017-01282 | 2016-12-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+2.0/096 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6)2016 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA 510(K):_THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4) | -15.0/+2.0/096 | OD |  | 07/22/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6498927&pc=MTA | 2023826-2017-00648 | 2016-12-30 | Injury | STAAR SURGICAL COMPANY | 2017-04-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.5/+1/076 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS NOT RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. REFERENCE MDR# 2023826-2017-00649 FOR PROBLEM RESOLUTION. Manufacturer Narrative: THE OUTCOMES ATTRIBUTED TO ADVERSE EVENT IN THE INITIAL MDR ARE NOT RELATED TO CONGENITAL ANOMALY/ BIRTH DEFECT. IT NEEDS TO BE UNCHECKED. THE DATE OF THE INITIAL REPORT IS NOT (B)(6) 2016, IT IS (B)(6) 2017. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN A LIQUID INSIDE A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS PRODUCT TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. CONCLUSION: OFF-LABEL USE (ACD <3.0MM). (B)(4). | -16.5/+1/076 | OD | Physician | 03/20/2017 | 2017 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6413665&pc=MTA | 2023826-2017-00425 | 2017-01-04 | Injury | STAAR SURGICAL COMPANY | 2017-03-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.0/4.0/78 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION-LENS WAS RETURNED IN A SMALL CENTRIFUGE VIAL. VISUAL INSPECTION FOUND STICKY CLEAR RESIDUE ON LENS SURFACE, AND SMALL TEAR IN THE OPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.0/4.0/78 | OD | Physician | 02/19/2017 | 2017 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6519083&pc=MTA | 2023826-2017-00700 | 2017-01-09 | Injury | STAAR SURGICAL COMPANY | 2017-04-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0/+2.0/92 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: CORRECTED DATA: IN MDR FOLLOW-UP #1, REFERRED TO CLAIM# (B)(4). SHOULD INSTEAD REFER TO CLAIM# (B)(4). (B)(4). Manufacturer Narrative: THE DEVICE WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISUAL DAMAGE TO THE LENS AND DEBRIS AND CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.0/+2.0/92 | OS | Physician | 03/30/2017 | 2017 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6522311&pc=MTA | 2023826-2017-00715 | 2017-01-09 | Injury | STAAR SURGICAL COMPANY | 2017-04-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+0.5/107 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: IN THE INITIAL MDR REPORT, THE CLAIM NUMBER CURRENTLY REFERENCED IS CLAIM# (B)(4). THE CORRECT CLAIM NUMBER SHOULD BE CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.5/+0.5/107 | OS | Physician | 03/30/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6300367&pc=MTA | 2023826-2017-00161 | 2017-01-10 | Injury | STAAR SURGICAL COMPANY | 2017-02-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.00/+2.00/76 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING (VAULT VALUE OF 1300 MICROMETERS). THE REPORTER STATED THE PATIENT DECIDED NOT TO BE IMPLANTED WITH A NEW ICL TEMPORARILY. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -17.00/+2.00/76 | OS | Physician | 01/12/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6311329&pc=MTA | 2023826-2017-00167 | 2017-01-10 | Injury | STAAR SURGICAL COMPANY | 2017-02-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/+3.50/096 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS 13.2 MM VTICMO13.2 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) IS 20/20 ON (B)(6) 2017. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) LENS NOT RETURNED. | -14.50/+3.50/096 | OS | Physician | 01/12/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6271794&pc=MTA | 2023826-2017-00108 | 2017-01-11 | Injury | STAAR SURGICAL COMPANY | 2017-01-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+3.5/90 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. DURING PATIENT'S LAST VISIT ON (B)(6) 2017, PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FIBERS ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -13.00/+3.5/90 | OD | Physician | 01/13/2017 | 2017 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8379529&pc=MTA | 2023826-2019-00306 | 2017-01-12 | Injury | STAAR SURGICAL COMPANY | 2019-02-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.5/+1.5/154 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO SIZING ISSUE. THE PATIENT HAD REDUCED NIGHT COMPLAINTS. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.5/+1.5/154 | OD | Physician | 01/29/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6855487&pc=MTA | 2023826-2017-01427 | 2017-01-17 | Injury | STAAR SURGICAL COMPANY | 2017-09-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -12.50/+1.00/092 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULTING. THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S POST-OP BCVA WAS 20/25 AND UCVA WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL AND HAD CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TO BE BENT AND CLEAR RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.50/+1.00/092 | OD | Physician | 08/14/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6785766&pc=MTA | 2023826-2017-01274 | 2017-01-19 | Injury | STAAR SURGICAL COMPANY | 2017-08-10 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+4.5/076 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/25. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -12.5/+4.5/076 | OS | Physician | 07/20/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6370471&pc=MTA | 2023826-2017-00304 | 2017-01-21 | Injury | STAAR SURGICAL COMPANY | 2017-03-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/1.5/089 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION-LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): SECONDARY SURGICAL INTERVENTION, LENS WAS EXCHANGED FOR SHORTER LENS. CONCLUSION: PATIENT IS UNDER THE AGE OF 21 AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS. (B)(4). LENS NOT RETURNED. | -12.5/1.5/089 | OS | Physician | 02/03/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6424132&pc=MTA | 2023826-2017-00453 | 2017-01-23 | Injury | STAAR SURGICAL COMPANY | 2017-03-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/2.5/080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION-LENS WAS RETURNED IN A SMALL CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS, AND STICKY CLEAR RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/2.5/080 | OD | Physician | 02/24/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6504956&pc=MTA | 2023826-2017-00664 | 2017-01-23 | Injury | STAAR SURGICAL COMPANY | 2017-04-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/089 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A -11.5/+1.5/009 DIOPTER LENS OF THE SAME LENGTH DUE TO EXCESSIVE VAULT. INITIAL VAULT WAS 2.0-2.5 CT AND VAULT AFTER EXCHANGE WAS 1.5 CT. PATIENT ALSO EXPERIENCED A SHALLOW ANTERIOR CHAMBER (10 UM). REPORTER STATED LENS WAS IMPLANTED VERTICALLY DUE TO THE "AXIS OF ASTIGMATISM." THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN LENS CASE/VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | -11.5/+1.5/089 | OD | Physician | 03/20/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6439125&pc=MTA | 2023826-2017-00509 | 2017-01-25 | Injury | STAAR SURGICAL COMPANY | 2017-03-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/95 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION-LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. CONCLUSION: OFF LABEL USE (ACD<3.00MM). (B)(4). | -12.5/+2.5/95 | OS | Physician | 02/28/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6669941&pc=MTA | 2023826-2017-01049 | 2017-01-28 | Injury | STAAR SURGICAL COMPANY | 2017-06-27 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/3.5/91 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/16. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON THE LENS. (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGICAL INTERVENTION, LENS WAS EXCHANGED FOR A LONGER LENS. (B)(4). | -13.5/3.5/91 | OS | Physician | 06/07/2017 | 2017 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6670651&pc=MTA | 2023826-2017-01051 | 2017-01-28 | Injury | STAAR SURGICAL COMPANY | 2017-06-27 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.5/2.0/80 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/16. Manufacturer Narrative: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: (B)(4). | -12.5/2.0/80 | OD | Physician | 06/07/2017 | 2017 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6668520&pc=MTA | 2023826-2017-01039 | 2017-02-01 | Injury | STAAR SURGICAL COMPANY | 2017-06-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+3.0/022 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND A MISSING PIECE OF HAPTIC AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+3.0/022 | OS | Physician | 06/01/2017 | 2017 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6507573&pc=MTA | 2023826-2017-00677 | 2017-02-08 | Injury | STAAR SURGICAL COMPANY | 2017-04-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.00/+1.5/81 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO LOW VAULT, AND THEN EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR AND WHITE RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/+1.5/81 | OD | Physician | 03/22/2017 | 2017 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6505004&pc=MTA | 2023826-2017-00665 | 2017-02-10 | Injury | STAAR SURGICAL COMPANY | 2017-04-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.5/092 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). | -10.0/+2.5/092 | OD | Physician | 03/20/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6463859&pc=MTA | 2023826-2017-00579 | 2017-02-12 | Injury | STAAR SURGICAL COMPANY | 2017-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/092 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A SMALL VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION: OFF LABEL USE (ACD<3.00MM). (B)(4). | -8.0/+1.0/092 | OD | Physician | 03/09/2017 | 2017 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6539256&pc=MTA | 2023826-2017-00745 | 2017-02-12 | Injury | STAAR SURGICAL COMPANY | 2017-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/098 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN A SMALL VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC MISSING A PIECE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+1.5/098 | OD | Physician | 04/04/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6410825&pc=MTA | 2023826-2017-00413 | 2017-02-13 | Injury | STAAR SURGICAL COMPANY | 2017-03-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.5/1.5/84 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.5/1.5/84 | OS | Physician | 02/19/2017 | 2017 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6410961&pc=MTA | 2023826-2017-00420 | 2017-02-13 | Injury | STAAR SURGICAL COMPANY | 2017-03-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/1.5/104 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: ADDITIONAL INFORMATION: PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY ON (B)(6) 2017 WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/1.5/104 | OD | Physician | 02/19/2017 | 2017 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6411150&pc=MTA | 2023826-2017-00426 | 2017-02-14 | Injury | STAAR SURGICAL COMPANY | 2017-03-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -8.5/2.0/098 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) /2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN, DEBRIS AND CLEAR AND BROWNISH RESIDUE ON LENS SURFACE, AND MISSING PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/2.0/098 | OD | Physician | 02/19/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6503651&pc=MTA | 2023826-2017-00662 | 2017-02-14 | Injury | STAAR SURGICAL COMPANY | 2017-04-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.5/+2.5/96 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: LENS WAS RETURNED DRY, IN SMALL VIAL. VISUAL INSPECTION FOUND MISSING PIECE OF HAPTIC, CLEAR, BROWNISH RESIDUE, AND DEBRIS ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/+2.5/96 | OD | Physician | 03/21/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6557196&pc=MTA | 2023826-2017-00785 | 2017-02-14 | Injury | STAAR SURGICAL COMPANY | 2017-05-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.0/80 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.5/+1.0/80 | OD | Physician | 04/18/2017 | 2017 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6557214&pc=MTA | 2023826-2017-00786 | 2017-02-14 | Injury | STAAR SURGICAL COMPANY | 2017-05-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.0/85 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION: OFF LABEL USE (BELOW 21 YEARS OF AGE AT THE TIME ON IMPLANT). (B)(4). | -13.5/+1.0/85 | OS | Physician | 04/18/2017 | 2017 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6721527&pc=MTA | 2023826-2017-01128 | 2017-02-14 | Injury | STAAR SURGICAL COMPANY AG | 2017-07-18 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICMO12.0 IMPLANTABLE COLLAMER LENS, -17.0/+4.0/93 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED WITH VTICMO13.2 AND THIS RESOLVED THE PROBLEM. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A MISSING PIECE OF THE HAPTIC AND CLEAR RESIDUE ON LENS. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -17.0/+4.0/93 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED WITH VTICMO13.2 AND THIS RESOLVED THE PROBLEM. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: (B)(4). | -17.0/+4.0/93 | OD | Physician | 06/28/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6522299&pc=MTA | 2023826-2017-00712 | 2017-02-15 | Injury | STAAR SURGICAL COMPANY | 2017-04-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.5/091 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THE DEVICE WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT, THE HAPTIC TORN, AND A CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.0/+2.5/091 | OD | Physician | 03/30/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6679638&pc=MTA | 2023826-2017-01101 | 2017-02-15 | Injury | STAAR SURGICAL COMPANY | 2017-06-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.0/+3.5/089 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR A LONGER LENS OF A SIMILAR DIOPTER ON (B)(6) 2017 DUE TO THE PATIENT EXPERIENCING LOW VAULT. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED AND UNCORRECTED VISUAL ACUITIES WERE 20/15 AS OF (B)(6) 2017. Manufacturer Narrative: THIS DEVICE IS NOT MARKETED IN THE U.S. (B)(4). | -13.0/+3.5/089 | OS | Physician | 06/07/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6619649&pc=MTA | 2023826-2017-00946 | 2017-02-20 | Injury | STAAR SURGICAL COMPANY | 2017-06-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/090 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (CLEAR AND BROWN RESIDUE) AND THE LENS MISSING A PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+1.5/090 | OD | Physician | 05/08/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6619661&pc=MTA | 2023826-2017-00947 | 2017-02-20 | Injury | STAAR SURGICAL COMPANY | 2017-06-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/089 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF HAPTIC AND OPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+1.5/089 | OS | Physician | 05/08/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6677388&pc=MTA | 2023826-2017-01071 | 2017-02-20 | Injury | STAAR SURGICAL COMPANY | 2017-06-29 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS WITH A -12.5/+1.0/085 DIOPTER IN THE RIGHT EYE (OD) OF THE (B)(6) YEARS OLD PATIENT. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS, A 12.1MM VICMO12.1 WITH A -12.5/+1.0/079 DIOPTER AND THE EXCHANGE RESOLVED THE PROBLEM. PATIENT'S ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) WAS 2.83. ON PATIENT'S LAST VISIT ON (B)(6) 2017 THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) IS 20/20 WITH A VAULT VALUE OF 1.5 MICROMETERS. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (WHITE, RED, AND CLEAR RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/085 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. \REPORT SOURCE: LITERATURE IS INCORRECT, SHOULD BE HEALTH PROFESSIONAL. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS NOT MARKETED IN THE U.S. (B)(4). CONCLUSION - PER DFU, REVIEW OF THIS FILE IS NOT INDICATED FOR VTICL WITH AN ANTERIOR ENDOTHELIUM CHAMBER DEPTH BELOW 3.0MM. ACD IS 2.83MM. (B)(4). | -12.5/+1.0/085 | OD | Physician | 06/09/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6677396&pc=MTA | 2023826-2017-01070 | 2017-02-20 | Injury | STAAR SURGICAL COMPANY | 2017-06-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS WITH A -14.5/+2.5/107 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6)2016. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS, 12.1MM VICMO12.1 WITH A -14.5/+2.5/111 DIOPTER AND THE EXCHANGE RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6)2017 HAD A BEST CORRECTED VISUAL ACUITY (BCVA) OF 20/20. Manufacturer Narrative: LENS WAS EXCHANGED DUE TO EXCESSIVE VAULTING. PREVIOUS COMMON DEVICE NAME VICMO12.6 IS INCORRECT; CORRECTED DATA IS PHAKIC INTRAOCULAR LENS FOR COMMON DEVICE NAME. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+2.5/107 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE) AND A TEAR ON HAPTIC. (B)(4). Manufacturer Narrative: PMA 510(K): PRODUCT NOT MARKETED IN THE U.S. (B)(4). ENDOTHELIUM CHAMBER DEPTH (ACD) DID NOT MEET REQUIREMENTS BECAUSE ANTERIOR ENDOTHELIUM CHAMBER DEPTH IS BELOW 3.0MM. (B)(4) | -14.5/+2.5/107 | OS | Physician | 06/09/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6752252&pc=MTA | 2023826-2017-01164 | 2017-02-20 | Injury | STAAR SURGICAL COMPANY | 2017-07-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.0/+2.0/88 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND THE LENS WAS EXCHANGED FOR A LENS WITH A DIFFERENT LENGTH AND SIMILAR DIOPTER. THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (RED AND CLEAR RESIDUE) AND THE HAPTICS WERE CURVED IN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.0/88 | OD | Physician | 07/10/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6764150&pc=MTA | 2023826-2017-01165 | 2017-02-20 | Injury | STAAR SURGICAL COMPANY | 2017-08-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -9.50/+2.0/88 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S POST-OP BCVA WAS 20/20 AND UCVA WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAS THE CURVED SHAPE, RED AND CLEAR RESIDUE ON THE PRODUCT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/+2.0/88 | OS | Physician | 07/10/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6658986&pc=MTA | 2023826-2017-01025 | 2017-02-21 | Injury | STAAR SURGICAL COMPANY | 2017-06-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/+2.5/079 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND CLEAR RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.0/+2.5/079 | OS | Physician | 05/26/2017 | 2017 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6721332&pc=MTA | 2023826-2017-01173 | 2017-02-21 | Injury | STAAR SURGICAL COMPANY | 2017-07-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/X119 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT YET BEEN RETURNED FOR EVALUATION. ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S, BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | -10.5/+1.0/119 | OS | Physician | 07/02/2017 | 2017 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6540648&pc=MTA | 2023826-2017-00736 | 2017-02-22 | Injury | STAAR SURGICAL COMPANY | 2017-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/2.0/081 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS BUT CLEAR RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -8.5/2.0/081 | OD | Physician | 04/04/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6721367&pc=MTA | 2023826-2017-01172 | 2017-02-22 | Injury | STAAR SURGICAL COMPANY | 2017-07-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/X055 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND WAS EXCHANGED FOR A LONGER LENS. THE EXCHANGE RESOLVED THE PROBLEM. ON (B)(6) 2017, THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. ON AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THIS PRODUCT IS MANUFACTURED IN THE U.S, BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.0/055 | OD |  | 07/02/2017 | 2017 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6498925&pc=MTA | 2023826-2017-00649 | 2017-02-26 | Injury | STAAR SURGICAL COMPANY | 2017-04-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+1.5/087 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CUSTOMER STATED ACCORDING TO THE DR. THE PATIENT IS WEARING GLASSES NOW, DID NOT LOSE BCVA, AND DOES NOT PLAN TO IMPLANT ANOTHER LENS. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. REFERENCE TO MDR# 2023826-2017-00648. Manufacturer Narrative: CORRECTED DATA: THE OUTCOMES ATTRIBUTED TO ADVERSE EVENT AS STATED IN THE INITIAL MDR DOES NOT INCLUDE CONGENITAL ANOMALY/BIRTH DEFECT. IT SHOULD BE LEFT UNCHECKED. THE DATE OF THE INITIAL REPORT WAS NOT (B)(6) 2016, BUT (B)(6) 2017. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. CONCLUSION: OFF-LABEL USE (ACD <3.0MM). (B)(4). | -16.5/+1.5/087 | OD | Physician | 03/20/2017 | 2017 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6458086&pc=MTA | 2023826-2017-00567 | 2017-03-01 | Injury | STAAR SURGICAL COMPANY | 2017-04-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.50/+1.5/143 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A SMALL VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS). CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | -6.50/+1.5/143 | OD | Physician | 03/08/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6463852&pc=MTA | 2023826-2017-00580 | 2017-03-01 | Injury | STAAR SURGICAL COMPANY | 2017-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/84 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCED EXCESSIVE VAULTING. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THAT THE LENS WAS WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THAT THE LENS WAS WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -12.5/+2.5/84 | OS | Physician | 03/08/2017 | 2017 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6723368&pc=MTA | 2023826-2017-01149 | 2017-03-02 | Injury | STAAR SURGICAL COMPANY | 2017-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11/2/090 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND WAS EXCHANGED FOR A SHORTER LENS. THE EXCHANGE RESOLVED THE PROBLEM. ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. (B)(4). Manufacturer Narrative: (SHOULD HAVE BEEN INCLUDED IN FOLLOW-UP #1) VISUAL INSPECTION FOUND LENS TORN AND CLEAR RESIDUE ON LENS. TORN PIECE WAS STUCK TO LENS. (B)(4). Manufacturer Narrative: ADDITONAL DATA: THIS PRODUCT IS MANUFACTURED IN THE U.S, BUT NOT MARKETED IN THE U.S. (B)(4). | -11/2/090 | OD | Physician | 07/05/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6723728&pc=MTA | 2023826-2017-01148 | 2017-03-02 | Injury | STAAR SURGICAL COMPANY | 2017-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14/1.5/091 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND WAS EXCHANGED FOR A SHORTER LENS. THE EXCHANGE RESOLVED THE PROBLEM. ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THIS PRODUCT IS MANUFACTURED IN THE U.S, BUT NOT MARKETED IN THE U.S. (B)(4). | -14/1.5/091 | OS | Physician | 07/05/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6680932&pc=MTA | 2023826-2017-01080 | 2017-03-06 | Injury | STAAR SURGICAL COMPANY | 2017-06-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/1.0/82 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR THE SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BENT AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR AND WHITE RESIDUE). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.5/1.0/82 | OD | Physician | 06/08/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6606359&pc=MTA | 2023826-2017-00898 | 2017-03-07 | Injury | STAAR SURGICAL COMPANY | 2017-06-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -18.0/+2.5/106 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS, DEBRIS ON THE LENS SURFACE. (B)(4). | -18.0/+2.5/106 | OD | Physician | 05/04/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6717887&pc=MTA | 2023826-2017-01198 | 2017-03-09 | Injury | STAAR SURGICAL COMPANY | 2017-07-17 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/1.0/083 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC WAS TORN AND PRESENCE OF CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/1.0/083 | OD | Physician | 06/21/2017 | 2017 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6491502&pc=MTA | 2023826-2017-00640 | 2017-03-10 | Injury | STAAR SURGICAL COMPANY | 2017-04-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -11.0/+2.0/086 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO THE PATIENT EXPERIENCING LOW VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR REPORTING THE LENS HAS NOT YET BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THERE WAS EVIDENCE OF DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THERE WAS EVIDENCE OF DEBRIS ON THE LENS. (B)(4). | -11.0/+2.0/086 | OS | Physician | 03/15/2017 | 2017 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6540637&pc=MTA | 2023826-2017-00735 | 2017-03-10 | Injury | STAAR SURGICAL COMPANY | 2017-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/1.0/077 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND CLEAR AND BROWN RESIDUE ON LENS SURFACE AND MISSING PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -15.00/1.0/077 | OD | Physician | 04/04/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6651809&pc=MTA | 2023826-2017-00995 | 2017-03-10 | Injury | STAAR SURGICAL COMPANY | 2017-06-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+3.0/090 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND CLEAR RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))]. (B)(4). | -12.5/+3.0/090 | OS | Physician | 05/24/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6521057&pc=MTA | 2023826-2017-00703 | 2017-03-13 | Injury | STAAR SURGICAL COMPANY | 2017-04-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -3.5/1.0/170 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE (DEBRIS, CLEAR AND BROWNISH RESIDUE). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR SHORTER LENS. (B)(4). | -3.5/1.0/170 | OS | Physician | 03/29/2017 | 2017 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6507567&pc=MTA | 2023826-2017-00675 | 2017-03-16 | Injury | STAAR SURGICAL COMPANY | 2017-04-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5/4.0/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THE DEVICE WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.5/4.0/085 | OS | Physician | 03/23/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6554052&pc=MTA | 2023826-2017-00780 | 2017-03-16 | Injury | STAAR SURGICAL COMPANY | 2017-05-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/+1.0/X095 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED WRAPPED IN GAUZE INSIDE A ZIPLOC BAG. VISUAL INSPECTION FOUND DEBRIS ON LENS SURFACE WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT AND PIECE OF HAPTIC MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.00/+1.0/095 | OD | Physician | 04/13/2017 | 2017 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7303802&pc=MTA | 2023826-2018-00297 | 2017-03-17 | Injury | STAAR SURGICAL COMPANY | 2018-02-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+4.5/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMAGE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+4.5/082 | OD | Physician | 01/29/2018 | 2018 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6669745&pc=MTA | 2023826-2017-01048 | 2017-03-20 | Injury | STAAR SURGICAL COMPANY | 2017-06-27 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.5/+2/092 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) LENS NOT RETURNED. | -11.5/+2/092 | OD | Physician | 06/08/2017 | 2017 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050124&pc=MTA | 2023826-2017-01767 | 2017-03-23 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.00/+2.5/077 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS CURVED IN AND PRESENCE OF DEBRIS AND CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.98MM)]. (B)(4). | -6.00/+2.5/077 | OD | Physician | 10/26/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6531855&pc=MTA | 2023826-2017-00728 | 2017-03-24 | Injury | STAAR SURGICAL COMPANY | 2017-04-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/089 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | -9.0/+1.5/089 | OD | Physician | 03/30/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6531870&pc=MTA | 2023826-2017-00729 | 2017-03-24 | Injury | STAAR SURGICAL COMPANY | 2017-04-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/087 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -8.5/+1.0/087 | OS | Physician | 03/30/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6673660&pc=MTA | 2023826-2017-01074 | 2017-03-26 | Injury | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.5/1.5/76 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION HAS BEEN COMPLETED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -13.5/1.5/76 | OD | Physician | 06/03/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6531867&pc=MTA | 2023826-2017-00730 | 2017-03-27 | Injury | STAAR SURGICAL COMPANY | 2017-04-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/082 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENTS POST-OP STATUS WAS NORMAL. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN THE LENS CONTAINER WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND LENS ARE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.5/+1.0/082 | OD | Physician | 03/30/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6637718&pc=MTA | 2023826-2017-00972 | 2017-04-01 | Injury | STAAR SURGICAL COMPANY | 2017-06-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0/+2.0/088 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON THE SAME DATE BUT NOT DURING THE SAME SURGERY, THE LENS WAS EXCHANGED FOR A SMALLER LENS DUE TO THE PATIENT EXPERIENCING EXCESSIVE VAULT. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A MISSING PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: PATIENT WEIGHT-UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NOT RETURNED TO MANUFACTURER. | -6.0/+2.0/088 | OD | Physician | 05/15/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6641758&pc=MTA | 2023826-2017-00980 | 2017-04-01 | Injury | STAAR SURGICAL COMPANY | 2017-06-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+2.5/079 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/+2.5/079 | OS | Physician | 05/15/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6786791&pc=MTA | 2023826-2017-01222 | 2017-04-06 | Injury | STAAR SURGICAL COMPANY | 2017-08-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.5/3.5/084 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. THERE WAS SOME MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (RED RESIDUE ON LENS) AND A PIECE OF HAPTIC WAS MISSING. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT NOT MARKETED IN THE U.S. (B)(4). | -16.5/3.5/084 | OS | Physician | 07/17/2017 | 2017 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6788891&pc=MTA | 2023826-2017-01223 | 2017-04-06 | Injury | STAAR SURGICAL COMPANY | 2017-08-11 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.5/4.0/085 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A PIECE OF HAPTIC MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -16.5/4.0/085 | OD | Physician | 07/17/2017 | 2017 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6959957&pc=MTA | 2023826-2017-01588 | 2017-04-06 | Injury | STAAR SURGICAL COMPANY | 2017-10-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+1.5/099 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR SHORTER LENS. (B)(4). | -17.5/+1.5/099 | OD | Physician | 09/23/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6684959&pc=MTA | 2023826-2017-01106 | 2017-04-07 | Injury | STAAR SURGICAL COMPANY | 2017-07-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -14.50/+2.5/170 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO THE PATIENT EXPERIENCING LOW VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TEARS, THE HAPTIC BENT AND DEFORMED, AND DEBRIS ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | -14.50/+2.5/170 | OD | Physician | 06/08/2017 | 2017 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6959996&pc=MTA | 2023826-2017-01589 | 2017-04-07 | Injury | STAAR SURGICAL COMPANY | 2017-10-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+1.5/086 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR SHORTER LENS. (B)(4). | -15.0/+1.5/086 | OS | Physician | 09/23/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6558494&pc=MTA | 2023826-2017-00787 | 2017-04-08 | Injury | STAAR SURGICAL COMPANY | 2017-05-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/109 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO THE PATIENT WAS EXPERIENCING EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND AN EXTREMELY SHALLOW A/C. THE PROBLEM IS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT AND FOREIGN MATERIAL (STAINS) ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NOT RETURNED TO MANUFACTURER. | -18.0/+1.5/109 | OD | Physician | 04/10/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6551457&pc=MTA | 2023826-2017-00778 | 2017-04-11 | Injury | STAAR SURGICAL COMPANY | 2017-05-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/X66 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES "SPECIALLY 12 HOURS", THE DR. CONCERNED. ON (B)(6) 2017 THE LENS WAS EXPLANTED. ACCORDING TO THE CUSTOMER, THE DR. PLANS TO IMPLANT ANOTHER LENS. THIS CASE IS UNDER MEDICAL CONSULT. AS OF THE DAY OF MDR SUBMISSION, THE LENS NOT BEEN RETURNED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION - PRODUCT EVALUATION FOUND LENS RETURNED DRY IN THE LENS CONTAINER WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN AND BENT. DIMENSIONAL INSPECTION FOUND LENS ARE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE SURGEON NOTED ANGLE CLOSURE WITHOUT ELEVATED IOP PRIOR TO LENS EXPLANTATION. THE PATIENT'S POST-OP BCVA WAS 20/25. CONCLUSION CODE: OFF LABEL USE (PATIENT HAD OCULAR SURGERY - CORNEAL RINGS FOR KERATOCONUS) (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/+1.5/66 | OS | Physician | 04/11/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6588672&pc=MTA | 2023826-2017-00881 | 2017-04-12 | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, DIOPTER -10.0/+1.5/077 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCING LOW VAULT. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT YET BEEN RETURNED. Manufacturer Narrative: ADDITIONAL INFORMATION: ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A LARGER LENS, SIMILAR DIOPTER. THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC CURLED UP. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT YET RETURNED. | -10.0/+1.5/077 | OD | Physician | 05/02/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6595515&pc=MTA | 2023826-2017-00888 | 2017-04-12 | Injury | STAAR SURGICAL COMPANY | 2017-05-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/2.0/098 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND CLEAR RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.5/2.0/098 | OS | Physician | 05/04/2017 | 2017 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6591451&pc=MTA | 2023826-2017-00883 | 2017-04-13 | Injury | STAAR SURGICAL COMPANY | 2017-05-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/077 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT YET BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT. (B)(4). Manufacturer Narrative: PATIENT WEIGHT-UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+2.0/077 | OS | Physician | 05/02/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7115054&pc=MTA | 2023826-2017-01891 | 2017-04-13 | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR SHORTER LENS. (B)(4). | -11.5/+1.0/085 | OS | Physician | 11/19/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7033593&pc=MTA | 2023826-2017-01722 | 2017-04-19 | Injury | STAAR SURGICAL COMPANY | 2017-11-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/171 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). | -18.0/+2.0/171 | OS | Physician | 10/19/2017 | 2017 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050899&pc=MTA | 2023826-2017-01704 | 2017-04-20 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -7.0/+2.5/175 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE A CURVED SHAPE AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4) SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, SECONDARY SURGERY, LENS EXCHANGE. (B)(4). | -7.0/+2.5/175 | OD | Physician | 10/24/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6682173&pc=MTA | 2023826-2017-00996 | 2017-04-22 | Injury | STAAR SURGICAL COMPANY | 2017-06-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, DIOPTER -0.5/+6/017 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCING EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | -0.5/+6/017 | OS | Physician | 06/02/2017 | 2017 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6658956&pc=MTA | 2023826-2017-01000 | 2017-04-24 | Injury | STAAR SURGICAL COMPANY | 2017-06-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/086 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE REPORTER STATED THAT APPROXIMATELY ONE MONTH PRIOR TO LENS EXCHANGE, IRIS OVALIZATION AND ANGLE CLOSURE OCCURRED. THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND A PIECE OF THE HAPTIC MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER | -11.5/+1.0/086 | OS | Physician | 05/25/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6677290&pc=MTA | 2023826-2017-01084 | 2017-04-24 | Injury | STAAR SURGICAL COMPANY | 2017-06-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -3.5/+2.0/092 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXCHANGED ON (B)(6) 2017 FOR A SHORTER LENS DUE TO THE PATIENT EXPERIENCING EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC MISSING A PIECE AND CURVED. (B)(4). Manufacturer Narrative: (B)(4). | -3.5/+2.0/092 | OD | Physician | 06/06/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6678897&pc=MTA | 2023826-2017-01093 | 2017-04-24 | Injury | STAAR SURGICAL COMPANY | 2017-06-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -4.0/+2.5/X106 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR SHORTER LENS OF A SIMILAR DIOPTER ON (B)(6) 2017 DUE TO THE PATIENT EXPERIENCING EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY ON (B)(6) 2017 WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE OPTIC TORN. (B)(4). Manufacturer Narrative: THIS DEVICE IS NOT MARKETED IN THE U.S. (B)(4). | -4.0/+2.5/106 | OS | Physician | 06/06/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6836408&pc=MTA | 2023826-2017-01385 | 2017-04-24 | Injury | STAAR SURGICAL COMPANY | 2017-08-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.0//+2.5/083 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS FOUND TO BE DISCOLORED AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0//+2.5/083 | OS | Physician | 08/04/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7115020&pc=MTA | 2023826-2017-01893 | 2017-04-25 | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50/+2.5/080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN, HAPTIC TORN, AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE AND DEBRIS ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR SHORTER LENS (B)(4). | -08.50/+2.5/080 | OD | Physician | 11/19/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7115021&pc=MTA | 2023826-2017-01892 | 2017-04-25 | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.00/+1.0/098 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS). (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) - SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR SHORTER LENS. (B)(4). | -6.00/+1.0/098 | OS | Physician | 11/19/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6645734&pc=MTA | 2023826-2017-00989 | 2017-04-26 | Injury | STAAR SURGICAL COMPANY | 2017-06-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/093 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SMALLER LENS DUE TO THE PATIENT EXPERIENCING EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NOT RETURNED TO MANUFACTURER. | -10.0/+1.5/093 | OD | Physician | 05/18/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6677377&pc=MTA | 2023826-2017-01059 | 2017-04-28 | Injury | STAAR SURGICAL COMPANY | 2017-06-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -12.5/+1/078 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2017. THE LENS WAS EXPLANTED ON (B)(6)2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4) Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). LENS NOT RETURNED. | -12.5/+1/078 | OD | Physician | 06/08/2017 | 2017 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6723384&pc=MTA | 2023826-2017-01146 | 2017-05-02 | Injury | STAAR SURGICAL COMPANY | 2017-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.00/+0.5/58 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND WAS EXCHANGED FOR A SHORTER LENS. THE EXCHANGE RESOLVED THE PROBLEM. ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20, AND THE PATIEN T'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE). (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THIS PRODUCT IS MANUFACTURED IN THE U.S, BUT NOT MARKETED IN THE U.S. (B)(4). | -8.00/+0.5/58 | OS | Physician | 07/04/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050110&pc=MTA | 2023826-2017-01766 | 2017-05-02 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.0/+1.5/169 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/DEFORMED, FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS), AND LENS WAS CAUGHT IN LENS VIAL CAP. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.0/+1.5/169 | OS | Physician | 10/26/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7245227&pc=MTA | 2023826-2018-00203 | 2017-05-02 | Injury | STAAR SURGICAL COMPANY | 2018-02-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL OBSERVATION FOUND THE LENS MISSING A PIECE OF HAPTIC AND RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -13.0/+1.0/089 | OD | Physician | 01/11/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7132129&pc=MTA | 2023826-2017-01854 | 2017-05-05 | Injury | STAAR SURGICAL COMPANY | 2017-12-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/028 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS). (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.0/028 | OS | Physician | 11/21/2017 | 2017 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7456986&pc=MTA | 2023826-2018-00594 | 2017-05-07 | Injury | STAAR SURGICAL COMPANY | 2018-04-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -13.5/+2.0/179 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS REMOVED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON THE LENS SURFACE. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | -13.5/+2.0/179 | OS | Physician | 03/27/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6684203&pc=MTA | 2023826-2017-01105 | 2017-05-10 | Injury | STAAR SURGICAL COMPANY | 2017-07-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/3.5/084 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A PIECE OF HAPTIC MISSING, AND OPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/3.5/084 | OS | Physician | 06/03/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050898&pc=MTA | 2023826-2017-01711 | 2017-05-10 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -11.50/+2.0/084 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FIBER AND STAINS ON THE LENS SURFACE. (B)(4). | -11.50/+2.0/084 | OD | Physician | 10/25/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6637214&pc=MTA | 2023826-2017-00974 | 2017-05-15 | Injury | STAAR SURGICAL COMPANY | 2017-06-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.5/+3.5/064 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ANGLE CLOSURE. ON (B)(6) 2017 THE LENS WAS EXPLANTED. THE SURGEON COMMENTED THAT HE IS PLANNING TO EXCHANGE A LENS FOR THE PATIENT. AS OF THE DAY OF MDR SUBMISSION THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: LENS WAS RETURNED DRY, IN A LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC BENT AND DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: FUNCTIONAL INSPECTION WAS NOT PERFORMED AS IT WASN'T REQUIRED FOR THIS CASE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.5/+3.5/064 | OD | Physician | 05/16/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6690912&pc=MTA | 2023826-2017-01116 | 2017-05-15 | Injury | STAAR SURGICAL COMPANY | 2017-07-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/050 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: PREVIOUS: LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. UPDATED: LENS HAS BEEN RETURNED AND EVALUATED. SHOULD HAVE LISTED 06/09/2017 IN PRIOR SUBMISSION. DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.5/050 | OD | Physician | 06/09/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6970593&pc=MTA | 2023826-2017-01602 | 2017-05-15 | Injury | STAAR SURGICAL COMPANY | 2017-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -17.00/+1.0/084 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ENLARGEMENT OF PI OR ADDITION OF NEW PI WAS PERFORMED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, SECONDARY SURGERY, ENLARGEMENT OF PI OR ADDITION OF NEW PI, LENS EXCHANGE. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | -17.00/+1.0/084 | OS | Physician | 09/27/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6970594&pc=MTA | 2023826-2017-01601 | 2017-05-15 | Injury | STAAR SURGICAL COMPANY | 2017-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -15.00/+2.5/089 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ENLARGEMENT OF PI OR ADDITION OF NEW PI WAS PERFORMED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4) SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, SECONDARY SURGERY, ENLARGEMENT OF PI OR ADDITION OF NEW PI, LENS EXCHANGE. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | -15.00/+2.5/089 | OD | Physician | 09/27/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6998012&pc=MTA | 2023826-2017-01683 | 2017-05-15 | Injury | STAAR SURGICAL COMPANY | 2017-11-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -7.0/+2.5/098 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF LENS HAPTIC BROKEN AND BENT. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/+2.5/098 | OD | Physician | 10/09/2017 | 2017 | 51
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6677391&pc=MTA | 2023826-2017-01050 | 2017-05-16 | Injury | STAAR SURGICAL COMPANY | 2017-06-29 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2MM IMPLANTABLE COLLAMER LENS WITH A -18.0/4.0/91 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2017. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6)2017 DUE TO EXCESSIVE VAULTING AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN, HAPTIC TORN, THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE), AND THE TEAR RUNS THROUGH THE HAPTIC AND OPTIC. (B)(4). Manufacturer Narrative: PMA 510(K):- THIS PRODUCT WAS NOT MANUFACTURED IN THE U.S. (B)(4) | -18.0/4.0/91 | OS | Physician | 06/04/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6677390&pc=MTA | 2023826-2017-01047 | 2017-05-17 | Injury | STAAR SURGICAL COMPANY | 2017-06-29 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS WITH A -12.5/2.0/92 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2017. THE LENS WAS EXPLANTED ON (B)(6)2017 DUE TO EXCESSIVE VAULTING. LENS WAS EXCHANGED ON (B)(6)2017 WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. ON PATIENT'S LAST VISIT ON (B)(6)2017 THE BEST CORRECTED VISUAL ACUITY IS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: PMA 510(K):- THIS PRODUCT WAS NOT MARKETED IN THE U.S. (B)(4). | -12.5/2.0/92 | OD | Physician | 06/04/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6677395&pc=MTA | 2023826-2017-01064 | 2017-05-17 | Injury | STAAR SURGICAL COMPANY | 2017-06-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS WITH A -12.0/1.50/85 DIOPTER IN THE PATIENT'S LEFT EYE (OD) ON (B)(6)2017. THE LENS WAS EXCHANGED FOR ANOTHER LENS WITH THE SAME SIZE AND DIOPTER AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ON PATIENT'S LAST VISIT ON (B)(6) 2017; PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. LENS WAS EXCHANGED DUE TO EXCESSIVE VAULTING. LENS WAS EXCHANGED FOR ANOTHER LENS WITH THE SAME SIZE AND DIOPTER IS INCORRECT; CORRECTED DATA IS THAT THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH SOME MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE, CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: ON PATIENT'S LAST VISIT ON (B)(6) 2017; PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. LENS WAS EXCHANGED DUE TO EXCESSIVE VAULTING. LENS WAS EXCHANGED FOR ANOTHER LENS WITH THE SAME SIZE AND DIOPTER IS INCORRECT; CORRECTED DATA IS THAT THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). | -12.0/1.50/85 | OS | Physician | 06/04/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7268606&pc=MTA | 2023826-2018-00223 | 2017-05-17 | Injury | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/1/92 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017 . THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A CENTRIFUGE VIAL. THERE WAS NO VISIBLE DAMAGE TO THE LENS. THERE WAS DEBRIS ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR A SHORTER LENS PATIENT CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). | -14.5/1/92 | OD | Physician | 01/17/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6658984&pc=MTA | 2023826-2017-01024 | 2017-05-19 | Injury | STAAR SURGICAL COMPANY | 2017-06-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.0/+3.5/087 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.0/+3.5/087 | OD | Physician | 05/25/2017 | 2017 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6661739&pc=MTA | 2023826-2017-01013 | 2017-05-19 | Injury | STAAR SURGICAL COMPANY | 2017-06-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -4.5/-1.5/089 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION - PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A SMALL SPECIMEN CUP AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND CLEAR RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -4.5/-1.5/089 | OS | Physician | 05/29/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6665392&pc=MTA | 2023826-2017-01014 | 2017-05-19 | Injury | STAAR SURGICAL COMPANY | 2017-06-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5/-1/090 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION - PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A SMALL SPECIMEN CUP AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND CLEAR RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5/-1/090 | OD | Physician | 05/29/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7113135&pc=MTA | 2023826-2017-01890 | 2017-05-19 | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/094 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+1.5/094 | OD | Physician | 11/19/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7131158&pc=MTA | 2023826-2017-01905 | 2017-05-20 | Injury | STAAR SURGICAL COMPANY | 2017-12-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -11.0/6.0/114 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL AND HAD CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | -11.0/6.0/114 | OD | Physician | 11/21/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6670140&pc=MTA | 2023826-2017-01032 | 2017-05-23 | Injury | STAAR SURGICAL COMPANY | 2017-06-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.5/+2.0/168 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN, THE HAPTIC BROKEN, AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.5/+2.0/168 | OD | Physician | 05/31/2017 | 2017 | 47
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6679382&pc=MTA | 2023826-2017-01073 | 2017-05-23 | Injury | STAAR SURGICAL COMPANY | 2017-06-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, DIOPTER -10.50/+1.50/091 IMPLANTABLE COLLAMER LENS (ICL), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: PREVIOUS: STATED THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. UPDATED: LENS HAS BEEN RETURNED AND EVALUATED. DEVICE EVALUATION: LENS WAS RETURNED DRY IN A SMALL VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT ((B)(6)), ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.99MM)]. (B)(4). | -10.50/+1.50/091 | OD | Physician | 06/02/2017 | 2017 | 47
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6847034&pc=MTA | 2023826-2017-01409 | 2017-05-23 | Injury | STAAR SURGICAL COMPANY | 2017-09-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -12.0/1.5/085 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. EXCESSIVE VAULT, PIGMENT DISPERSION, AND SHALLOW AC WAS OBSERVED. A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: THE LENS WILL REMAIN IMPLANT AND UNDER MONITORING. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/25. (B)(4). Manufacturer Narrative: TWO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: "DISABLILITY OR PERMANENT DAMAGE" WAS NOT REPORTED. IT IS NOT AN OUTCOME ATTRIBUTED TO AN ADVERSE EVENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/1.5/085 | OS | Physician | 08/10/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6657537&pc=MTA | 2023826-2017-01017 | 2017-05-24 | Injury | STAAR SURGICAL COMPANY | 2017-06-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+1.0/90 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE CUSTOMER STATED THE PATIENT DOES NOT WANT TO IMPLANT ANOTHER LENS AT THIS MOMENT. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: \DEVICE EVALUATION: UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO THE HAPTIC BEING MISSING. CLAIM #(B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN/BENT AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (FIBERS). DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CONCLUSION: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.0/+1.0/90 | OD | Physician | 05/24/2017 | 2017 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6658306&pc=MTA | 2023826-2017-01010 | 2017-05-24 | Injury | STAAR SURGICAL COMPANY | 2017-06-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/097 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL, WITH THICK RESIDUE ON LENS. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (THICK RESIDUE). (B)(4). Manufacturer Narrative: CONCLUSION: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+2.5/097 | OD | Physician | 05/24/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6679349&pc=MTA | 2023826-2017-01072 | 2017-05-24 | Injury | STAAR SURGICAL COMPANY | 2017-06-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, DIOPTER -9.50/+4/080 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO THE PATIENT EXPERIENCING EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: PREVIOUSLY STATED THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. UPDATED: THE LENS HAS BEEN RETURNED AND EVALUATED. DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAS A CURVED SHAPE. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | -9.50/+4/080 | OS | Physician | 06/01/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6670027&pc=MTA | 2023826-2017-01031 | 2017-05-25 | Injury | STAAR SURGICAL COMPANY | 2017-06-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/3.0/097 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE WITH SOME MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.5/3.0/097 | OS | Physician | 06/01/2017 | 2017 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6677100&pc=MTA | 2023826-2017-01081 | 2017-05-25 | Injury | STAAR SURGICAL COMPANY | 2017-06-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.50/1.5/072 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND EXCHANGED FOR A LONGER LENS. THE EXCHANGE RESOLVED THE PROBLEM. AS OF THE DATE OF WRITING THIS MDR, THE LENS HAS NOT YET BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN, FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE), AND THE LENS WAS IN A CURVED SHAPE. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: WEIGHT - UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S, BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -9.50/1.5/072 | OD | Physician | 06/03/2017 | 2017 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7115038&pc=MTA | 2023826-2017-01894 | 2017-05-27 | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+1.5/085 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC SMASHED AND RESIDUE ON THE LENS. IN ADDITION, THE LENS WAS CAUGHT IN THE LENS VIAL CAP. CORRECTION: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+1.5/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4)SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR SHORTER LENS. (B)(4). | -14.0/+1.5/085 | OD | Physician | 11/19/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7301499&pc=MTA | 2023826-2018-00295 | 2017-05-28 | Injury | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/105 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO VAULT WAS TOO HIGH AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A VICMO12.1 LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+2.0/105 | OD | Physician | 01/29/2018 | 2018 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6956701&pc=MTA | 2023826-2017-01587 | 2017-05-29 | Injury | STAAR SURGICAL COMPANY | 2017-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/082 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND EXCHANGED FOR A SHORTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20-2. Manufacturer Narrative: THE LENS WAS RETURNED IN A LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE HAPTICS BENT. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR SHORTER LENS. CONCLUSION CODE: OFF LABEL USE (ACD<3.0MM). (B)(4). | -12.0/+1.5/082 | OS | Physician | 09/18/2017 | 2017 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6955574&pc=MTA | 2023826-2017-01582 | 2017-06-05 | Injury | STAAR SURGICAL COMPANY | 2017-10-17 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A VTICM5_12.6 DIOPTER -10.50/+0.50/X002 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). CONCLUSION: 24-OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0 MM (2.87 MM)]. (B)(4). | -10.50/+0.50/002 | OD | Physician | 09/21/2017 | 2017 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6890298&pc=MTA | 2023826-2017-01471 | 2017-06-06 | Injury | STAAR SURGICAL COMPANY | 2017-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -9.50/1.0/079 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND A TEAR ON THE HAPTIC AND CLEAR RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.0/079 | OS | Physician | 08/29/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6950080&pc=MTA | 2023826-2017-01575 | 2017-06-06 | Injury | STAAR SURGICAL COMPANY | 2017-10-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+0.5/095 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND SHALLOW ANTERIOR CHAMBER DEPTH. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. THE LENS WAS PLACED VERTICALLY. Manufacturer Narrative: CORRECTED INFORMATION: PATIENT ID: (B)(6). (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+0.5/095 | OD | Physician | 09/18/2017 | 2017 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7308348&pc=MTA | 2023826-2018-00310 | 2017-06-06 | Injury | STAAR SURGICAL COMPANY | 2018-03-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.00/+2.5/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. EVENT DATE: (B)(6) 2017. THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.00/+2.5/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. IMPLANT DATE: (B)(\6) 2017. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LENS CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4). | -16.00/+2.5/080 | OS | Physician | 02/05/2018 | 2018 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6950084&pc=MTA | 2023826-2017-01574 | 2017-06-07 | Injury | STAAR SURGICAL COMPANY | 2017-10-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+2.0/095 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND SHALLOW ANTERIOR CHAMBER DEPTH. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/60. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN, FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS), TORN PIECE OF HAPTIC IS MISSING, AND THE LENS HAS A CURVED SHAPE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.0/+2.0/095 | OD | Physician | 09/18/2017 | 2017 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6691794&pc=MTA | 2023826-2017-01124 | 2017-06-08 | Injury | STAAR SURGICAL COMPANY | 2017-07-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/092 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE, AND CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICED NOT RETURNED TO MANUFACTURER. | -11.0/+1.5/092 | OS | Physician | 06/09/2017 | 2017 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6715652&pc=MTA | 2023826-2017-01136 | 2017-06-08 | Injury | STAAR SURGICAL COMPANY | 2017-07-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -11.0/1.5/094 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND TORN. THERE WAS CLEAR RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/1.5/094 | OS | Physician | 06/29/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6725266&pc=MTA | 2023826-2017-01129 | 2017-06-08 | Injury | STAAR SURGICAL COMPANY | 2017-07-19 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/102 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SMALLER LENS AND THIS RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/16.6 AND UNCORRECTED VISUAL ACUITY WAS 20/16.6. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: (B)(4). | -11.5/+1.5/102 | OD | Paramedic | 06/28/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6795665&pc=MTA | 2023826-2017-01317 | 2017-06-08 | Injury | STAAR SURGICAL COMPANY | 2017-08-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -3.5/3.5/166 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBELM WAS RESOLVED, AND THE DR. COMEMENTED THAT THE PATIENT IS DOING GOOD. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL AND THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND FIBERS PRESENT ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -3.5/3.5/166 | OD | Physician | 07/18/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594267&pc=MTA | 2023826-2018-00882 | 2017-06-10 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/109 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A LONGER LENS. THE VAULT VALUE PRIOR TO EXCHANGE IS REPORTED AS 120UM; AFTER EXCHANGE IS REPORTED AS 230UM. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4) | -8.0/+1.0/109 | OD | Physician | 05/15/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6832513&pc=MTA | 2023826-2017-01296 | 2017-06-13 | Injury | STAAR SURGICAL COMPANY AG | 2017-08-30 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A VTICM5_12.6 WITH A -9.5/+1.5/90 DIOPTER IN THE PATIENT ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, SO THE SURGEON REPLACED THE LENS WITH A SMALLER ONE AND A SIMILAR DIOPTER ON THE SAME DAY. IT IS UNKNOWN IF THIS OCCURRED DURING THE SAME SURGERY. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC TORN/BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). THIS PRODUCT IS CONTRAINDICATED FOR USE IN PATIENTS WITH AN ANTERIOR ENDOTHELIUM CHAMBER DEPTH LESS THAN 3.0 MM (2.99 MM). (B)(4). | -9.5/+1.5/90 | NA | Physician | 08/02/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7068476&pc=MTA | 2023826-2017-01794 | 2017-06-14 | Injury | STAAR SURGICAL COMPANY | 2017-11-29 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A VTICM5_13.2 DIOPTER -4.5/+3.0/X172 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, EXCESSIVE VAULT AND REDUCTION OF IRIDO-CORNEAL ANGLES (ICA) WAS REPORTED. A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: UPDATED INFORMATION: THE LENS HAS BEEN EXPLANTED; PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -4.5/+3.0/172 | OD | Physician | 10/31/2017 | 2017 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7068490&pc=MTA | 2023826-2017-01795 | 2017-06-14 | Injury | STAAR SURGICAL COMPANY | 2017-11-29 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A VTICM5_13.2 DIOPTER -4.5/+3.5/X175 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, EXCESSIVE VAULT AND REDUCTION OF IRIDO-CORNEAL ANGLES (ICA) WAS REPORTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: -UPDATED: LENS EXPLANTED; PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TO BE BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO CODE AVAILABLE (SIGNIFICANT REDUCTION OF ICA). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -4.5/+3.5/175 | OS | Physician | 10/31/2017 | 2017 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7301483&pc=MTA | 2023826-2018-00273 | 2017-06-15 | Injury | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/081 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED DUE EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.5/081 | OD | Physician | 01/29/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6719743&pc=MTA | 2023826-2017-01203 | 2017-06-16 | Injury | STAAR SURGICAL COMPANY | 2017-07-17 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.50/2/90 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TO BE BENT ALONG WITH DEBRIS AND CLEAR RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS. (B)(4). | -10.50/2/90 | OD | Physician | 07/06/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6719627&pc=MTA | 2023826-2017-01202 | 2017-06-21 | Injury | STAAR SURGICAL COMPANY | 2017-07-17 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/3.0/93 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) - SECONDARY SURGICAL INTERVENTION, LENS WAS EXCHANGED FOR SHORTER LENS. CLAIM #(B)(4). | -16.5/3.0/93 | OD | Physician | 07/05/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7213902&pc=MTA | 2023826-2018-00128 | 2017-06-23 | Injury | STAAR SURGICAL COMPANY AG | 2018-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2009. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND TWO LENSES RETURNED IN THE SAME VIAL. BOTH LENSES WERE RETURNED DRY WITH RESIDUE/DEBRIS ON THE LENSES. VISUAL INSPECTION FOUND THE HAPTIC OF THE FIRST LENS BENT AND DEBRIS AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC OF THE SECOND LENS BROKEN, A PIECE OF THE HAPTIC MISSING, AND DEBRIS AND RED RESIDUE ON THE LENS. REPORT SOURCE CHANGED FROM USER FACILITY TO (B)(6). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. (B)(4). | -12.00/+1.5/086 | OD | Physician | 01/15/2018 | 2018 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7516947&pc=MTA | 2023826-2018-00731 | 2017-06-27 | Injury | STAAR SURGICAL COMPANY | 2018-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/071 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND A TEAR IN THE LENS HAPTIC AND PRESENCE OF RESIDUE ON THE PRODUCT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.0/071 | OD | Physician | 04/16/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6720246&pc=MTA | 2023826-2017-01195 | 2017-06-28 | Injury | STAAR SURGICAL COMPANY | 2017-07-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -10/+2/X091 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND WAS EXCHANGED FOR A SHORTER LENS OF SIMILAR DIOPTER. THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL AND CLEAR RESIDUE WAS FOUND ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND BROKEN. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -10/+2/091 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. MODEL NO.: IT IS CORRECTED TO " VTICM5_13.2". (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S, BUT NOT MARKETED IN THE U.S. (B)(4). | -10/+2/091 | OD | Physician | 06/30/2017 | 2017 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6818715&pc=MTA | 2023826-2017-01332 | 2017-06-29 | Injury | STAAR SURGICAL COMPANY | 2017-08-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -11.5/+0.5/78 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHTT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED IN LIQUID IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/+0.5/78 | OD | Physician | 07/31/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050904&pc=MTA | 2023826-2017-01719 | 2017-07-02 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -11.0/+1.0/46 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT, FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS), AND THE LENS MISSING A PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/+1.0/46 | OD | Physician | 10/27/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6752238&pc=MTA | 2023826-2017-01151 | 2017-07-05 | Injury | STAAR SURGICAL COMPANY | 2017-07-28 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/3.0/086 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS OF THE SAME MODEL AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS CURVED IN AND PRESENCE OF CLEAR RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS BELOW THE AGE OF 21 YEARS AT THE TIME OF IMPLANTATION. (B)(4). | -11.5/3.0/086 | OD | Physician | 07/06/2017 | 2017 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6795069&pc=MTA | 2023826-2017-01308 | 2017-07-05 | Injury | STAAR SURGICAL COMPANY | 2017-08-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/2.5/088 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/100. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR AND BROWN RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -14.5/2.5/088 | OD | Physician | 08/10/2017 | 2017 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6850297&pc=MTA | 2023826-2017-01311 | 2017-07-05 | Injury | STAAR SURGICAL COMPANY | 2017-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/078 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE AND CLEAR SURGICAL RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF LABEL USE (PATIENT BELOW 21 YEARS OF AGE AT TIME OF IMPLANT). (B)(4). | -13.5/+1.5/078 | OS | Physician | 08/09/2017 | 2017 | 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6769393&pc=MTA | 2023826-2017-01270 | 2017-07-06 | Injury | STAAR SURGICAL COMPANY | 2017-08-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WAS RETURNED DRY, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTICS CURVED IN, CLEAR SURGICAL RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.5/090 | OS | Physician | 07/09/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6786780&pc=MTA | 2023826-2017-01220 | 2017-07-06 | Injury | STAAR SURGICAL COMPANY | 2017-08-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -7.0/2.0/066 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR A LARGER LENS ON SAME DAY DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTROINDICATED FOR IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | -7.0/2.0/066 | OD | Physician | 07/17/2017 | 2017 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6786787&pc=MTA | 2023826-2017-01221 | 2017-07-06 | Injury | STAAR SURGICAL COMPANY | 2017-08-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -5.0/1.5/106 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6)2017. THE LENS WAS EXCHANGED FOR A LARGER LENS ON SAME DAY DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS). (B)(4). Manufacturer Narrative: PMA 510)(K): THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | -5.0/1.5/106 | OS | Physician | 07/17/2017 | 2017 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7126717&pc=MTA | 2023826-2017-01910 | 2017-07-06 | Injury | STAAR SURGICAL COMPANY | 2017-12-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: REPORT SOURCE: (B)(6) IS INCORRECT, SHOULD BE (B)(6). (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.0/095 | OD | Physician | 12/04/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7144249&pc=MTA | 2023826-2017-01911 | 2017-07-06 | Injury | STAAR SURGICAL COMPANY | 2017-12-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT CONDITION POST-OP IS PUPIL IS SLIGHTLY DECENTERED INFERIOR. Manufacturer Narrative: REPORT SOURCE: (B)(6) IS INCORRECT, SHOULD BE (B)(6). (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR SHORTER LENS. (B)(4). SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. (B)(4). THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). | -12.0/+1.0/092 | OS | Physician | 12/04/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6850695&pc=MTA | 2023826-2017-01407 | 2017-07-10 | Injury | STAAR SURGICAL COMPANY | 2017-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS DIOPTER -12.0/+1.0/X072, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE PRESENT ON IT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.0/072 | OS | Physician | 08/09/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978043&pc=MTA | 2023826-2017-01621 | 2017-07-10 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -13.0/+2.5/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND NARROW ANGLES. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, NARROW ANGLES). (B)(4). | -13.0/+2.5/090 | OD | Physician | 09/29/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6872810&pc=MTA | 2023826-2017-01450 | 2017-07-11 | Injury | STAAR SURGICAL COMPANY | 2017-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -8.50/+3.0/92 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHTT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND WHITE DEBRIS ON LENS. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/+3.0/92 | OD | Physician | 08/23/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6800911&pc=MTA | 2023826-2017-01281 | 2017-07-15 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/112 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. DUE TO THE CILIARY BODY CYSTS AT 9 O'CLOCK, THE PATIENT'S VAULT WAS TOO HIGH. THE LENS WAS EXCHANGED FOR A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -10.5 DIOPTER ON (B)(6) 2017. AFTER REPLACEMENT OF ICL TO VERTICAL POSITION THE PROBLEM WAS RESOLVED AND THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS AT 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: PMA 510(K): THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4) | -11.0/+1.0/112 | OD | Physician | 07/23/2017 | 2017 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6918179&pc=MTA | 2023826-2017-01537 | 2017-07-15 | Injury | STAAR SURGICAL COMPANY | 2017-10-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -13.50/+4.0/092 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | -13.50/+4.0/092 | OS | Physician | 09/07/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6796035&pc=MTA | 2023826-2017-01305 | 2017-07-17 | Injury | STAAR SURGICAL COMPANY | 2017-08-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.00/+1.00/096 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE LENS EXCHANGE RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.00/+1.00/096 | OS | Physician | 07/21/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7051208&pc=MTA | 2023826-2017-01779 | 2017-07-17 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+2.5/088 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE REPORTER INDICATED THE 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 AND THE PATIENT IS IN GOOD CONDITION. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON LENS AND PRESENCE OF FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. (B)(4) Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+2.5/088 | OD | Physician | 10/30/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7051434&pc=MTA | 2023826-2017-01780 | 2017-07-17 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/+2.5/093 DIOPTER, IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. THE LENS WAS LOST AT THE FACILITY. Manufacturer Narrative: THE REPORTER INDICATED THE 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON 03/30/2017. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 AND THE PATIENT IS IN GOOD CONDITION. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/+2.5/093 | OS | Physician | 10/30/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7232876&pc=MTA | 2023826-2018-00156 | 2017-07-17 | Injury | STAAR SURGICAL COMPANY | 2018-01-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -16.0/+3.5/096 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A LONGER LENS, THE SURGEON NOTING LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTION: THE LENS WAS EXPLANTED WITH A LONGER LENS, THE SURGEON NOTED LOW VAULTING. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO THE LENS, LENS IS CURVED IN SHAPE, AND CONTAINS FOREIGN MATERIALS ON LENS SURFACE (RESIDUE ON LENS). THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -16.0/+3.5/096 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO LOW VAULT. POST-OPERATIVELY A LONGER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | -16.0/+3.5/096 | OS | Physician | 01/10/2018 | 2018 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6872816&pc=MTA | 2023826-2017-01442 | 2017-07-18 | Injury | STAAR SURGICAL COMPANY | 2017-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -11.50/+1.5/X085 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHTT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PATIENT POST-OP BCVA WAS 20/20 AND UCVA WAS 20/20 ON (B)(6) 2017. PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/+1.5/085 | OD | Physician | 08/21/2017 | 2017 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7131175&pc=MTA | 2023826-2017-01903 | 2017-07-18 | Injury | STAAR SURGICAL COMPANY | 2017-12-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICM5_13.7, -3.0/+4.5/79 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE AND CLEAR RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -3.0/+4.5/79 | OS | Physician | 11/21/2017 | 2017 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6813304&pc=MTA | 2023826-2017-01292 | 2017-07-20 | Injury | STAAR SURGICAL COMPANY | 2017-08-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.0/077 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -13.5/+1.0/077 | OD | Physician | 07/30/2017 | 2017 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6918164&pc=MTA | 2023826-2017-01538 | 2017-07-20 | Injury | STAAR SURGICAL COMPANY | 2017-10-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.50/+2.5/107 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.50/+2.5/107 | OS | Physician | 09/07/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7122689&pc=MTA | 2023826-2017-01896 | 2017-07-20 | Injury | STAAR SURGICAL COMPANY | 2017-12-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/135 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS), THE LENS MISSING A PIECE OF HAPTIC, AND THE LENS HAS A CURVED SHAPE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR LONGER LENS. PATIENT CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). | -8.0/+1.5/135 | OS | Physician | 11/27/2017 | 2017 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6812330&pc=MTA | 2023826-2017-01355 | 2017-07-21 | Injury | STAAR SURGICAL COMPANY | 2017-08-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+2.5/084 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE ON IT. THERE WAS ALSO CLEAR SURGICAL RESIDUE/DEBRIS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF HAPTIC AND CLEAR RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))]. (B)(4). | -13.5/+2.5/084 | OS | Physician | 07/30/2017 | 2017 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7150038&pc=MTA | 2023826-2017-01980 | 2017-07-24 | Injury | STAAR SURGICAL COMPANY | 2017-12-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/177 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS). (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -8.0/+1.5/177 | OD | Physician | 12/07/2017 | 2017 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7150076&pc=MTA | 2023826-2017-01981 | 2017-07-24 | Injury | STAAR SURGICAL COMPANY | 2017-12-27 | MTA | IMPLANATABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/003 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: PATIENT ACD IS LESS THAN 3.0MM AND PATIENT AGE WAS BELOW 21 YEARS OLD AT THE TIME OF IMPLANTATION. (B)(4). | -9.5/+1.5/003 | OS | Physician | 12/07/2017 | 2017 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6812369&pc=MTA | 2023826-2017-01356 | 2017-07-26 | Injury | STAAR SURGICAL COMPANY | 2017-08-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/084 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT, A PIECE OF HAPTIC MISSING AND CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))]. (B)(4). | -10.0/+1.5/084 | OD | Physician | 07/30/2017 | 2017 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6865940&pc=MTA | 2023826-2017-01455 | 2017-07-26 | Injury | STAAR SURGICAL COMPANY | 2017-09-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+3.0/102 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+3.0/102 | OD | Physician | 08/18/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6899558&pc=MTA | 2023826-2017-01486 | 2017-07-26 | Injury | STAAR SURGICAL COMPANY | 2017-09-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5/+0.5/165 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP CONDITION WAS: UCVA - 20/20 -2, INTRAOCULAR PRESSURE (IOP) WAS NORMAL AND VISION IS GOOD. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MINI CENTRIFUGE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS OPTIC TORN AND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE OTHER HAPTICS WERE BENT. THERE WAS EVIDENCE OF RED AND WHITE RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.5/+0.5/165 | OS | Physician | 09/01/2017 | 2017 | 53
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6956708&pc=MTA | 2023826-2017-01585 | 2017-07-26 | Injury | STAAR SURGICAL COMPANY | 2017-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/088 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (DEBRIS ON LENS) AND THE LENS MISSING A PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR SHORTER LENS. (B)(4). | -12.5/+2.0/088 | OD | Physician | 09/18/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6867156&pc=MTA | 2023826-2017-01435 | 2017-07-27 | Injury | STAAR SURGICAL COMPANY | 2017-09-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -2.5/+4.0/089 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH LIQUID ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -2.5/+4.0/089 | OD | Physician | 08/20/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6899539&pc=MTA | 2023826-2017-01507 | 2017-07-28 | Injury | STAAR SURGICAL COMPANY | 2017-09-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/090 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND CLEAR RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))] (B)(4). | -12.5/+2.5/090 | OS | Physician | 08/31/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7073558&pc=MTA | 2023826-2017-01822 | 2017-07-30 | Injury | STAAR SURGICAL COMPANY | 2017-11-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, DIOPTER -13.0/+0.5/116 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A LONGER LENS, THE SURGEON NOTING LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME LIQUID. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4) CONCLUSION CODE: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6)), ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.99MM)] (B)(4). | -13.0/+0.5/116 | OS | Physician | 11/02/2017 | 2017 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6899538&pc=MTA | 2023826-2017-01508 | 2017-07-31 | Injury | STAAR SURGICAL COMPANY | 2017-09-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+3.5/082 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WAS RETURNED A MICROCENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+3.5/082 | OD | Physician | 08/31/2017 | 2017 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7081601&pc=MTA | 2023826-2017-01848 | 2017-07-31 | Injury | STAAR SURGICAL COMPANY | 2017-12-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/092 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+2.5/092 | OD | Physician | 11/10/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7081606&pc=MTA | 2023826-2017-01849 | 2017-07-31 | Injury | STAAR SURGICAL COMPANY | 2017-12-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+3.0/088 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+3.0/088 | OS | Physician | 11/10/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7021007&pc=MTA | 2023826-2017-01699 | 2017-08-01 | Injury | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+2.0/088 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/33.3. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE ON IT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF ITS HAPTIC AND CLEAR RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.5/+2.0/088 | OD | Physician | 10/15/2017 | 2017 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6926695&pc=MTA | 2023826-2017-01546 | 2017-08-02 | Injury | STAAR SURGICAL COMPANY | 2017-10-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/096 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A LIQUID INSIDE A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+2.0/096 | OS | Physician | 09/11/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6970599&pc=MTA | 2023826-2017-01604 | 2017-08-03 | Injury | STAAR SURGICAL COMPANY | 2017-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICM5_13.7, -4.5/+0.5/X031 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, SECONDARY SURGERY, LENS EXCHANGE. (B)(4). | -4.5/+0.5/031 | OS | Physician | 09/27/2017 | 2017 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050098&pc=MTA | 2023826-2017-01768 | 2017-08-03 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/086 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6)2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS. SURGEON NOTES A VAULT VALUE OF 3.0 CT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: CONFIRMED LENS EXCHANGE WAS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND PRESENCE OF CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF ITS HAPTIC AND PRESENCE OF DEBRIS AND RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+2.0/086 | OS | Physician | 10/26/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978403&pc=MTA | 2023826-2017-01631 | 2017-08-04 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+1.5/088 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND THE LENS MISSING A PIECE OF ITS HAPTIC. THERE WAS PRESENCE OF CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF LABEL USE (PATIENT BELOW 21 YEARS OF AGE AT TIME OF IMPLANT) (B)(4). | -15.00/+1.5/088 | OD | Physician | 10/09/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978412&pc=MTA | 2023826-2017-01632 | 2017-08-04 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+2.0/096 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS CURVED IN AND WHITE RESIDUE PRESENT ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF LABEL USE (PATIENT BELOW 21 YEARS OF AGE AT TIME OF IMPLANT). (B)(4). | -14.5/+2.0/096 | OS | Physician | 10/09/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7213870&pc=MTA | 2023826-2018-00124 | 2017-08-04 | Injury | STAAR SURGICAL COMPANY AG | 2018-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0 MM TICM120V4 IMPLANTABLE COLLAMER LENS, -20.5/+3.5/079 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION OF THE LENS FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: PER DFU FOR THIS LENS MODEL, ICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. (B)(4). | -20.5/+3.5/079 | OD | Physician | 01/08/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7126804&pc=MTA | 2023826-2017-01908 | 2017-08-08 | Injury | STAAR SURGICAL COMPANY | 2017-12-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+5.0/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0/ +5.0/ 100 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO- CORNEAL ANGLES. A NEW SHORTER LENS WAS IMPLANTED LATER AND THE PROBLEM WAS RESOLVED. REPORT SOURCE: THE COUNTRY THIS REPORT ORIGINATED FROM IS CORRECTED FROM "(B)(6) " TO "(B)(6)" IN THE INITIAL MDR. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+5.0/100 | OD | Physician | 12/04/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7126854&pc=MTA | 2023826-2017-01909 | 2017-08-08 | Injury | STAAR SURGICAL COMPANY | 2017-12-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+4.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/ 4.0/ 88 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO- CORNEAL ANGLES. A NEW SHORTER LENS WAS IMPLANTED LATER AND THE PROBLEM WAS RESOLVED. REPORT SOURCE: THE COUNTRY THIS REPORT ORIGINATED FROM IS CORRECTED FROM "(B)(6)" TO "(B)(6)" IN THE INITIAL MDR. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+4.0/088 | OS | Physician | 12/04/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6926688&pc=MTA | 2023826-2017-01545 | 2017-08-09 | Injury | STAAR SURGICAL COMPANY | 2017-10-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.0/+3.0/075 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/10. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION - OFF LABEL USE (PATIENT BELOW 21 YEARS OF AGE AT TIME OF IMPLANT). (B)(4). | -6.0/+3.0/075 | OS | Physician | 09/11/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7113125&pc=MTA | 2023826-2017-01888 | 2017-08-09 | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.5/090 | OS | Physician | 11/19/2017 | 2017 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7001403&pc=MTA | 2023826-2017-01680 | 2017-08-10 | Injury | STAAR SURGICAL COMPANY | 2017-11-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/079 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF LABEL USE (PATIENT BELOW 21 YEARS OF AGE AT TIME OF IMPLANT). (B)(4). | -12.00/+2.0/079 | OD | Physician | 10/15/2017 | 2017 | 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7001500&pc=MTA | 2023826-2017-01681 | 2017-08-10 | Injury | STAAR SURGICAL COMPANY | 2017-11-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/105 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF ITS HAPTIC AND PRESENCE OF DEBRIS AND CLEAR RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF LABEL USE (PATIENT BELOW 21 YEARS OF AGE AT TIME OF IMPLANT) (B)(4). | -13.0/+2.0/105 | OS | Physician | 10/15/2017 | 2017 | 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135979&pc=MTA | 2023826-2017-01934 | 2017-08-10 | Injury | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.0/+1.50/84 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN, FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND RESIDUE ON LENS), AND THE LENS MISSING A PIECE OF HAPTIC. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -16.0/+1.50/84 | OS | Physician | 11/29/2017 | 2017 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7579088&pc=MTA | 2023826-2018-00862 | 2017-08-10 | Injury | STAAR SURGICAL COMPANY | 2018-06-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VITCMO12.1, -11.5/+1.0/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS DUE TO OBSERVATION OF LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION - LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | -11.5/+1.0/098 | OD | Physical Therapist | 05/09/2018 | 2018 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6950048&pc=MTA | 2023826-2017-01572 | 2017-08-14 | Injury | STAAR SURGICAL COMPANY | 2017-10-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0/+3.5/083 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF DEBRIS AND THE LENS HAD AN ORANGE TONE. CONCLUSION CODE: OFF LABEL USE (PATIENT BELOW 21 YEARS OF AGE AT TIME OF IMPLANT), (ACD<3.0MM) (B)(4). | -10.0/+3.5/083 | OS | Physician | 09/18/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6869684&pc=MTA | 2023826-2017-01437 | 2017-08-15 | Injury | STAAR SURGICAL COMPANY | 2017-09-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0/+0.5/106 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (RED RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+0.5/106 | OS | Physician | 08/20/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6937606&pc=MTA | 2023826-2017-01551 | 2017-08-16 | Injury | STAAR SURGICAL COMPANY | 2017-10-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -12.0/2/85 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4) SECONDARY SURGERY, LENS EXCHANGE. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT UNDER THE AGE OF 21 YEARS OLD. (B)(4). | -12.0/2/85 | OS | Physician | 09/13/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6867299&pc=MTA | 2023826-2017-01436 | 2017-08-17 | Injury | STAAR SURGICAL COMPANY | 2017-09-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/086 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN; FOREIGN MATERIAL ON LENS SURFACE (RED AND CLEAR RESIDUE); AND THE LENS MISSING A PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+1.5/086 | OD | Physician | 08/20/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7259941&pc=MTA | 2023826-2018-00171 | 2017-08-17 | Injury | STAAR SURGICAL COMPANY | 2018-02-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/6.0/103 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/6.0/103 | OS | Physician | 01/10/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978045&pc=MTA | 2023826-2017-01623 | 2017-08-18 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/X085 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON THE SAME DATE THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING, TWO PIECES OF THE HAPTIC TORN, A PIECE OF THE HAPTIC BENT, AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). (B)(4). | -13.0/+1.5/085 | OD | Physician | 09/28/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135227&pc=MTA | 2023826-2017-01921 | 2017-08-19 | Injury | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: RETURNED PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.00/+1.5/093 | OD | Physician | 11/24/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135230&pc=MTA | 2023826-2017-01922 | 2017-08-19 | Injury | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE PRODUCT WAS RETURNED IN LIQUID, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS WITH A HAPTIC TEAR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.0/092 | OS | Physician | 11/24/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6970596&pc=MTA | 2023826-2017-01599 | 2017-08-23 | Injury | STAAR SURGICAL COMPANY | 2017-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -13.5/2.0/94 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RED AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. | -13.5/2.0/94 | OS | Physician | 09/26/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7081258&pc=MTA | 2023826-2017-01813 | 2017-08-23 | Injury | STAAR SURGICAL COMPANY | 2017-12-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+3.5/091 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON 08(B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS DUE TO THE PATIENT DID NOT WANT A SECOND OPERATION. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND. (B)(4) Manufacturer Narrative: A DIMENSIONAL INSPECTION WAS PERFORMED. THE OVERALL LENS DIAMETER WAS THEN FOUND OUT OF TOLERANCE. THE LENS SN: (B)(4) WAS A 13.2MM IMPLANTABLE LENS THAT IS 0.27MM LARGER THAN THE THEORETICAL; THE TOLERANCE IS +/-0.20MM. SINCE THE WORK ORDER SEARCH DID NOT FIND ADDITIONAL SIMILAR ERRORS TO THIS COMPLAINT EVENT, IT CAN BE CONCLUDED THAT THE OOS CONDITION IS NOT LIKELY DUE TO MANUFACTURING OR PACKAGING. (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+3.5/091 | OD | Physician | 11/05/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7068473&pc=MTA | 2023826-2017-01798 | 2017-08-24 | Injury | STAAR SURGICAL COMPANY | 2017-11-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -13.5/+1.5/071 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: THE LENS EXCHANGE DID NOT RESOLVE THE PROBLEM. (B)(4) Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). CONCLUSION: 24-OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0 MM (2.97 MM)]. (B)(4). | -13.5/+1.5/071 | OD | Physician | 11/01/2017 | 2017 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7081264&pc=MTA | 2023826-2017-01814 | 2017-08-24 | Injury | STAAR SURGICAL COMPANY | 2017-12-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+4.0/089 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS DUE TO THE PATIENT DID NOT WANT A SECOND OPERATION. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN/BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+4.0/089 | OS | Physician | 11/05/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7081275&pc=MTA | 2023826-2017-01816 | 2017-08-26 | Injury | STAAR SURGICAL COMPANY | 2017-12-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/088 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS DUE TO THE PATIENT DID NOT WANT A SECOND OPERATION. Manufacturer Narrative: THE PATIENT IS IN NORMAL CONDITION AFTER REMOVAL OF THE LENS AND THERE IS NO COMPLAINT. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS). DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+2.0/088 | OD | Physician | 11/05/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7073568&pc=MTA | 2023826-2017-01823 | 2017-08-27 | Injury | STAAR SURGICAL COMPANY | 2017-11-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -12.0/+1.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: 24-OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))] (B)(4). | -12.0/+1.0/093 | OD | Physician | 11/02/2017 | 2017 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7081268&pc=MTA | 2023826-2017-01815 | 2017-08-27 | Injury | STAAR SURGICAL COMPANY | 2017-12-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/081 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS DUE TO THE PATIENT DID NOT WANT A SECOND OPERATION. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL ALONG WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND A TEAR IN THE LENS HAPTIC, LENS MISSING A PIECE OF ITS HAPTIC AND PRESENCE OF DEBRIS AND RESIDUE ON THE LENS SURFACE. METHOD CODE: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+2.0/081 | OS | Physician | 11/05/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050894&pc=MTA | 2023826-2017-01710 | 2017-08-28 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -14.0/+3.5/87 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. PATIENT CODE: (B)(4) -SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, SECONDARY SURGERY, LENS EXCHANGE. (B)(4). | -14.0/+3.5/87 | OS | Physician | 10/25/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6965643&pc=MTA | 2023826-2017-01610 | 2017-08-29 | Injury | STAAR SURGICAL COMPANY | 2017-10-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13..0/+1.50/X091 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED WITH A SHORTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: METHOD: THE PRODUCT WAS RETURNED DRY IN A LENS CASE/ VIAL AND HAD CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+1.50/091 | OS | Physician | 09/21/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978093&pc=MTA | 2023826-2017-01618 | 2017-09-01 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A VTICM5_12.6 DIOPTER -10.5/+2.5/090 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL AND HAD CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -10.5/+2.5/090 | OS | Physician | 09/28/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978103&pc=MTA | 2023826-2017-01617 | 2017-09-01 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A VTICM5_12.6 DIOPTER -10.5/+2.5/090 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF ITS HAPTIC AND PRESENCE OF CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -10.5/+2.5/090 | OD | Physician | 09/28/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7069418&pc=MTA | 2023826-2017-01806 | 2017-09-01 | Injury | STAAR SURGICAL COMPANY | 2017-11-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/085 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND A PIECE OF THE HAPTIC MISSING. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+2.0/085 | OD | Physician | 11/05/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6908139&pc=MTA | 2023826-2017-01491 | 2017-09-04 | Injury | STAAR SURGICAL COMPANY | 2017-10-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -9.00/+1.0/59 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC TORN/BENT/BROKEN/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4) SECONDARY SURGERY, LENS EXCHANGE. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | -9.00/+1.0/59 | OD | Physician | 09/06/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135616&pc=MTA | 2023826-2017-01931 | 2017-09-05 | Injury | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -9.00/+1.50/173 DIOPTER, IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL AND HAD CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -9.00/+1.50/173 | OS | Physician | 11/27/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7149219&pc=MTA | 2023826-2017-01955 | 2017-09-05 | Injury | STAAR SURGICAL COMPANY | 2017-12-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.0/+1.0/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS). (B)(4). Manufacturer Narrative: INITIAL MDR DATE IS INCORRECT, SHOULD BE 07 DEC 2017. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.0/+1.0/083 | OD | Physician | 12/07/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7156013&pc=MTA | 2023826-2017-01969 | 2017-09-05 | Injury | STAAR SURGICAL COMPANY | 2017-12-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+1.50/099 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE AND DEBRIS ON LENS). (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.89MM)]. (B)(4). | -17.0/+1.50/099 | OS | Physician | 12/06/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7021356&pc=MTA | 2023826-2017-01666 | 2017-09-07 | Injury | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -7.5/+2.5/092 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION: 24- PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | -7.5/+2.5/092 | OS | Physician | 10/17/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7069468&pc=MTA | 2023826-2017-01804 | 2017-09-07 | Injury | STAAR SURGICAL COMPANY | 2017-11-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+2.0/096 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR LONGER LENS. (B)(4). | -15.5/+2.0/096 | OD | Physician | 11/05/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6997295&pc=MTA | 2023826-2017-01653 | 2017-09-10 | Injury | STAAR SURGICAL COMPANY | 2017-11-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -8.0/+1.5/92 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND, NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF CLEAR AND RED RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.0/+1.5/92 | OS | Physician | 10/11/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7169829&pc=MTA | 2023826-2018-00009 | 2017-09-10 | Injury | STAAR SURGICAL COMPANY | 2018-01-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.5/+2.0/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS). (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). (B)(4). | -16.5/+2.0/086 | OD | Physician | 12/10/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7211359&pc=MTA | 2023826-2018-00114 | 2017-09-10 | Injury | STAAR SURGICAL COMPANY | 2018-01-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICM5_12.1, -11.0/+1.0/82 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND CLEAR RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/+1.0/82 | OS | Physician | 12/26/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6938513&pc=MTA | 2023826-2017-01550 | 2017-09-12 | Injury | STAAR SURGICAL COMPANY | 2017-10-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -14.5/1.5/70 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND LENS CURVED IN AND PRESENCE OF DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.5/1.5/70 | OD | Physician | 09/13/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7156018&pc=MTA | 2023826-2017-01971 | 2017-09-12 | Injury | STAAR SURGICAL COMPANY | 2017-12-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS MISSING TWO PIECES OF HAPTIC WITH A CLEAR AND RED RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.0/095 | OS | Physician | 12/07/2017 | 2017 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7516760&pc=MTA | 2023826-2018-00743 | 2017-09-13 | Injury | STAAR SURGICAL COMPANY | 2018-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+2.0/160 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ON (B)(6) 2017 A SHORTER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.0/+2.0/160 | OD | Physician | 04/20/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050897&pc=MTA | 2023826-2017-01709 | 2017-09-15 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -14.0/+3.5/097 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED. (B)(4). | -14.0/+3.5/097 | OD | Physician | 10/25/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7051442&pc=MTA | 2023826-2017-01783 | 2017-09-15 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.0/+2.00/111 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND A TEAR IN THE HAPTIC, THE HAPTICS BENT AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+2.00/111 | OD | Physician | 10/30/2017 | 2017 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7065663&pc=MTA | 2023826-2017-01785 | 2017-09-15 | Injury | STAAR SURGICAL COMPANY | 2017-11-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.50/+2.00/061 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A PIECE OF THE LENS MISSING AND DEBRIS ON LENS. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.50/+2.00/061 | OS | Physician | 10/30/2017 | 2017 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6952446&pc=MTA | 2023826-2017-01565 | 2017-09-17 | Injury | STAAR SURGICAL COMPANY | 2017-10-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -12.50/2/94 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (WHITE AND CLEAR RESIDUE ON LENS). (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM AND PATIENTS UNDER THE AGE OF 21 YEARS OLD. (B)(4). | -12.50/2/94 | OD | Physician | 09/19/2017 | 2017 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7507551&pc=MTA | 2023826-2018-00689 | 2017-09-17 | Injury | STAAR SURGICAL COMPANY | 2018-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.0/+1.0/84 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(46 2017. THE LENS WAS REMOVED ON (B)(6) 2018 DUE TO OBSERVATION OF EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: RETURNED TO MANUFACTURER: 27-APRIL-2018. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: LENS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS SURFACE. (B)(4). | -10.0/+1.0/84 | OD | Physician | 04/11/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7021458&pc=MTA | 2023826-2017-01671 | 2017-09-19 | Injury | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -6.5/2.0/090 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND CURVED IN AND CLEAR AND RED RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT UNDER THE AGE OF 21 YEARS OF AGE. (B)(4). | -6.5/2.0/090 | OS | Physician | 10/17/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7026723&pc=MTA | 2023826-2017-01739 | 2017-09-19 | Injury | STAAR SURGICAL COMPANY | 2017-11-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/093 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL AND HAD CLEAR SURGICAL RESIDUE/ DEBRIS ON IT. VISUAL INSPECTION FOUND A PIECE OF LENS HAPTIC MISSING AND PRESENCE OF CLEAR RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR SHORTER LENS. (B)(4). | -12.5/+2.5/093 | OD | Physician | 10/16/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7070534&pc=MTA | 2023826-2017-01805 | 2017-09-19 | Injury | STAAR SURGICAL COMPANY | 2017-11-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+3.5/099 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: \DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+3.5/099 | OD | Physician | 11/05/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7484279&pc=MTA | 2023826-2018-00651 | 2017-09-19 | Injury | STAAR SURGICAL COMPANY | 2018-05-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SAME LENGTH, DIFFERENT MODEL LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) (B)(4). | -12.5/+1.0/080 | OS | Physician | 04/06/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7026709&pc=MTA | 2023826-2017-01738 | 2017-09-21 | Injury | STAAR SURGICAL COMPANY | 2017-11-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/089 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN, A PIECE OF THE HAPTIC MISSING, AND CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+2.5/089 | OS | Physician | 10/16/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7033553&pc=MTA | 2023826-2017-01726 | 2017-09-22 | Injury | STAAR SURGICAL COMPANY | 2017-11-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, DIOPTER -9.5/+2.5/095 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR AND WHITE RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION: 24-OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0 MM (2.84 MM)]. CLAIM# (B)(4). | -9.5/+2.5/095 | OS | Physician | 10/20/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7001473&pc=MTA | 2023826-2017-01677 | 2017-09-23 | Injury | STAAR SURGICAL COMPANY | 2017-11-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.5/072 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO IMPLANTABLE COLLAMER LENS -13.5/2.5/072 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXCHANGED ON (B)(6) 2017 DUE TO SIGNIFICANT REDUCTION IN IRIDO CORNEAL ANGLES AND "ARCH TOO HIGH". THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF DEBRIS, RED AND CLEAR RESIDUE ON LENS SURFACE. (B)(4). METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+2.5/072 | OD | Physician | 10/15/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7144283&pc=MTA | 2023826-2017-01914 | 2017-09-23 | Injury | STAAR SURGICAL COMPANY | 2017-12-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -8.0/+5.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR A SHORTER LENS. (B)(4). THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). | -8.0/+5.0/088 | OS | Physician | 12/04/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978086&pc=MTA | 2023826-2017-01619 | 2017-09-25 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A VTICM5_13.7 DIOPTER -5.5/+1.0/090 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, SIGNIFICANT REDUCTION OF ICA). (B)(4). | -5.5/+1.0/090 | OD | Physician | 09/28/2017 | 2017 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6993831&pc=MTA | 2023826-2017-01656 | 2017-09-25 | Injury | STAAR SURGICAL COMPANY | 2017-11-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/073 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | -9.0/+1.0/073 | OD | Physician | 10/05/2017 | 2017 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6993832&pc=MTA | 2023826-2017-01657 | 2017-09-25 | Injury | STAAR SURGICAL COMPANY | 2017-11-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/123 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | -9.5/+2.5/123 | OS | Physician | 10/05/2017 | 2017 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7219324&pc=MTA | 2023826-2018-00129 | 2017-09-25 | Injury | STAAR SURGICAL COMPANY | 2018-01-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.00/+2.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND PRESENCE OF RED RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LENS. CLAIM # (B)(4). | -6.00/+2.0/094 | OD | Physician | 01/15/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7286012&pc=MTA | 2023826-2018-00244 | 2017-09-25 | Injury | STAAR SURGICAL COMPANY | 2018-02-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -13.5/1/170 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANT ON (B)(6) 2017 DUE TO LOW VAULTING. THERE IS NO PLAN TO IMPLANT ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND A LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: DEVICE MANUFACTURE DATE: IT IS CORRECTED FROM "25-AUG-2016" TO "24-AUG-2016" IN THE INITIAL MDR. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/1/170 | OD | Physician | 01/23/2018 | 2018 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7035808&pc=MTA | 2023826-2017-01745 | 2017-09-26 | Injury | STAAR SURGICAL COMPANY | 2017-11-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+3.0/081 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) AND UNCORRECTED VISUAL ACUITY (UCVA) WERE FOUND TOP BE 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS) AND THE LENS MISSING A PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED IN THE US. (B)(4). | -8.5/+3.0/081 | OS | Physician | 10/16/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7197905&pc=MTA | 2023826-2018-00076 | 2017-09-26 | Injury | STAAR SURGICAL COMPANY | 2018-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -11.0/+2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO LOW VAULTING. THERE IS NO PLANS TO IMPLANT ANOTHER LENS IN THE PATIENT'S EYE. Manufacturer Narrative: DATE OF EVENT: IT IS CORRECTED FROM "(B)(6) 2017" TO "(B)(6) 2017". DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+2.5/093 | OD | Physician | 12/19/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7197910&pc=MTA | 2023826-2018-00077 | 2017-09-26 | Injury | STAAR SURGICAL COMPANY | 2018-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -11.0/+1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO LOW VAULTING. THERE IS NO PLANS TO IMPLANT ANOTHER LENS IN THE PATIENT'S EYE. Manufacturer Narrative: DATE OF EVENT: IT IS CORRECTED FROM "(B)(6) 2017" TO "(B)(6) 2017". DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+1.5/093 | OS | Physician | 12/19/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7109987&pc=MTA | 2023826-2017-01876 | 2017-09-27 | Injury | STAAR SURGICAL COMPANY | 2017-12-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/091 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON 01-NOV-2017 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT/DEFORMED; FOREIGN MATERIAL ON LENS SURFACE (RED AND CLEAR RESIDUE ON LENS) AND THE LENS MISSING A PIECE OF HAPTIC. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.5/091 | OD | Physician | 11/20/2017 | 2017 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7129291&pc=MTA | 2023826-2017-01928 | 2017-09-27 | Injury | STAAR SURGICAL COMPANY | 2017-12-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VITCM5_13.2 IMPLANTABLE COLLAMER LENS, -7.0/2.0/97 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL AND HAD CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF HAPTIC AND PRESENCE OF CLEAR AND RED RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LENS. CLAIM# (B)(4). | -7.0/2.0/97 | OS | Physician | 11/20/2017 | 2017 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7211132&pc=MTA | 2023826-2018-00113 | 2017-09-28 | Injury | STAAR SURGICAL COMPANY | 2018-01-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -10.0/+3.5/85 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHTT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING AND CLEAR RESIDUE ON LENS. (B)(4) Manufacturer Narrative: IN SUPPLEMENTAL MDR #1, THE DATE RECEIVED BY MANUFACTURER IS 01 FEB 2018. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT UNDER THE AGE OF 21 YEARS OLD. (B)(4). | -10.0/+3.5/85 | OD | Physician | 12/26/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7211142&pc=MTA | 2023826-2018-00112 | 2017-09-28 | Injury | STAAR SURGICAL COMPANY | 2018-01-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -9.50/+3.5/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND CLEAR RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT UNDER THE AGE OF 21 YEARS OLD. (B)(4). | -9.50/+3.5/94 | OS | Physician | 12/26/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7020569&pc=MTA | 2023826-2017-01694 | 2017-09-29 | Injury | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+3.5/089 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+3.5/089 | OS | Physician | 10/15/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7034166&pc=MTA | 2023826-2017-01741 | 2017-09-29 | Injury | STAAR SURGICAL COMPANY | 2017-11-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/1.0/105 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) AND UNCORRECTED VISUAL ACUITY (UCVA) WAS FOUND TO BE 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL AND HAD CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND A TEAR IN THE LENS HAPTIC AND PRESENCE OF CLEAR RESIDUE AND DEBRIS ON THE PRODUCT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.00/1.0/105 | OD | Physician | 10/16/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7034989&pc=MTA | 2023826-2017-01742 | 2017-09-29 | Injury | STAAR SURGICAL COMPANY | 2017-11-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50/1.5/090 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) AND UNCORRECTED VISUAL ACUITY (UCVA) WERE FOUND TO BE 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED IN THE US. PATIENT CODE (B)(4): SECONDARY SURGICAL INTERVENTION; LENS EXCHANGE. (B)(4). | -13.50/1.5/090 | OS | Physician | 10/16/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6987332&pc=MTA | 2023826-2017-01635 | 2017-09-30 | Injury | STAAR SURGICAL COMPANY | 2017-10-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -12.50/+1.0/102 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS). (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT UNDER THE AGE OF 21 YEARS OLD. (B)(4). | -12.50/+1.0/102 | OD | Physician | 10/04/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7211140&pc=MTA | 2023826-2018-00111 | 2017-09-30 | Injury | STAAR SURGICAL COMPANY | 2018-01-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -14.0/+1.5/63 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL AND HAD CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF CLEAR RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | -14.0/+1.5/63 | OD | Physician | 12/26/2017 | 2017 | 95
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7346773&pc=MTA | 2023826-2018-00423 | 2017-10-01 | Injury | STAAR SURGICAL COMPANY | 2018-03-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.50/+2.5/099 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON LENS. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF HAPTIC AND RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.50/+2.5/099 | OD | Physician | 02/17/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320727&pc=MTA | 2023826-2018-00315 | 2017-10-02 | Injury | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+0.5/052 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM IMPLANTABLE COLLAMER LENS OF -7.0/+0.5/052 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. A SHORTER LENS WAS IMPLANTED LATER AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))]. (B)(4). | -7.0/+0.5/052 | OS | Physician | 02/08/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7021088&pc=MTA | 2023826-2017-01698 | 2017-10-10 | Injury | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+4.0/094 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE LENS WAS TORN IN TWO PIECES. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF LABEL USE (PATIENT BELOW 21 YEARS OF AGE AT TIME OF IMPLANT). (B)(4). | -12.0/+4.0/094 | OS | Physician | 10/15/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7144271&pc=MTA | 2023826-2017-01913 | 2017-10-10 | Injury | STAAR SURGICAL COMPANY | 2017-12-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -7.5/+1.0/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A VICMO13.7 LENS (-7.5 DIOPTER) AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR ANOTHER LENS PATIENT CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). | -7.5/+1.0/080 | OS | Physician | 12/04/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320763&pc=MTA | 2023826-2018-00330 | 2017-10-10 | Injury | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+3.5/076 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND NARROW ANTERIOR CHAMBER DEPTH. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: (B)(6) THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, NARROW ANTERIOR CHAMBER DEPTH). (B)(4). | -15.5/+3.5/076 | OS | Physician | 02/08/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7481808&pc=MTA | 2023826-2018-00629 | 2017-10-10 | Injury | STAAR SURGICAL COMPANY | 2018-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -10.50/1.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED ON 20 OCT 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/1.5/085 | OS | Physician | 04/05/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7579219&pc=MTA | 2023826-2018-00861 | 2017-10-13 | Injury | STAAR SURGICAL COMPANY | 2018-06-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VITCMO12.6, -15.5/+2.0/067 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS DUE TO OBSERVATION OF LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: ON (B)(6) 2017 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS DUE TO OBSERVATION OF LOW VAULT. PROBLEM WAS NOT RESOLVED. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | -15.5/+2.0/067 | OS | Physical Therapist | 05/09/2018 | 2018 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7113157&pc=MTA | 2023826-2017-01889 | 2017-10-16 | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/031 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF HAPTIC WITH CLEAR RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.5/031 | OD | Physician | 11/19/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7484285&pc=MTA | 2023826-2018-00652 | 2017-10-18 | Injury | STAAR SURGICAL COMPANY | 2018-05-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND TEARS IN THE LENS HAPTIC, HAPTIC BENT AND PRESENCE OF RESIDUE ON THE LENS SURFACE. IMPLANTED DATE: IT IS CORRECTED TO (B)(6) 2017. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) (B)(4). | -11.5/+1.5/086 | OD | Physician | 04/06/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7720844&pc=MTA | 2023826-2018-01123 | 2017-10-18 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LENS. (B)(4). | -13.0/+2.0/092 | OD | Physician | 06/25/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7073159&pc=MTA | 2023826-2017-01817 | 2017-10-19 | Injury | STAAR SURGICAL COMPANY | 2017-11-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.00/+1.00/159 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 AND THE PATIENT'S EYE WAS NORMAL. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A LIQUID INSIDE A LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE - (B)(4)- SECONDARY SURGICAL INTERVENTION, EXPLANTED AND EXCHANGED FOR A SHORTER LENS. (B)(4). | -16.00/+1.00/159 | OD | Physician | 11/06/2017 | 2017 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7089427&pc=MTA | 2023826-2017-01832 | 2017-10-23 | Injury | STAAR SURGICAL COMPANY | 2017-12-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1, -9.5/1.5/083 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | -9.5/1.5/083 | OS | Physician | 11/07/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135208&pc=MTA | 2023826-2017-01915 | 2017-10-23 | Injury | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+3.5/093 SPHERE/CYLINDAR/AXIS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THERE ARE NO KNOWN PLANS TO IMPLANT ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC SPLIT. DIMENSIONAL INSPECTION UNABLE TO BE PERFORMED. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.0/+3.5/093 | OD | Physician | 11/23/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050153&pc=MTA | 2023826-2017-01770 | 2017-10-25 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+2.00/150 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR AND RED RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.00/+2.00/150 | OD | Physician | 10/27/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7114686&pc=MTA | 2023826-2017-01861 | 2017-10-31 | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -13.0/+1.5/86 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND A TEAR IN THE HAPTIC, THE LENS MISSING A PIECE OF ITS HAPTIC, AND PRESENCE OF CLEAR RESIDUE ON THE LENS SURFACE. DATE RECEIVED BY MANUFACTURER: INITIAL MDR STATED 27-NOV-1989. IT IS CORRECTED TO 14-NOV-2017. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.0/+1.5/86 | OD | Physician | 11/14/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7481827&pc=MTA | 2023826-2018-00638 | 2017-10-31 | Injury | STAAR SURGICAL COMPANY | 2018-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7, -11.0/+3.0/062 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO CORNEAL ANGLES WAS NOTICED. ATTEMPTS AT OBTAINING ADDITIONAL INFORMATION HAVE NOT BEEN FORTHCOMING. Manufacturer Narrative: IT IS CORRECTED TO"ADVERSE EVENT". ADDITIONALLY, ANGLE CLOSURE WAS ALSO NOTICED IN THE LOWER QUADRANT OF THE EYE. IT IS CORRECTED TO "SERIOUS INJURY". (B)(4). Manufacturer Narrative: CORRECTED DATA: IT IS CORRECTED TO "PRODUCT PROBLEM". IT IS CORRECTED TO "MALFUNCTION". CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+3.0/062 | OS | Physician | 04/05/2018 | 2018 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7085063&pc=MTA | 2023826-2017-01819 | 2017-11-03 | Injury | STAAR SURGICAL COMPANY | 2017-12-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.5/+2.5/091 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/22.2 AND INTRAOCULAR PRESSURE WAS 18.3MM HG. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE WITH MOISTURE ON THE LENS.VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+2.5/091 | OS | Physician | 11/06/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7122681&pc=MTA | 2023826-2017-01895 | 2017-11-06 | Injury | STAAR SURGICAL COMPANY | 2017-12-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.0/110 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. AN ADDITIONAL YAG WAS PERFORMED FOR ENLARGEMENT/ADDITION OF PERIPHERAL IRIDECTOMIES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THE ANGLES AND VAULT WERE OK. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS IS CURVED AND DEBRIS AND CLEAR RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | -18.0/+1.0/110 | OD | Physician | 11/27/2017 | 2017 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7515936&pc=MTA | 2023826-2018-00698 | 2017-11-07 | Injury | STAAR SURGICAL COMPANY | 2018-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICM5_12.6, -10.50/+1.50/88 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). REPORT SOURCE: USER FACILITY IS INCORRECT, SHOULD BE DISTRIBUTOR. (B)(4). Manufacturer Narrative: CONCLUSION CODE: PREVIOUS CODE IS NOT APPLICABLE, THIS LENS MODEL IS NOT CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM.(B)(4). Manufacturer Narrative: (B)(4). CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). | -10.50/+1.50/88 | OD | Physician | 04/19/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7081599&pc=MTA | 2023826-2017-01847 | 2017-11-08 | Injury | STAAR SURGICAL COMPANY | 2017-12-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, -7.0/+1.0/X172 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE SURGEON REPORTS HIGH VAULT, SIGNIFICANT SHALLOW ANGLES AND ACD, AND IRIS ATROPHY AT THE IMPLANT INCISION SITE. SURGEON ALSO STATES PATIENT IS "PRONE" TO IRIS PROLAPSE. LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL SECONDARY SURGICAL INTERVENTION NOTED BY SURGEON: ENLARGEMENT OF/ADDITION OF NEW PI USING YAG LASER. SURGEON HAS NO PLAN TO EXCHANGE THE LENS, WILL MONITOR CORNEA; STATES "REMAINING ACD ICL-CORNEA: 1MM." (B)(4). CONCLUSION: PREVIOUSLY STATED ACD OF 2.80MM; CORRECTED VALUE IS 2.78MM. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT ((B)(6)), ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.80MM)] (B)(4). | -7.0/+1.0/172 | OS | Physician | 11/10/2017 | 2017 | 51
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7122696&pc=MTA | 2023826-2017-01897 | 2017-11-08 | Injury | STAAR SURGICAL COMPANY | 2017-12-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMAGE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+2.0/085 | OS | Physician | 11/27/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7213977&pc=MTA | 2023826-2018-00134 | 2017-11-08 | Injury | STAAR SURGICAL COMPANY | 2018-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))] (B)(4). | -8.5/+1.5/089 | OS | Physician | 12/27/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7245210&pc=MTA | 2023826-2018-00201 | 2017-11-08 | Injury | STAAR SURGICAL COMPANY | 2018-02-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -15.5/+3.0/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED IN A CONTACT CASE WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMAGE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.5/+3.0/095 | OD | Physician | 01/11/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7114683&pc=MTA | 2023826-2017-01862 | 2017-11-09 | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -9.0/+1.5/84 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISUAL DAMAGE TO LENS. DATE OF THIS REPORT: IT IS CORRECTED FROM 29-APR-1987 TO 14-NOV-2017 IN THE INITIAL MDR. DEVICE MANUFACTURE DATE: IT IS CORRECTED FROM 30-JAN-2020 TO 27-JULY-2017 IN THE INITIAL MDR. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.0/+1.5/84 | OS | Physician | 11/14/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7499545&pc=MTA | 2023826-2018-00703 | 2017-11-09 | Injury | STAAR SURGICAL COMPANY | 2018-05-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -13.5/3.5/106 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS REMOVED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | -13.5/3.5/106 | OD | Physician | 04/11/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7113123&pc=MTA | 2023826-2017-01887 | 2017-11-10 | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+1.5/006 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. THERE WAS CLEAR SURGICAL RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.5/+1.5/006 | OD | Physician | 11/19/2017 | 2017 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7399224&pc=MTA | 2023826-2018-00505 | 2017-11-10 | Injury | STAAR SURGICAL COMPANY | 2018-04-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICM5_12.6,-10.5/+4.5/97 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMAGE AS WELL AS RESIDUE AND DEBRIS ON LENS SURFACE. (B)(4) Manufacturer Narrative: CONCLUSION CODE: PREVIOUS CODE IS NOT APPLICABLE. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/+4.5/97 | OD | Physician | 03/07/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7114685&pc=MTA | 2023826-2017-01869 | 2017-11-13 | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -10.5/+5.0/91 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL OBSERVATION FOUND THE LENS WITH AN OPTIC TEAR AND CLEAR RESIDUE ON LENS. IN ADDITION, THE LENS WAS CAUGHT IN THE LENS VIAL, AND THE HAPTIC WAS SMASHED. CONCLUSION CODE: CORRECTED TO, "PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATION OF A LENS IN A PERSON LESS THAN TWENTY ONE YEARS OF AGE." (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/+5.0/91 | OS | Physician | 11/15/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7300929&pc=MTA | 2023826-2018-00291 | 2017-11-15 | Injury | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS), THE LENS HAS A CURVED SHAPE, AND THE LENS MISSING A PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.5/087 | OD | Physician | 02/07/2018 | 2018 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7321824&pc=MTA | 2023826-2018-00354 | 2017-11-15 | Injury | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/92 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017 THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT, FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS), LENS HAS A CURVED SHAPE, AND THE LENS MISSING A PIECE OF HAPTIC. CONCLUSION: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS OVER THE AGE OF 45 YEARS AND FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.5/92 | OS | Physician | 02/07/2018 | 2018 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7149217&pc=MTA | 2023826-2017-01954 | 2017-11-17 | Injury | STAAR SURGICAL COMPANY | 2017-12-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.50/+2.0/105 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMAGE AND CLEAR RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.50/+2.0/105 | OD | Physician | 12/08/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7487541&pc=MTA | 2023826-2018-00670 | 2017-11-20 | Injury | STAAR SURGICAL COMPANY | 2018-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+5.0/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2018, THE VAULT WAS 1.2CT. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND THE TORN PIECE MISSING. THERE WAS PRESENCE OF RESIDUE ON THE PRODUCT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LENS. CLAIM # (B)(4). | -14.0/+5.0/096 | OD | Physician | 04/14/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8164823&pc=MTA | 2023826-2018-01732 | 2017-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -13.00/1.5/092 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2017 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATED BY MFR -DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -13.00/1.5/092 | OD |  | 11/20/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7268477&pc=MTA | 2023826-2018-00193 | 2017-11-22 | Injury | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -11.00/+2.5/+80 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS REMOVED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMAGE AND FOREIGN MATERIAL ON THE LENS SURFACE (DEBRIS ON LENS). (B)(4) Manufacturer Narrative: PATIENT CODE: (B)(4) SECONDARY SURGERY, LENS EXCHANGE. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT OVER THE AGE OF 45 YEARS OLD. (B)(4). | -11.00/+2.5/+80 | OS | Physician | 01/16/2018 | 2018 | 47
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7268420&pc=MTA | 2023826-2018-00221 | 2017-11-23 | Injury | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13/2.5/89 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017 . THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION:LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. DEBRIS AND RESIDUE ON LENS. NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR A LONGER LENS. (B)(4). | -13/2.5/89 | OS | Physician | 01/17/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7232877&pc=MTA | 2023826-2018-00157 | 2017-11-25 | Injury | STAAR SURGICAL COMPANY | 2018-01-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/174 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+1.0/174 | OD | Physician | 01/10/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7188217&pc=MTA | 2023826-2018-00063 | 2017-11-28 | Injury | STAAR SURGICAL COMPANY | 2018-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0/+0.5/096 (SPHERE/CYLINDER/AXIS) , IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS LATER EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED , A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -6.0/+0.5/096 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS LATER EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017 DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. EXPLANTED DATA: IT IS CORRECTED FROM UNK TO (B)(6) 2017. CORRECTED DATA: AGE AT THE TIME OF IMPLANTATION IS (B)(6) YEARS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.0/+0.5/096 | OS | Physician | 12/13/2017 | 2017 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7201650&pc=MTA | 2023826-2018-00089 | 2017-11-29 | Injury | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+1.5/096 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.5/+1.5/096 | OS | Physician | 12/21/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7204549&pc=MTA | 2023826-2018-00132 | 2017-11-29 | Injury | STAAR SURGICAL COMPANY | 2018-01-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+1.5/096 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMAGE, AS WELL AS RESIDUE AND DEBRIS ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.5/+1.5/096 | OS | Physician | 12/21/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7242991&pc=MTA | 2023826-2018-00153 | 2017-11-30 | Injury | STAAR SURGICAL COMPANY | 2018-02-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -11.5/+1.5/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+1.5/078 | OS | Physician | 01/09/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7378417&pc=MTA | 2023826-2018-00472 | 2017-11-30 | Injury | STAAR SURGICAL COMPANY | 2018-03-28 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5/+5.0/071 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF ITS HAPTIC AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, REDUCTION OF ICA) (B)(4). | -9.5/+5.0/071 | OD | Physician | 03/07/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7265290&pc=MTA | 2023826-2018-00185 | 2017-12-04 | Injury | STAAR SURGICAL COMPANY | 2018-02-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.0/+4.5/103 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE ELNS AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR A LONGER LENS ; THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). | -13.0/+4.5/103 | OD | Physician | 01/16/2018 | 2018 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7265440&pc=MTA | 2023826-2018-00220 | 2017-12-04 | Injury | STAAR SURGICAL COMPANY | 2018-02-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6/1/91 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION:LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. LEN HAS A CURVED SHAPE. RESIDUE AND DEBRIS ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR A SHORTER LENS (B)(4). | -6/1/91 | OD | Physician | 01/17/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7213883&pc=MTA | 2023826-2018-00125 | 2017-12-11 | Injury | STAAR SURGICAL COMPANY | 2018-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+1.0/092 | OD | Physician | 01/15/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8028305&pc=MTA | 2023826-2018-01589 | 2017-12-11 | Injury | STAAR SURGICAL COMPANY | 2018-11-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -8.5/+1.0/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: THIS IS A DUPLICATE FILE OF MDR # 2023826-2018-00125. PLEASE REFER TO MDR # 2023826-2018-00125 FOR FURTHER DETAILS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/+1.0/92 | OD | Physician | 10/03/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7281446&pc=MTA | 2023826-2018-00254 | 2017-12-12 | Injury | STAAR SURGICAL COMPANY | 2018-02-19 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A VTICM5_13.2 IMPLANTABLE COLLAMER LENS DIOPTER -10.5/+2.5/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: RETURNED PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND THE LENS HAD A CURVED SHAPE AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -10.5/+2.5/092 | OS | Physician | 01/25/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7281449&pc=MTA | 2023826-2018-00252 | 2017-12-12 | Injury | STAAR SURGICAL COMPANY | 2018-02-19 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A VTICM5_13.2 DIOPTER -9.0/+2.0/093 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE A CURVED SHAPE AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -9.0/+2.0/093 | OD | Physician | 01/25/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721258&pc=MTA | 2023826-2018-01086 | 2017-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/174 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA/ DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. THIS REPORT IS BEING SUBMITTED PER FDA REQUEST. THIS REPORT WAS PREVIOUSLY FILED WITH INCORRECT MFR NUMBER. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: THIS REPORT IS BEING SUBMITTED PER FDA REQUEST. THIS REPORT WAS PREVIOUSLY FILED WITH INCORRECT MFR NUMBER. CLAIM# (B)(4). Manufacturer Narrative: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.0/174 | OS | Physician | 07/02/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7268520&pc=MTA | 2023826-2018-00222 | 2017-12-15 | Injury | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -7.5/1.5/102 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017 . THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A CENTRIFUGE VIAL. THERE WAS NO VISIBLE DAMAGE TO THE LENS. THERE WAS DEBRIS ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR A SHORTER LENS. (B)(4). | -7.5/1.5/102 | OD | Physician | 01/17/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7260409&pc=MTA | 2023826-2018-00172 | 2017-12-19 | Injury | STAAR SURGICAL COMPANY | 2018-02-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHNAGE FOR SHORTER LENS. PATIENT CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4). | -10.0/+1.0/092 | OD | Physician | 01/10/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320762&pc=MTA | 2023826-2018-00332 | 2017-12-19 | Injury | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.5/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND NARROW ANTERIOR CHAMBER. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+2.5/089 | OS | Physician | 02/08/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7213978&pc=MTA | 2023826-2018-00136 | 2017-12-20 | Injury | STAAR SURGICAL COMPANY | 2018-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+2.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND PIECES OF THE HAPTIC MISSING AND RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+2.5/093 | OD | Physician | 12/29/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7300921&pc=MTA | 2023826-2018-00280 | 2017-12-20 | Injury | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMAGE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+2.5/095 | OD | Physician | 01/31/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7300923&pc=MTA | 2023826-2018-00279 | 2017-12-20 | Injury | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+3.0/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: RETURNED PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT, A PIECE OF THE HAPTIC MISSING, AND DEBRIS ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.0/+3.0/087 | OS | Physician | 01/31/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7516777&pc=MTA | 2023826-2018-00748 | 2017-12-20 | Injury | STAAR SURGICAL COMPANY | 2018-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+1.0/057 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+1.0/057 | OD | Physician | 04/27/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7211139&pc=MTA | 2023826-2018-00117 | 2017-12-21 | Injury | STAAR SURGICAL COMPANY | 2018-01-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -16.0/3.0/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, SHALLOW AC, AND THE IRIS BEING PUSHED FORWARD. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS VIAL. VISUAL INSPECTION FOUND NO TEARS; HOWEVER THE LENS APPEARED STRETCHED. CORRECTED INFORMATION: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, SHALLOW ACD, AND THE IRIS BEING PUSHED FORWARD. DATE RECEIVED BY MANUFACTURER CHANGED FROM 27-DEC-2071 TO 27-DEC-2017. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT UNDER THE AGE OF 21 YEARS OLD. (B)(4). | -16.0/3.0/94 | OD | Physician | 12/27/2017 | 2017 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7335040&pc=MTA | 2023826-2018-00337 | 2017-12-27 | Injury | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICM5_12.1, -7.0/0.5/113, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHTT EYE (OD) ON (B)(6) 2017. THE LENS WAS REMOVED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID INSIDE A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.0/0.5/113 | OD | Physician | 02/13/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7378431&pc=MTA | 2023826-2018-00473 | 2017-12-28 | Injury | STAAR SURGICAL COMPANY | 2018-03-28 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL OBSERVATION FOUND THE LENS WITH AN OPTIC TEAR, HAPTIC TEAR, AND RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.5/090 | OD | Physician | 03/07/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7510779&pc=MTA | 2023826-2018-00724 | 2017-12-28 | Injury | STAAR SURGICAL COMPANY | 2018-05-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -6.5/+3.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THE PRESENCE OF FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.5/+3.0/094 | OD | Physician | 04/16/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7510794&pc=MTA | 2023826-2018-00725 | 2017-12-28 | Injury | STAAR SURGICAL COMPANY | 2018-05-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.0/+4.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+4.0/094 | OS | Physician | 04/16/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7213986&pc=MTA | 2023826-2018-00126 | 2017-12-29 | Injury | STAAR SURGICAL COMPANY | 2018-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+2.5/015 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.0/+2.5/015 | OD | Physician | 01/14/2018 | 2018 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7219298&pc=MTA | 2023826-2018-00127 | 2017-12-29 | Injury | STAAR SURGICAL COMPANY | 2018-01-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+2.5/007 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.0/+2.5/007 | OS | Physician | 01/14/2018 | 2018 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7456984&pc=MTA | 2023826-2018-00592 | 2017-12-30 | Injury | STAAR SURGICAL COMPANY | 2018-04-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -11.5/+3.0/83 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS REMOVED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/+3.0/83 | OS | Physician | 03/27/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7281192&pc=MTA | 2023826-2018-00240 | 2018-01-02 | Injury | STAAR SURGICAL COMPANY | 2018-02-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: VTHE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.5/095 | OD | Physician | 01/21/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7267546&pc=MTA | 2023826-2018-00215 | 2018-01-04 | Injury | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1, -11.5/+1.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS, THE SURGEON NOTING LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMAGE AND FOREIGN MATERIALS ON LENS SURFACE (RESIDUE ON LENS). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.5/090 | OS | Physician | 01/19/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7405067&pc=MTA | 2023826-2018-00519 | 2018-01-05 | Injury | STAAR SURGICAL COMPANY | 2018-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.0/+1.0/094 | OD | Physician | 03/10/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438481&pc=MTA | 2023826-2018-00540 | 2018-01-05 | Injury | STAAR SURGICAL COMPANY | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -18/2/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS REMOVED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND THE TORN PIECE MISSING. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4) -SECONDARY SURGERY, LENS EXCHANGE. (B)(4). | -18/2/90 | OS | Physician | 03/20/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438810&pc=MTA | 2023826-2018-00561 | 2018-01-05 | Injury | STAAR SURGICAL COMPANY | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING AND RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) CONCLUSION CODE: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))] (B)(4). | -11.0/+2.5/090 | OS | Physician | 03/21/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7740005&pc=MTA | 2023826-2018-01146 | 2018-01-06 | Injury | STAAR SURGICAL COMPANY | 2018-08-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5/+0.5/125 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+0.5/125 | OD | Physician | 07/06/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7261778&pc=MTA | 2023826-2018-00217 | 2018-01-08 | Injury | STAAR SURGICAL COMPANY | 2018-02-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+2.5/064 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH LIQUID ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LENS. (B)(4). | -13.5/+2.5/064 | OS | Physician | 01/21/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7261786&pc=MTA | 2023826-2018-00218 | 2018-01-08 | Injury | STAAR SURGICAL COMPANY | 2018-02-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/108 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH LIQUID ON LENS. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LENS (B)(4). | -11.0/+2.0/108 | OS | Physician | 01/21/2018 | 2018 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7261813&pc=MTA | 2023826-2018-00219 | 2018-01-08 | Injury | STAAR SURGICAL COMPANY | 2018-02-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH LIQUID ON LENS. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+1.5/088 | OD | Physician | 01/21/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7275448&pc=MTA | 2023826-2018-00238 | 2018-01-08 | Injury | STAAR SURGICAL COMPANY | 2018-02-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/081 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THE DOCTOR REPLIED THAT THE PATIENT DID NOT WANT THE OPERATION FOR THE TIME BEING, AND THE PATIENT WAS IN GOOD CONDITION. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. | -11.5/+1.5/081 | OD | Physician | 01/21/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7515357&pc=MTA | 2023826-2018-00699 | 2018-01-08 | Injury | STAAR SURGICAL COMPANY | 2018-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICM5_12.6, -14.0/+2.5/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2017. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TORN HAPTIC (TORN PIECE IS MISSING) AND RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: CORRECTED DATA: DATE RETURNED TO MANUFACTURER: 16-MAY-2018. DEVICE EVALUATION: THE LENS WAS RETURNED DRY. NOT INCLUDED IN THE SUPPLEMENTAL#1 MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -14.0/+2.5/91 | OD | Physician | 04/19/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7268469&pc=MTA | 2023826-2018-00192 | 2018-01-10 | Injury | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -11.00/+2.0/+106 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE MANUFACTURE DATE: IT IS CORRECTED FROM "18-SEP-2018" TO "18-SEP-2017". (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4) SECONDARY SURGERY, LENS EXCHANGE. (B)(4). | -11.00/+2.0/+106 | OD | Physician | 01/16/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7378435&pc=MTA | 2023826-2018-00482 | 2018-01-10 | Injury | STAAR SURGICAL COMPANY | 2018-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.5/+2.0/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.5/+2.0/087 | OS | Physician | 03/16/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7516774&pc=MTA | 2023826-2018-00750 | 2018-01-10 | Injury | STAAR SURGICAL COMPANY | 2018-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/070 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED, FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS), AND THE HAPTIC CAUGHT WITH VIAL CAP. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.0/070 | OD | Physician | 04/27/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721242&pc=MTA | 2023826-2018-01079 | 2018-01-10 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.5/+2.0/094 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4) Manufacturer Narrative: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.5/+2.0/094 | OD | Physician | 06/29/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721244&pc=MTA | 2023826-2018-01081 | 2018-01-10 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+2.0/091 | OS | Physician | 06/29/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7516763&pc=MTA | 2023826-2018-00740 | 2018-01-11 | Injury | STAAR SURGICAL COMPANY | 2018-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND THE LENS MISSING THE TORN PIECE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE); CONCLUSION CODE: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))]. (B)(4). | -13.5/+1.5/089 | OS | Physician | 04/20/2018 | 2018 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7516764&pc=MTA | 2023826-2018-00738 | 2018-01-11 | Injury | STAAR SURGICAL COMPANY | 2018-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.0/+3.0/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -15.0/+3.0/097 | OD | Physician | 04/20/2018 | 2018 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7281447&pc=MTA | 2023826-2018-00253 | 2018-01-16 | Injury | STAAR SURGICAL COMPANY | 2018-02-19 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A VTICM5_13.2 DIOPTER -9.5/+2.5/088 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND THE LENS MISSING A PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -9.5/+2.5/088 | OD | Physician | 01/25/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7465317&pc=MTA | 2023826-2018-00618 | 2018-01-16 | Injury | STAAR SURGICAL COMPANY | 2018-04-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/112 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND THE TORN PIECE MISSING. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.5/112 | OD | Physician | 04/06/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7557823&pc=MTA | 2023826-2018-00822 | 2018-01-16 | Injury | STAAR SURGICAL COMPANY | 2018-05-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CONCLUSION CODE: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))]. (B)(4). | -12.0/+1.0/095 | OS | Physician | 05/08/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7499541&pc=MTA | 2023826-2018-00704 | 2018-01-17 | Injury | STAAR SURGICAL COMPANY | 2018-05-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -12/3/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS REMOVED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS) AND LENS HAS A CURVED SHAPE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12/3/92 | OS | Physician | 04/11/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7499552&pc=MTA | 2023826-2018-00702 | 2018-01-17 | Injury | STAAR SURGICAL COMPANY | 2018-05-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -11/1.5/80 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS REMOVED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMGE TO THE LENS. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO "EXCESSIVE VAULTING". A NEW LENS OF LONGER LENGTH WAS IMPLANTED INSTEAD AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAVE BEEN FORTHCOMING. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. PATIENT CODE: (B)(4)-SECONDARY SURGERY, LENS EXCHANGE. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT UNDER THE AGE OF 21 YEARS OLD. (B)(4). | -11/1.5/80 | OD | Physician | 04/11/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7612277&pc=MTA | 2023826-2018-00935 | 2018-01-19 | Injury | STAAR SURGICAL COMPANY | 2018-06-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -5/5.5/49 (SPHERE,CYLINDER,AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION:LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS, RESIDUE AND DEBRIS ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-ANGLES. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4) | -5/5.5/49 | OD | Physician | 05/23/2018 | 2018 | 67
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7612293&pc=MTA | 2023826-2018-00936 | 2018-01-19 | Injury | STAAR SURGICAL COMPANY | 2018-06-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -3/6/133 (SPHERE,CYLINDER,AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MIRO-CENTRIFUGE VIAL AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. (B)(4). Manufacturer Narrative: (B)(6) NOT INCLUDED IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR A SHORTER LENGTH LENS, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. (B)(4). | -3/6/133 | OS | Physician | 05/23/2018 | 2018 | 67
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7401372&pc=MTA | 2023826-2018-00302 | 2018-01-20 | Injury | STAAR SURGICAL COMPANY | 2018-04-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -17.0/+1.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT, MISSING A PIECE OF HAPTIC AND RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4) SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, SECONDARY SURGERY, LENS EXCHANGE. THIS REPORT WAS INITIALLY SUBMITTED ON 06-MAR-2018 WITH THE INCORRECT MDR NUMBER 2023382-2018-00302. THE REPORT IS BEING RESUBMITTED PER FDA REQUEST UNDER CORRECT MDR NUMBER 2023826-2018-00302. (B)(4). | -17.0/+1.0/109 | OD | Physician | 02/06/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7720731&pc=MTA | 2023826-2018-01117 | 2018-01-22 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+2.5/094 | OS | Physician | 06/25/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7301464&pc=MTA | 2023826-2018-00275 | 2018-01-25 | Injury | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/+2.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. AT ONE DAY POST-OP THE INTRAOCULAR PRESSURE WAS 20MMHG. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.0/+2.0/094 | OD | Physician | 01/29/2018 | 2018 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7378392&pc=MTA | 2023826-2018-00478 | 2018-01-25 | Injury | STAAR SURGICAL COMPANY | 2018-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.0/093 | OD | Physician | 03/09/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7378398&pc=MTA | 2023826-2018-00477 | 2018-01-26 | Injury | STAAR SURGICAL COMPANY | 2018-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.0/+1.5/092 | OD | Physician | 03/09/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7378404&pc=MTA | 2023826-2018-00475 | 2018-01-26 | Injury | STAAR SURGICAL COMPANY | 2018-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+1.0/093 | OS | Physician | 03/09/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7335038&pc=MTA | 2023826-2018-00336 | 2018-01-27 | Injury | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -7.5/+0.5/003 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS REMOVED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL OBSERVATION FOUND THE LENS MISSING A PIECE OF HAPTIC AND RESIDUE ON LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4) -SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, SECONDARY SURGERY, LENS EXCHANGE. (B)(4). | -7.5/+0.5/003 | OD | Physician | 02/13/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7346844&pc=MTA | 2023826-2018-00424 | 2018-01-29 | Injury | STAAR SURGICAL COMPANY | 2018-03-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.50/+2.0/090 | OD | Physician | 02/18/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7347092&pc=MTA | 2023826-2018-00425 | 2018-01-30 | Injury | STAAR SURGICAL COMPANY | 2018-03-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+1.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4). | -13.00/+1.5/091 | OS | Physician | 02/18/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7378436&pc=MTA | 2023826-2018-00480 | 2018-01-30 | Injury | STAAR SURGICAL COMPANY | 2018-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WASEXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN, FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS) AND THE HAPTIC TORN PIECE IS MISSING. (B)(4). Manufacturer Narrative: PATIENT ID: (B)(6). THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE : NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) (B)(4). | -10.5/+1.5/088 | OD | Physician | 03/09/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7300912&pc=MTA | 2023826-2018-00283 | 2018-01-31 | Injury | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.00/+1.5/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. HOWEVER, THE PROBLEM WAS NOT RESOLVED AS THE SURGEON MISTAKENLY IMPLANTED AN INCORRECT LENS (IMPLANTED THE LENS FOR THE OS EYE). SEE CLAIM# (B)(4) FOR CASE RESOLUTION (NO MDR# ISSUED AS OF YET). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON LENS. (B)(4) Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS FOUND. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -8.00/+1.5/080 | OD | Physician | 02/01/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7300938&pc=MTA | 2023826-2018-00285 | 2018-01-31 | Injury | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.00/+1.50/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE PRODUCT WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL OBSERVATION FOUND THE LENS WITH AN OPTIC TEAR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -6.00/+1.50/095 | OS | Physician | 02/01/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7378434&pc=MTA | 2023826-2018-00481 | 2018-01-31 | Injury | STAAR SURGICAL COMPANY | 2018-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5/+3.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.5/+3.0/092 | OS | Physician | 03/13/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7378440&pc=MTA | 2023826-2018-00483 | 2018-01-31 | Injury | STAAR SURGICAL COMPANY | 2018-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/096 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.0/096 | OD | Physician | 03/13/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7403690&pc=MTA | 2023826-2018-00511 | 2018-02-01 | Injury | STAAR SURGICAL COMPANY | 2018-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.5/+1.5/102 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON THE SAME DATE THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. IT IS UNKNOWN IF THIS WAS PERFORMED INTRAOPERATIVELY OR NOT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.86MM)]. (B)(4) | -7.5/+1.5/102 | OD | Physician | 03/09/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7456991&pc=MTA | 2023826-2018-00590 | 2018-02-01 | Injury | STAAR SURGICAL COMPANY | 2018-04-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -9.0/+3.0/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND A TEAR IN THE LENS HAPTIC AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: DATE OF THIS REPORT: IT IS CORRECTED FROM "02-APR-2018" TO "30-MAR-2018" IN THE INITIAL MDR. DATE RECEIVED BY MANUFACTURER: IT IS CORRECTED FROM "02-APR-2018" TO "30-MAR-2018" IN THE INITIAL MDR. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT UNDER THE AGE OF 21 YEARS OLD. (B)(4). | -9.0/+3.0/91 | OS | Physician | 03/30/2018 | 2018 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7364286&pc=MTA | 2023826-2018-00452 | 2018-02-03 | Injury | STAAR SURGICAL COMPANY | 2018-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+3.5/078 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH PRESENCE OF RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+3.5/078 | OS | Physician | 02/25/2018 | 2018 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7660078&pc=MTA | 2023826-2018-00992 | 2018-02-04 | Injury | STAAR SURGICAL COMPANY | 2018-07-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+2.0/011 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). CORRECTED DATA: DATE RECEIVED BY MANUFACTURER: 05 JUN 2018 IS THE CORRECT DATE IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S. (B)(4). | -15.5/+2.0/011 | OD | Physician | 06/05/2018 | 2018 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7360095&pc=MTA | 2023826-2018-00446 | 2018-02-05 | Injury | STAAR SURGICAL COMPANY | 2018-03-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR LONGER LENS. (B)(4). | -10.5/+1.5/088 | OD | Physician | 02/25/2018 | 2018 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8164876&pc=MTA | 2023826-2018-01742 | 2018-02-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/1.0/087 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -10.50/1.0/087 | OD |  | 11/20/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7359032&pc=MTA | 2023826-2018-00358 | 2018-02-07 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICM5_13.7, -11.50/+1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS NOTICED. LENS REMAINS IMPLANTED. A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: CORRECTED DATA: IT IS CORRECTED FROM "ADVERSE EVENT" TO "PRODUCT PROBLEM". TYPE OF REPORTABLE EVENT: IT IS CORRECTED FROM "SERIOUS INJURY" TO "MALFUNCTION". CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER STATED THAT THE LENS WILL REMAIN IN THE EYE AND WILL NOT BE EXCHANGED. THE POST-OP UCVA WAS 20/15 WITH 900 UM VAULT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/+1.5/094 | OS | Physician | 02/20/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7395778&pc=MTA | 2023826-2018-00500 | 2018-02-07 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 VTICM5_13.2,-11.0/+2.5/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. AN EXCESSIVE VAULT WAS NOTICED WITHIN THE SAME SURGERY (INTRAOPERATIVELY) AND THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENS.THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN, RESIDUE ON LENS AND MISSING PIECE OF HAPTIC. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -11.0/+2.5/92 | OD | Physician | 03/07/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8164877&pc=MTA | 2023826-2018-01743 | 2018-02-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.0/2.5/082 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED WITH MOISTURE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND; NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/2.5/082 | OS |  | 11/20/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438811&pc=MTA | 2023826-2018-00567 | 2018-02-08 | Injury | STAAR SURGICAL COMPANY | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -8.5/+0.5/085 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. Manufacturer Narrative: PREVIOUSLY STATED THE VALUES OF THE SPHERE/CYLINDER/AXIS AS -8.5/+0.5/085. THE ACTUAL VALUES ARE -8.5/+0.5/086. (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.5/+0.5/085 | OS | Physician | 03/23/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7710886&pc=MTA | 2023826-2018-01061 | 2018-02-09 | Injury | STAAR SURGICAL COMPANY | 2018-07-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5/+3.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED AND PRESENCE OF RESIDUE ON LENS SURFACE. LENS HAPTIC CAUGHT WITH VIAL CAP. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.5/+3.0/091 | OD | Physician | 06/25/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7710924&pc=MTA | 2023826-2018-01062 | 2018-02-09 | Injury | STAAR SURGICAL COMPANY | 2018-07-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5/+2.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND PRESENCE OF RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.5/+2.5/093 | OS | Physician | 06/25/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7465298&pc=MTA | 2023826-2018-00609 | 2018-02-10 | Injury | STAAR SURGICAL COMPANY | 2018-04-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+3.5/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING AND RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.5/+3.5/095 | OD | Physician | 04/06/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7465300&pc=MTA | 2023826-2018-00611 | 2018-02-10 | Injury | STAAR SURGICAL COMPANY | 2018-04-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+3.5/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS) AND THE LENS MISSING A PIECE OF HAPTIC. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+3.5/080 | OS | Physician | 04/06/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7557826&pc=MTA | 2023826-2018-00821 | 2018-02-11 | Injury | STAAR SURGICAL COMPANY | 2018-05-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/081 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE AND DEBRIS ON LENS). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). (B)(4). | -9.5/+1.5/081 | OS | Physician | 05/08/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7557905&pc=MTA | 2023826-2018-00819 | 2018-02-11 | Injury | STAAR SURGICAL COMPANY | 2018-05-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/068 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL AND THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/ BENT, AND PRESENCE OF RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+1.5/068 | OD | Physician | 05/08/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7347102&pc=MTA | 2023826-2018-00426 | 2018-02-12 | Injury | STAAR SURGICAL COMPANY | 2018-03-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(4) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) | -12.00/+2.0/082 | OS | Physician | 02/18/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7499549&pc=MTA | 2023826-2018-00706 | 2018-02-13 | Injury | STAAR SURGICAL COMPANY | 2018-05-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -10.5/2/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: INITIAL MDR DATE RECEIVED BY MANUFACTURER SHOULD BE 10 APR 2018. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/2/91 | OD | Physician | 04/10/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7716161&pc=MTA | 2023826-2018-01105 | 2018-02-16 | Injury | STAAR SURGICAL COMPANY | 2018-07-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.0/084 | OD | Physician | 06/25/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7571621&pc=MTA | 2023826-2018-00852 | 2018-02-20 | Injury | STAAR SURGICAL COMPANY | 2018-06-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: REPORT SOURCE - USER FACILITY IS INCORRECT, REPORT SOURCE SHOULD BE DISTRIBUTOR. ADDITIONAL INFORMATION: DEVICE EVALUATION - LENS WAS RETURNED DRY IN MICRO-CENTRIFUGE VIAL, WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT AND FOREIGN MATERIAL (RESIDUE) ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) (B)(4). | -12.0/+1.0/087 | OS | Physician | 05/11/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7571624&pc=MTA | 2023826-2018-00851 | 2018-02-20 | Injury | STAAR SURGICAL COMPANY | 2018-06-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.0/+1.0/073 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS; FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) (B)(4). | -6.0/+1.0/073 | OD | Physician | 05/11/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7405079&pc=MTA | 2023826-2018-00520 | 2018-02-22 | Injury | STAAR SURGICAL COMPANY | 2018-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.0/+3.5/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.0/+3.5/080 | OS | Physician | 03/10/2018 | 2018 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7516769&pc=MTA | 2023826-2018-00752 | 2018-02-24 | Injury | STAAR SURGICAL COMPANY | 2018-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+0.5/006 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -9.5/+0.5/006 | OD | Physician | 04/27/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8204879&pc=MTA | 2023826-2018-01937 | 2018-02-24 | Injury | STAAR SURGICAL COMPANY | 2018-12-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.5/+1.5/071 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.5/+1.5/071 | OS | Physician | 12/03/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8204874&pc=MTA | 2023826-2018-01936 | 2018-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0/+2.0/109 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATED BY MFR - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.0/+2.0/109 | OD |  | 12/03/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7660644&pc=MTA | 2023826-2018-01010 | 2018-02-25 | Injury | STAAR SURGICAL COMPANY | 2018-07-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9/0.5/152 SPHERE, CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED, FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS) AND THE LENS HAPTIC CAUGHT IN VIAL CAP. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9/0.5/152 | OD | Physician | 06/05/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7660676&pc=MTA | 2023826-2018-01011 | 2018-02-26 | Injury | STAAR SURGICAL COMPANY | 2018-07-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.5/2/84 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DATE OF EVENT IS CORRECTED TO (B)6) 2018 IN THE INITIAL MDR.(B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN, HAPTIC BENT AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.5/2/84 | OD | Physician | 06/05/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7408944&pc=MTA | 2023826-2018-00525 | 2018-02-27 | Injury | STAAR SURGICAL COMPANY | 2018-04-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5/+2.0/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS PRESENCE OF RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.5/+2.0/096 | OS | Physician | 03/10/2018 | 2018 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7409118&pc=MTA | 2023826-2018-00524 | 2018-02-27 | Injury | STAAR SURGICAL COMPANY | 2018-04-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC DEFORMED AND PRESENCE OF RESIDUE ON THE LENS SURFACE. DEVICE MANUFACTURE DATE: IT IS CORRECTED TO 12-AUG-2017. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) | -9.0/+1.0/090 | OS | Physician | 03/10/2018 | 2018 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7660066&pc=MTA | 2023826-2018-00994 | 2018-02-27 | Injury | STAAR SURGICAL COMPANY | 2018-07-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/074 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: IT IS CORRECTED TO 05-JUN-2018 IN THE INITIAL MDR. IT IS CORRECTED TO 05-JUN-2018 IN THE INITIAL MDR. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+1.0/074 | OS | Physician | 06/05/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7853009&pc=MTA | 2023826-2018-01344 | 2018-02-27 | Injury | STAAR SURGICAL COMPANY | 2018-09-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+1.0/093 | OS | Physician | 08/10/2018 | 2018 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7567638&pc=MTA | 2023826-2018-00795 | 2018-02-28 | Injury | STAAR SURGICAL COMPANY | 2018-06-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND LONGER LENS IMPLANTED. (B)(4). | -10.5/+1.5/091 | OD | Physician | 05/06/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8048741&pc=MTA | 2023826-2018-01573 | 2018-02-28 | Injury | STAAR SURGICAL COMPANY | 2018-11-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.5/096 | OD | Physician | 10/08/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8048764&pc=MTA | 2023826-2018-01574 | 2018-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/057 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATED BY MFR - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC AND OPTIC TORN. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.5/057 | OS |  | 10/08/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7794417&pc=MTA | 2023826-2018-01194 | 2018-03-01 | Injury | STAAR SURGICAL COMPANY | 2018-08-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS. CORRECTED DATA: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) - SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LENS. (B)(4). | -13.0/+2.5/074 | OS | Physician | 07/22/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7465296&pc=MTA | 2023826-2018-00614 | 2018-03-02 | Injury | STAAR SURGICAL COMPANY | 2018-04-26 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.5/+1.0/031 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO LOW VAULT. IT IS NOTED THAT THE PATIENT DOES NOT WANT TO UNDERGO ANOTHER OPERATION. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.0/031 | OS | Physician | 04/06/2018 | 2018 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7405082&pc=MTA | 2023826-2018-00521 | 2018-03-07 | Injury | STAAR SURGICAL COMPANY | 2018-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/077 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. (B)(4) Manufacturer Narrative: CORRECTED INFORMATION: DEVICE EVALUATED BY MFR? - DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.0/+2.0/077 | OS | Physician | 03/10/2018 | 2018 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7507909&pc=MTA | 2023826-2018-00690 | 2018-03-07 | Injury | STAAR SURGICAL COMPANY | 2018-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.0/+2.0/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED ON 04/03/2018 DUE TO OBSERVATION OF EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: -RETURNED TO MANUFACTURER: 27-APRIL-2018. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS ANF RESIDUE ON THE LENS SURFACE. (B)(4). | -12.0/+2.0/95 | OS | Physician | 04/11/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7527900&pc=MTA | 2023826-2018-00756 | 2018-03-07 | Injury | STAAR SURGICAL COMPANY | 2018-05-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/104 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE INCISION WAS ENLARGED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. AT A POST-OP VISIT, THE PATIENTS EYE HAD LOST 1200 CELLS/MM2 OF ECC. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH THICK RESIDUE ON LENS. VISUAL INSPECTION FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LENS. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, ENLARGEMENT OF INCISION. PATIENT CODE: (B)(4) SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. (B)(4). | -13.0/+1.5/104 | OD | Physician | 04/23/2018 | 2018 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7503531&pc=MTA | 2023826-2018-00679 | 2018-03-08 | Injury | STAAR SURGICAL COMPANY | 2018-05-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -13.5/1.5/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4 Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -13.5/1.5/89 | OD | Physician | 04/11/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7631766&pc=MTA | 2023826-2018-00940 | 2018-03-08 | Injury | STAAR SURGICAL COMPANY | 2018-06-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+2.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON PRODUCT). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND SHORTER LENS IMPLANTED. CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). | -15.0/+2.5/093 | OS | Physician | 06/04/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7465303&pc=MTA | 2023826-2018-00608 | 2018-03-09 | Injury | STAAR SURGICAL COMPANY | 2018-04-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+4.0/085 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.5/+4.0/085 | OD | Physician | 04/06/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7514451&pc=MTA | 2023826-2018-00729 | 2018-03-10 | Injury | STAAR SURGICAL COMPANY | 2018-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL.THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR LONGER LENS CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). | -11.5/+1.5/094 | OD | Physician | 04/16/2018 | 2018 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7514479&pc=MTA | 2023826-2018-00730 | 2018-03-10 | Injury | STAAR SURGICAL COMPANY | 2018-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/113 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGE=ICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. (B)(4). Manufacturer Narrative: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. "SURGICAL" WAS SPELLED INCORRECTLY IN PREVIOUS SUBMISSION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.5/113 | OS | Physician | 04/16/2018 | 2018 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7518823&pc=MTA | 2023826-2018-00732 | 2018-03-10 | Injury | STAAR SURGICAL COMPANY | 2018-05-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. AT 3 HOURS POST-OP, THE VAULT WAS EXCESSIVE, SO THE SURGEON EXPLANTED THE LENS AND EXCHANGED FOR A SHORTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. (B)(4). | -13.0/+1.5/090 | OD | Physician | 04/17/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7639505&pc=MTA | 2023826-2018-00980 | 2018-03-10 | Injury | STAAR SURGICAL COMPANY | 2018-06-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -10.5/+2.0/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/+2.0/103 | OD | Physician | 05/29/2018 | 2018 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7597086&pc=MTA | 2023826-2018-00892 | 2018-03-12 | Injury | STAAR SURGICAL COMPANY | 2018-06-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/176 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(4) 2018. ON (B)(4) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, THIS WAS DUE TO USER ERROR. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+2.0/176 | OS | Physician | 05/18/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7516773&pc=MTA | 2023826-2018-00749 | 2018-03-14 | Injury | STAAR SURGICAL COMPANY | 2018-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.0/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC DEFORMED AND PRESENCE OF RESIDUE ON THE LENS SURFACE. THE HAPTIC WAS CAUGHT WITH THE VIAL CAP. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CONCLUSION CODE: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))]. (B)(4). | -18.0/+1.0/088 | OD | Physician | 04/27/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7660088&pc=MTA | 2023826-2018-01001 | 2018-03-14 | Injury | STAAR SURGICAL COMPANY | 2018-07-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4) Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -8.0/+1.5/097 | OS | Physician | 06/12/2018 | 2018 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7885129&pc=MTA | 2023826-2018-01400 | 2018-03-14 | Injury | STAAR SURGICAL COMPANY | 2018-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.0/087 | OS | Physician | 08/23/2018 | 2018 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7378428&pc=MTA | 2023826-2018-00487 | 2018-03-16 | Injury | STAAR SURGICAL COMPANY | 2018-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/097 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) (B)(4). | -13.0/+2.5/097 | OS | Physician | 03/16/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7378429&pc=MTA | 2023826-2018-00485 | 2018-03-16 | Injury | STAAR SURGICAL COMPANY | 2018-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -14.0/+1.5/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VAIL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON THE PRODUCT. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+1.5/087 | OD | Physician | 03/16/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7636117&pc=MTA | 2023826-2018-00969 | 2018-03-16 | Injury | STAAR SURGICAL COMPANY | 2018-06-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -3.5/+1.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE WITH MOISTURE. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -3.5/+1.0/091 | OD | Physician | 06/10/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7647598&pc=MTA | 2023826-2018-00986 | 2018-03-18 | Injury | STAAR SURGICAL COMPANY | 2018-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.5/+3.0/103 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.5/+3.0/103 | OD | Physician | 06/17/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7465314&pc=MTA | 2023826-2018-00615 | 2018-03-20 | Injury | STAAR SURGICAL COMPANY | 2018-04-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+2.0/094 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+2.0/094 | OS | Physician | 04/06/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721245&pc=MTA | 2023826-2018-01083 | 2018-03-20 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -14.5/+2.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2008. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE CAUSE IS LISTED AS A PATIENT-RELATED FACTOR. REPORTEDLY, THE PATIENT'S NATURAL LENS GREW AFTER THE INITIAL LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED AFTER THE LENS EXCHANGE. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4) Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) (B)(4). | -14.5/+2.0/095 | OD | Physician | 07/02/2018 | 2018 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721247&pc=MTA | 2023826-2018-01085 | 2018-03-20 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -14.5/+2.0/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2008. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE CAUSE IS LISTED AS A PATIENT-RELATED FACTOR. REPORTEDLY, THE PATIENT'S NATURAL LENS GREW AFTER THE INITIAL LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED AFTER THE LENS EXCHANGE. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -14.5/+2.0/085 | OS | Physician | 07/02/2018 | 2018 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7943014&pc=MTA | 2023826-2018-01463 | 2018-03-20 | Injury | STAAR SURGICAL COMPANY | 2018-10-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.00/1.5/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND A LONGER LENGTH LENS WAS LATER IMPLANTED DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL AND THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE AND DEBRIS ON LENS. (B)(4). Manufacturer Narrative: PT INFO:- UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.00/1.5/078 | OS | Physician | 09/25/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7557850&pc=MTA | 2023826-2018-00824 | 2018-03-23 | Injury | STAAR SURGICAL COMPANY | 2018-05-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/039 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN MICRO-CENTRIFUGE. VISUAL INSPECTION FOUND HAPTIC TEAR AND LENS CAUGHT WITH VIAL CAP. CLAIM #(B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) (B)(4). | -10.5/+1.0/039 | OD | Physician | 05/08/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7696905&pc=MTA | 2023826-2018-01046 | 2018-03-23 | Injury | STAAR SURGICAL COMPANY | 2018-07-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0/+2.0/081 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE SURGEON STATES THERE IS A DIFFERENCE BETWEEN THE RIGHT AND LEFT EYE VISION, HOWEVER, THE PATIENT'S VISION IS GOOD AND NO FURTHER TREATMENT IS PLANNED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4) Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -8.0/+2.0/081 | OS | Physician | 06/26/2018 | 2018 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7696922&pc=MTA | 2023826-2018-01044 | 2018-03-23 | Injury | STAAR SURGICAL COMPANY | 2018-07-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.5/+2.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE SURGEON STATES THERE IS A DIFFERENCE BETWEEN THE RIGHT AND LEFT EYE VISION, HOWEVER, THE PATIENT'S VISION IS GOOD AND NO FURTHER TREATMENT IS PLANNED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -8.5/+2.5/093 | OD | Physician | 06/26/2018 | 2018 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7675290&pc=MTA | 2023826-2018-01025 | 2018-03-25 | Injury | STAAR SURGICAL COMPANY | 2018-07-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.50/1.5/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS SURFACE.(B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. SECONDARY SURGICAL INTERVENTION, LENS REMOVED AND EXCHANGED FOR A LONGER LENS. CLAIM#: (B)(4). | -7.50/1.5/088 | OS | Physician | 06/12/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7557878&pc=MTA | 2023826-2018-00828 | 2018-03-26 | Injury | STAAR SURGICAL COMPANY | 2018-05-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/112 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CONCLUSION CODE: OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))] (B)(4). | -11.0/+1.0/112 | OD | Physician | 05/08/2018 | 2018 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7557879&pc=MTA | 2023826-2018-00830 | 2018-03-26 | Injury | STAAR SURGICAL COMPANY | 2018-05-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+2.0/084 | OS | Physician | 05/08/2018 | 2018 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7575699&pc=MTA | 2023826-2018-00872 | 2018-03-26 | Injury | STAAR SURGICAL COMPANY | 2018-06-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/011 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TOP THE LENS. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+1.5/011 | OD | Physician | 05/07/2018 | 2018 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7469552&pc=MTA | 2023826-2018-00584 | 2018-03-26 | Injury | STAAR SURGICAL COMPANY AG | 2018-04-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | -13.0/+1.5/091 | OD | Physician | 04/09/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7466260&pc=MTA | 2023826-2018-00624 | 2018-03-27 | Injury | STAAR SURGICAL COMPANY | 2018-04-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 VTICM5_13.7, -10.5/+6.0/028 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. IT WAS REPORTED THAT DURING ICL IMPLANTATION, THE ICL AND IRIS WERE BULKING UP, IRIS WAS TENDING TO RUN OUT IN THE MAIN WOUND, THEREFORE A SUTURE WAS PLACED. ON (B)(6) 2018 THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. THE REPORTED STATED THAT IN HIS OPINION THE CAUSE OF EVENT WAS DUE TO A PATIENT RELATED FACTOR AND INABILITY TO PRECISELY PREDICT THE OPTIMAL ICL DIAMETER. THE PATIENT HAD A PKP (PENETRATING KERATOPLASTY) ABOUT 30 YEARS AGO BECAUSE OF KERATOCONUS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND THE LENS MISSING THE TORN PIECE OF HAPTIC. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE (B)(4): SECONDARY SURGICAL INTERVENTION, SUTURE, EXPLANT AND EXCHANGED. (B)(4). | -10.5/+6.0/028 | OD | Physician | 03/29/2018 | 2018 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7857666&pc=MTA | 2023826-2018-01360 | 2018-04-01 | Injury | STAAR SURGICAL COMPANY | 2018-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -10/+4/136 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4) -SECONDARY SURGICAL INTERVENTION, EXCHANGE (B)(4). | -10/+4/136 | OD | Physician | 08/08/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7857734&pc=MTA | 2023826-2018-01350 | 2018-04-01 | Injury | STAAR SURGICAL COMPANY | 2018-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -10/+3.5/54 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4)-SECONDARY SURGICAL INTERVENTION, EXCHANGE (B)(4). | -10/+3.5/54 | OS | Physician | 08/08/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7484268&pc=MTA | 2023826-2018-00641 | 2018-04-02 | Injury | STAAR SURGICAL COMPANY | 2018-05-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/108 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND ADMINISTRATION OF ACETAZOLAMIDE 250 MG. THE STATUS OF THE EYE ON (B)(6) 2018 IS LISTED AS VERY NARROW ANGLES, THE ICL IS PUSHING FORWARD THE IRIS, NO FLARE OR PIGMENT IN ANTERIOR CHAMBER, THE PATIENT IS ASYMPTOMATIC AND HAPPY WITH HER VISION. IN THE SURGEON'S OPINION, THE CAUSE OF THIS EVENT IS THAT THE ICL IS "BIG FOR THIS PATIENT." REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED (B)(6) 2018. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.0/108 | OS | Physician | 04/04/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7660089&pc=MTA | 2023826-2018-01000 | 2018-04-02 | Injury | STAAR SURGICAL COMPANY | 2018-07-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/066 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED, FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS) AND THE HAPTIC WAS CAUGHT IN VIAL CAP. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+1.5/066 | OS | Physician | 06/12/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8092513&pc=MTA | 2023826-2018-01707 | 2018-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+5.0/X078 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY; IN LENS CASE/VIAL. VISUAL INSPECTION FOUND ONE HAPTIC BROKEN. CLAIM# (B)(4). Manufacturer Narrative: CLAIM #(B)(4). Manufacturer Narrative: PT INFO: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+5.0/078 | OS |  | 10/22/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7629812&pc=MTA | 2023826-2018-00938 | 2018-04-03 | Injury | STAAR SURGICAL COMPANY | 2018-06-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/075 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION: - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4) | -11.0/+1.5/075 | OS | Physician | 06/04/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7612230&pc=MTA | 2023826-2018-00934 | 2018-04-05 | Injury | STAAR SURGICAL COMPANY | 2018-06-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -5.5/2.5/88 (SPHERE,CYLINDER,AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL AND THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-ANGLES. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR A SHORTER LENGTH LENS. (B)(4). | -5.5/2.5/88 | OS | Physician | 05/23/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7533623&pc=MTA | 2023826-2018-00771 | 2018-04-06 | Injury | STAAR SURGICAL COMPANY | 2018-05-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -14.5/+1.5/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -14.5/+1.5/89 | OD | Other | 04/25/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7534316&pc=MTA | 2023826-2018-00772 | 2018-04-06 | Injury | STAAR SURGICAL COMPANY | 2018-05-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -15.5/+1.5/87 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TORN HAPTIC (TORN PIECE IS MISSING) AND RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.5/+1.5/87 | OS | Other | 04/25/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7571308&pc=MTA | 2023826-2018-00838 | 2018-04-06 | Injury | STAAR SURGICAL COMPANY | 2018-06-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.00/+3.5/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.00/+3.5/082 | OD | Physician | 05/07/2018 | 2018 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7833174&pc=MTA | 2023826-2018-01227 | 2018-04-06 | Injury | STAAR SURGICAL COMPANY | 2018-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -09.00/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -09.00/1.5/089 | OS | Physician | 08/01/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7535076&pc=MTA | 2023826-2018-00774 | 2018-04-07 | Injury | STAAR SURGICAL COMPANY | 2018-05-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.0/+1.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CORRECTED DATA: (B)(6) IS THE WRONG DATE OF BIRTH, CORRECT DATE OF BIRTH IS (B)(6). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/+1.0/090 | OS | Physician | 04/27/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7575677&pc=MTA | 2023826-2018-00799 | 2018-04-11 | Injury | STAAR SURGICAL COMPANY | 2018-06-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS AND PRESENCE OF RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND SHORTER LENS IMPLANTED CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). | -10.5/+1.0/074 | OD | Physician | 05/07/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594269&pc=MTA | 2023826-2018-00886 | 2018-04-11 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/028 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. CORRECTED DATA: CORRECT DATE IS (B)(6) 2018, (B)(6) 2018 IS INCORRECT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+1.5/028 | OD | Physician | 05/15/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7896089&pc=MTA | 2023826-2018-01416 | 2018-04-12 | Injury | STAAR SURGICAL COMPANY | 2018-09-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -08.5/+2/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUG VIAL, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A TORN HAPTIC AND RESIDUE ON LENS. LENS HAS A CURVED SHAPE. (B)(4). | -08.5/+2/104 | OD | Physician | 08/22/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7896224&pc=MTA | 2023826-2018-01417 | 2018-04-12 | Injury | STAAR SURGICAL COMPANY | 2018-09-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -09.5/+2/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE HAS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: INITIAL REPORT LISTS LENS DIOPTER AS -09.5/+2//91 (SPHERE/CYLINDER/AXIS); CORRECT DIOPTER IS -9.5/+2.5/91 (SPHERE/CYLINDER/AXIS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. (B)(4). | -09.5/+2/91 | OS | Physician | 08/22/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8133035&pc=MTA | 2023826-2018-01784 | 2018-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -18.0/-2.0/088, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED ON (B)(6) 2018 DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. A REPLACEMENT LENS, SHORTER LENGTH WAS IMPLANTED BUT IT DID NOT RESOLVE THE PROBLEM. "NEEDS EVEN FURTHER LENS SIZE REDUCTION" WAS REPORTED. THE REPORTER STATED THAT THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATED BY MFR - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: "-18.0/-2.0/088" SHOULD BE CORRECTED TO "-18.0/2.0/088" IN THE ORIGINAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.0/2.0/088 | OD |  | 08/11/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7579748&pc=MTA | 2023826-2018-00863 | 2018-04-15 | Injury | STAAR SURGICAL COMPANY | 2018-06-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 VTICM5_13.2, -13.5/+2.0/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.5/+2.0/95 | OD | Physician | 05/09/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7589629&pc=MTA | 2023826-2018-00876 | 2018-04-15 | Injury | STAAR SURGICAL COMPANY | 2018-06-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND WAS EXCHANGED FOR A SHORTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR SHORTER LENS. CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). | -11.0/+1.0/091 | OD | Physician | 05/14/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7589637&pc=MTA | 2023826-2018-00877 | 2018-04-15 | Injury | STAAR SURGICAL COMPANY | 2018-06-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND WAS EXCHANGED FOR A SHORTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4) | -11.5/+1.0/084 | OS | Physician | 05/14/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7542996&pc=MTA | 2023826-2018-00794 | 2018-04-16 | Injury | STAAR SURGICAL COMPANY | 2018-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+1.0/102 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: CORRECT AWARE DATE IS (B)(6) 2018 NOT (B)(6) 2018. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LENS PATIENT CODE: (B)(4)- SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). | -17.0/+1.0/102 | OD | Physician | 05/05/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7554910&pc=MTA | 2023826-2018-00807 | 2018-04-17 | Injury | STAAR SURGICAL COMPANY | 2018-05-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -12.5/+1.0/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TORN HAPTIC (TORN PIECE IS MISSING) AND RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.5/+1.0/90 | OD | Physician | 05/03/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7720835&pc=MTA | 2023826-2018-01122 | 2018-04-18 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/081 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LENS (B)(4). | -9.5/+1.0/081 | OD | Physician | 06/25/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7660601&pc=MTA | 2023826-2018-01009 | 2018-04-19 | Injury | STAAR SURGICAL COMPANY | 2018-07-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.5/1.5/93 (SPHERE,CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.5/1.5/93 | OD | Physician | 06/05/2018 | 2018 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7696915&pc=MTA | 2023826-2018-01038 | 2018-04-20 | Injury | STAAR SURGICAL COMPANY | 2018-07-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TEAR AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.5/093 | OD | Physician | 06/22/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7696920&pc=MTA | 2023826-2018-01040 | 2018-04-20 | Injury | STAAR SURGICAL COMPANY | 2018-07-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA:H3-DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MIRCO-CENTRIFUGE VIAL AND THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE.(B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) (B)(4). | -12.5/+1.5/091 | OS | Physician | 06/22/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7700085&pc=MTA | 2023826-2018-01064 | 2018-04-20 | Injury | STAAR SURGICAL COMPANY | 2018-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -16/1.5/69 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DATE OF EVENT CORRECTED TO (B)(6) 2018. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL AND THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN US. (B)(4). | -16/1.5/69 | OD | Physician | 06/19/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7639431&pc=MTA | 2023826-2018-00979 | 2018-04-23 | Injury | STAAR SURGICAL COMPANY | 2018-06-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -14.0/+2.5/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL AND THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.0/+2.5/93 | OS | Physician | 05/29/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7647519&pc=MTA | 2023826-2018-00971 | 2018-04-23 | Injury | STAAR SURGICAL COMPANY | 2018-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.5/+3.0/103 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE LENS HAPTIC TORN. THERE WAS RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+3.0/103 | OD | Physician | 06/12/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8028723&pc=MTA | 2023826-2018-01542 | 2018-04-24 | Injury | STAAR SURGICAL COMPANY | 2018-11-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.50/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN ALONG WITH RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -15.50/2.0/095 | OS | Physician | 10/16/2018 | 2018 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7710876&pc=MTA | 2023826-2018-01056 | 2018-04-25 | Injury | STAAR SURGICAL COMPANY | 2018-07-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/077 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.5/077 | OS | Physician | 06/25/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7660382&pc=MTA | 2023826-2018-01007 | 2018-04-26 | Injury | STAAR SURGICAL COMPANY | 2018-07-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18/3.5/76 SPHERE, CYLINDER/AXIS, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18/3.5/76 | OS | Physician | 06/05/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7724404&pc=MTA | 2023826-2018-01097 | 2018-04-27 | Injury | STAAR SURGICAL COMPANY | 2018-07-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8/1/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE LENS DAMAGE AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, EXCHANGE (B)(4). | -8/1/89 | OS | Physician | 06/27/2018 | 2018 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7587831&pc=MTA | 2023826-2018-00911 | 2018-04-29 | Injury | STAAR SURGICAL COMPANY | 2018-06-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -4.5/2/88 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND EXCHANGED FOR SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN AND RESIDUE ON LENS SURFACE. CLAIM #:(B)(4). Manufacturer Narrative: CORRECTED DATA: DEVICE EVALUATION: VISUAL INSPECTION CORRECTED TO- "HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE" IN THE SUPPLEMENTAL MDR#1. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -4.5/2/88 | OD | Physician | 05/16/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7720826&pc=MTA | 2023826-2018-01121 | 2018-04-30 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND SGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN AND CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LENS. PATIENT CODE: (B)(4)- SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. (B)(4). | -8.5/+1.0/095 | OD | Physician | 06/25/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7676121&pc=MTA | 2023826-2018-01029 | 2018-05-02 | Injury | STAAR SURGICAL COMPANY | 2018-07-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+1.5/069 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC BENT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) | -6.5/+1.5/069 | OS | Physician | 06/13/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7676126&pc=MTA | 2023826-2018-01027 | 2018-05-02 | Injury | STAAR SURGICAL COMPANY | 2018-07-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.0/+1.5/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, SIGNIFICANT REDUCTION OF ICA) (B)(4). | -5.0/+1.5/089 | OD | Physician | 06/13/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7612191&pc=MTA | 2023826-2018-00933 | 2018-05-03 | Injury | STAAR SURGICAL COMPANY | 2018-06-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -13.5/3.0/75 (SPHERE,CYLINDER,AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS DUE TO OBSERVATION OF EXCESSIVE VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR A SHORTER LENGTH LENS. (B)(4). | -13.5/3.0/75 | OS | Physician | 05/23/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7739732&pc=MTA | 2023826-2018-01132 | 2018-05-03 | Injury | STAAR SURGICAL COMPANY | 2018-08-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO 13.7, -11.5/2/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION- THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/2/102 | OD | Physician | 07/04/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7716190&pc=MTA | 2023826-2018-01100 | 2018-05-06 | Injury | STAAR SURGICAL COMPANY | 2018-07-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+1.0/084 | OD | Physician | 06/25/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7972131&pc=MTA | 2023826-2018-01510 | 2018-05-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-10-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+0.5/102 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND THE LENS WAS TORN IN HALF, UNABLE TO IDENTIFY LENS MODEL WITH THE LENS PIECE THAT WAS RETURNED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+0.5/102 | OD |  | 09/16/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7897311&pc=MTA | 2023826-2018-01418 | 2018-05-08 | Injury | STAAR SURGICAL COMPANY | 2018-09-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2, -10.00/+2.00/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). DEVICE EVALUATION: LENS RETURNED DRY WITH RESIDE IN A MICRO-CENRIFUGE VIAL. VISUAL INSPECTION FOUND A TORN HAPTIC. (B)(4). | -10.00/+2.00/079 | OS | Physician | 08/22/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7629774&pc=MTA | 2023826-2018-00919 | 2018-05-09 | Injury | STAAR SURGICAL COMPANY | 2018-06-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+3.0/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC DEFORMED, LENS CAUGHT WITH VIAL CAP, AND FOUND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. CORRECTED DATA: "SECONDARY SURGICAL INTERVENTION, EXCHANGE" IS CORRECT. THE STATEMENT "LENS EXPLANTED" IS INCORRECT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION: - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4) | -10.5/+3.0/085 | OD | Physician | 05/28/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7629785&pc=MTA | 2023826-2018-00920 | 2018-05-10 | Injury | STAAR SURGICAL COMPANY | 2018-06-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.5/+4.0/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED, FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS), AND LENS WAS CAUGHT IN VIAL CAP. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -09.5/+4.0/082 | OS | Physician | 05/28/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7598934&pc=MTA | 2023826-2018-00881 | 2018-05-11 | Injury | STAAR SURGICAL COMPANY | 2018-06-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/135 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN; FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS); AND THE TORN HAPTIC PIECE IS MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND SHORTER LENS IMPLANTED (B)(4). | -13.0/+1.5/135 | OS | Physician | 05/14/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7598990&pc=MTA | 2023826-2018-00880 | 2018-05-11 | Injury | STAAR SURGICAL COMPANY | 2018-06-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+1.0/070 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. THIS IS A RESUBMISSION TO CORRECT THE MFR#. THE MFR# WAS REPORTED INCORRECTLY (MFR#2023628-2018-00880). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+1.0/070 | OD | Physician | 05/14/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7888974&pc=MTA | 2023826-2018-01313 | 2018-05-11 | Injury | STAAR SURGICAL COMPANY | 2018-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/1.5/057 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND TEARS IN THE HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/1.5/057 | OS | Physician | 08/28/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7720739&pc=MTA | 2023826-2018-01118 | 2018-05-15 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -8.5/+2.0/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON 05/16/2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+2.0/100 | OD | Physician | 06/25/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7720750&pc=MTA | 2023826-2018-01119 | 2018-05-15 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/117 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WAS RETURNED IN A MICRO CENTRIGUE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.5/+1.5/117 | OD | Physician | 06/25/2018 | 2018 | 47
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7720786&pc=MTA | 2023826-2018-01120 | 2018-05-15 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+1.0/007 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE, VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR LONGER LENS. (B)(4). | -14.5/+1.0/007 | OD | Physician | 06/25/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7720855&pc=MTA | 2023826-2018-01124 | 2018-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+0.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+0.5/095 | OD |  | 06/25/2018 | 2018 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594281&pc=MTA | 2023826-2018-00893 | 2018-05-16 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.5/+1.0/137 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE AND THERE WAS CLEAR RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE AND DEBRIS ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.0/137 | OS | Physician | 05/18/2018 | 2018 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7631781&pc=MTA | 2023826-2018-00944 | 2018-05-16 | Injury | STAAR SURGICAL COMPANY | 2018-06-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MIRCO-CENTRIFUGE VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (RESIDE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LENS. PATIENT CODE: (B)(4) SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. (B)(4). | -13.5/+1.0/095 | OS | Physician | 06/04/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7871083&pc=MTA | 2023826-2018-01371 | 2018-05-16 | Injury | STAAR SURGICAL COMPANY | 2018-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/X044 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING, WITH SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2018. THE LENS WAS EXCHANGED FOR SHORTER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. THERE WAS RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S., (B)(4). | -12.5/+2.5/044 | OS | Physician | 08/13/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7943019&pc=MTA | 2023826-2018-01464 | 2018-05-16 | Injury | STAAR SURGICAL COMPANY | 2018-10-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.50/3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED ON (B)(6) 2018 AND A SHORTER LENGTH LENS WAS LATER IMPLANTED DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.50/3.0/093 | OD | Physician | 09/25/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7639926&pc=MTA | 2023826-2018-00957 | 2018-05-17 | Injury | STAAR SURGICAL COMPANY | 2018-06-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CONCLUSION CODE: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))] (B)(4). | -10.0/+2.5/090 | OD | Physician | 06/01/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7647539&pc=MTA | 2023826-2018-00970 | 2018-05-17 | Injury | STAAR SURGICAL COMPANY | 2018-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+3.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS RESIDUE AND DEBRIS ON THE LENS. CORRECTED DATA: CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+3.5/091 | OS | Physician | 06/12/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7704629&pc=MTA | 2023826-2018-01048 | 2018-05-17 | Injury | STAAR SURGICAL COMPANY | 2018-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.0/+1.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND CLEAR SURGICAL RESIDUE ON THE PRODUCT. DIMENSIONAL AND FUNCTIONAL TESTING WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED; PATIENT CODE: (B)(4)- SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.0/+1.5/089 | OS | Physician | 06/19/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7704634&pc=MTA | 2023826-2018-01049 | 2018-05-17 | Injury | STAAR SURGICAL COMPANY | 2018-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.0/+2.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS DUE TO IRIS PROLAPSE DURING EXPLANTATION DUE TO A NOT PERFECT INCISIONAL TUNNEL. THE PATIENT WILL BE RE-EVALUATED ONCE THE EYE IS COMPLETELY CALM. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DATE OF THIS REPORT: -IT IS CORRECTED TO 15-JUN-2018 IN THE INITIAL MDR. DATE RECEIVED BY MFR: IT IS CORRECTED TO 15-JUN-2018 IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND ONE HAPTIC BROKEN AND CLEAR SURGICAL RESIDUE ON THE PRODUCT. DIMENSIONAL AND FUNCTIONAL TESTING WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED; PATIENT CODE: (B)(4)- SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, IRIS PROLAPSE. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.0/+2.5/094 | OD | Physician | 06/15/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8002611&pc=MTA | 2023826-2018-01567 | 2018-05-23 | Injury | STAAR SURGICAL COMPANY | 2018-10-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.5/097 | OD | Physician | 09/24/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8237935&pc=MTA | 2023826-2019-00082 | 2018-05-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.50/+4.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.50/+4.5/088 | OS |  | 12/11/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7853000&pc=MTA | 2023826-2018-01332 | 2018-05-26 | Injury | STAAR SURGICAL COMPANY | 2018-09-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -8.0/+3.0/009 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -8.0/+3.0/009 | OS | Physician | 08/09/2018 | 2018 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7853005&pc=MTA | 2023826-2018-01329 | 2018-05-26 | Injury | STAAR SURGICAL COMPANY | 2018-09-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -8.5/+4.0/002 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4) | -8.5/+4.0/002 | OD | Physician | 08/09/2018 | 2018 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7827644&pc=MTA | 2023826-2018-01230 | 2018-05-28 | Injury | STAAR SURGICAL COMPANY | 2018-08-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -07.25/-0.75/010 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -07.25/-0.75/010 | OD | Physician | 08/01/2018 | 2018 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8066017&pc=MTA | 2023826-2018-01660 | 2018-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/073 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: -UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+1.0/073 | OS |  | 10/18/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8255329&pc=MTA | 2023826-2019-00108 | 2018-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0/+2.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: ADDITIONAL DATA DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+2.5/094 | OD |  | 12/17/2018 | 2018 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7884846&pc=MTA | 2023826-2018-01264 | 2018-05-29 | Injury | STAAR SURGICAL COMPANY | 2018-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -15.00/1/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION DEVICE EVALUATION- THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.00/1/090 | OS | Physician | 08/21/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7884848&pc=MTA | 2023826-2018-01263 | 2018-05-29 | Injury | STAAR SURGICAL COMPANY | 2018-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.00/1/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4) - SECONDARY SURGICAL INTERVENTION, EXCHANGE (B)(4). | -15.00/1/090 | OD | Physician | 08/21/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8107288&pc=MTA | 2023826-2018-01725 | 2018-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATED BY MFR -DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MIRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS, AND PRESENCE OF RESIDUE AND DEBRIS ON LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -8.0/+1.0/095 | OS |  | 11/01/2018 | 2018 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7907435&pc=MTA | 2023826-2018-01425 | 2018-05-31 | Injury | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+2.5/061 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.50/+2.5/061 | OS | Physician | 08/27/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7696907&pc=MTA | 2023826-2018-01039 | 2018-06-05 | Injury | STAAR SURGICAL COMPANY | 2018-07-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.5/090 | OS | Physician | 06/22/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7696910&pc=MTA | 2023826-2018-01037 | 2018-06-05 | Injury | STAAR SURGICAL COMPANY | 2018-07-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.0/+3.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL AND THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+3.0/090 | OD | Physician | 06/22/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721251&pc=MTA | 2023826-2018-01074 | 2018-06-05 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON THE SAME DATE, IN A SEPARATE SURGERY, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.5/090 | OD | Physician | 06/28/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7749376&pc=MTA | 2023826-2018-01156 | 2018-06-05 | Injury | STAAR SURGICAL COMPANY | 2018-08-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICMO12.6, -8.5/1/87 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO LOW VAULT. IT WAS REPORTED THAT THE PATIENT IS IN GOOD CONDITION AFTER THE SURGERY. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THAT THE PATIENT HAS HAD A REPLACEMENT (IOL). PATIENT IS IN GOOD CONDITION. (B)(4) Manufacturer Narrative: "CAPSULAR TEAR" NOT INCLUDED IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, EXPLANT (B)(4). | -8.5/1/87 | OS | Physician | 07/04/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8149314&pc=MTA | 2023826-2018-01809 | 2018-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/107 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13/+2.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+2.5/107 | OS |  | 11/20/2018 | 2018 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8149325&pc=MTA | 2023826-2018-01805 | 2018-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CORRECTION DATA: THE DIOPTER VALUE IS CORRECTED TO -13.5/+1.5/089 FOR THE INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+2.5/091 | OD |  | 11/20/2018 | 2018 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7947566&pc=MTA | 2023826-2018-01488 | 2018-06-12 | Injury | STAAR SURGICAL COMPANY | 2018-10-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -9.00/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: CORRECTED TO: "SERIOUS INJURY" IN THE INITIAL MDR. ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/1.5/095 | OD | Physician | 10/02/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7927071&pc=MTA | 2023826-2018-01472 | 2018-06-13 | Injury | STAAR SURGICAL COMPANY | 2018-10-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.5/+1.5/089 | OD | Physician | 09/07/2018 | 2018 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7838243&pc=MTA | 2023826-2018-01323 | 2018-06-14 | Injury | STAAR SURGICAL COMPANY | 2018-08-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -11/1/72 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. IT WAS REPORTED THAT AN EXCHANGE FOR A SHORTER LENGTH LENS WAS PERFORMED AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. RETURNED PRODUCT INSPECTION FOUND NO VISIBLE DAMAGE TO LENS WITH FOREIGN MATERIAL (RESIDUE) ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11/1/72 | OD | Physician | 08/01/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8164891&pc=MTA | 2023826-2018-01744 | 2018-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -08.50/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. "EXCESIVE" VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS WAS OBSERVED, THE LENS WAS REMOVED AND EXCHANGED WITH A SHORTER LENS ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -08.50/1.5/093 | OD |  | 11/20/2018 | 2018 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8181675&pc=MTA | 2023826-2018-01842 | 2018-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/1.0/020 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VAIL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.50/1.0/020 | OS |  | 11/20/2018 | 2018 | 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7832625&pc=MTA | 2023826-2018-01224 | 2018-06-20 | Injury | STAAR SURGICAL COMPANY | 2018-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -5.5/3/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.5/3/088 | OD | Physician | 08/01/2018 | 2018 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7716175&pc=MTA | 2023826-2018-01098 | 2018-06-21 | Injury | STAAR SURGICAL COMPANY | 2018-07-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS AND THE PROBLEM WAS NOT RESOLVED. THE SURGEON CHOSE THE SAME LENS MODEL BECAUSE SHE IMPLANTED THE SECOND LENS IN OBLIQUE POSITION (DIFFERENT AXIS) AND SHE BELIEVED THAT COULD DECREASE THE VAULT. SEE MFR. REPORT 2023826-2018-01106 FOR THE EXCHANGED LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.0/+2.5/088 | OD | Physician | 06/24/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7830175&pc=MTA | 2023826-2018-01242 | 2018-06-21 | Injury | STAAR SURGICAL COMPANY | 2018-08-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -05.5/2/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -05.5/2/091 | OS | Physician | 08/01/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7837801&pc=MTA | 2023826-2018-01320 | 2018-06-21 | Injury | STAAR SURGICAL COMPANY | 2018-08-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -6/2.5/88 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON 06/21/2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A TORN, BENT HAPTIC AND PRESENCE OF RESIDUE ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4) | -6/2.5/88 | OD | Physician | 08/01/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8002653&pc=MTA | 2023826-2018-01564 | 2018-06-21 | Injury | STAAR SURGICAL COMPANY | 2018-10-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -7.0/+2.0/103 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/+2.0/103 | OD | Physician | 09/24/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8002647&pc=MTA | 2023826-2018-01562 | 2018-06-22 | Injury | STAAR SURGICAL COMPANY | 2018-10-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0/+3.0/099 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR ANOTHER SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+3.0/099 | OD | Physician | 09/24/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8149318&pc=MTA | 2023826-2018-01800 | 2018-06-23 | Injury | STAAR SURGICAL COMPANY | 2018-12-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+4.0/075 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+4.0/075 | OD | Physician | 11/14/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721263&pc=MTA | 2023826-2018-01091 | 2018-06-26 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/+2.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) (B)(4). | -16.0/+2.5/093 | OD | Physician | 07/02/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8149326&pc=MTA | 2023826-2018-01810 | 2018-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/025 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND CLEAR SURGICAL RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.5/025 | OD |  | 11/20/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8149330&pc=MTA | 2023826-2018-01814 | 2018-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.5/+3.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: DATE OF REPORT TO: "19-NOV-2018" IN THE INITIAL MDR, (B)(4). Manufacturer Narrative: CORRECTED DATA: DATE OF REPORT-CORRECTED FROM (B)(6)2018 TO (B)(6)2018 IN THE SUPPLEMENTAL MDR#2. CLAIM #(B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.5/+3.0/093 | OD |  | 11/20/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721270&pc=MTA | 2023826-2018-01092 | 2018-06-27 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+2.0/127 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED DUE TO LOW VAULT. THE REPLACEMENT LENS IS THE SAME SIZE AS THE ORIGINAL, BUT A DIFFERENT AXIS. ACCORDING TO THE SURGEON, ROTATING THE LENS CAN INCREASE THE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) (B)(4). | -8.5/+2.0/127 | OS | Physician | 07/02/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7726685&pc=MTA | 2023826-2018-01128 | 2018-06-27 | Injury | STAAR SURGICAL COMPANY | 2018-07-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A13.2MM VTICM5_13.2, -13.5/3.5/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS). THE LENS WAS LATER EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THAT THE EXCHANGE RESOLVED THE PROBLEM AND THE CAUSE OF EVENT WAS USER ERROR. (B)(4) Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/3.5/89 | OS | Physician | 06/27/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8002381&pc=MTA | 2023826-2018-01561 | 2018-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-10-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+1.0/087 | OS |  | 09/24/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8149313&pc=MTA | 2023826-2018-01813 | 2018-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/145 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATED BY MFR -THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. METHOD CODE ADDED. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.5/145 | OS |  | 11/20/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8149335&pc=MTA | 2023826-2018-01817 | 2018-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATED BY MFR -THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. METHOD CODE ADDED CLAIM #(B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.5/+2.5/091 | OS |  | 11/20/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8272677&pc=MTA | 2023826-2019-00149 | 2018-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+1.0/096 | OD |  | 12/26/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8272679&pc=MTA | 2023826-2019-00151 | 2018-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.5/095 | OS |  | 12/26/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7747625&pc=MTA | 2023826-2018-01163 | 2018-07-02 | Injury | STAAR SURGICAL COMPANY | 2018-08-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.00/1.0/023(SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4) Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.00/1.0/023 | OD | Physician | 07/04/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7832893&pc=MTA | 2023826-2018-01225 | 2018-07-02 | Injury | STAAR SURGICAL COMPANY | 2018-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -13.00/1.5/115 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS SURFACE. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/1.5/115 | OD | Physician | 08/01/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7886303&pc=MTA | 2023826-2018-01409 | 2018-07-02 | Injury | STAAR SURGICAL COMPANY | 2018-09-18 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -14.0/+2.0/105 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+2.0/105 | OD | Physician | 08/24/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824058&pc=MTA | 2023826-2018-01273 | 2018-07-03 | Injury | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+1.5/093 | OS | Physician | 08/02/2018 | 2018 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824064&pc=MTA | 2023826-2018-01271 | 2018-07-03 | Injury | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.5/+1.5/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DIOPTER CORRECTED TO: "-7.50/+1.0/095". (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: (B)(4). THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.5/+1.5/095 | OD | Physician | 08/02/2018 | 2018 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8164861&pc=MTA | 2023826-2018-01739 | 2018-07-04 | Injury | STAAR SURGICAL COMPANY | 2018-12-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.50/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/1.0/088 | OS | Physician | 11/20/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8164862&pc=MTA | 2023826-2018-01738 | 2018-07-04 | Injury | STAAR SURGICAL COMPANY | 2018-12-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.50/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/1.0/092 | OD | Physician | 11/20/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8066018&pc=MTA | 2023826-2018-01659 | 2018-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.5/+1.0/094 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS, THE SURGEON NOTING EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE IS LISTED AS UNKNOWN. ATTEMPTS TO OBTAIN CLARIFICATION REGARDING CLAIM DETAILS HAVE NOT BEEN SUCCESSFUL.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5/+1.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS, THE SURGEON NOTING EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE IS LISTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: THIS CLAIM IS A DUPLICATE OF MFR REPORT #: 2023826-2018-01745. SEE ORIGINAL REPORT FOR CLAIM DETAILS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.5/+1.0/094 | OS |  | 10/18/2018 | 2018 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8164897&pc=MTA | 2023826-2018-01745 | 2018-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -07.50/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS WAS OBSERVED, THE LENS WAS REMOVED AND EXCHANGED WITH A SHORTER LENS ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. CLAIM#: (B)(4). Manufacturer Narrative: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -07.50/1.5/093 | OS |  | 11/20/2018 | 2018 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8002545&pc=MTA | 2023826-2018-01559 | 2018-07-09 | Injury | STAAR SURGICAL COMPANY | 2018-10-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -6.0/+1.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.0/+1.0/090 | OD | Physician | 09/24/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824060&pc=MTA | 2023826-2018-01278 | 2018-07-10 | Injury | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, REDUCTION OF ICA). CLAIM #: (B)(4). | -12.5/+1.0/090 | OD | Physician | 08/02/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8149329&pc=MTA | 2023826-2018-01815 | 2018-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/072 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATED BY MFR -THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. METHOD CODE ADDED. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.0/072 | OS |  | 11/20/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8066008&pc=MTA | 2023826-2018-01661 | 2018-07-11 | Injury | STAAR SURGICAL COMPANY | 2018-11-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.0/+1.0/106 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.0/+1.0/106 | OS | Physician | 10/18/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8066023&pc=MTA | 2023826-2018-01658 | 2018-07-11 | Injury | STAAR SURGICAL COMPANY | 2018-11-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.0/+1.5/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.0/+1.5/086 | OD | Physician | 10/18/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824497&pc=MTA | 2023826-2018-01259 | 2018-07-12 | Injury | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/077 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. THERE WAS RESIDUE ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.5/077 | OD | Physician | 07/30/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824801&pc=MTA | 2023826-2018-01260 | 2018-07-12 | Injury | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -1.50/+6.0/014 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND TRACES OF PINK COLOR ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND LENS WITHIN SPECIFICATIONS. CLAIM #: (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -1.50/+6.0/014 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -1.50/+6.0/014 | OS | Physician | 07/29/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8011514&pc=MTA | 2023826-2018-01538 | 2018-07-12 | Injury | STAAR SURGICAL COMPANY | 2018-10-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+2.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. THERE WAS RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+2.0/094 | OD | Physician | 09/27/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8044007&pc=MTA | 2023826-2018-01598 | 2018-07-12 | Injury | STAAR SURGICAL COMPANY | 2018-11-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -13.00/2.0/112 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND RESIDUE AND DEBRIS ON LENS. (B)(4). Manufacturer Narrative: PT INFO - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/2.0/112 | OS | Physician | 10/12/2018 | 2018 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8255342&pc=MTA | 2023826-2019-00114 | 2018-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+2.0/094 | OD |  | 12/20/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8043745&pc=MTA | 2023826-2018-01597 | 2018-07-13 | Injury | STAAR SURGICAL COMPANY | 2018-11-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, --9.50/1.0/060 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BENT AND RESIDUE ON THE LENS. PREVIOUS CARTRIDGE MODEL NOT APPLICABLE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | --9.50/1.0/060 | OD | Physician | 10/12/2018 | 2018 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8148502&pc=MTA | 2023826-2018-01792 | 2018-07-14 | Injury | STAAR SURGICAL COMPANY | 2018-12-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+3.5/081 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+3.5/081 | OS | Physician | 11/10/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7814794&pc=MTA | 2023826-2018-01220 | 2018-07-15 | Injury | STAAR SURGICAL COMPANY | 2018-08-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -13.5/+3.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4) - SECONDARY SURGICAL INTERVENTION, EXCHANGE (B)(4). | -13.5/+3.0/092 | OS | Physician | 07/31/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8203652&pc=MTA | 2023826-2018-01890 | 2018-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/1.0/147 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: DIOPTER CORRECTED TO: "-11.00/ 1.0/ 150 (SPHERE/CYLINDER/AXIS)" IN THE INITIAL MDR. ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -11.00/1.0/147 | OS |  | 11/28/2018 | 2018 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7870370&pc=MTA | 2023826-2018-01380 | 2018-07-17 | Injury | STAAR SURGICAL COMPANY | 2018-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+4.0/096 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018, IN A SEPARATE SURGERY, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, IRIDOCORNEAL TOUCH 360, LOSS OF ANTERIOR CHAMBER SPACE (AC), AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM RESOLVED WITH THE LENS EXPLANT, THE EYE "RETURNED TO NORMAL WITH A CLEAR CORNEA." THE SURGEON STATES THIS WAS A "VERY UNUSUAL CASE WITH EXPECTED SULCUS ANOMALIES" AND THAT HE "SUGGEST[S] SULCUS ANOMALIES." Manufacturer Narrative: ADDITIONAL INFORMATION: THE SURGEON STATES THE LENS WAS IMPLANTED AND IT WAS "VERY BIG." ALSO STATED IS THAT "THE PATIENT WAS CONSIDERED WITH SOME CONGENITAL ABNORMALITIES SO HE IS NOT A CANDIDATE FOR ICL." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+4.0/096 | OS | Physician | 08/16/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7906069&pc=MTA | 2023826-2018-01374 | 2018-07-17 | Injury | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT WHEN INITIALLY CALCULATING THE LENS, THE PATIENT'S ASTIGMATISM AXIS WAS INCORRECTLY ENTERED INTO THE OCOS. THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6., -5.00/+1.0/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. AFTER THE FIRST SURGERY, STAFF FOUND THAT HE INPUT THE WRONG AXIS ALSO VAULT WAS INADEQUATE. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND DIFFERENT AXIS. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.00/+1.0/099 | OD | Physician | 09/04/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7794386&pc=MTA | 2023826-2018-01197 | 2018-07-18 | Injury | STAAR SURGICAL COMPANY | 2018-08-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/142 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE SURGEON CHANGED THE AXIS OF THE LENS AND CHANGING THE ORIENTATION OF THE LENS CAN INCREASE THE VAULT. Manufacturer Narrative: CORRECTED DATA: (B)(4). ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) | -8.5/+1.5/142 | OS | Physician | 07/22/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7949804&pc=MTA | 2023826-2018-01460 | 2018-07-18 | Injury | STAAR SURGICAL COMPANY | 2018-10-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/104 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+1.5/104 | OD | Physician | 09/10/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7949817&pc=MTA | 2023826-2018-01461 | 2018-07-19 | Injury | STAAR SURGICAL COMPANY | 2018-10-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/069 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC DEFORMED. THERE WAS RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+1.5/069 | OS | Physician | 09/10/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8122320&pc=MTA | 2023826-2018-01765 | 2018-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.5/+1.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATED BY MFR: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND THE HAPTICS BENT. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.5/+1.5/094 | OD |  | 11/08/2018 | 2018 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824083&pc=MTA | 2023826-2018-01279 | 2018-07-20 | Injury | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.5/+2.0/083 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON THE SAME DATE, IN A SEPARATE SURGERY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CLAIM# (B)(4). | -8.5/+2.0/083 | OS | Physician | 08/03/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8182301&pc=MTA | 2023826-2018-01839 | 2018-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.00/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -10.00/1.5/095 | OS |  | 11/20/2018 | 2018 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8107280&pc=MTA | 2023826-2018-01726 | 2018-07-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. INTRAOPERATIVELY, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THIS IS EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+1.5/080 | OD |  | 11/02/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8002714&pc=MTA | 2023826-2018-01568 | 2018-07-24 | Injury | STAAR SURGICAL COMPANY | 2018-10-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -4.5/+1.5/118 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS RESIDUE ON LENS. (B)(4). Manufacturer Narrative: CORRECTED DATA: DEVICE EVALUATION: THERE WAS "NO RESIDUE ON LENS SURFACE" IN THE SUPPLEMENTAL MDR#1 CLAIM# (B)(4). Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. (B)(4). | -4.5/+1.5/118 | OS | Physician | 09/24/2018 | 2018 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7976069&pc=MTA | 2023826-2018-01525 | 2018-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -10.5/+1.5/085 | OS |  | 09/17/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8028759&pc=MTA | 2023826-2018-01548 | 2018-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -16.50/4.0/080 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -16.50/4.0/080 | OS |  | 10/16/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7870362&pc=MTA | 2023826-2018-01377 | 2018-07-31 | Injury | STAAR SURGICAL COMPANY | 2018-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.0/+3.0/125 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, IRIDOCORNEAL TOUCH 360, LOSS OF ANTERIOR CHAMBER SPACE (AC), AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PATIENT HAS NOT RETURNED SINCE THE LENS WAS EXPLANTED FOR FURTHER EVALUATION. THE SURGEON STATES THIS WAS A "VERY UNUSUAL CASE WITH EXPECTED SULCUS ANOMALIES." Manufacturer Narrative: ADDITIONAL INFORMATION: THE SURGEON STATES THE LENS IS "VERY BIG" AND THAT THE PATIENT "WAS CONSIDERED WITH SOME CONGENITAL ANOMALIES SO HE IS NOT A CANDIDATE FOR ICL." DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.0/+3.0/125 | OD | Physician | 08/16/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7949769&pc=MTA | 2023826-2018-01496 | 2018-07-31 | Injury | STAAR SURGICAL COMPANY | 2018-10-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+2.0/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.5/+2.0/082 | OD | Physician | 09/10/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8066005&pc=MTA | 2023826-2018-01654 | 2018-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -9.0/+1.0/103 | OD |  | 10/18/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8066011&pc=MTA | 2023826-2018-01657 | 2018-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -13.0/+2.5/092 | OS |  | 10/18/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8076072&pc=MTA | 2023826-2018-01653 | 2018-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATED BY MFR - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+2.5/094 | OS |  | 10/18/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8076089&pc=MTA | 2023826-2018-01645 | 2018-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.0/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATED BY MFR -DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.00/+1.0/097 | OD |  | 10/18/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8312989&pc=MTA | 2023826-2019-00196 | 2018-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-06 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+1.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN AND RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.0/+1.5/094 | OS |  | 01/10/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7949773&pc=MTA | 2023826-2018-01458 | 2018-08-01 | Injury | STAAR SURGICAL COMPANY | 2018-10-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.5/+1.5/173 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.5/+1.5/173 | OS | Physician | 09/10/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7898147&pc=MTA | 2023826-2018-01433 | 2018-08-02 | Injury | STAAR SURGICAL COMPANY | 2018-09-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -09.00/+2.5/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WAS OBSERVED, THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. (B)(4) | -09.00/+2.5/95 | OD | Physician | 08/23/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7860757&pc=MTA | 2023826-2018-01221 | 2018-08-03 | Injury | STAAR SURGICAL COMPANY | 2018-09-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM 513.2 -11.5/+2.5/98 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO PATIENTS RIGHT EYE ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN ALONG WITH RESIDUE ON THE LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/+2.5/98 | OD | Physician | 08/14/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7830767&pc=MTA | 2023826-2018-01298 | 2018-08-04 | Injury | STAAR SURGICAL COMPANY | 2018-08-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/141 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4)SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED (B)(4). | -11.5/+2.0/141 | OS | Physician | 08/06/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8194056&pc=MTA | 2023826-2018-01891 | 2018-08-07 | Injury | STAAR SURGICAL COMPANY | 2018-12-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPIC TORN. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -12.5/+1.5/090 | OS | Physician | 11/28/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7976084&pc=MTA | 2023826-2018-01524 | 2018-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE DEVICE WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+1.5/073 | OS |  | 09/17/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7857645&pc=MTA | 2023826-2018-01365 | 2018-08-10 | Injury | STAAR SURGICAL COMPANY | 2018-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO WRONG SIZING RECOMMENDED. AT LAST POST-OP VISIT, THE RIGHT EYE IS RED BUT FEELS WELL. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) | -12.0/+1.5/090 | OD | Physician | 08/12/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8207006&pc=MTA | 2023826-2018-01942 | 2018-08-10 | Injury | STAAR SURGICAL COMPANY | 2018-12-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+1.5/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.0/+1.5/083 | OS | Physician | 12/03/2018 | 2018 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8204876&pc=MTA | 2023826-2018-01938 | 2018-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+3.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION- THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+3.0/092 | OS |  | 12/03/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7886288&pc=MTA | 2023826-2018-01407 | 2018-08-13 | Injury | STAAR SURGICAL COMPANY | 2018-09-18 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.5/+2.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS USER ERROR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+2.5/093 | OS | Physician | 08/23/2018 | 2018 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7886290&pc=MTA | 2023826-2018-01404 | 2018-08-13 | Injury | STAAR SURGICAL COMPANY | 2018-09-18 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS USER ERROR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+2.0/094 | OD | Physician | 08/23/2018 | 2018 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8323169&pc=MTA | 2023826-2019-00205 | 2018-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+4.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION- THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+4.0/087 | OD |  | 01/10/2019 | 2019 | 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7947544&pc=MTA | 2023826-2018-01489 | 2018-08-17 | Injury | STAAR SURGICAL COMPANY | 2018-10-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.00/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WAS RETURNED DRY IN A MICRO-CENTRUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH FOREIGN MATERIAL (RESIDUE) ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/2.5/093 | OD | Physician | 10/02/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8048831&pc=MTA | 2023826-2018-01576 | 2018-08-17 | Injury | STAAR SURGICAL COMPANY | 2018-11-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+2.5/092 | OD | Physician | 10/08/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8048863&pc=MTA | 2023826-2018-01577 | 2018-08-17 | Injury | STAAR SURGICAL COMPANY | 2018-11-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+2.5/102 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+2.5/102 | OD | Physician | 10/08/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8148525&pc=MTA | 2023826-2018-01794 | 2018-08-17 | Injury | STAAR SURGICAL COMPANY | 2018-12-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+4.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+4.0/090 | OD | Physician | 11/10/2018 | 2018 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7890056&pc=MTA | 2023826-2018-01413 | 2018-08-20 | Injury | STAAR SURGICAL COMPANY | 2018-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON DECIDED TO IMPLANT AN 13.2 MM VTICM5_13.2, -12.50/+5.0/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMERLENS LENS INTO PAIENTS RIGHT EYE ON (B)(6) 2018. THE LENGTH WAS TO SHORT CAUSING LOW VAULT. IT WAS HIS FAULT NOT TO SELECT THE 13.7MM LENS LENGTH WHICH IS NOW IN THE PAITENTS EYE (AFTER EXCHANGE ON (B)(6) 2018) AND SHOWS A GOOD VAULTING OF ~400M. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE ON LENS, IN MICRO-CENTRIFUGE VIAL, CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TEAR AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, EXCHANGE (B)(4). | -12.50/+5.0/107 | OD | Physician | 08/21/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7976094&pc=MTA | 2023826-2018-01514 | 2018-08-20 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+2.0/075 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTICS DEFORMED , WITH RESIDUE ON THE LENS. CLAIM (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+2.0/075 | OD | Physician | 09/17/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7976238&pc=MTA | 2023826-2018-01522 | 2018-08-20 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/081 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+2.0/081 | OD | Physician | 09/17/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7924355&pc=MTA | 2023826-2018-01478 | 2018-08-21 | Injury | STAAR SURGICAL COMPANY | 2018-10-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -16.00/4.0/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL STUCK ON A PIECE OF TAPE. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT; DEBRIS AND RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.00/4.0/076 | OD | Physician | 09/05/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7945361&pc=MTA | 2023826-2018-01502 | 2018-08-21 | Injury | STAAR SURGICAL COMPANY | 2018-10-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+2.5/092 | OS | Physician | 09/14/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7976196&pc=MTA | 2023826-2018-01520 | 2018-08-21 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.0/103 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+1.0/103 | OD | Physician | 09/17/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7976200&pc=MTA | 2023826-2018-01521 | 2018-08-21 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTICS BENT AND DEFORMED. THERE WAS RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+2.5/093 | OS | Physician | 09/17/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8073848&pc=MTA | 2023826-2018-01688 | 2018-08-21 | Injury | STAAR SURGICAL COMPANY | 2018-11-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.00/4.0/089 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS ON (B)(6) 2018 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/4.0/089 | OD | Physician | 10/18/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8073856&pc=MTA | 2023826-2018-01692 | 2018-08-21 | Injury | STAAR SURGICAL COMPANY | 2018-11-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.00/4.0/085 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEBRIS ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/4.0/085 | OS | Physician | 10/18/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7976009&pc=MTA | 2023826-2018-01519 | 2018-08-22 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+3.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -11.5/+3.0/094 | OS | Physician | 09/17/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8073879&pc=MTA | 2023826-2018-01689 | 2018-08-22 | Injury | STAAR SURGICAL COMPANY | 2018-11-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.00/2.5/095 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WAS RETURNED DRY, IN A MICRO-CENTRFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -14.00/2.5/095 | OD | Physician | 10/18/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8075922&pc=MTA | 2023826-2018-01639 | 2018-08-22 | Injury | STAAR SURGICAL COMPANY | 2018-11-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.50/+1.5/001 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN AND HAPTIC TORN AND BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.50/+1.5/001 | OD | Physician | 10/18/2018 | 2018 | 47
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7924401&pc=MTA | 2023826-2018-01479 | 2018-08-23 | Injury | STAAR SURGICAL COMPANY | 2018-10-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -16.5/4.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL STUCK ON A PIECE OF TAPE. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND RESIDUE AND DEBRIS ON LENS, LENS STUCK ON A PIECE OF TAPE. (B)(4). Manufacturer Narrative: CORRECTED DATA: IMPLANT DATE CORRECTED TO: "(B)(6) 2018" IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.5/4.0/093 | OS | Physician | 09/05/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8208764&pc=MTA | 2023826-2019-00007 | 2018-08-23 | Injury | STAAR SURGICAL COMPANY | 2019-01-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -16.0/+3.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA DATE RECEIVED BY MANUFACTURER CORRECTED FROM "17-JAN-2018" TO "17-JAN-2019" FOR FOLLOW-UP# 1. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | -16.0/+3.0/090 | OS | Physician | 12/07/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8122405&pc=MTA | 2023826-2018-01764 | 2018-08-25 | Injury | STAAR SURGICAL COMPANY | 2018-12-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.5/+2.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.5/+2.0/091 | OD | Physician | 11/04/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7906084&pc=MTA | 2023826-2018-01393 | 2018-08-27 | Injury | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.5/+6/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -11.5/+6/085 | OS | Physician | 09/05/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8028747&pc=MTA | 2023826-2018-01546 | 2018-08-28 | Injury | STAAR SURGICAL COMPANY | 2018-11-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -17.50/1.5/091 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17.50/1.5/091 | OS | Physician | 10/16/2018 | 2018 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8090046&pc=MTA | 2023826-2018-01695 | 2018-08-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+2.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATED BY MFR THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC BENT. CLAIM #(B)(4). Manufacturer Narrative: CLAIM #(B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -15.5/+2.0/093 | OS |  | 10/25/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8149312&pc=MTA | 2023826-2018-01812 | 2018-08-29 | Injury | STAAR SURGICAL COMPANY | 2018-12-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+1.5/096 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/+1.5/096 | OD | Physician | 11/20/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7976168&pc=MTA | 2023826-2018-01518 | 2018-08-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS WAS CAUGHT IN THE VIAL CAP AND HAD A DEFORMED SHAPE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.5/098 | OS |  | 09/17/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8073878&pc=MTA | 2023826-2018-01691 | 2018-08-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/1.0/114 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATED BY MFR-DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.0/114 | OS |  | 10/18/2018 | 2018 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7976245&pc=MTA | 2023826-2018-01523 | 2018-08-30 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+2.5/094 | OS | Physician | 09/17/2018 | 2018 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8075930&pc=MTA | 2023826-2018-01649 | 2018-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.5/+3.5/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.5/+3.5/098 | OS |  | 10/18/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8404309&pc=MTA | 2023826-2019-00376 | 2018-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -7.5/1.5/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. CORRECTED DATA: COMMON DEVICE NAME SHOULD BE CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS IN ORIGINAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.5/1.5/079 | OS |  | 02/14/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7976120&pc=MTA | 2023826-2018-01515 | 2018-09-05 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.5/085 | OS | Physician | 09/17/2018 | 2018 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8075916&pc=MTA | 2023826-2018-01650 | 2018-09-05 | Injury | STAAR SURGICAL COMPANY | 2018-11-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+1.5/108 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: LENS WAS RETURNED DRY IN MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.50/+1.5/108 | OD | Physician | 10/18/2018 | 2018 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7976130&pc=MTA | 2023826-2018-01516 | 2018-09-06 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.5/+3.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+3.0/088 | OS | Physician | 09/17/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7976021&pc=MTA | 2023826-2018-01517 | 2018-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+1.5/094 | OS |  | 09/17/2018 | 2018 | 47
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8075954&pc=MTA | 2023826-2018-01643 | 2018-09-07 | Injury | STAAR SURGICAL COMPANY | 2018-11-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.0/093 | OD | Physician | 10/18/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8075992&pc=MTA | 2023826-2018-01644 | 2018-09-07 | Injury | STAAR SURGICAL COMPANY | 2018-11-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.0/084 | OS | Physician | 10/18/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8307845&pc=MTA | 2023826-2019-00218 | 2018-09-07 | Injury | STAAR SURGICAL COMPANY | 2019-02-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TEAR IN THE LENS HAPTIC AND RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+2.0/092 | OD | Physician | 01/10/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8117912&pc=MTA | 2023826-2018-01758 | 2018-09-10 | Injury | STAAR SURGICAL COMPANY | 2018-11-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.5/+3.5/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE SURGEON DOES NOT PLAN TO IMPLANT ANOTHER ICL LENS. Manufacturer Narrative: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. DIMENSIONAL INSPECTION FOUND LENS TO BE SPECIFICATION. (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/+3.5/080 | OD | Physician | 11/01/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8117933&pc=MTA | 2023826-2018-01759 | 2018-09-10 | Injury | STAAR SURGICAL COMPANY | 2018-11-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -9.5/+4.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE SURGEON DOES NOT PLAN TO IMPLANT ANOTHER ICL LENS. Manufacturer Narrative: LENS WAS RETURNED DRY IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. DIMESIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/+4.0/089 | OS | Physician | 11/01/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8127056&pc=MTA | 2023826-2018-01774 | 2018-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-03 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/052 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. STATUS OF THE EYE IS REPORTED AS "OK." THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATED BY MFR -THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.5/052 | OS |  | 11/08/2018 | 2018 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8176867&pc=MTA | 2023826-2018-01860 | 2018-09-12 | Injury | STAAR SURGICAL COMPANY | 2018-12-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.00/+1.50/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.00/+1.50/100 | OS | Physician | 11/19/2018 | 2018 | 51
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8176883&pc=MTA | 2023826-2018-01861 | 2018-09-12 | Injury | STAAR SURGICAL COMPANY | 2018-12-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.00/+1.50/104 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF DEBRIS ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.00/+1.50/104 | OD | Physician | 11/19/2018 | 2018 | 51
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8028734&pc=MTA | 2023826-2018-01543 | 2018-09-14 | Injury | STAAR SURGICAL COMPANY | 2018-11-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.00/2.0/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISUTRE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. ON (B)(6) 2018" SHOULD BE CORRECTED TO "ON (B)(6) 2018" (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.00/2.0/100 | OD | Physician | 10/16/2018 | 2018 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8028735&pc=MTA | 2023826-2018-01544 | 2018-09-14 | Injury | STAAR SURGICAL COMPANY | 2018-11-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.50/2.5/141 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.50/2.5/141 | OS | Physician | 10/16/2018 | 2018 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8002534&pc=MTA | 2023826-2018-01558 | 2018-09-20 | Injury | STAAR SURGICAL COMPANY | 2018-10-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.5/+1.0/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.5/+1.0/073 | OD | Physician | 09/24/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8148455&pc=MTA | 2023826-2018-01790 | 2018-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.5/+2.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATED BY MFR - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.5/+2.0/095 | OD |  | 11/10/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8148488&pc=MTA | 2023826-2018-01791 | 2018-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -16.0/+1.5/078 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATED BY MFR - DEVICE EVALUATION - THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.0/+1.5/078 | OS |  | 11/10/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8090068&pc=MTA | 2023826-2018-01700 | 2018-09-27 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -8.5/+1.5/092 | OD | Physician | 10/26/2018 | 2018 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8172144&pc=MTA | 2023826-2018-01846 | 2018-09-28 | Injury | STAAR SURGICAL COMPANY | 2018-12-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO LOW VAULT. ON (B)(6) 2018 A LONGER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -12.5/+2.5/085 | OS | Physician | 11/23/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8200342&pc=MTA | 2023826-2018-01917 | 2018-09-28 | Injury | STAAR SURGICAL COMPANY | 2018-12-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. THE LENS IS NOTED TO HAVE A CURVED SHAPE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.5/090 | OS | Physician | 11/28/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8200063&pc=MTA | 2023826-2018-01916 | 2018-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.50/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -17.50/1.5/095 | OD |  | 11/28/2018 | 2018 | 49
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8194058&pc=MTA | 2023826-2018-01893 | 2018-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0/+2.5/123 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.0/+2.5/123 | OS |  | 11/28/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8200926&pc=MTA | 2023826-2018-01915 | 2018-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.50/2.0/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6 )2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -10.50/2.0/069 | OD |  | 11/28/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8200771&pc=MTA | 2023826-2018-01918 | 2018-10-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6.2MM VTICMO12.6, -13.50/3.0/089 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: LENS SIZE CORRECTED TO: "12.6MM" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.50/3.0/089 | OS |  | 11/28/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8073838&pc=MTA | 2023826-2018-01627 | 2018-10-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -11.50/5.0/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT WAS OBSERVED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: RESULT CODE CORRECTED TO: "213" IN THE INITIAL MDR, PREVIOUS CODE NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/5.0/100 | OD |  | 10/16/2018 | 2018 | 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8271026&pc=MTA | 2023826-2019-00169 | 2018-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -11.50/1.0/028 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLMAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. THE FOLLOWING ADDITIONAL INFORMATION WAS REPORTED: IT WAS VERY LOW VAULT IN YOUNG PATIENT, NEEDED TO EXCHANGE WITH LARGER LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/1.0/028 | OD |  | 12/26/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8133039&pc=MTA | 2023826-2018-01785 | 2018-10-09 | Injury | STAAR SURGICAL COMPANY | 2018-12-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6 MM VTICMO12.6, -17.50/2.0/081, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018 AS A REPLACEMENT LENS BUT IT DID NOT RESOLVE THE EXCESSIVE VAULT. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO THE VAULT AND A SHORTER LENGTH LENS WAS IMPLANTED. IT WAS REPORTED THAT THIS REPLACEMENT RESOLVED THE PROBLEM. PATIENT RELATED FACTOR WAS REPORTED AS CAUSE OF EVENT. Manufacturer Narrative: DATE OF THIS REPORT CORRECTED TO: "08-NOV-2018" IN THE INITIAL MDR. LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17.50/2.0/081 | OD | Physician | 08/11/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8176996&pc=MTA | 2023826-2018-01865 | 2018-10-09 | Injury | STAAR SURGICAL COMPANY | 2018-12-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. RETURNED PRODUCT EVALUATION FOUND THE OPTIC AND HAPTIC TORN ALONG WITH RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+2.5/095 | OD | Physician | 11/19/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8051065&pc=MTA | 2023826-2018-01601 | 2018-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -02.50/0.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS WERE OBSERVED. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN DRY; IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN, BENT AND DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -02.50/0.5/089 | OD |  | 10/16/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8051070&pc=MTA | 2023826-2018-01602 | 2018-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -03.00/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS WERE OBSERVED, THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE AND PRESENCE OF FIBERS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -03.00/1.0/087 | OS |  | 10/16/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8158160&pc=MTA | 2023826-2018-01797 | 2018-10-10 | Injury | STAAR SURGICAL COMPANY | 2018-12-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+2.0/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+2.0/084 | OS | Physician | 11/14/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8194050&pc=MTA | 2023826-2018-01896 | 2018-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.5/089 | OS |  | 11/28/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8378733&pc=MTA | 2023826-2019-00340 | 2018-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.0/2.5/001 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. IT WAS REPORTED THAT ON (B)(6) 2019 A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONLY A PORTION OF LENS WAS RETURNED. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.0/2.5/001 | OD |  | 01/30/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8203471&pc=MTA | 2023826-2018-01919 | 2018-10-12 | Injury | STAAR SURGICAL COMPANY | 2018-12-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.00/1.0/092 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -8.00/1.0/092 | OD | Physician | 11/28/2018 | 2018 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8107290&pc=MTA | 2023826-2018-01724 | 2018-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATED BY MFR -DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.0/+1.5/070 | OS |  | 11/01/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8194046&pc=MTA | 2023826-2018-01895 | 2018-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/096 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED; PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+1.0/096 | OD |  | 11/28/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8073881&pc=MTA | 2023826-2018-01629 | 2018-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -8.00/1.5/088 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATED BY MFR -DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -8.00/1.5/088 | OD |  | 11/06/2018 | 2018 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8336988&pc=MTA | 2023826-2019-00241 | 2018-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+3.0/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+3.0/100 | OD |  | 01/14/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8203332&pc=MTA | 2023826-2018-01911 | 2018-10-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/2.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATED BY MFR - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND RESIDUE ON LENS. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -9.50/2.0/082 | OS |  | 11/28/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8201021&pc=MTA | 2023826-2018-01910 | 2018-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -7.50/1.5/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DEFORMED; PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -7.50/1.5/098 | OS |  | 11/28/2018 | 2018 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8176971&pc=MTA | 2023826-2018-01863 | 2018-10-23 | Injury | STAAR SURGICAL COMPANY | 2018-12-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+1.5/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+1.5/082 | OD | Physician | 11/19/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8122408&pc=MTA | 2023826-2018-01766 | 2018-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/075 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATED BY MFR - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+2.0/075 | OD |  | 11/08/2018 | 2018 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8176976&pc=MTA | 2023826-2018-01864 | 2018-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATED BY MFR -DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+2.0/096 | OS |  | 11/19/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8323013&pc=MTA | 2023826-2019-00203 | 2018-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+3.0/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION DEVICE EVALUATED BY MFR : DEVICE EVALUATION- THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS OPTIC TORN. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+3.0/082 | OD |  | 01/10/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8296789&pc=MTA | 2023826-2019-00135 | 2018-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/1.5/069 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/1.5/069 | OS |  | 01/10/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8129050&pc=MTA | 2023826-2018-01778 | 2018-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+3.0/139 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR AN ALTERNATE SAME SIZE LENS PLACED ON A DIFFERENT AXIS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. THE SURGEON STATES, "CHANGING ORIENTATION TO MAKE VAULT HIGHER." Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATED BY MFR - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+3.0/139 | OD |  | 11/09/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8149319&pc=MTA | 2023826-2018-01802 | 2018-10-29 | Injury | STAAR SURGICAL COMPANY | 2018-12-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/105 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. Manufacturer Narrative: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DEFORMED; AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -9.0/+1.0/105 | OS | Physician | 11/14/2018 | 2018 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8229434&pc=MTA | 2023826-2019-00028 | 2018-10-31 | Injury | STAAR SURGICAL COMPANY | 2019-01-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+0.5/179 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA DEVICE EVALUATION- THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+0.5/179 | OS | Physician | 12/11/2018 | 2018 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8401914&pc=MTA | 2023826-2019-00373 | 2018-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -6.5/1.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC HAD A TEAR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.5/1.5/097 | OD |  | 02/14/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8403993&pc=MTA | 2023826-2019-00374 | 2018-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -5.5/1.0/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.5/1.0/104 | OS |  | 02/14/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8129042&pc=MTA | 2023826-2018-01780 | 2018-11-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+2.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATED BY MFR -DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT; AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.0/+2.0/090 | OD |  | 11/09/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8409313&pc=MTA | 2023826-2019-00386 | 2018-11-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/129 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISUAL DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.0/129 | OD |  | 02/15/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8237899&pc=MTA | 2023826-2019-00077 | 2018-11-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.0/083 | OD |  | 12/11/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8208765&pc=MTA | 2023826-2019-00001 | 2018-11-13 | Injury | STAAR SURGICAL COMPANY | 2019-01-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0/+3.0/101 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND UNREACTIVE (FIXED) PUPIL. THE PROBLEM IS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+3.0/101 | OD | Physician | 12/07/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8208779&pc=MTA | 2023826-2019-00004 | 2018-11-13 | Injury | STAAR SURGICAL COMPANY | 2019-01-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0/+4.0/085 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND UNREACTIVE (FIXED) PUPIL. THE PROBLEM IS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+4.0/085 | OS | Physician | 12/07/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8237884&pc=MTA | 2023826-2019-00029 | 2018-11-13 | Injury | STAAR SURGICAL COMPANY | 2019-01-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+2.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE LENS OPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.5/+2.0/087 | OS | Physician | 12/11/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8203419&pc=MTA | 2023826-2018-01889 | 2018-11-14 | Injury | STAAR SURGICAL COMPANY | 2018-12-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.00/1.5/148 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -16.00/1.5/148 | OS | Physician | 11/28/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8227433&pc=MTA | 2023826-2019-00046 | 2018-11-15 | Injury | STAAR SURGICAL COMPANY | 2019-01-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -2.0/+4.0/052 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS ATTEMPTED TO BE EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS NOT RESOLVED. SEE MFR REPORT 2023826-2019-00049 FOR CASE CONTINUATION. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -2.0/+4.0/052 | OS | Physician | 12/13/2018 | 2018 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8227419&pc=MTA | 2023826-2019-00052 | 2018-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/077 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS.-METHOD CODE ADDED CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -13.5/+2.0/077 | OS |  | 12/14/2018 | 2018 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8194052&pc=MTA | 2023826-2018-01898 | 2018-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -14.0/+1.5/114 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. METHOD CODE ADDED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+1.5/114 | OD |  | 11/30/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8194496&pc=MTA | 2023826-2018-01901 | 2018-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED FOR A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. METHOD CODE ADDED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.0/070 | OS |  | 11/30/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8287618&pc=MTA | 2023826-2019-00181 | 2018-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.5/+1.0/087 | OD |  | 01/03/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8194045&pc=MTA | 2023826-2018-01899 | 2018-11-22 | Injury | STAAR SURGICAL COMPANY | 2018-12-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.0/+0.5/096 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. "THE SURGEON AND THE PATIENT ARE HAPPY WITH THE OUTCOME." Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE/DEBRIS ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -6.0/+0.5/096 | OS | Physician | 11/30/2018 | 2018 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8227416&pc=MTA | 2023826-2019-00051 | 2018-11-23 | Injury | STAAR SURGICAL COMPANY | 2019-01-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -12.5/+2.5/090 | OS | Physician | 12/14/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8164901&pc=MTA | 2023826-2018-01824 | 2018-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2, -11.50/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BENT AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/2.0/093 | OD |  | 12/04/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8164912&pc=MTA | 2023826-2018-01825 | 2018-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 -12.00/2.5/093, (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATED BY MFR - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -12.00/2.5/093 | OS |  | 12/04/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8273971&pc=MTA | 2023826-2019-00131 | 2018-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -10.00/1.0/002 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/1.0/002 | OS |  | 01/10/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8356039&pc=MTA | 2023826-2019-00284 | 2018-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION- THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+2.5/094 | OD |  | 12/27/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8194049&pc=MTA | 2023826-2018-01897 | 2018-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -14.5/+1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.5/+1.5/094 | OD |  | 11/30/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8194051&pc=MTA | 2023826-2018-01894 | 2018-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/077 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THE LENS HAS A CURVED SHAPE. METHOD CODE ADDED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+1.5/077 | OS |  | 11/30/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8322753&pc=MTA | 2023826-2019-00199 | 2018-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+3.0/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION- THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+3.0/096 | OS |  | 01/10/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8313004&pc=MTA | 2023826-2019-00197 | 2018-12-02 | Injury | STAAR SURGICAL COMPANY | 2019-02-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+0.5/179 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+0.5/179 | OS | Physician | 01/10/2019 | 2019 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8404797&pc=MTA | 2023826-2019-00352 | 2018-12-02 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.5/+3.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS REMAINS IMPLANTED. EXPLANTATION OF THE LENS IS NOT PLANNED TO DATE. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -6.5/+3.0/088 | OD | Physician | 02/07/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8322887&pc=MTA | 2023826-2019-00200 | 2018-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.5/093 | OD |  | 01/10/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8273963&pc=MTA | 2023826-2019-00129 | 2018-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/1.5/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/1.5/079 | OD |  | 01/10/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8337028&pc=MTA | 2023826-2019-00245 | 2018-12-06 | Injury | STAAR SURGICAL COMPANY | 2019-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -0.5/+1.0/031 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PIGMENT DISPERSION. ANOTHER ICL LENS WAS NOT IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS WAS DISPOSED OF AT THE FACILITY AND WILL NOT BE RETURNED. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -0.5/+1.0/031 | OD | Physician | 01/14/2019 | 2019 | 67
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8287617&pc=MTA | 2023826-2019-00179 | 2018-12-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019, THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+1.0/086 | OD |  | 01/04/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8287620&pc=MTA | 2023826-2019-00182 | 2018-12-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -7.5/+1.0/104 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019, THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.5/+1.0/104 | OS |  | 01/04/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8354669&pc=MTA | 2023826-2019-00294 | 2018-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATE THAT THE SURGEON IMPLANTED A 12.1MM, VTICM5_12.1, -11.0/2.5/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION : LENS WAS RETURNED IN DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. PRESENCE OF RESIDUE ON LENS SURFACE WAS ALSO OBSERVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/2.5/107 | OD |  | 01/23/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8255315&pc=MTA | 2023826-2019-00112 | 2018-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.5/+2.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.5/+2.0/089 | OD |  | 12/17/2018 | 2018 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8255307&pc=MTA | 2023826-2019-00111 | 2018-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/065 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018, DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.0/065 | OS |  | 12/17/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8255324&pc=MTA | 2023826-2019-00113 | 2018-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.0/+2.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA CLAIM# (B')(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.0/+2.0/093 | OS |  | 12/17/2018 | 2018 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8275003&pc=MTA | 2023826-2019-00171 | 2018-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM,VTICMO12.6, -11.00/1.5/098 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: THIS IS A DUPLICATE FILE OF MDR # 2023826-2019-00154. PLEASE REFER TO MDR # 2023826-2019-00154 FOR FURTHER DETAILS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/1.5/098 | OD |  | 12/27/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8287333&pc=MTA | 2023826-2019-00132 | 2018-12-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.00/3.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION- THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/3.0/087 | OS |  | 01/10/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8352027&pc=MTA | 2023826-2019-00288 | 2018-12-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A12.1MM, VTICM5_12.1, -8.00/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: LENS WAS RETURNED IN DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND TEARS IN THE OPTIC AND A HAPTIC WAS TORN. PRESENCE OF RESIDUE ON LENS SURFACE WAS ALSO OBSERVED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.00/2.0/089 | OD |  | 01/23/2019 | 2019 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8356224&pc=MTA | 2023826-2019-00296 | 2019-01-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATE THAT THE SURGEON IMPLANTED A 13.2MM, VTICM5_13.2, -12.00/2.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS DUE TO HIGH VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. PRESENCE OF RESIDUE ON LENS SURFACE WAS OBSERVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/2.5/084 | OS |  | 01/23/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372522&pc=MTA | 2023826-2019-00328 | 2019-01-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION DEVICE EVALUATED BY MFR : DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.0/+1.5/103 | OD |  | 02/01/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8409326&pc=MTA | 2023826-2019-00391 | 2019-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -6.5/+1.5/105 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.5/+1.5/105 | OD |  | 02/15/2019 | 2019 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8710342&pc=MTA | 2023826-2019-01097 | 2019-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -18.0/+1.0/115 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). Manufacturer Narrative: PHAKIC INTRAOCULAR LENS, PRODUCT CODE: MTA SHOULD BE CHANGED TO " PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB " IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL, DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN AND FIBERS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: "DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS", SHOULD HAVE BEEN INCLUDED WITH MEDWATCH SUPPLEMENTAL #2. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+1.0/115 | OD |  | 05/23/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6041624&pc=MTA | 2023826-2016-01494 |  | Injury | STAAR SURGICAL COMPANY | 2016-10-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS - -11.50/+3.0/093 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.50/+3.0/093 | OD | Physician | 09/26/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6091604&pc=MTA | 2023826-2016-01602 |  | Injury | STAAR SURGICAL COMPANY | 2016-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS, -9.5/+2.50/091 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL (NOT POSSIBLE TO SPECIFY) ON LENS SURFACE. CORRECTED DATA: REPORT SOURCE IS DISTRIBUTOR, NOT USER FACILITY. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+2.50/091 | OS | Physician | 10/13/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6094684&pc=MTA | 2023826-2016-01586 |  | Injury | STAAR SURGICAL COMPANY | 2016-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50/+2.0/086 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED AND FIBERS ON LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.50/+2.0/086 | OD | Physician | 10/13/2016 | 2016 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6234140&pc=MTA | 2023826-2017-00037 |  | Injury | STAAR SURGICAL COMPANY | 2017-01-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+4.5/111 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE VAULT VALUE WENT FROM 1215 MICROMETERS TO 452 MICROMETERS. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DRIED, DISCOLORED MATERIAL ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -18.0/+4.5/111 | OD | Physician | 12/14/2016 | 2016
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6246839&pc=MTA | 2023826-2017-00048 |  | Injury | STAAR SURGICAL COMPANY | 2017-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+0.5/59 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ON (B)(6) 2015, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE LENS HAS BEEN RETURNED HOWEVER, AS OF REPORT SUBMISSION DATE, IT HAS NOT YET BEEN EVALUATED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+0.5/59 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT (VAULT VALUE: 1182 UM) AND WAS EXCHANGED ON (B)(6) 2016 FOR A SHORTER SAME MODEL, SIMILAR DIOPTER LENS. PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS REPORTED AS 20/25 WITH VAULT VALUE = 450 UM. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.5/+0.5/59 | OS | Physician | 12/14/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6432222&pc=MTA | 2023826-2017-00438 |  | Injury | STAAR SURGICAL COMPANY | 2017-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11/1.0/139 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11/1.0/139 | OS | Physician | 02/23/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6527428&pc=MTA | 2023826-2017-00722 |  | Injury | STAAR SURGICAL COMPANY | 2017-04-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/76 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, PIGMENT DISPERSION, AND UNDERWENT THE SECONDARY YAG. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH OCULAR HYPERTENSION IN EITHER EYE. (B)(4). Manufacturer Narrative: (B)(4). | -13.5/+1.5/76 | OS | Physician | 03/30/2017 | 2017 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6546101&pc=MTA | 2023826-2017-00767 |  | Injury | STAAR SURGICAL COMPANY | 2017-05-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/X008 IN THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED EXCESSIVE VAULTING, NARROWING OF THE IRIDO-CORNEAL ANGLES. THE DOCTOR PLANS TO EXCHANGE TO A SMALLER LENS AND SAID THAT THE PROBLEM IS DUE TO W-T-W MEASUREMENT. AT THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. PATIENT DOES NOT UNDERSTAND WHY EXCHANGE IS NECESSARY WHEN VISION IS GOOD. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/+1.0/x008 | OS | Physician | 04/05/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6811211&pc=MTA | 2023826-2017-01341 |  | Injury | STAAR SURGICAL COMPANY | 2017-08-22 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 , -7.5/+1.5/112 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO PATIENT EXPERIENCED EXCESSIVE VAULT. ATTEMPTS TO GET MORE INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.5/+1.5/112 | NA | Physician | 07/24/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6836404&pc=MTA | 2023826-2017-01384 |  | Injury | STAAR SURGICAL COMPANY | 2017-08-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/161 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED LENS HAD EXCESSIVE VAULTING, CROWDED ANGLES AND THE ANTERIOR CHAMBER WAS VERY SHALLOW. THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/161 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED LENS HAD EXCESSIVE VAULTING, CROWDED ANGLES AND THE ANTERIOR CHAMBER WAS VERY SHALLOW. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PATIENT IS HAPPY NOW. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEBRIS ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+1.5/161 | OD | Physician | 08/02/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6998023&pc=MTA | 2023826-2017-01684 |  | Injury | STAAR SURGICAL COMPANY | 2017-11-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -0.5/+4.5/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF LENS HAPTIC BROKEN AND BENT. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -0.5/+4.5/085 | OS | Physician | 10/09/2017 | 2017 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7211131&pc=MTA | 2023826-2018-00118 |  | Injury | STAAR SURGICAL COMPANY | 2018-01-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -7.5/+1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, SHALLOW AC, AND THE IRIS BEING PUSHED FORWARD. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT UNDER THE AGE OF 21 YEARS OLD. (B)(4). | -7.5/+1.0/090 | OS | Physician | 12/27/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7629864&pc=MTA | 2023826-2018-00948 |  | Injury | STAAR SURGICAL COMPANY | 2018-06-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE REPORTER INDICATED THEY FELT THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) | -11.00/+1.5/097 | OS | Physician | 05/28/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7786610&pc=MTA | 2023826-2018-01217 |  | Injury | STAAR SURGICAL COMPANY | 2018-08-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/5.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO ASYMMETRICAL VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DRY, CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. CLAIM# (B)(4). | -10.00/5.5/091 | OS | Physician | 07/17/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7903841&pc=MTA | 2023826-2018-01391 |  | Injury | STAAR SURGICAL COMPANY | 2018-09-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.00/+3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18.00/+3.0/093 | OS | Physician | 08/28/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8471358&pc=MTA | 2023826-2019-00425 |  | Injury | STAAR SURGICAL COMPANY | 2019-04-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.00/2.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WAS OBSERVED, THE LENS WAS REMOVED ON (B)(6) 2018 AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.00/2.5/081 | OD | Physician | 03/05/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4922988&pc=MTA | 2023826-2015-00820 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-07-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/13. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -10.0/+1.5/155 | OS | Physician | 06/18/2015 | 2015 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7253318&pc=MTA | 2023826-2018-00150 | 2017-12-04 | Injury | STAAR SURGICAL COMPANY | 2018-02-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+4.0/105 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017. THE LENS EXCHANGE RESOLVED THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER STATED THAT "THE DOCTOR'S REPLY WAS A SPARE LENS FOR THE SAME OPERATION AND RESERVED." THE LENS WAS EXPLANTED AND EXCHANGED PRESUMABLY DUE TO OBSERVATION OF AN EXCESSIVE VAULT. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT, FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS), AND THE LENS MISSING A PIECE OF HAPTIC. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.50/+4.0/105 | OD | Physician | 01/10/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7253307&pc=MTA | 2023826-2018-00149 | 2017-12-11 | Injury | STAAR SURGICAL COMPANY | 2018-02-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+1.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017. THE LENS EXCHANGE RESOLVED THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXPLANTED DUE TO "HIGH ARCHES AND THERE WERE NO OTHER COMPLICATIONS. THE PATIENT WAS IN GOOD CONDITION AFTER THE REPLACEMENT." (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/+1.5/087 | OS | Physician | 01/10/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5830544&pc=MTA | 2023826-2016-00989 | 2012-07-06 | Injury | STAAR SURGICAL COMPANY AG | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.0/+2.5/083 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND REMAINS IMPLANTED. THE PATIENT EXPERIENCED A REFRACTIVE CHANGE OVERTIME. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/16. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4). | -10.0/+2.5/083 | OS | Physician | 07/12/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4923018&pc=MTA | 2023826-2015-00821 | 2013-03-04 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/13. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). LENS NOT RETURNED. | -9.5/+2.5/035 | OD | Physician | 06/18/2015 | 2015 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4786542&pc=MTA | 2023826-2015-00570 | 2013-08-17 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-20 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2014 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS DRY WITH NO VISIBLE DAMAGE. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND WITHIN THE WORKORDER. MEDICAL REVIEW: ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSION - BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4) SECONDARY SURGERY. NEW INCISION. EXPLANTED. VAULTING. (B)(4). | -6.50/1.50/096 | OS | Physician | 02/05/2015 | 2015
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4954776&pc=MTA | 2023826-2015-00963 | 2013-11-07 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). | -12.5/+1.5/005 | OS | Physician | 07/14/2015 | 2015 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4806954&pc=MTA | 2023826-2015-00584 | 2014-01-11 | Injury | STAAR SURGICAL COMPANY | 2015-05-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -16.5/+3.5/078 | OS | Physician | 02/17/2015 | 2015 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4786529&pc=MTA | 2023826-2015-00569 | 2014-05-12 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-20 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. EXCESSIVE VAULT AND CLOSURE OF IRIDO-CORNEAL ANGLES ASSOCIATED WITH ELEVATED IOP WERE NOTED, THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2014, AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST KNOWN UCVA WAS 20/20 AS OF (B)(6) 2014. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW & WORK ORDER SEARCH. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITION MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY OF THESE COMPLICATIONS FOR OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. THERE IS NOT ENOUGH INFORMATION TO DETERMINE THE EXACT CAUSE OF THIS EVENT, HOWEVER, WE CANNOT RULE OUT EXCESSIVE VAULTING SECONDARY TO A LONG LENS AS A CONTRIBUTING FACTOR. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -16.5/+3.0/091 | OS | Physician | 02/03/2015 | 2015 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4599024&pc=MTA | 2023826-2015-00308 | 2014-06-02 | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH A TORN HAPTIC AND THERE WAS A CLEAR SURGICAL RESIDUE ON THE LENS. A LENS WORK ORDER SEARCH REVEALED THERE WAS ONCE SIMILAR COMPLAINT WITHIN THE WORK ORDER. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. MEDICAL REVIEW - REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -16.00/+2.0/088 | OS | Physician | 02/10/2015 | 2015 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4615063&pc=MTA | 2023826-2015-00331 | 2014-06-21 | Injury | STAAR SURGICAL COMPANY | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE, RESULTING IN AN EXPLANT OF THE LENS ON (B)(6) 2014 AND EXCHANGE FOR A LENS WITH A SMALLER VAULT. THE ISSUE WAS RESOLVED WITH THE IMPLANTATION OF THE NEW LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4) - (SECONDARY SURGERY); (NEW INCISION); SIZE INCORRECT FOR PATIENT; (EXPLANTED); (VAULTING). DEVICE EVALUATED BY THE MANUFACTURER? NO. LENS WAS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -10.0/+2.5/089 | OD | Physician | 02/17/2015 | 2015 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5164296&pc=MTA | 2023826-2015-01268 | 2014-08-30 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/X079 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2008. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. PATIENT VISIT ON (B)(6) 2015 WAS 20/16. SEE MFR. #2023826-2015-01274 FOR RIGHT EYE. Manufacturer Narrative: (B)(4). BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) | -13.5/+1.5/079 | OS | Physician | 09/22/2015 | 2015 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5164307&pc=MTA | 2023826-2015-01274 | 2014-08-30 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -14.5/+1.5/X094 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2008. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. PATIENT VISIT ON (B)(6) 2015 WAS 20/13. SEE MFR. #2023826-2015-01268 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4): BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) | -14.5/+1.5/094 | OD | Physician | 09/22/2015 | 2015 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4736021&pc=MTA | 2023826-2015-00480 | 2014-09-12 | Injury | STAAR SURGICAL COMPANY | 2015-04-29 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4) (SECONDARY SURGERY); (NEW INCISION); SIZE INCORRECT FOR PATIENT; (EXPLANTED); (VAULTING). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -9.5/1.0/10 | OS | Physician | 04/02/2015 | 2015 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5082837&pc=MTA | 2023826-2015-01224 | 2014-09-24 | Injury | STAAR SURGICAL COMPANY | 2015-09-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 11.5MM TICM115V4, -12.0/+3.0/X176 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -12.0/+3.0/X176. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -12.0/+3.0/176 | OS | Physician | 08/20/2015 | 2015 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4784694&pc=MTA | 2023826-2015-00591 | 2014-09-27 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 AND LOW VAULT WAS NOTED IN THE POST-OPERATIVE PERIOD. THE LENS WAS EXCHANGED FOR A LONGER LENGTH ON (B)(6) 2014 AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS DRY WITH NO VISIBLE DAMAGE. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND WITHIN THE WORKORDER. MEDICAL REVIEW: ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -1.3/+2/084 | OS | Physician | 02/05/2015 | 2015 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4958310&pc=MTA | 2023826-2015-00979 | 2014-10-01 | Injury | STAAR SURGICAL COMPANY | 2015-07-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4). | - 7.5/5.5/152 | OD | Physician | 07/08/2015 | 2015 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5168232&pc=MTA | 2023826-2015-01271 | 2014-10-01 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -16.0/+4.5/X097 DIOPTER IN THE PATIENT'S RIGHT EYE (OD). LOW VAULT AND LENS ROTATION WAS OBSERVED ON (B)(6) 2014. THE LENS WAS REPOSITIONED ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. PATIENT VISIT ON (B)(6) 2015 - UCVA WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. CONCLUSION: BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -16.0/+4.5/097 | OD | Physician | 09/22/2015 | 2015 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5209471&pc=MTA | 2023826-2015-01435 | 2014-10-04 | Injury | STAAR SURGICAL COMPANY | 2015-11-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -09.00/+06.00/X088 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED AND THE ICL WAS EXPLANTED ON (B)(6) 2014. NO OTHER LENS WAS IMPLANTED. SEE MFR. #2023829-2015-01431 FOR RIGHT EYE. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXCHANGED (DATE UNKNOWN) FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE LENS WAS DEFORMED. HOWEVER, THERE WAS NO OTHER VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: DIMENSIONAL INSPECTION OF THE RETURNED LENS FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -09.00/+06.00/X088. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -09.00/+06.00/088 | OS | Physician | 10/13/2015 | 2015 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5210090&pc=MTA | 2023826-2015-01431 | 2014-10-04 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -8.50/+6.00/X094 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED AND THE ICL WAS EXPLANTED ON (B)(6) 2014. NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4): BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. THE LENS WAS NOT RETURNED. | -8.50/+6.00/094 | OD | Physician | 10/13/2015 | 2015 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4745849&pc=MTA | 2023826-2015-00500 | 2014-10-06 | Injury | STAAR SURGICAL COMPANY | 2015-05-04 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT AGE AND WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OS THE RETURNED LENS SHOWED THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -13.0/+2.0/98 | OD | Physician | 04/04/2015 | 2015
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4806959&pc=MTA | 2023826-2015-00586 | 2014-10-11 | Injury | STAAR SURGICAL COMPANY | 2015-05-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(6): EVALUATION: METHOD - WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). CONCLUSIONS - BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -8.00/+2.5/092 | OD | Physician | 03/03/2015 | 2015 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4600865&pc=MTA | 2023826-2015-00313 | 2014-10-21 | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE ICL WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DATE OF BIRTH - UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -11.00/+3.5/101 | OD | Other | 02/13/2015 | 2015 | 47
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5210650&pc=MTA | 2023826-2015-01430 | 2014-11-10 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, 8.00/+2.50/X103 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LOW VAULT WAS NOTED AND THE ICL WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2014. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT IS NOT MARKETED IN THE U.S. DIOPTER: -8.0/+2.5/X103. (B)(4). LENS NOT RETURNED. | -8.0/+2.5/x103 | OS | Physician | 10/13/2015 | 2015 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5079600&pc=MTA | 2023826-2015-01226 | 2014-11-18 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4, -20.00/+2.50/X097 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. SEE MFR #2023826-2015-01225 FOR LEFT EYE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT NOT MARKETED IN THE U.S. DIOPTER: -20.00/+2.5/X097. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -20.00/+2.50/097 | OD | Physician | 08/20/2015 | 2015 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5088287&pc=MTA | 2023826-2015-01225 | 2014-11-18 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4, -22.00/+3.00/X083 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. POST-OP UNCORRECTED VISUAL ACUITY WAS 20/25. SEE MFR #2023826-2015-01226 FOR RIGHT EYE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT NOT MARKETED IN THE U.S. DIOPTER: -22.00/+3.00/X083. DEVICE EVALUATED BY MANUFACTURER? NO-LENS NOT RETURNED. (B)(4). METHOD CODE: EVALUATION METHOD: LENS WORK ORDER SEARCH RESULTS CODE: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -22.00/+3.00/083 | OS | Physician | 08/20/2015 | 2015 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4578608&pc=MTA | 2023826-2015-00259 | 2014-11-20 | Injury | STAAR SURGICAL COMPANY | 2015-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST BCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW & LENS WORK ORDER SEARCH. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED DRY AND A HAPTIC WAS TORN. LENS WORK ORDER SEARCH REVEALED THAT THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -12.5/42.0/080 | OS | Physician | 02/04/2015 | 2015 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4806336&pc=MTA | 2023826-2015-00581 | 2014-11-26 | Injury | STAAR SURGICAL COMPANY | 2015-05-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION CODES: METHOD - (OTHER): WORK ORDER SEARCH, MEDICAL REVIEW RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS - (OTHER): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -6.0/+1.5/163 | OS | Physician | 02/10/2015 | 2015 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4786483&pc=MTA | 2023826-2015-00564 | 2014-11-26 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-20 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WERE NOTED, THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2014 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH AND MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED NO VISIBLE DAMAGE WAS OBSERVED. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND WITHIN THE WORKORDER. MEDICAL REVIEW: ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP, ECT.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -11/1/015 | OS | Physician | 01/07/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5178938&pc=MTA | 2023826-2015-01329 | 2014-12-05 | Injury | STAAR SURGICAL COMPANY | 2015-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER -13.0/+2.0/X103 DIOPTER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -13.0/+2.0/X103. DEVICE EVALUATED BY MANUFACTURER? LENS NOT RETURNED. (B)(4) NEW INCISION. SECONDARY SURGERY. LENS EXPLANTED. VAULTING. EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -13.0/+2.0/103 | OD | Physician | 03/12/2015 | 2015
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5287996&pc=MTA | 2023826-2015-01608 | 2014-12-08 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/072 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | -13.0/+1.5/072 | OS | Physician | 10/27/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4641288&pc=MTA | 2023826-2015-00398 | 2014-12-13 | Injury | STAAR SURGICAL COMPANY | 2015-03-30 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LOW VAULT WAS NOTED AND THE LENS WAS REPOSITIONED ON (B)(6) 2014. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THIS RESOLVED THE PROBLEM. PATIENT'S LAST KNOW UCVA IS 20/22 AS OF (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT:UNK (B)(4). DEVICE EVALUATED BY THE MANUFACTURER? NO. LENS WAS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED DRY AND A HAPTIC WAS TORN. MEDICAL REVIEW - CLINICAL STUDIES HAVE SHOWN THAT TORIC ICL ROTATION OCCURS IN (B)(4) OF PATIENTS WHERE A TICL HAS BEEN IMPLANTED. SEVERAL FACTORS MAY CONTRIBUTE TO THIS PHENOMENON (I.E. ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION,ECT). THERE IS NOT ENOUGH CLINICAL INFORMATION TO DETERMINE THE EXACT ROOT CAUSE; HOWEVER, IT SEEMS THE ICL WAS SHORT FOR THE EYE. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -11.0/+3.0/091 | OS | Physician | 03/03/2015 | 2015 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5079611&pc=MTA | 2023826-2015-01223 | 2014-12-13 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4, -13.5/+1.5/X081 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT NOT MARKETED IN THE U.S. DIOPTER: -13.5/+1.5/X081. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -13.5/+1.5/081 | OS | Physician | 08/20/2015 | 2015 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4954612&pc=MTA | 2023826-2015-00965 | 2014-12-14 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4). | -19.0/+4.0/095 | OS | Physician | 07/14/2015 | 2015 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4488989&pc=MTA | 2023826-2015-00110 | 2014-12-16 | Injury | STAAR SURGICAL COMPANY | 2015-02-05 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) (NEW INCISION); SIZE INCORRECT FOR PATIENT; (EXPLANTED); (VAULTING). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -17.5/+2.5/076 | OS | Physician | 01/07/2015 | 2015 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4488990&pc=MTA | 2023826-2015-00109 | 2014-12-16 | Injury | STAAR SURGICAL COMPANY | 2015-02-05 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) - NO CODE AVAILABLE (NEW INCISION). SIZE INCORRECT FOR PATIENT; (EXPLANTED); (VAULTING). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -18.0/+3.0/108 | OD | Physician | 01/07/2015 | 2015 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4816577&pc=MTA | 2023826-2015-00637 | 2014-12-22 | Injury | STAAR SURGICAL COMPANY AG | 2015-06-03 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2006. LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON 03/17/2015 AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST KNOWN BCVA WAS 20/16 ON (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGERY. NEW INCISION. EXPLANTED. VAULTING. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED DRY AND THERE WAS NO VISIBLE DAMAGE OBSERVED. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -12.5/+3.0/5 | OS | Other | 05/08/2015 | 2015 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5089551&pc=MTA | 2023826-2015-01235 | 2014-12-29 | Injury | STAAR SURGICAL COMPANY | 2015-09-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/073 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -13.5/+2.0/073 | OD | Physician | 08/20/2015 | 2015 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4529096&pc=MTA | 2023826-2015-00171 | 2015-01-02 | Injury | STAAR SURGICAL COMPANY | 2015-02-19 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -16.5/+2.0/088 | OD | Physician | 01/28/2015 | 2015 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5082645&pc=MTA | 2023826-2015-01227 | 2015-01-12 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4, -13.0/+1.5/X100 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT NOT MARKETED IN THE U.S. DIOPTER: -13.0/+1.5/X100. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -13.0/+1.5/100 | OS | Physician | 08/20/2015 | 2015 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4802612&pc=MTA | 2023826-2015-00616 | 2015-01-16 | Injury | STAAR SURGICAL COMPANY | 2015-05-28 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 ON (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. IMPLANT DATE: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH AND MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP, ACT.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, PRODUCT EVALUATION AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -11/+0.5/97 | OD | Other | 04/28/2015 | 2015 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4786506&pc=MTA | 2023826-2015-00568 | 2015-01-17 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-20 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LOW VAULT WAS NOTED, THE LENS WAS EXPLANTED ON (B)(6) 2015, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT:UNK. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS DRY AND A HAPTIC WAS TORN. THE LENS WAS RE-HYDRATED AND BOTH THE LENGTH AND WIDTH WERE REMEASURE AND DETERMINED TO BE WIHIN ORIGINAL DESIGN SPECIFICATIONS. THEREFORE, IT CAN BE JUSTIFIED THE COMPLAINT WAS NOT PRODUCT RELATED. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND WITHIN THE WORK ORDER. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). STAAR RECOMMENDS THE ICL TO BE LEFT IMPLANTED IF NO CATARACTS ARE NOTED OR IN CASES WHERE THERE IS A TRACE ANTERIOR SUB-CAPSULAR CATARACT WITH NO PROGRESSION. STAAR BELIEVES THAT THE ACTION TO REMOVE AND/OR REPLACE THE ICL INCREASE THE RISK FOR ANTERIOR SUB-CAPSULAR CATARACT. CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -19/12.5/143 | OD | Physician | 02/02/2015 | 2015 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5252018&pc=MTA | 2023826-2015-01552 | 2015-01-21 | Injury | STAAR SURGICAL COMPANY | 2015-11-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/037 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD A LOW VAULT AND A REFRACTIVE SURPRISE. THE SURGEON IS PLANNING TO EXPLANT AND EXCHANGE FOR A LONGER LENS. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THE HAPTIC TO BE TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -8.0/+1.0/037 | OD | Physician | 11/11/2015 | 2015 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5084309&pc=MTA | 2023826-2015-01228 | 2015-01-22 | Injury | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4, -14.0/+2.0/X095 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -14.0/+2.0/X095; DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | -14.0/+2.0/095 | OS | Physician | 08/20/2015 | 2015 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4842136&pc=MTA | 2023826-2015-00677 | 2015-02-05 | Injury | STAAR SURGICAL COMPANY | 2015-06-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. REFRACTIVE SURPRISE AND EXCESSIVE VAULT WAS NOTED. CUSTOMER IS CURRENTLY UNDER CONSULTATION. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). LENS REMAINS IMPLANTED. (B)(4). Manufacturer Narrative: CORRECTED DATA/ADDITIONAL INFORMATION: CORRECTED TO ADVERSE EVENT. STATED EXCESSIVE VAULTING. SUPPLEMENTAL INFORMATION RECEIVED STATED NO EXCESSIVE VAULT OCCURRED. PREVIOUS "OTHER" SECTION: STATED INCORRECT DIOPTER. ACTUAL DIOPTER IS -6/+2.5/10. PREVIOUS EXP DATE IS INCORRECT. PREVIOUS MANUFACTURE DATE INCORRECT. (B)(4). | -6/+2.5/10 | OS | Physician | 05/22/2015 | 2015 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4669884&pc=MTA | 2023826-2015-00418 | 2015-02-13 | Injury | STAAR SURGICAL COMPANY | 2015-04-08 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED, AS WELL AS SIGNIFICANT REDUCTION OF ICA (IRIDO-CORNEAL ANGLES), THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4): EVALUATION: METHOD: WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) | -11/2.5/099 | OD | Physician | 03/12/2015 | 2015 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5525083&pc=MTA | 2023826-2016-00337 | 2015-02-14 | Injury | STAAR SURGICAL COMPANY | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 -16.50/+2.00/X053 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2015. EXCESSIVE VAULTING WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES NOTED. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. ON (B)(6) 2015, PATIENT'S LAST VISIT; UCVA WAS 20/22. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -16.50/+2.00/X053, DEVICE EVALUATED BY MANUFACTURER? NO - DEVICE NOT RETURNED. (B)(4). | -16.50/+2.00/053 | OD | Physician | 02/26/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4669900&pc=MTA | 2023826-2015-00419 | 2015-02-15 | Injury | STAAR SURGICAL COMPANY | 2015-04-08 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED, AS WELL AS SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4) - (SECONDARY SURGERY); (NEW INCISION); (REDUCTION OF IRIDO-CORNEAL ANGLES). SIZE INCORRECT FOR PATIENT; (EXPLANTED); (VAULTING). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED DRY WITH A TORN HAPTIC. MEDICAL REVIEW - REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP, ECT.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -18/2.5/011 | OD | Physician | 03/12/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4803648&pc=MTA | 2023826-2015-00618 | 2015-02-25 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-28 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYES (OD) AND PATIENT EXPERIENCED AN UNEXPECTED REFRACTIVE OUTCOME AND LOW VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A SMALLER SIZE AND THIS RESOLVED THE ISSUES.Event Description: ADDITIONAL INFORMATION RECEIVED - THE SURGEON NOTICED ROTATION AFTER IMPLANTATION. THE LENS WAS EXCHANGED FOR ANOTHER SAME SIZE LENS BUT DIFFERENT DIOPTER. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION METHOD: LENS WORK ORDER SEARCH. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. (B)(4). Manufacturer Narrative: (B)(4). METHOD (PROCESS EVALUATION): MEDICAL REVIEW. RESULTS (EVALUATION RESULT): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH, WIDTH, DIOPTER AND SPHERE WERE MEASURED AND THE RESULTS OF THE MEASUREMENTS WERE COMPARED AGAINST THE ORIGINAL VALUES AND THE LENS WAS FOUND TO BE IN SPECIFICATION. PER MEDICAL REVIEW - THE MOST LIKELY CAUSE OF THE RESIDUAL REFRACTION IS A COMBINATION OF SELECTING A LENS POWER THAT WAS OVER-CORRECTING THE REFRACTIVE ERROR AND THE OFF-AXIS POSITION OF THE TORIC LENS SECONDARY TO ROTATION OR MISALIGNMENT. WE CANNOT COMPLETELY RULE OUT LOW VAULT SECONDARY TO A SHORT LENS AS CONTRIBUTING FACTOR. CLINICAL STUDIES HAVE SHOWN THAT TORIC ICL ROTATION OCCURS IN 0.5% OF PATIENTS WHERE A TICL HAS BEEN IMPLANTED. SEVERAL FACTORS MAY CONTRIBUTE TO THIS PHENOMENON (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). SEVERAL FACTORS MAY BE INVOLVED IN A POST-OPERATIVE REFRACTIVE SURPRISE, INCLUDING INAPPROPRIATE REFRACTIVE SURGICAL PLANNING, A WRONG LENS CALCULATION/SELECTION, PREOPERATIVE INACCURATE DETERMINATION OF THE EYE'S MEASUREMENTS TO DETERMINE THE POWER OF THE ICL, INDUCED REFRACTIVE CHANGE BY INCISIONS, UNSTABLE CORNEA AND/OR REFRACTION PRIOR TO IMPLANTATION, CL-INDUCED WARPAGE, UPSIDE-DOWN OR TILTED ICL, ETC. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSION CODE 17: (EVALUATION CONCLUSION): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. CLAIM # 707642 | -12.5/1.5/89 | OD | Other | 04/30/2015 | 2015 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5124743&pc=MTA | 2023826-2015-01293 | 2015-02-27 | Injury | STAAR SURGICAL COMPANY | 2015-10-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.5/+2.5/X087 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | -13.5/+2.5/087 | OD | Physician | 09/01/2015 | 2015 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4724365&pc=MTA | 2023826-2015-00453 | 2015-03-05 | Injury | STAAR SURGICAL COMPANY | 2015-04-23 | MTA | VISIAN ICL IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4) SECONDARY SURGERY. NEW INCISION. EXPLANTED. VAULTING. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: MEDICAL REVIEW - REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). IF ROTATION, SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ECT.). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -12.5/2.5/78 | OS | Physician | 03/25/2015 | 2015 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5167030&pc=MTA | 2023826-2015-01373 | 2015-03-05 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.00/+3.0/X173 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | -8.00/+3.0/173 | OS | Physician | 09/23/2015 | 2015 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4724363&pc=MTA | 2023826-2015-00454 | 2015-03-10 | Injury | STAAR SURGICAL COMPANY | 2015-04-23 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED AND THE LENS WAS EXPLANTED ON (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). EVALUATION CODES: METHOD: (OTHER) WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED IN LIQUID WITH NO VISIBLE DAMAGE. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. MEDICAL REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. (B)(4). | -4.5/1.5/80 | OS | Physician | 03/25/2015 | 2015 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5171642&pc=MTA | 2023826-2015-01382 | 2015-03-12 | Injury | STAAR SURGICAL COMPANY | 2015-10-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -6.0/+3.5/X111 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -6.0/+3.5/111 | OS | Physician | 09/23/2015 | 2015 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4724348&pc=MTA | 2023826-2015-00452 | 2015-03-13 | Injury | STAAR SURGICAL COMPANY | 2015-04-23 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS). LOW VAULT WAS NOTED AND THE LENS WAS EXPLANTED ON (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). EVALUATION, METHOD: (OTHER) WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. MEDICAL REVIEW - REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE).OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. (B)(4). | -11/1.5/83 | OS | Physician | 03/25/2015 | 2015 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4806337&pc=MTA | 2023826-2015-00589 | 2015-03-20 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4) NEW INCISION. SECONDARY SURGERY. EXPLANTED. VAULTING. METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). CONCLUSIONS - 68 (OTHER): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -9.5/+1.0/087 | OS | Physician | 03/23/2015 | 2015 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4743957&pc=MTA | 2023826-2015-00504 | 2015-03-21 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: (B)(4). | -12.0/+1.5/100 | OD | Other | 04/02/2015 | 2015 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4770082&pc=MTA | 2023826-2015-00541 | 2015-03-21 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS REMOVED AND NOT REPLACED. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OS THE RETURNED LENS SHOWED THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. THEREFORE, IT CAN BE JUSTIFIED THAT THE COMPLAINT WAS NOT PRODUCT RELATED. MEDICAL REVIEW - THERE IS NOT ENOUGH CLINICAL DATA TO DETERMINE A EXACT CAUSE OF THE EVENT, HOWEVER, WE CANNOT RULE OUT LOW VAULTING SECONDARY TO A SHORT LENS AS A CONTRIBUTING FACTOR. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -12/+1.5/100 | OD | Physician | 04/15/2015 | 2015 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5043434&pc=MTA | 2023826-2015-01087 | 2015-03-24 | Injury | STAAR SURGICAL COMPANY | 2015-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.00/+1.50/101 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND PATIENT HAD A SHALLOW CHAMBER. NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S BUT NOT MARKETED IN THE U.S. DIOPTER: -14.00/+1.50/101. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS WAS NOT RETURNED. | -14.00/+1.50/101 | OD | Physician | 08/05/2015 | 2015 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4791922&pc=MTA | 2023826-2015-00603 | 2015-03-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-22 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND EXCESSIVE VAULT WAS NOTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(6). (B)(4). (OTHER): LENS REMAINS IMPLANTED. (B)(4). LENS REMAINS IMPLANTED. Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH AND MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP, ACT.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. THE LENS WAS NOT RETURNED. (B)(4). | -4/+3.5/82 | OD | Other | 04/27/2015 | 2015 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5043693&pc=MTA | 2023826-2015-01088 | 2015-03-25 | Injury | STAAR SURGICAL COMPANY | 2015-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.00/+1.00/97 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND PATIENT HAD A SHALLOW CHAMBER. NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S BUT NOT MARKETED IN THE U.S. DIOPTER: -14.00/+1.00/97. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS WAS NOT RETURNED. | -14.00/+1.00/97 | OS | Physician | 08/05/2015 | 2015 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5565337&pc=MTA | 2023826-2016-00396 | 2015-03-25 | Injury | STAAR SURGICAL COMPANY | 2016-04-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 -18.00/+1.00/X089 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED THE NEXT DAY. THE LENS WAS EXPLANTED AND WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. PATIENT'S VISIT ON (B)(6) 2015, UCVA WAS 20/30. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND LENS HAPTIC TORN/BROKEN/BENT/DEFORMED, AND FOREIGN MATERIAL ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -18.00/+1.00/X089 DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). METHOD CODE(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODE(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -18.00/+1.00/089 | OD | Physician | 03/11/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4743827&pc=MTA | 2023826-2015-00498 | 2015-03-27 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-01 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO ELEVATED IOP ASSOCIATED WITH EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. SEE MFR REPORT# 2023826-2015-00498 FOR THE OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4) NEW INCISION. LENS EXPLANTED. VAULTING. (B)(4). Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH & MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE OPTIC WAS TORN AND THE LENS WAS RETURNED DRY. THE WORK ORDER SEARCH SHOWED NO OTHER COMPLAINT RECORDED RELATED TO THE SAME WORK ORDER. MEDICAL REVIEW - REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -6.0/3.0/004 | OD | Physician | 04/02/2015 | 2015 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5579343&pc=MTA | 2023826-2016-00478 | 2015-04-06 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.00/+2.50/086 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT'S UNCORRECTED VISUAL ACUITY IS 20/25. DEVICE EVALUATION - THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. PERFORMED VISUAL INSPECTION AND FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.00/+2.50/086 | OD | Physician | 03/15/2016 | 2016 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5054920&pc=MTA | 2023826-2015-01161 | 2015-04-08 | Injury | STAAR SURGICAL COMPANY | 2015-09-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18/1.5/105 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES ASSOCIATED WITH EXCESSIVE VAULTING WAS OBSERVED. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE EYE IS QUIET AND DEEP AC NOW. PATIENT IS DOING WELL AND EXTREMELY HAPPY NOW. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). | -18/1.5/105 | OD | Physician | 08/12/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4933038&pc=MTA | 2023826-2015-00911 | 2015-04-20 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) AND LOW VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. SEE MFR REPORT# 2023826-2015-00910 FOR THE OTHER EYE. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. ANTICIPATED BUT NOT YET BEGUN. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -13.0/3.0/83 | OS | Other | 06/25/2015 | 2015 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5405799&pc=MTA | 2023826-2016-00046 | 2015-04-23 | Injury | STAAR SURGICAL COMPANY | 2016-02-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -14.50/+4.00/X080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULTING WAS NOTED AND THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2015, FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. SEE MFR.#2023826-2016-00045 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER:-14.50/+4.00/X080 (B)(4). METHOD CODE(S): (EVALUATION METHOD): WORK ORDER SEARCH. RESULTS CODE(S): (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -14.50/+4.00/080 | OD | Physician | 12/11/2015 | 2015 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5535802&pc=MTA | 2023826-2016-00319 | 2015-04-25 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/+1.50/X088 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2015. EXCESSIVE VAULT AND SHALLOW CHAMBER WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. UCVA IS 20/24. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -10.00/+1.50/X088 (B)(4). | -10.00/+1.50/088 | OS | Physician | 02/29/2016 | 2016 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4930378&pc=MTA | 2023826-2015-00910 | 2015-04-27 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (ODS) ON (B)(6) 2013 AND LOW VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST BCVA WAS 20/16. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO, ANTICIPATED BUT NOT YET BEGUN. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: PATIENT'S LAST POST-OP BEST-CORRECTED AND UNCORRECTED VISUAL ACUITIES WERE 20/16. (B)(4). | -13.0/3.0/87 | OD | Other | 06/25/2015 | 2015 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4919456&pc=MTA | 2023826-2015-00837 | 2015-04-28 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-16 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 AND EXCESSIVE VAULT WAS OBSERVED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2015 AND THIS RESOLVED THE PROBLEM. PATIENT'S LAST BCVA WAS 20/20. SEE MFR REPORT# 2023826-2015-00836 FOR THE OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. BASED ON THE COMPLAINT INFORMATION AND LENS WORK ORDER SEARCH, WE WERE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4). | -8/+3.5/89 | OS | Physician | 06/16/2015 | 2015 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5084418&pc=MTA | 2023826-2015-01134 | 2015-05-02 | Injury | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -8.0/+1.5/X086 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. LAST DOCTOR VISIT ON (B)(6) 2015, UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -8.0/+1.5/X086. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THIS EVENT COULD NOT BE DETERMINED. (B)(4). | -8.0/+1.5/086 | OS | Physician | 08/20/2015 | 2015 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5047744&pc=MTA | 2023826-2015-01112 | 2015-05-04 | Injury | STAAR SURGICAL COMPANY | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -10.50/+1.0/080 | OD | Physician | 08/04/2015 | 2015 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5525075&pc=MTA | 2023826-2016-00330 | 2015-05-05 | Injury | STAAR SURGICAL COMPANY | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+0.5/097 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | -12.5/+0.5/097 | OD | Physician | 02/25/2016 | 2016 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5525077&pc=MTA | 2023826-2016-00331 | 2015-05-05 | Injury | STAAR SURGICAL COMPANY | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+3.0/070DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | -15.5/+3.0/070 | OS | Physician | 02/25/2016 | 2016 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5408506&pc=MTA | 2023826-2016-00096 | 2015-05-10 | Injury | STAAR SURGICAL COMPANY | 2016-02-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+0.50/X109 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULTING WAS NOTED AND THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2015 FOR A SHORTER LENS WITH A SIMILAR DIOPTER. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DIOPTER:-12.00/+0.50/X109. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -12.00/+0.50/109 | OD | Physician | 12/11/2015 | 2015 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4949679&pc=MTA | 2023826-2015-00895 | 2015-05-12 | Injury | STAAR SURGICAL COMPANY | 2015-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -11.5/+3.0/075 | OD | Physician | 07/01/2015 | 2015 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218059&pc=MTA | 2023826-2015-01496 | 2015-05-12 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+0.5/174 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | -17.0/+0.5/174 | OD | Physician | 10/13/2015 | 2015 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5154937&pc=MTA | 2023826-2015-01355 | 2015-05-20 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+1.0/X081 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | -15.0/+1.0/081 | OS | Physician | 09/21/2015 | 2015 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5408918&pc=MTA | 2023826-2016-00123 | 2015-05-23 | Injury | STAAR SURGICAL COMPANY | 2016-02-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2M VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.00/+2.00/005 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -6.00/+2.00/005 | OS | Physician | 01/05/2016 | 2016
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5386394&pc=MTA | 2023826-2016-00056 | 2015-05-26 | Injury | STAAR SURGICAL COMPANY | 2016-01-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS, -16.5/+3.0/085 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.5/+3.0/085 | OD | Physician | 12/11/2015 | 2015 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4975056&pc=MTA | 2023826-2015-00972 | 2015-05-27 | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING AND THE LENS DISLOCATED/SUBLUXED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -16.0/+1.50/102 | OD | Physician | 07/14/2015 | 2015 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4975057&pc=MTA | 2023826-2015-00973 | 2015-05-27 | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING AND THE LENS DISLOCATED/SUBLUXED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -12.0/+1.00/094 | OS | Physician | 07/14/2015 | 2015 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5412037&pc=MTA | 2023826-2016-00044 | 2015-05-28 | Injury | STAAR SURGICAL COMPANY | 2016-02-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS, -16.50/+2.00/X085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULTING WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2015 FOR A LONGER LENGTH LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DIOPTER:-16.50/+2.00/X085. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -16.50/+2.00/085 | OS | Physician | 12/11/2015 | 2015 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4975030&pc=MTA | 2023826-2015-00970 | 2015-06-02 | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING EXTREMELY WELL AND HAPPY. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | -13.0/+0.5/070 | OS | Physician | 07/14/2015 | 2015 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5224357&pc=MTA | 2023826-2015-01458 | 2015-06-04 | Injury | STAAR SURGICAL COMPANY | 2015-11-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/+1.50/X086 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WITH LENS ROTATION WAS NOTED. REPOSITIONING OF THE LENS WAS PERFORMED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED FOR LONGER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, BCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -17.50/+1.50/X086. (B)(4). METHOD CODE(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODE(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -17.50/+1.50/086 | OS | Physician | 10/13/2015 | 2015 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5400115&pc=MTA | 2023826-2016-00075 | 2015-06-10 | Injury | STAAR SURGICAL COMPANY | 2016-01-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/081 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD(PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -13.0/+1.5/081 | OS | Physician | 12/11/2015 | 2015 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119169&pc=MTA | 2023826-2015-01302 | 2015-06-15 | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/082 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -10.0/+1.0/082 | OD | Physician | 09/01/2015 | 2015 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5178917&pc=MTA | 2023826-2015-01323 | 2015-06-18 | Injury | STAAR SURGICAL COMPANY | 2015-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER -9.0/+1.0/X018 DIOPTER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SMALLER LENGTH AND SIMILAR DIOPTER LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -9.0/+1.0/X018. (B)(4). METHOD CODES: EVALUATION METHOD: LENS WORK ORDER SEARCH RESULTS CODES: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -9.0/+1.0/018 | OD | Physician | 09/29/2015 | 2015 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5642732&pc=MTA | 2023826-2016-00650 | 2015-06-24 | Injury | STAAR SURGICAL COMPANY | 2016-05-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4, -15.50/+3.0/X076 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SMALLER LENGTH LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. SEE MFR. #2023826-2016-00649 FOR RIGHT EYE. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. CONCLUSION: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL BELOW THE AGE OF 21 YEARS IS CONTRAINDICATED. (B)(4). NO NEW INCISION WAS MADE. THIS CODE IS NOT APPLICABLE UNDER DEVICE CODE; IT SHOULD BE LISTED UNDER PATIENT CODE INSTEAD. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.50/+3.0/076 | OS | Physician | 04/11/2016 | 2016 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5642727&pc=MTA | 2023826-2016-00649 | 2015-06-24 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4, -16.00/+4.00/X089 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SMALLER LENGTH LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. SEE MFR. #2023826-2016-00650 FOR LEFT EYE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THE HAPTIC TO BE TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL ON THE LENS SURFACE (NOT POSSIBLE TO SPECIFY). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). | -16.00/+4.00/089 | OD | Physician | 04/11/2016 | 2016 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4933376&pc=MTA | 2023826-2015-00901 | 2015-07-02 | Injury | STAAR SURGICAL COMPANY | 2015-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2006. THE LENS HAD AN INADEQUATE VAULT AND REFRACTIVE CHANGE OVER TIME. THE SURGEON IS PLANNING TO EXPLANT AND EXCHANGE THE LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4). | -11.0/+2.0/083 | OS | Physician | 07/06/2015 | 2015 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4933377&pc=MTA | 2023826-2015-00900 | 2015-07-02 | Injury | STAAR SURGICAL COMPANY | 2015-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON 09/18/2006. THE LENS HAD AN INADEQUATE VAULT AND REFRACTIVE CHANGE OVER TIME. THE SURGEON IS PLANNING TO EXPLANT AND EXCHANGE THE LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). LENS REMAINS IMPLANTED. (B)(4). | -12.0/+2.0/084 | OD | Physician | 07/06/2015 | 2015 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5371920&pc=MTA | 2023826-2016-00063 | 2015-07-06 | Injury | STAAR SURGICAL COMPANY | 2016-01-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+4.0/100 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). LENS NOT RETURNED. | -9.5/+4.0/100 | OD | Physician | 12/19/2015 | 2015 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5171706&pc=MTA | 2023826-2015-01287 | 2015-07-07 | Injury | STAAR SURGICAL COMPANY | 2015-10-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER -10.5/+1.0/X109 DIOPTER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULTING WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND WAS EXCHANGED FOR A SMALLER LENS. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015 - UCVA WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS . (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -10.5/+1.0/X109. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -10.5/+1.0/109 | OS | Physician | 09/23/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5525076&pc=MTA | 2023826-2016-00332 | 2015-07-08 | Injury | STAAR SURGICAL COMPANY | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/094 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: CORRECTED INFORMATION: (B)(4). ADDITIONAL INFORMATION: CONCLUSION CODE: AS PER THE DFU OF THIS LENS MODEL, IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0MM IS CONTRADICTED. THIS PATIENT HAD AN ACD OF 2.92MM AT THE TIME OF IMPLANTATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | -12.5/+1.5/094 | OD | Physician | 02/27/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5325786&pc=MTA | 2023826-2015-01715 | 2015-07-13 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/X113 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -11.0/+2.0/113 | OD | Physician | 11/16/2015 | 2015 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5325934&pc=MTA | 2023826-2015-01716 | 2015-07-13 | Injury | STAAR SURGICAL COMPANY | 2015-12-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+0.5/X083 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) | -11.0/+0.5/083 | OS | Physician | 11/16/2015 | 2015 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5099291&pc=MTA | 2023826-2015-01173 | 2015-07-14 | Injury | STAAR SURGICAL COMPANY | 2015-09-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/+2.0/X040 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ON (B)(6) 2015, EXCESSIVE VAULTING WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED. ON (B)(6) 2015. THE LENS WAS EXPLANTED AND WAS EXCHANGED FOR A SMALLER LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -12.50/+2.0/X040. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -12.50/+2.0/040 | OS | Physician | 08/25/2015 | 2015 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5792720&pc=MTA | 2023826-2016-00889 | 2015-07-18 | Injury | STAAR SURGICAL COMPANY | 2016-07-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -12.5/1.00/094 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULT. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL (BLOOD) WAS PRESENT ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DOB-UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | -12.5/1.00/094 | OD | Physician | 06/11/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5362051&pc=MTA | 2023826-2016-00006 | 2015-07-22 | Injury | STAAR SURGICAL COMPANY | 2016-01-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -18.00/+2.50/X079 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS. THIS RESOLVED THE PROBLEM. SEE MFR. #2023826-2016-00007 FOR RIGHT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -18.00/+2.50/X079. DEVICE EVALUATED BY MANUFACTURER? NO. DEVICE NOT RETURNED. EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -18.00/+2.50/079 | OS | Physician | 12/14/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5209473&pc=MTA | 2023826-2015-01436 | 2015-07-23 | Injury | STAAR SURGICAL COMPANY | 2015-11-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -15.00/+2.00/X142 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS NOTED AND THE ICL WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/22. SEE MFR. #2023829-2015-01432 FOR RIGHT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(6). DIOPTER: -15.00/+2.00/X142. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -15.00/+2.00/142 | OS | Physician | 10/13/2015 | 2015 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5167014&pc=MTA | 2023826-2015-01372 | 2015-07-24 | Injury | STAAR SURGICAL COMPANY | 2015-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/080 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -10.0/+2.0/080 | OS | Physician | 09/23/2015 | 2015 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5210640&pc=MTA | 2023826-2015-01432 | 2015-07-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -14.50/+3.00/X005 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT WAS NOTED AND THE ICL WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/22. SEE MFR. #2023829-2015-01436 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT IS NOT MARKETED IN THE U.S. (B)(6). DIOPTER: -14.50/+3.00/X005. (B)(4). CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -14.50/+3.00/005 | OD | Physician | 10/13/2015 | 2015 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218099&pc=MTA | 2023826-2015-01499 | 2015-07-28 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5/+1.5/083 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -7.5/+1.5/083 | OS | Physician | 10/13/2015 | 2015 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218094&pc=MTA | 2023826-2015-01497 | 2015-07-29 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/114 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -12.5/+1.5/114 | OS | Physician | 10/13/2015 | 2015 | 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218820&pc=MTA | 2023826-2015-01501 | 2015-07-29 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/094 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -12.0/+1.5/094 | OD | Physician | 10/13/2015 | 2015 | 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218971&pc=MTA | 2023826-2015-01492 | 2015-07-29 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -11.5/+1.0/085 | OS | Physician | 10/13/2015 | 2015 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5362048&pc=MTA | 2023826-2016-00007 | 2015-07-29 | Injury | STAAR SURGICAL COMPANY AG | 2016-01-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -18.00/+4.00/X095 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS. THIS RESOLVED THE PROBLEM. SEE MFR. #2023826-2016-00006 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. DIOPTER: -18.00/+4.00/X095 DEVICE EVALUATED BY MANUFACTURER? NO. DEVICE NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD CODE(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODE(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). DEVICE NOT RETURNED. | -18.00/+4.00/095 | OD | Physician | 12/14/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218992&pc=MTA | 2023826-2015-01493 | 2015-07-30 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/093 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -11.0/+1.5/093 | OD | Physician | 10/13/2015 | 2015 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5638606&pc=MTA | 2023826-2016-00591 | 2015-07-30 | Injury | STAAR SURGICAL COMPANY | 2016-05-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+3.5/127 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED. PATIENT HAD KERATOCONUS AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH KERATOCONUS. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/+3.5/127 | OD | Physician | 04/08/2016 | 2016
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5151468&pc=MTA | 2023826-2015-01242 | 2015-07-31 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/X093 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS REMOVED ON THE SAME DAY DUE TO EXCESSIVE VAULTING AND THE PATIENT HAD SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS WITH A SIMILAR DIOPTER. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015 - UCVA: 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -14.00/+2.0/X093. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4) SHALLOW CHAMBER. VAULTING. EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -14.00/+2.0/093 | OD | Physician | 09/15/2015 | 2015 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5154908&pc=MTA | 2023826-2015-01344 | 2015-08-01 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -14.5/+5.5/092 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | -14.5/+5.5/092 | OD | Physician | 09/14/2015 | 2015 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5243470&pc=MTA | 2023826-2015-01533 | 2015-08-08 | Injury | STAAR SURGICAL COMPANY | 2015-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -07.50/1.0/X101 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED (B)(6) 2015. THE ICL WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR SHORTER ICL WITH A SIMILAR DIOPTER SIZE. THE RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/16.6. SEE MFR. #2023826-2015-01532 FOR RIGHT EYE. Manufacturer Narrative: (B)(4). THE REPORTER INDICATE THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.50/1.0/X101 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER SAME MODEL LENS WITH SIMILAR DIOPTER AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST VISIT WAS ON (B)(6) 2015 WITH POST-OP UNCORRECTED VISUAL ACUITY (UCVA) 20/16.6. SEE MFR. #2023826-2015-01532 FOR RIGHT EYE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -07.50/1.0/X101. (B)(4). EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -07.50/1.0/101 | OS | Physician | 10/28/2015 | 2015 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5243666&pc=MTA | 2023826-2015-01532 | 2015-08-08 | Injury | STAAR SURGICAL COMPANY | 2015-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -07.50/1.50/X103 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED (B)(6) 2015. THE ICL WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR SHORTER ICL WITH A SIMILAR DIOPTER SIZE. THE RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/16.6. SEE MFR. #2023826-2015-01533 FOR LEFT EYE. Manufacturer Narrative: (B)(4). THE REPORTER INDICATE THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.50/1.50/X103 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER SAME MODEL LENS WITH SIMILAR DIOPTER AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST VISIT WAS ON (B)(6) 2015 WITH POST-OP UNCORRECTED VISUAL ACUITY (UCVA) 20/16.6. SEE MFR. #2023826-2015-01533 FOR LEFT EYE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -07.50/1.5/X103. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -07.50/1.50/103 | OD | Physician | 10/28/2015 | 2015 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5146312&pc=MTA | 2023826-2015-01240 | 2015-08-12 | Injury | STAAR SURGICAL COMPANY | 2015-10-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+2.0/X068 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 015 DUE TO EXCESSIVE VAULTING AND THE PATIENT HAD A SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS WITH A SIMILAR DIOPTER. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015 - UCVA: 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -11.00/+2.0/X068. (B)(4). A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -11.00/+2.0/068 | OD | Physician | 09/15/2015 | 2015 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5409015&pc=MTA | 2023826-2016-00093 | 2015-08-15 | Injury | STAAR SURGICAL COMPANY | 2016-02-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.50/+2.00/X090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULTING WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2015 FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DIOPTER:-8.50/+2.00/X090. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -8.50/+2.00/090 | OS | Physician | 12/11/2015 | 2015 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5219011&pc=MTA | 2023826-2015-01503 | 2015-08-17 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/095 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/16. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE US. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -12.5/+1.5/095 | OD | Physician | 10/13/2015 | 2015 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5219058&pc=MTA | 2023826-2015-01504 | 2015-08-17 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.5/+2.0/107 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4) METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -16.5/+2.0/107 | OD | Physician | 10/13/2015 | 2015 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5247885&pc=MTA | 2023826-2015-01556 | 2015-08-17 | Injury | STAAR SURGICAL COMPANY | 2015-11-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+2.0/X011 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -15.5/+2.0/011 | OD | Physician | 11/05/2015 | 2015 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218087&pc=MTA | 2023826-2015-01494 | 2015-08-18 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/089 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) . THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -15.5/+1.5/089 | OD | Physician | 10/13/2015 | 2015 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5325833&pc=MTA | 2023826-2015-01710 | 2015-08-25 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.0/+1.5/X090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -14.0/+1.5/090 | OS | Physician | 11/16/2015 | 2015 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5213886&pc=MTA | 2023826-2015-01423 | 2015-08-27 | Injury | STAAR SURGICAL COMPANY | 2015-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.00/+1.50/X085 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS NOTED ON (B)(6) 2015. THE ICL WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/20. SEE MFR. #2023826-2015-01422 FOR RIGHT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -13.00/+1.50/X085. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | -13.00/+1.50/085 | OS | Physician | 10/13/2015 | 2015 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5869089&pc=MTA | 2023826-2016-01041 | 2015-09-01 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/127 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DRIED DISCOLORED MATERIAL ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -10.5/+2.0/127 | OD | Physician | 08/03/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5308710&pc=MTA | 2023826-2015-01663 | 2015-09-03 | Injury | STAAR SURGICAL COMPANY | 2015-12-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0/+1.5/112 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -6.0/+1.5/112 | OS | Physician | 11/20/2015 | 2015 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5325832&pc=MTA | 2023826-2015-01708 | 2015-09-10 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -17.0/+2.5/X094 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -17.0/+2.5/094 | OS | Physician | 11/16/2015 | 2015 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5212793&pc=MTA | 2023826-2015-01459 | 2015-09-21 | Injury | STAAR SURGICAL COMPANY | 2015-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -1.00/+4.00/X086 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS NOTED ON (B)(6) 2015. THE ICL WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, BCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -1.00/+4.00/X086. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -1.00/+4.00/086 | OS | Physician | 10/13/2015 | 2015 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4578662&pc=MTA | 2023826-2015-00260 | 2015-09-22 | Injury | STAAR SURGICAL COMPANY | 2015-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND THE LENS WAS EXPLANTED THE NEXT DAY DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST UCVA WAS 20/60. Manufacturer Narrative: (B)(4). METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. MEDICAL REVIEW -ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -10.5/1.5/90 | OD | Physician | 02/04/2015 | 2015 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5325780&pc=MTA | 2023826-2015-01711 | 2015-09-23 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/X075 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -9.5/+1.5/075 | OD | Physician | 11/16/2015 | 2015 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5287698&pc=MTA | 2023826-2015-01635 | 2015-09-24 | Injury | STAAR SURGICAL COMPANY | 2015-12-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/X080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -13.5/+1.5/080 | OD | Physician | 11/10/2015 | 2015 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5325785&pc=MTA | 2023826-2015-01713 | 2015-09-24 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+0.5/X087 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/16. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -13.5/+0.5/087 | OD | Physician | 11/16/2015 | 2015 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5247874&pc=MTA | 2023826-2015-01557 | 2015-09-30 | Injury | STAAR SURGICAL COMPANY | 2015-11-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.5/+2.5/X119 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | -5.5/+2.5/119 | OD | Physician | 11/05/2015 | 2015 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5193763&pc=MTA | 2023826-2015-01445 | 2015-10-01 | Injury | STAAR SURGICAL COMPANY | 2015-11-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+2.0/092 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND A REFRACTIVE SURPRISE. THE SURGEON IS PLANNING ON EXPLANTING THE LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -8.5/+2.0/092 | OS | Physician | 10/02/2015 | 2015 | 49
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5312164&pc=MTA | 2023826-2015-01676 | 2015-10-08 | Injury | STAAR SURGICAL COMPANY | 2015-12-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+2.5/088 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -9.0/+2.5/088 | OD | Physician | 11/26/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5572299&pc=MTA | 2023826-2016-00436 | 2015-10-10 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -20.0/+2.0/088 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC TORN. CORRECTED DATA: PATIENT CODE: (NEW INCISION) IS INCORRECT. NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | -20.0/+2.0/088 | OD | Physician | 02/18/2016 | 2016 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5649010&pc=MTA | 2023826-2016-00569 | 2015-10-10 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -18.5/+2.0/X087 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CORRECTED DATA: (B)(4). NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | -18.5/+2.0/087 | OS | Physician | 02/18/2016 | 2016 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5242709&pc=MTA | 2023826-2015-01528 | 2015-10-15 | Injury | STAAR SURGICAL COMPANY | 2015-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/-1.50/X090 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT AND LENS ROTATION WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS WITH A SIMILAR DIOPTER. PROBLEM WAS RESOLVES. POST VISIT ON (B)(6) 2015. UCVA WAS 20/20. Manufacturer Narrative: CORRECTED DATA: PREVIOUS MODEL VICMO13.2 IS INCORRECT, CORRECT DATA IS VTICMO13.2. (B)(4). ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC TORN/SPLIT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -10.50/+1.50/X090 DEVICE EVALUATED BY MANUFACTURER? NO, DEVICE NOT RETURNED. (B)(4). CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). DEVICE NOT RETURNED. | -10.50/-1.50/090 | OD | Physician | 10/28/2015 | 2015 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5252473&pc=MTA | 2023826-2015-01554 | 2015-10-21 | Injury | STAAR SURGICAL COMPANY | 2015-11-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -3.0/+4.5/019 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -3.0/+4.5/019 | OS | Physician | 11/11/2015 | 2015 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5247889&pc=MTA | 2023826-2015-01555 | 2015-10-22 | Injury | STAAR SURGICAL COMPANY | 2015-11-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/X082 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS NOT RETURNED. | -11.0/+1.0/082 | OD | Physician | 11/04/2015 | 2015 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5349511&pc=MTA | 2023826-2016-00003 | 2015-10-27 | Injury | STAAR SURGICAL COMPANY | 2016-01-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 -7.50/+2.00/X107 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. EXCESSIVE VAULTING WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -7.50/+2.00/X107. DEVICE EVALUATED BY MANUFACTURER? NO - (OTHER): DEVICE NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). DEVICE NOT RETURNED. | -7.50/+2.00/107 | OS | Physician | 12/10/2015 | 2015 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5247870&pc=MTA | 2023826-2015-01545 | 2015-10-30 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -2.5/+2.5/X011 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD A LOW VAULT AND REFRACTIVE SURPRISE. THE SURGEON IS PLANNING TO EXPLANT AND EXCHANGE FOR A LONGER LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. (B)(4). | -2.5/+2.5/011 | OD | Physician | 11/02/2015 | 2015 | 54
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5408763&pc=MTA | 2023826-2016-00090 | 2015-11-04 | Injury | STAAR SURGICAL COMPANY | 2016-02-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+2.00/X094 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULTING WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2015 FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. PATIENT'S LAST OFFICE VISIT ON (B)(6) 2015, UCVA WAS 20/20. Manufacturer Narrative: DIOPTER:-12.00/+2.50/X094. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -12.00/+2.00/094 | OS | Physician | 12/22/2015 | 2015 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5296541&pc=MTA | 2023826-2015-01631 | 2015-11-09 | Injury | STAAR SURGICAL COMPANY | 2015-12-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -12.50/+3.0/X095 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED AND THE LENS WAS EXPLANTED ON (B)(6) 2015. THE EXCHANGED LENS, A SHORTER LENS WAS IMPLANTED ON (B)(6) 2015. Manufacturer Narrative: AGE AT TIME OF EVENT IS INCORRECT. DATE OF BIRTH PROVIDED. (B)(4). THE PROBLEM WAS RESOLVED. DEVICE EVALUATION : PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THERE WAS CLEAR, SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. HOWEVER, THERE WAS NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -12.50/+3.00/X095. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -12.50/+3.0/095 | OS | Physician | 11/17/2015 | 2015 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5505912&pc=MTA | 2023826-2016-00275 | 2015-11-09 | Injury | STAAR SURGICAL COMPANY | 2016-03-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/071 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -10.0/+1.0/071 | OD | Physician | 01/25/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5349510&pc=MTA | 2023826-2016-00002 | 2015-11-10 | Injury | STAAR SURGICAL COMPANY | 2016-01-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -8.50/+2.0/X094 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. EXCESSIVE VAULTING WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. PATIENT IS DOING WELL AND HAPPY. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -8.50/+2.00/X094, DEVICE EVALUATED BY MANUFACTURER? NO, DEVICE NOT RETURNED. (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -8.50/+2.0/094 | OS | Physician | 12/10/2015 | 2015 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5649011&pc=MTA | 2023826-2016-00568 | 2015-11-11 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -19.5/+2.5/X088 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4) IS INCORRECT. NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | -19.5/+2.5/088 | OD | Physician | 02/18/2016 | 2016 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5506283&pc=MTA | 2023826-2016-00274 | 2015-11-12 | Injury | STAAR SURGICAL COMPANY | 2016-03-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/077 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN THE CASE/VIAL; VISUAL INSPECTION FOUND UNSPECIFIED FOREIGN MATERIAL ON SURFACE BUT NO VISIBLE DAMAGE (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -12.0/+1.5/077 | OS | Physician | 01/22/2016 | 2016
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5349526&pc=MTA | 2023826-2016-00004 | 2015-11-19 | Injury | STAAR SURGICAL COMPANY | 2016-01-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 -13.00/+2.50X094 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. EXCESSIVE VAULTING WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -13.0/+2.50/X094. DEVICE EVALUATED BY MANUFACTURER: DEVICE NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -13.00/+2.50094 | OD | Physician | 12/10/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5816259&pc=MTA | 2023826-2016-00969 | 2015-11-26 | Injury | STAAR SURGICAL COMPANY AG | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.0/+3.5/175 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/25. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED INFORMATION: (B)(4) SHOULD BE INCLUDED. (B)(4) CODE NOT APPLICABLE, SHOULD BE UNDER PATIENT CODE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE (B)(4) BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -6.0/+3.5/175 | OS | Physician | 05/12/2016 | 2016 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5676333&pc=MTA | 2023826-2016-00712 | 2015-11-27 | Injury | STAAR SURGICAL COMPANY | 2016-05-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.5/+1.0/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: LENS NOT RETURNED. (B)(4). | -16.5/+1.0/085 | OS | Physician | 04/27/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5480832&pc=MTA | 2023826-2016-00191 | 2015-11-27 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 -7.50/+2.50/X118 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. ON (B)(6) 20105, EXCESSIVE VAULTING WAS NOTED WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES. ON (B)(6) 2015, THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. SEE MFR. REPORT #2023826-2016-00192 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -7.50/+2.50/X118, (B)(4). | -7.50/+2.50/118 | OD | Physician | 12/28/2015 | 2015 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5481076&pc=MTA | 2023826-2016-00192 | 2015-11-27 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 -4.0/+4.00/X031 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. ON (B)(6) 20105, EXCESSIVE VAULTING WAS NOTED WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES. ON (B)(6) 2015, THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. SEE MFR. REPORT #2023826-2016-00191 FOR RIGHT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -4.0/+4.0/X031. (B)(4). | -4.0/+4.00/031 | OS | Physician | 12/28/2015 | 2015 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5446118&pc=MTA | 2023826-2016-00171 | 2015-12-04 | Injury | STAAR SURGICAL COMPANY | 2016-02-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+2.5/130 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | -15.0/+2.5/130 | OD | Physician | 01/20/2016 | 2016 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5577550&pc=MTA | 2023826-2016-00448 | 2015-12-04 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.0MM TICM130V4 -14.00/+3.50/X106 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2009. LOW VAULT WAS OBSERVED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED WITH A DIFFFERENT SIZE LENS ON (B)(6) 2016. PATIENT'S LAST VISIT ON (B)(6) 2016, UCVA WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. CONCLUSION - AS PER DFU FOR THIS LENS MODEL, TICLS ARE CONTRAINDICATED FOR PATIENTS OVER 45 YEARS OLD. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. DIOPTER: -14.00/+3.50/X106. DEVICE NOT RETURNED. (B)(4). | -14.00/+3.50/106 | OD | Physician | 03/16/2016 | 2016 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5593294&pc=MTA | 2023826-2016-00483 | 2015-12-09 | Injury | STAAR SURGICAL COMPANY | 2016-04-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -1.5/+5.5/096 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: PREVIOUS REPORT SOURCE, (B)(6), WAS INCORRECT. CORRECT REPORT SOURCE IS (B)(6). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN AND FOREIGN MATERIAL ON LENS SURFACE. (B)(4). CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4). | -1.5/+5.5/096 | NA | Physician | 01/25/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5577775&pc=MTA | 2023826-2016-00445 | 2015-12-11 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 -08.50/+2.0/X080 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WITH SHALLOW CHAMBER WAS NOTED. THE ICL WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2016, UCVA WAS 20/20. Manufacturer Narrative: DIOPTER: -08.50/+2.00/X080. (B)(4). | -08.50/+2.0/080 | OS | Physician | 03/15/2016 | 2016 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5429379&pc=MTA | 2023826-2016-00124 | 2015-12-13 | Injury | STAAR SURGICAL COMPANY | 2016-02-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/088 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | -11.5/+1.0/088 | OS | Physician | 01/12/2016 | 2016 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5830684&pc=MTA | 2023826-2016-00574 | 2015-12-15 | Injury | STAAR SURGICAL COMPANY | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/100 DIOPTER, IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WEIGHT: UNK. EVENT PROBLEM AND EVALUATION CODES: (B)(4). | -8.5/+1.5/100 | OD | Physician | 02/04/2016 | 2016 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5495680&pc=MTA | 2023826-2016-00259 | 2015-12-17 | Injury | STAAR SURGICAL COMPANY | 2016-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+1.0/034 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE ANTERIOR CHAMBER IS DEEP AND THE PATIENT IS HAPPY. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -13.5/+1.0/034 | OS | Physician | 12/28/2015 | 2015 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5638603&pc=MTA | 2023826-2016-00597 | 2015-12-22 | Injury | STAAR SURGICAL COMPANY | 2016-05-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/X119 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4). | -18.0/+2.0/119 | OD | Physician | 04/08/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5878874&pc=MTA | 2023826-2016-01068 | 2015-12-25 | Injury | STAAR SURGICAL COMPANY | 2016-08-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -12.00/1.0/075 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN AND FOREIGN MATERIAL ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). DEVICE NOT RETURNED. | -12.00/1.0/075 | OD | Physician | 07/21/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5676346&pc=MTA | 2023826-2016-00711 | 2015-12-30 | Injury | STAAR SURGICAL COMPANY | 2016-05-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.0/+1.5/125 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -16.0/+1.5/125 | OD | Physician | 04/27/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5830651&pc=MTA | 2023826-2016-00577 | 2015-12-30 | Injury | STAAR SURGICAL COMPANY | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/024 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.9MM AT THE TIME OF IMPLANTATION. CORRECTED DATA: REPORT SOURCE IS DISTRIBUTOR, NOT USER FACILITY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+2.0/024 | OS | Physician | 02/11/2016 | 2016 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5676394&pc=MTA | 2023826-2016-00715 | 2016-01-11 | Injury | STAAR SURGICAL COMPANY | 2016-05-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.0/+4.0/091 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT AND THAT THERE WAS DRIED, DISCOLORED MATERIAL ON THE LENS SURFACE. CORRECTED DATA: CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM IS CONTRAINDICATED. THIS PATIENT HAS AN ACD OF 2.97MM. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -14.0/+4.0/091 | OD | Physician | 04/25/2016 | 2016 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5565079&pc=MTA | 2023826-2016-00394 | 2016-01-13 | Injury | STAAR SURGICAL COMPANY | 2016-04-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 -12.00/+2.00/X049 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2015. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED ON (B)(6) 2016. THE LENS WAS EXPLANTED AND WAS EXCHANGED FOR A SHORTER LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS ALSO CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -12.00/2.00/X049. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -12.00/+2.00/049 | OS | Physician | 03/11/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5621835&pc=MTA | 2023826-2016-00533 | 2016-01-14 | Injury | STAAR SURGICAL COMPANY | 2016-04-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.50/-0.5/075 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID AND IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN, OPTIC TORN/SPLIT AND PRESENCE OF FOREIGN MATERIAL (SPOT/BLOCK) ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.50/-0.5/075 | OD | Physician | 02/16/2016 | 2016 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5621848&pc=MTA | 2023826-2016-00534 | 2016-01-14 | Injury | STAAR SURGICAL COMPANY | 2016-04-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.00/+1.0/080 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.00/+1.0/080 | OS | Physician | 02/16/2016 | 2016 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536849&pc=MTA | 2023826-2016-00309 | 2016-01-15 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -12.5/+2.5/90 DIOPTER IMPLANTABLE COLLAMER LENS ON (B)(6) 2016 IN TO THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2016 AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -12.5/+2.5/90 (B)(4) SECONDARY SURGERY. NEW INCISION. EXPLANTED. VAULTING. METHOD: WORK ORDER SEARCH. RESULTS: A WORK ORDER SEARCH FOUND ONE SIMILAR COMPLAINT. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -12.5/+2.5/90 | OD | Physician | 02/29/2016 | 2016 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5648608&pc=MTA | 2023826-2016-00667 | 2016-01-17 | Injury | STAAR SURGICAL COMPANY | 2016-05-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+4.0/077 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC AND THE OPTIC WERE BOTH BENT. CORRECTED DATA: CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM IS CONTRAINDICATED. THIS PATIENT HAS AN ACD OF 2.82MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4). LENS NOT RETURNED. | -8.5/+4.0/077 | OS | Physician | 04/08/2016 | 2016 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5900788&pc=MTA | 2023826-2016-00620 | 2016-01-22 | Injury | STAAR SURGICAL COMPANY | 2016-08-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15/+2/090 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND UNREACTIVE PUPIL. THE PATIENT'S INCISION WAS ENLARGED AND THE ANTERIOR CHAMBER WAS IRRIGATED OF REMAINING VISCO FLUIDS. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). DEVICE NOT RETURNED. | -15/+2/090 | OS | Physician | 03/17/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5831903&pc=MTA | 2023826-2016-00929 | 2016-01-24 | Injury | STAAR SURGICAL COMPANY | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED AND THERE WAS FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS SPOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). CLAIM # (B)(4). NOT RETURNED. | -10.0/+2.0/085 | OS | Physician | 06/28/2016 | 2016 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5525072&pc=MTA | 2023826-2016-00339 | 2016-01-26 | Injury | STAAR SURGICAL COMPANY | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 -1.50/+2.5X095 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. ON (B)(6) 2016, PATIENT'S LAST VISIT; UCVA WAS 20/20. Manufacturer Narrative: CONCLUSION: THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0 MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0 MM. CORRECTED DATA: (B)(4). CONCLUSION: CONCLUSION WAS FOUND AND UNABLE TO CONFIRM COMPLAINT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -1.50/+2.50/X095. (B)(4). | -1.50/+2.5095 | OD | Physician | 02/26/2016 | 2016 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5903748&pc=MTA | 2023826-2016-01139 | 2016-01-27 | Injury | STAAR SURGICAL COMPANY | 2016-08-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5 MM TICM125V4 IMPLANTABLE COLLAMER LENS, -20.0/+3.0/79 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -20.0/+3.0/79 | OD | Physician | 08/12/2016 | 2016 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5831914&pc=MTA | 2023826-2016-00934 | 2016-01-27 | Injury | STAAR SURGICAL COMPANY AG | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS, -14.0/+1.00/079 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL PRESENT ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | -14.0/+1.00/079 | OS | Physician | 06/30/2016 | 2016 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5464649&pc=MTA | 2023826-2016-00201 | 2016-01-29 | Injury | STAAR SURGICAL COMPANY | 2016-02-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+0.5/135 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD AN EXCESSIVE VAULT AND A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CONCLUSION CODE: (B)(4)-PER DFU OF THIS LENS MODEL, IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIAL TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD LESS THAN 3.0MM AT THE TIME OF IMPLANTATION. PER DFU FOR THIS LENS MODEL, VTICMOS ARE CONTRAINDICATED FOR PATIENTS ABOVE THE AGE OF 45 YEARS OLD. THIS PATIENT WAS OVER 45 AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -9.50/+0.5/135 | OD | Physician | 02/05/2016 | 2016 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5525102&pc=MTA | 2023826-2016-00336 | 2016-02-01 | Injury | STAAR SURGICAL COMPANY | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 -4.00/+4.00/X087 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2016. LOW VAULT WAS NOTED. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. ON (B)(6) 2016, PATIENT'S LAST VISIT; UCVA WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM IS CONTRAINDICATED. THIS PATIENT HAS AN ACD OF 2.83MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -4.00/+4.00/X087. (B)(4). | -4.00/+4.00/087 | OD | Physician | 02/26/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5542874&pc=MTA | 2023826-2016-00362 | 2016-02-01 | Injury | STAAR SURGICAL COMPANY | 2016-04-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2, -11.50/+1.00/X091 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2016 AND EXCESSIVE VAULTING WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS WITH A SIMILAR DIOPTER SIZE ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN AND FOREIGN MATERIAL ON LENS SURFACE. (B)(4). Manufacturer Narrative: DIOPTER: -11.50/+1.00/X091. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -11.50/+1.00/091 | OS | Physician | 03/02/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5638252&pc=MTA | 2023826-2016-00627 | 2016-02-02 | Injury | STAAR SURGICAL COMPANY | 2016-05-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO 13.7, -12.00/+3.00/083 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2016. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXPLANTED AND WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2016; UCVA WAS 20/25. Manufacturer Narrative: CORRECTION TO DATE OF EVENT: (B)(6) 2016. CORRECTION TO DESCRIBE EVENT OR PROBLEM: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.00/+3.00/083 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXPLANTED AND WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2016, BCVA WAS 20/25. CORRECTION TO IMPLANT DATE: (B)(6) 2016. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THE HAPTIC TO BE TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL ON THE LENS SURFACE (SPOT). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED. | -12.00/+3.00/083 | OS | Physician | 04/06/2016 | 2016 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5672071&pc=MTA | 2023826-2016-00645 | 2016-02-05 | Injury | STAAR SURGICAL COMPANY | 2016-05-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.00/+2.0/077 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, ANTERIOR ENDOTHELIUM CHAMBER DEPTH BELOW 3.0MM FOR VICL/VTICL. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN, AND THERE WAS FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS HAIR ON HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -8.00/+2.0/077 | OS | Physician | 04/07/2016 | 2016 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5664128&pc=MTA | 2023826-2016-00695 | 2016-02-09 | Injury | STAAR SURGICAL COMPANY | 2016-05-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+3.0/180 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THERE WAS NO VISIBLE DAMAGE TO THE LENS BUT THERE WAS FOREIGN MATERIAL ON THE LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -18.0/+3.0/180 | OS | Physician | 05/09/2016 | 2016 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5577767&pc=MTA | 2023826-2016-00446 | 2016-02-10 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -16.00/+3.50/X022 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED. THE ICL WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS. THIS RESOLVED THE PROBLEM. PATIENT CONDITION ON LAST VISIT, STERILE UVEITIS. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: DIOPTER: -16.00/+3.50/X022. (B)(4). | -16.00/+3.50/022 | OS | Physician | 03/16/2016 | 2016 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5621632&pc=MTA | 2023826-2016-00539 | 2016-02-19 | Injury | STAAR SURGICAL COMPANY | 2016-04-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -6.50/+1.50/X133 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. LOW VAULTING WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENGTH LENS (B)(6) 2016. THIS RESOLVED THE PROBLEM. THE PATIENT'S LAST VISIT ON (B)(6) 2016, BCVA:20/25. THE ICL WAS DISCARDED BY THE FACILITY. Manufacturer Narrative: CORRECTED DATA: CORRECT DIOPTER OF THE LENS IS -6.50/+1.50/X113. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE WAS DISCARDED BY THE FACILITY. | -6.50/+1.50/133 | OD | Physician | 03/30/2016 | 2016 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5540148&pc=MTA | 2023826-2016-00377 | 2016-02-25 | Injury | STAAR SURGICAL COMPANY | 2016-03-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -3.5/+3.5/098 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS HAD AN EXCESSIVE VAULT AND THE PUPIL WAS UNREACTIVE (FIXED). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED. THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -3.5/+3.5/098 | OS | Physician | 03/01/2016 | 2016 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5724404&pc=MTA | 2023826-2016-00789 | 2016-02-29 | Injury | STAAR SURGICAL COMPANY | 2016-06-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -9.00/+1.0/64 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CORRECTED DATA: (B)(4). NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4) SECONDARY SURGERY. NEW INCISION. EXPLANTED. VAULTING. (B)(4). | -9.00/+1.0/64 | OD | Physician | 05/20/2016 | 2016 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5549238&pc=MTA | 2023826-2016-00369 | 2016-03-05 | Injury | STAAR SURGICAL COMPANY | 2016-04-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -14.00/1.50/X171 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2016. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES. ICL REMAINS IMPLANTED. SURGERY IS PENDING TO EXPLANT AND REPLACE WITH ANOTHER ICL. Manufacturer Narrative: PRODUCT EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: DIOPTER: -14.00/1.50/X171. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -14.00/1.50/171 | OS | Physician | 03/08/2016 | 2016 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5799978&pc=MTA | 2023826-2016-00926 | 2016-03-12 | Injury | STAAR SURGICAL COMPANY | 2016-07-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/098 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT OPTIC WAS TORN/ SPLIT, HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL (SPOT) ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -13.5/+1.5/098 | OD | Physician | 06/16/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5799970&pc=MTA | 2023826-2016-00925 | 2016-03-13 | Injury | STAAR SURGICAL COMPANY | 2016-07-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/087 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE - SPOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | -13.5/+1.5/087 | OS | Physician | 06/16/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5676420&pc=MTA | 2023826-2016-00716 | 2016-03-15 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -19.0/+1.0/012 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) IS 20/20. Manufacturer Narrative: CORRECTED DATA: AGE AT TIME OF EVENT: 37 TO 40 YEARS. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. A RESIDUE WAS PRESENT ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -19.0/+1.0/012 | OD | Physician | 04/29/2016 | 2016 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5629473&pc=MTA | 2023826-2016-00540 | 2016-03-16 | Injury | STAAR SURGICAL COMPANY | 2016-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -3.00/+3.50/X129 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. EXCESSIVE VAULTING WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2016, UCVA WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS WITH A KERATOCONUS EYE. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED. | -3.00/+3.50/129 | OD | Physician | 04/04/2016 | 2016 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5651122&pc=MTA | 2023826-2016-00662 | 2016-03-19 | Injury | STAAR SURGICAL COMPANY | 2016-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.5/072 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: CORRECTED DATA: DEVICE MANUFACTURE DATE 02/05/2016 IS INCORRECT, CORRECT DATA IS 01/20/2016. (B)(4) - CONCLUSION: REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH WAS BELOW 3.0MM FOR VICL/VTICL. ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN/SPLIT, THE HAPTIC TORN/BROKEN/BENT/DEFORMED, AND THERE WAS FOREIGN MATERIAL ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -10.0/+2.5/072 | OS | Physician | 04/15/2016 | 2016 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5940525&pc=MTA | 2023826-2016-01251 | 2016-03-19 | Injury | STAAR SURGICAL COMPANY | 2016-09-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+3.0/098 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FIBERS ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+3.0/098 | OS | Physician | 08/10/2016 | 2016 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5940533&pc=MTA | 2023826-2016-01252 | 2016-03-19 | Injury | STAAR SURGICAL COMPANY | 2016-09-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/096 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT FIBERS WERE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+1.5/096 | OD | Physician | 08/10/2016 | 2016 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5817531&pc=MTA | 2023826-2016-00917 | 2016-03-23 | Injury | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/073 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20.Event Description: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND NO VISIBLE DAMAGE (B)(4). Manufacturer Narrative: UCVA, INSTEAD OF BCVA, WAS REPORTED. PREVIOUS DEVICE CODE (B)(4) WAS INCORRECT; (B)(4) (SECONDARY SURGERY) WAS A PATIENT CODE INSTEAD OF A DEVICE CODE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+1.5/073 | OD | Physician | 06/22/2016 | 2016 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5655217&pc=MTA | 2023826-2016-00705 | 2016-03-30 | Injury | STAAR SURGICAL COMPANY | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+3.0/083 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT. AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WHO IS LESS THAN 21 YEARS OF AGE IS CONTRAINDICATED. THIS PATIENT IS (B)(6). CORRECTED INFORMATION: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | -15.5/+3.0/083 | OD | Physician | 04/13/2016 | 2016 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5756722&pc=MTA | 2023826-2016-00790 | 2016-03-30 | Injury | STAAR SURGICAL COMPANY | 2016-06-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -10.5/+0.5/062 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE PATIENTS POST OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: DIMENSIONAL INSPECTION PERFORMED FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -10.5/+0.5/062 | OD | Physician | 05/25/2016 | 2016 | 47
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5926678&pc=MTA | 2023826-2016-00680 | 2016-03-31 | Injury | STAAR SURGICAL COMPANY | 2016-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/096 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: CORRECTED DATA: (B)(4). ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS, BUT THERE WAS FOREIGN MATERIAL ON LENS SURFACE SPECIFIED HAIR. (B)(4). Manufacturer Narrative: (B)(4). | -12.5/+1.5/096 | OD | Physician | 04/01/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5884733&pc=MTA | 2023826-2016-01121 | 2016-04-02 | Injury | STAAR SURGICAL COMPANY | 2016-08-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+3.0/89 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT' POST-OP UNCORRECTED VISUAL ACUITY IS 20/20. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED INFORMATION: "USER FACILITY" TO "DISTRIBUTOR". (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4). | -12.5/+3.0/89 | OD | Physician | 06/16/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5662053&pc=MTA | 2023826-2016-00719 | 2016-04-04 | Injury | STAAR SURGICAL COMPANY | 2016-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/+1.50/X079 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2016. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXPLANTED AND WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2016; UCVA WAS 20/20. Manufacturer Narrative: (B)(4). ADDITIONAL DATA: WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT EVENT WITHIN ASSOCIATED LOTS WAS FOUND. CONCLUSION CODE: AS PER THE DFU OF THIS LENS MODEL, IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIAL TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD BELOW 3.0MM AT THE TIME OF IMPLANTATION. PER DFU FOR THIS LENS MODEL, VTICMOS ARE CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS OLD. THIS PATIENT WAS BELOW AGE 21 AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL WITH SOME LIQUID IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND CLEAR RESIDUE ON IT. THE LIQUID IN VIAL WAS THE SALINE SOLUTION DRYING UP. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED. | -12.00/+1.50/079 | OS | Physician | 04/18/2016 | 2016 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5918359&pc=MTA | 2023826-2016-01183 | 2016-04-04 | Injury | STAAR SURGICAL COMPANY | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/87 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL WAS OBSERVED TO BE ON THE LENS (DRIED MATERIAL). UPON DIMENSIONAL INSPECTION, THE LENS WAS FOUND TO BE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). DEVICE NOT RETURNED. | -9.5/+2.0/87 | OS | Physician | 08/02/2016 | 2016 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5655206&pc=MTA | 2023826-2016-00704 | 2016-04-05 | Injury | STAAR SURGICAL COMPANY | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+0.5/097 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | -14.0/+0.5/097 | OD | Physician | 04/13/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5938270&pc=MTA | 2023826-2016-01245 | 2016-04-05 | Injury | STAAR SURGICAL COMPANY | 2016-09-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+4.0/120 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. A SHORTER LENS WAS IMPLANTED ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DRIED, DISCOLORED MATERIAL ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | -10.0/+4.0/120 | OD | Physician | 08/09/2016 | 2016 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5741917&pc=MTA | 2023826-2016-00841 | 2016-04-12 | Injury | STAAR SURGICAL COMPANY | 2016-06-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.50/+2.5/123 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TO BE TORN/BROKEN/BENT/DEFORMED. FOREIGN MATERIAL IN THE FORM OF SPOTS AND FIBERS WERE PRESENT ON THE LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -5.50/+2.5/123 | OS | Physician | 05/30/2016 | 2016 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5662096&pc=MTA | 2023826-2016-00718 | 2016-04-14 | Injury | STAAR SURGICAL COMPANY | 2016-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2, -10.50/+2.50/X083 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2016. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXPLANTED AND WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2016; UCVA WAS 20/16. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: THE EXPLANT DATE, (B)(6) 2016, IS INCORRECT. THE CORRECT EXPLANT IS (B)(6) 2016. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE) AND A PIECE OF HAPTIC MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED. | -10.50/+2.50/083 | OD | Physician | 04/18/2016 | 2016 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5685207&pc=MTA | 2023826-2016-00754 | 2016-04-20 | Injury | STAAR SURGICAL COMPANY | 2016-05-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+2.0/089 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL BELOW THE AGE OF 21 YEARS IS CONTRAINDICATED. THIS PATIENT WAS (B)(6) YEARS OLD AT THE TIME OF THE EVENT. DEVICE CODE: (EXPLANTED) - THIS CODE IS NOT APPLICABLE UNDER DEVICE CODE; IT SHOULD BE LISTED UNDER PATIENT CODE INSTEAD. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -6.5/+2.0/089 | OD | Physician | 05/03/2016 | 2016 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5752027&pc=MTA | 2023826-2016-00794 | 2016-04-21 | Injury | STAAR SURGICAL COMPANY | 2016-06-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -17.00/3.5/85 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | -17.00/3.5/85 | OS | Physician | 05/25/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5724446&pc=MTA | 2023826-2016-00785 | 2016-04-28 | Injury | STAAR SURGICAL COMPANY | 2016-06-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/72 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS ALSO CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4). | -11.5/+1.5/72 | OS | Physician | 05/20/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5724383&pc=MTA | 2023826-2016-00788 | 2016-05-04 | Injury | STAAR SURGICAL COMPANY | 2016-06-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/106 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). CONCLUSION: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED (B)(4). | -10.5/+1.5/106 | OS | Physician | 05/16/2016 | 2016 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5923608&pc=MTA | 2023826-2016-01205 | 2016-05-07 | Injury | STAAR SURGICAL COMPANY | 2016-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -17.5/+3.5/074 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SHALLOW CHAMBER. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION-LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND OPTIC TORN, HAPTIC TORN AND DEFORMED, AND FIBERS ON LENS SURFACES. DIMENSIONAL INSPECTION CONCLUDE THAT THE LENS IS IN SPECIFICATIONS. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.5/+3.5/074 | OD | Physician | 08/02/2016 | 2016 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5902495&pc=MTA | 2023826-2016-01147 | 2016-05-21 | Injury | STAAR SURGICAL COMPANY | 2016-08-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -2.5/+2.0/175 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER INDICATED THE LENS HAD A LOW VAULT AND A REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED AND REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/10. Manufacturer Narrative: THE REPORTER INDICATED THAT THE LENS HAS SINCE BEEN EXCHANGED. CONCLUSION: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS WITH PREEXISTING KERATOCONUS EYE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -2.5/+2.0/175 | OS | Physician | 07/25/2016 | 2016 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5957333&pc=MTA | 2023826-2016-01264 | 2016-05-27 | Injury | STAAR SURGICAL COMPANY | 2016-09-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.5/+5.5/025 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND IRIS-ENDOTHELIUM TOUCH. THE SURGEON WILL PERFORM EXPLORATION OF RETROPUPILLARY SPACE TO TRY TO LIBERATE POSSIBLE POSTERIOUS PERIPHERIC SYNECHIAE. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/80. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -5.5/+5.5/025 | OD | Physician | 08/18/2016 | 2016 | 65
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5850908&pc=MTA | 2023826-2016-01023 | 2016-05-31 | Injury | STAAR SURGICAL COMPANY | 2016-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+1.0/092 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO AN EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE REPORTER STATED THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A SHORTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -15.5/+1.0/092 | OS | Physician | 07/26/2016 | 2016 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5870503&pc=MTA | 2023826-2016-01059 | 2016-06-01 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+2.5/088 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND WILL BE EXPLANTED. THE PUPIL WAS LARGE AND SLUGGISH. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: ON (B)(6) 2016 THE LENS WAS EXCHANGED WITH A SMALLER LENS. AFTER THIS SECONDARY SURGICAL INTERVENTION, PUPIL WAS MID-DILATED AND NON-REACTIVE. DEVICE EVALUATION - THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. PATIENT CODE - (B)(4) SECONDARY INTERVENTION, EXCHANGED LENS. MID-DILATED AND NON-REACTIVE. WORK ORDER SEARCH - NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE CORRECT REPORT SOURCE COUNTRY IS THE (B)(4) . (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). LENS NOT RETURNED. | -14.0/+2.5/088 | OS | Physician | 07/12/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5877567&pc=MTA | 2023826-2016-01075 | 2016-06-01 | Injury | STAAR SURGICAL COMPANY | 2016-08-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -7.0/+4.0/092 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/23. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CONCLUSION: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) BELOW 3.0MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.875MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -7.0/+4.0/092 | OS | Physician | 07/17/2016 | 2016 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5900906&pc=MTA | 2023826-2016-00893 | 2016-06-03 | Injury | STAAR SURGICAL COMPANY | 2016-08-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -5.5/1/142 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAD NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -5.5/1/142 | OD | Physician | 06/14/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5836860&pc=MTA | 2023826-2016-01004 | 2016-06-04 | Injury | STAAR SURGICAL COMPANY | 2016-08-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/097 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO AN EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -12.5/+2.5/097 | OS | Physician | 07/18/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5867041&pc=MTA | 2023826-2016-01014 | 2016-06-05 | Injury | STAAR SURGICAL COMPANY | 2016-08-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+3/112 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: ON (B)(6) 2016, THE PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/20. IMPLANTED (B)(6) 2016 . DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL/RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. CONCLUSION CODE: PATIENT IS UNDER THE AGE OF 21 AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -13.5/+3/112 | OS | Physician | 07/11/2016 | 2016 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5960051&pc=MTA | 2023826-2016-01167 | 2016-06-05 | Injury | STAAR SURGICAL COMPANY | 2016-09-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+3/80 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED. | -13.5/+3/80 | OS | Physician | 08/15/2016 | 2016 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5870300&pc=MTA | 2023826-2016-01053 | 2016-06-06 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/097 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND EXPLANTED WITHIN THE SAME DAY WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PREVIOUS CLAIM # (B)(4) WAS INCORRECT; THE CORRECT CLAIM # (B)(4). (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -8.0/+1.5/097 | OS | Physician | 07/12/2016 | 2016 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5886814&pc=MTA | 2023826-2016-01126 | 2016-06-07 | Injury | STAAR SURGICAL COMPANY | 2016-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/102 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED WITH THE EXCHANGED LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -9.5/+1.5/102 | OD | Physician | 07/19/2016 | 2016 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873723&pc=MTA | 2023826-2016-01072 | 2016-06-08 | Injury | STAAR SURGICAL COMPANY | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/083 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: (B)(4). CONCLUSION CODE: 24-PER DFU FOR THIS LENS MODEL, VTCMOS ARE CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS OLD. THIS PATIENT WAS BELOW 21 AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | -9.0/+1.5/083 | OS | Physician | 07/15/2016 | 2016 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5951518&pc=MTA | 2023826-2016-01256 | 2016-06-10 | Injury | STAAR SURGICAL COMPANY | 2016-09-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/082 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY IS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -12.0/+2.0/082 | OS | Physician | 08/15/2016 | 2016 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5923150&pc=MTA | 2023826-2016-01215 | 2016-06-13 | Injury | STAAR SURGICAL COMPANY | 2016-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS, -7.00/+1.0/095 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016 THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT HAPTIC WAS BROKEN AND BENT AND FOREIGN MATERIAL AND SCRATCHES ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION: AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS MODEL IS CONTRAINDICATED IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH(ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM. THE PATIENT HAD AN ACD OF 2.986MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -7.00/+1.0/095 | OD | Physician | 08/03/2016 | 2016 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5905516&pc=MTA | 2023826-2016-01173 | 2016-06-20 | Injury | STAAR SURGICAL COMPANY | 2016-08-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/108 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/16.5. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FIBER AND SMALL PARTICULATES ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WEIGHT: UNK. (B)(4). LENS NOT RETURNED. | -12.0/+1.0/108 | OS | Physician | 07/26/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5909200&pc=MTA | 2023826-2016-01172 | 2016-06-20 | Injury | STAAR SURGICAL COMPANY | 2016-08-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/090 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/16.5. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FIBER AND SMALL PARTICULATES ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -11.5/+1.0/090 | OD | Physician | 07/26/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5845333&pc=MTA | 2023826-2016-01020 | 2016-06-22 | Injury | STAAR SURGICAL COMPANY | 2016-08-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+3.5/86 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT DEVELOPED A REFRACTIVE SURPRISE AND THE LENS REMAINS IN THE PATIENTS EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4). | -12.0/+3.5/86 | OS | Physician | 07/06/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5938261&pc=MTA | 2023826-2016-01244 | 2016-06-22 | Injury | STAAR SURGICAL COMPANY | 2016-09-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+3.5/051 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN/BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -8.5/+3.5/051 | OS | Physician | 08/09/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5831924&pc=MTA | 2023826-2016-00949 | 2016-06-28 | Injury | STAAR SURGICAL COMPANY | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.5./X094 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (HAIR) ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: REPORT SOURCE IS DISTRIBUTOR, NOT USER FACILITY (B)(4). Manufacturer Narrative: (B)(4). | -12.0/+2.5./094 | OD | Physician | 06/30/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5834829&pc=MTA | 2023826-2016-00933 | 2016-06-28 | Injury | STAAR SURGICAL COMPANY | 2016-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/093 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT; VISUAL INSPECTION FOUND HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -13.0/+2.0/093 | OS | Physician | 07/04/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5870296&pc=MTA | 2023826-2016-01054 | 2016-06-29 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+3.0/100 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO AN EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED, BUT THERE WAS FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS SPOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | -18.0/+3.0/100 | OD | Physician | 07/17/2016 | 2016 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5951519&pc=MTA | 2023826-2016-01259 | 2016-06-29 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -15.0/+4.5/068 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). CONCLUSION: PER DFU FOR THIS LENS MODEL, VTICMOS ARE CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS OLD. THIS PATIENT WAS UNDER 21 AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4). LENS NOT RETURNED. | -15.0/+4.5/068 | OS | Physician | 08/15/2016 | 2016 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5971272&pc=MTA | 2023826-2016-01292 | 2016-07-01 | Injury | STAAR SURGICAL COMPANY | 2016-09-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM 13.2 IMPLANTABLE COLLAMER LENS, -12.5/+4.5/110 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN THE LENS CASE/VIAL AND IN LIQUID. VISUAL INSPECTION FOUND THE LENS TO HAPTIC TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -12.5/+4.5/110 | OD | Physician | 08/09/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5870077&pc=MTA | 2023826-2016-01013 | 2016-07-04 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/99 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL WAS OBSERVED TO BE ON THE LENS (SPOT). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). CONCLUSION: THIS LENS MODEL IS CONTRAINDICATED DUE TO THE AGE OF THE PATIENT BEING HIGHER THAN WHAT WAS INDICATED IN THE DFU. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+2.0/99 | OD | Physician | 07/12/2016 | 2016 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5879317&pc=MTA | 2023826-2016-01067 | 2016-07-04 | Injury | STAAR SURGICAL COMPANY | 2016-08-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7/+1/38 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT, HAPTIC TORN/BROKEN, AND FOREIGN MATERIAL ON LENS SURFACE. (B)(4) SHOULD BE INCLUDED. SHOULD NOT BE LISTED AS DEVICE CODE. (B)(4) SHOULD BE INCLUDED IN PATIENT CODE, NOT DEVICE CODE. CONCLUSION CODE: LENS RETURNED AND EVALUATED, CAN'T CONFIRM COMPLAINT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -7/+1/38 | OD | Physician | 07/22/2016 | 2016 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5902506&pc=MTA | 2023826-2016-01164 | 2016-07-06 | Injury | STAAR SURGICAL COMPANY | 2016-08-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/081 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+1.5/081 | OD | Physician | 07/26/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5902509&pc=MTA | 2023826-2016-01165 | 2016-07-06 | Injury | STAAR SURGICAL COMPANY | 2016-08-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/091 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.5/091 | OS | Physician | 07/26/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5886826&pc=MTA | 2023826-2016-01127 | 2016-07-11 | Injury | STAAR SURGICAL COMPANY | 2016-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/085 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON 07/11/2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). CONCLUSION: PER DFU FOR THIS LENS MODEL, VTICMOS ARE CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS OLD. THIS PATIENT WAS UNDER 21 AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4). | -12.5/+1.0/085 | OD | Physician | 07/20/2016 | 2016 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873728&pc=MTA | 2023826-2016-01073 | 2016-07-13 | Injury | STAAR SURGICAL COMPANY | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/095 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/ VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BROKEN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -11.0/+1.0/095 | OD | Physician | 07/15/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5817545&pc=MTA | 2023826-2016-00928 | 2016-07-15 | Injury | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+1.0/076 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN, OPTIC TORN/SPLIT AND PRESENCE OF FOREIGN MATERIAL (SPOT) ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -13.0/+1.0/076 | OS | Physician | 06/23/2016 | 2016 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5955099&pc=MTA | 2023826-2016-01263 | 2016-07-19 | Injury | STAAR SURGICAL COMPANY | 2016-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/077 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. PATIENT AGE BELOW 21 OR OVER 45 YEARS. PRODUCT EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN AND BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -12.0/+2.0/077 | OS | Physician | 08/17/2016 | 2016 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5902549&pc=MTA | 2023826-2016-01166 | 2016-07-21 | Injury | STAAR SURGICAL COMPANY | 2016-08-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/100 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL DATA: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). CORRECTIONAL DATA: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS), THE LENS MISSING A PIECE OF HAPTIC AND THE LENS WAS CAUGHT WITH VIAL CAP. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -18.0/+1.5/100 | OS | Physician | 07/26/2016 | 2016 | 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5926650&pc=MTA | 2023826-2016-01216 | 2016-07-22 | Injury | STAAR SURGICAL COMPANY | 2016-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/093 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE PRESENCE OF FOREIGN MATERIAL (FIBERS) ON THE LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -13.0/+2.5/093 | OD | Physician | 08/03/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5869130&pc=MTA | 2023826-2016-01037 | 2016-08-01 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/084 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A EXCESSIVE VAULT AND EXPLANTED WITHIN THE SAME DAY WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -18.0/+1.5/084 | OD | Physician | 08/04/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5951511&pc=MTA | 2023826-2016-01254 | 2016-08-04 | Injury | STAAR SURGICAL COMPANY | 2016-09-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -6.5/+1.0/076 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. A LONGER LENS WAS IMPLANTED ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -6.5/+1.0/076 | OS | Physician | 08/15/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5955107&pc=MTA | 2023826-2016-01257 | 2016-08-05 | Injury | STAAR SURGICAL COMPANY | 2016-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/083 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND NO VISIBLE DAMAGE (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -9.5/+1.5/083 | OS | Physician | 08/19/2016 | 2016 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5535804&pc=MTA | 2023826-2016-00318 | 2016-10-07 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.00/+2.00/X098 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2015. EXCESSIVE VAULT AND SHALLOW CHAMBER WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015; UCVA WAS 20/13. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -13.00/+2.00/X098. (B)(4). A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -13.00/+2.00/098 | OD | Physician | 02/29/2016 | 2016 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6091600&pc=MTA | 2023826-2016-01604 | 2016-10-12 | Injury | STAAR SURGICAL COMPANY | 2016-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/099 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR ANOTHER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE OPTIC WAS TORN AND THE HAPTIC WAS TORN AND BENT. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. THERE WERE FIBERS ON THE LENS SURFACE. CORRECTED DATA: OFF-LABEL USE INDICATED AS PATIENT IS PREGNANT AND ACD IS 2.88MM, BOTH OF WHICH ARE CONTRAINDICATIONS TO THE USE OF THIS LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+1.0/099 | OD | Physician | 10/12/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4488993&pc=MTA | 2023826-2015-00114 |  | Injury | STAAR SURGICAL COMPANY | 2015-02-05 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON 03/27/2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NEW INCISION. EXPLANTED. VAULTING. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -11.5/+2.0/090 | OS | Physician | 01/09/2015 | 2015 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4629122&pc=MTA | 2023826-2015-00384 |  | Injury | STAAR SURGICAL COMPANY | 2015-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAS AN EXCESSIVE VAULT AND REMAINS IMPLANTED. THE REPORTER INDICATED THE SURGEON PLANS TO EXPLANT THE LENS AND EXCHANGE FOR A SHORTER LENS. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS IMPLANTED. Manufacturer Narrative: MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -16.50/+4.0/090 | OS | Physician | 03/06/2015 | 2015 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4754658&pc=MTA | 2023826-2015-00508 |  | Injury | STAAR SURGICAL COMPANY | 2015-05-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE ICL WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. DATE OF EVENT: UNK. (B)(4). EVALUATION, METHOD - WORK ORDER SEARCH. RESULTS : A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE NOT AVAILABLE. (B)(4). | -17.5/+3.0/080 | OD | Other | 04/10/2015 | 2015 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4759147&pc=MTA | 2023826-2015-00520 |  | Injury | STAAR SURGICAL COMPANY | 2015-05-08 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED, THE LENS WAS REPLACED FOR A SHORTER LENS ON (B)(6) 2014 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). (SECONDARY SURGERY); (NEW INCISION); SIZE INCORRECT FOR PATIENT; (EXPLANTED); (VAULTING). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE LENS SHOWED IT WAS RETURNED DRY WITH NO VISIBLE DAMAGE. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW: REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP, ACT.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -14.5/+2.0/88 | OD | Other | 04/09/2015 | 2015
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4975014&pc=MTA | 2023826-2015-00964 |  | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -19.5/+1.5/100 | OD | Physician | 07/14/2015 | 2015 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5089549&pc=MTA | 2023826-2015-01234 |  | Injury | STAAR SURGICAL COMPANY | 2015-09-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -19.0/2.0/092 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -19.0/2.0/092 | OS | Physician | 08/20/2015 | 2015 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5089559&pc=MTA | 2023826-2015-01236 |  | Injury | STAAR SURGICAL COMPANY | 2015-09-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -23.0/+2.5/071 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -23.0/+2.5/071 | OS | Physician | 08/20/2015 | 2015 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5171646&pc=MTA | 2023826-2015-01383 |  | Injury | STAAR SURGICAL COMPANY | 2015-10-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+2.5/X080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -14.0/+2.5/080 | OD | Physician | 09/23/2015 | 2015 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5224354&pc=MTA | 2023826-2015-01506 |  | Injury | STAAR SURGICAL COMPANY | 2015-11-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.5/+3.50/X073 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -13.5/+3.50/073 | OD | Physician | 10/15/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5224355&pc=MTA | 2023826-2015-01507 |  | Injury | STAAR SURGICAL COMPANY | 2015-11-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+1.0/101 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -9.50/+1.0/101 | OD | Physician | 10/15/2015 | 2015
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5371776&pc=MTA | 2023826-2016-00064 |  | Injury | STAAR SURGICAL COMPANY | 2016-01-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/088 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS RE-POSITIONED ON (B)(6) 2015. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). | -10.5/+1.0/088 | OD | Physician | 12/18/2015 | 2015 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5518777&pc=MTA | 2023826-2016-00296 |  | Injury | STAAR SURGICAL COMPANY | 2016-03-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -3.00/+3.0/170 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISUAL DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. | -3.00/+3.0/170 | OD | Physician | 01/25/2016 | 2016
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5638607&pc=MTA | 2023826-2016-00592 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/102 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT, HAPTIC TORN/BROKEN/BENT/DEFORMED, AND FOREIGN MATERIAL ON LENS SURFACE (NOT SPECIFIED). (B)(4). CORRECTED DATA: REPORT SOURCE: INITIAL REPORT IS INCORRECT, SHOULD BE DISTRIBUTOR, NOT USER FACILITY. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4). | -12.0/+2.0/102 | OD | Physician | 04/08/2016 | 2016 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5667073&pc=MTA | 2023826-2016-00575 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/096 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE LENS WAS EXCHANGED ON (B)(6) 2016 FOR A LONGER LENS. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC BROKEN AND BENT. (B)(4). CORRECTED DATA: (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -13.0/+1.5/096 | OD | Physician | 02/08/2016 | 2016 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5673195&pc=MTA | 2023826-2016-00635 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.50/+1.0/093 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -6.50/+1.0/093 | OS | Physician | 03/18/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5724421&pc=MTA | 2023826-2016-00787 |  | Injury | STAAR SURGICAL COMPANY | 2016-06-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -11.0/+20/093 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION:THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -11.0/+20/093 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS UNDER 21 YEARS OF AGE. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED (B)(4). | -11.0/+20/093 | OD | Physician | 05/16/2016 | 2016 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5724456&pc=MTA | 2023826-2016-00786 |  | Injury | STAAR SURGICAL COMPANY | 2016-06-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM TICM 125V4 IMPLANTABLE COLLAMER LENS, -22.00/+4.5/081 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT AND THE HAPTIC TORN/BROKEN/BENT/DEFORMED. CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, TICMS ARE CONTRAINDICATED FOR PATIENTS UNDER 21 YEARS OF AGE. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. EVENT DATE: UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | -22.00/+4.5/081 | OD | Physician | 05/16/2016 | 2016 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5780973&pc=MTA | 2023826-2016-00887 |  | Injury | STAAR SURGICAL COMPANY | 2016-07-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+3.0/095 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) IS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PREVIOUSLY, IT WAS REPORTED THAT VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE. THIS IS INCORRECT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BROKEN AND BENT AND FOREIGN MATERIAL ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -15.5/+3.0/095 | OS | Physician | 06/11/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5867033&pc=MTA | 2023826-2016-01015 |  | Injury | STAAR SURGICAL COMPANY | 2016-08-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/62 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (SPOT) ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). DEVICE NOT RETURNED. | -15.5/+1.5/62 | OD | Physician | 07/11/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5909285&pc=MTA | 2023826-2016-01174 |  | Injury | STAAR SURGICAL COMPANY | 2016-08-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/101 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS, WITH FIBER AND SMALL PARTICULATES ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -11.5/+2.0/101 | OS | Physician | 07/26/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4736044&pc=MTA | 2023826-2015-00479 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-04-29 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE (B)(4) BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -8.5/1.5/138 | OD | Physician | 04/02/2015 | 2015 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4954563&pc=MTA | 2023826-2015-00962 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-07-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4). | -20.5/+4.5/120 | OD | Physician | 07/14/2015 | 2015
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5189897&pc=MTA | 2023826-2015-01406 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-10-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, -3.5/+3.0/X078 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | -3.5/+3.0/078 | OS | Physician | 09/29/2015 | 2015 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5648761&pc=MTA | 2023826-2016-00567 |  | Injury | STAAR SURGICAL COMPANY AG | 2016-05-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -15.0/+3.5/086 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: (B)(4). NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4). | -15.0/+3.5/086 | OS | Physician | 02/18/2016 | 2016 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6021746&pc=MTA | 2023826-2016-01443 | 2013-05-31 | Injury | STAAR SURGICAL COMPANY AG | 2016-10-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+4.0/012 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2013. THE LENS HAD BEEN REPOSITIONED TWICE WITH RE-ROTATION DUE TO LOW VAULT. THE LENS WAS REMOVED ON (B)(6) 2016, AND WILL BE EXCHANGED FOR A LENS OF LONGER LENGTH. Manufacturer Narrative: EVENT DATE: (B)(6) 2013. LENS REMOVED (B)(6) 2016; CORRECTED - LENS WAS EXCHANGED ON (B)(6) 2016 AND REPLACED WITH THE SAME SIZE LENS. PROBLEM WAS RESOLVED. LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. (PREVIOUS)-(B)(6) 2016, (CORRECTED)-(B)(6) 2016. (PREVIOUS)-NO, (CURRENT)-YES. (PREVIOUS)-LENS NOT RETURNED, (CURRENT)-LENS RETURNED BUT NOT YET EVALUATED. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -15.0/+4.0/012 | OD | Physician | 10/03/2016 | 2016 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6404638&pc=MTA | 2023826-2017-00396 | 2013-08-27 | Injury | STAAR SURGICAL COMPANY AG | 2017-03-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -15.0/+6.0/99 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE PATIENT EXPERIENCED LOW VAULT AND LENS ROTATION. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED IN SMALL CENTRIFUGE VIAL AND THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. THERE IS DEBRIS AND CLEAR RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.0/+6.0/99 | OD | Physician | 02/15/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6059288&pc=MTA | 2023826-2016-01529 | 2014-02-07 | Injury | STAAR SURGICAL COMPANY | 2016-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS - -15.5/+3.0/86 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT NOT MARKETED IN THE U.S (B)(4). | -15.5/+3.0/86 | OD | Physician | 10/11/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6041201&pc=MTA | 2023826-2016-01091 | 2015-01-30 | Injury | STAAR SURGICAL COMPANY | 2016-10-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+4.50/081 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND ROTATION OF THE LENS. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN AND DRIED DISCOLORATED RESIDUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.50/+4.50/081 | OS | Physician | 05/17/2016 | 2016 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6041198&pc=MTA | 2023826-2016-01092 | 2015-02-08 | Injury | STAAR SURGICAL COMPANY AG | 2016-10-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+3.0/085 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND ROTATION OF THE LENS. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND MULTIPLE DRIED, DISCOLORED PARTICULATES ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. (B)(4). | -11.00/+3.0/085 | OD | Physician | 05/17/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6112596&pc=MTA | 2023826-2016-01681 | 2015-07-07 | Injury | STAAR SURGICAL COMPANY | 2016-11-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+3.5/094 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING WITH ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL; SURGICAL RESIDUE WAS CLEAR; VISUAL INSPECTION FOUND FIBERS ON LENS SURFACE, BUT NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -09.50/+3.5/094 | OD | Physician | 11/06/2015 | 2015 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6509080&pc=MTA | 2023826-2017-00565 | 2015-10-06 | Injury | STAAR SURGICAL COMPANY | 2017-04-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM13.2, DIOPTER -17.0/+1.5/093 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ON (B)(6) 2016, THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO THE PATIENT EXPERIENCING LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: IN THE INITIAL MDR, THE LENS IS INCORRECTLY DESCRIBED AS A VTICM13.2. THE LENS IS A VTICMO13.2. IN THE INITIAL MDR, THE LENS IS INCORRECTLY DESCRIBED AS A VTICM13.2. THE LENS IS A VTICMO13.2. (B)(4). Manufacturer Narrative: CORRECTED DATA: IN FOLLOW-UP #1, THE DATE IS NOT LISTED. IT IS (B)(6) 2017. (B)(4). Manufacturer Narrative: PATIENT WEIGHT: UNK. THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION-LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND BENT. (B)(4). | -17.0/+1.5/093 | OD | Physician | 04/01/2016 | 2016 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6024720&pc=MTA | 2023826-2016-01457 | 2016-02-15 | Injury | STAAR SURGICAL COMPANY | 2016-10-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/089 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND ROTATION OF THE LENS. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PATIENT'S POST-OP (B)(6) 2016) UNCORRECTED VISUAL ACUITY WAS 20/20. THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: (B)(4). | -12.0/+2.0/089 | OS | Physician | 04/19/2016 | 2016 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6036774&pc=MTA | 2023826-2016-01472 | 2016-02-15 | Injury | STAAR SURGICAL COMPANY | 2016-10-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.50/+2.0/094 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND ROTATION OF THE LENS. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.98MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.50/+2.0/094 | OD | Physician | 04/19/2016 | 2016 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5984997&pc=MTA | 2023826-2016-01324 | 2016-05-07 | Injury | STAAR SURGICAL COMPANY | 2016-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.5/+3.5/72 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION FOUND FOREIGN MATERIAL ON SURFACE AND SPOTTY OPTIC (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -15.5/+3.5/72 | OS | Physician | 08/30/2016 | 2016 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5987394&pc=MTA | 2023826-2016-01339 | 2016-05-20 | Injury | STAAR SURGICAL COMPANY | 2016-09-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS - -7.50/+2.0/060 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED LOW VAULT AND LENS ROTATION "NOT" ASSOCIATED WITH LOW VAULT. THE LENS WAS REPOSITIONED IN THE EYE. THE LENS REMAINS IN THE EYE. THE SURGEON PLANS TO EXPLANT THE LENS. THE PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/80. Manufacturer Narrative: LENS EXCHANGED WAS PERFORMED, PROBLEM WAS RESOLVED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.50/+2.0/060 | OS | Physician | 08/31/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7783462&pc=MTA | 2023826-2018-01213 | 2016-06-07 | Injury | STAAR SURGICAL COMPANY | 2018-08-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+4.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO_13.2 IMPLANTABLE COLLAMER LENS -14.5/+4.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. REPORTEDLY, LENS REPOSITIONING WAS PREVIOUSLY ATTEMPTED BUT DID NOT REACH DESIRED OUTCOME. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE EXCHANGE FOR THE LONGER LENS RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS WITH FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -14.5/+4.5/087 | OS | Physician | 12/12/2017 | 2017 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6253455&pc=MTA | 2023826-2017-00078 | 2016-06-13 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.00/+1.5/088 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CORRECTED DATA: THERE IS LENS ROTATION ASSOCIATED WITH LOW VAULT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -5.00/+1.5/088 | OS | Physician | 01/03/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6216619&pc=MTA | 2023826-2016-01852 | 2016-06-20 | Injury | STAAR SURGICAL COMPANY | 2016-12-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -4.50/+6.0/003 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS REPOSITIONED. ON (B)(6) 2016 THE LENS WAS EXPLANTED DUE TO LOW VAULT AND ROTATION. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS IS IN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION: THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. CORRECTIONAL DATA: PREVIOUS EXPLANT DATE NOT APPLICABLE, ON 12/05/2016 THE LENS WAS EXPLANTED DUE TO LOW VAULT AND ROTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4), (LENS REPOSITIONING, LENS EXPLANT, SECONDARY SURGERY). (B)(4). DEVICE NOT RETURNED. | -4.50/+6.0/003 | OD | Physician | 12/05/2016 | 2016 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6620720&pc=MTA | 2023826-2017-00936 | 2016-06-28 | Injury | STAAR SURGICAL COMPANY | 2017-06-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -12.50/+1.0/080 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO THE PATIENT EXPERIENCING LOW VAULT AND LENS DISLOCATION/SUBLUXATION. THE PROBLEM HAS BEEN RESOLVED. Manufacturer Narrative: THE FOLLOWING INFORMATION WAS NOT INCLUDED IN THE INITIAL REPORT: THE PATIENT ALSO EXPERIENCED ENLARGEMENT OF INCISION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION-LENS WAS RETURNED DRY IN A SMALL VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). | -12.50/+1.0/080 | OS | Physician | 05/12/2017 | 2017 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6392833&pc=MTA | 2023826-2017-00358 | 2016-09-05 | Injury | STAAR SURGICAL COMPANY | 2017-03-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.50/133 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED LOW VAULT, AND LENS DISLOCATION/SUBLUXATION. THE LENS WAS REPOSITIONED BUT IT ROTATED AGAIN AFTER REPOSITIONING. ON (B)(6) 2017 THE LENS WAS THEN EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET BEEN EVALUATED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION-LENS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND PIECE OF HAPTIC MISSING. (B)(4). Manufacturer Narrative: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS WITH NARROW ANTERIOR CHAMBER ANGLES (I.E. LESS THAN GRADE III AS DETERMINED BY GONIOSCOPIC EXAM). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.5/+1.50/133 | OD | Physician | 02/10/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6531949&pc=MTA | 2023826-2017-00731 | 2016-09-05 | Injury | STAAR SURGICAL COMPANY | 2017-04-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+3.5/074 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR LONGER LENS. (B)(4). | -17.0/+3.5/074 | OS | Physician | 10/28/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6124382&pc=MTA | 2023826-2016-01671 | 2016-09-28 | Injury | STAAR SURGICAL COMPANY | 2016-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.50/+1.5/093 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT. CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM IS CONTRAINDICATED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.50/+1.5/093 | OD | Physician | 10/25/2016 | 2016 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6407222&pc=MTA | 2023826-2017-00405 | 2016-11-21 | Injury | STAAR SURGICAL COMPANY | 2017-03-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+3.5/99 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE PATIENT EXPERIENCED LOW VAULT AND LENS ROTATION. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A SMALL PIECE OF HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.5/+3.5/99 | OD | Physician | 02/15/2017 | 2017 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6404194&pc=MTA | 2023826-2017-00387 | 2016-12-02 | Injury | STAAR SURGICAL COMPANY | 2017-03-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+2.50/089 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED VERY LOW VAULT, AND LENS ROTATION 60 DEGREES CW. ON (B)(6) 2017 THE LENS WAS REPOSITIONED BUT IT ROTATED AGAIN AFTER REPOSITIONING. SURGEON PLANS TO EXCHANGE LENS. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR THE UNITS WITHIN THE SAME LOT. OFF LABEL USE (OVER 45 YEARS OF AGE AT THE TIME OF SURGERY, ACD<3.00MM) (B)(4). | -10.50/+2.50/089 | OD | Physician | 02/12/2017 | 2017 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6589767&pc=MTA | 2023826-2017-00877 | 2016-12-19 | Injury | STAAR SURGICAL COMPANY AG | 2017-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/088 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT, IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT, IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). | -12.0/+1.0/088 | OS | Physician | 04/28/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6539264&pc=MTA | 2023826-2017-00748 | 2017-02-02 | Injury | STAAR SURGICAL COMPANY | 2017-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/094 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND CLEAR RESIDUE AND DEBRIS ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+1.5/094 | OS | Physician | 04/04/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6442669&pc=MTA | 2023826-2017-00522 | 2017-02-15 | Injury | STAAR SURGICAL COMPANY | 2017-03-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED LOW VAULT, AND LENS ROTATION. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR LONGER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: METHOD CODE: CODE IS NOT APPLICABLE, THERE WAS INSPECTION OF OBJECTS OR SURFACES REQUIRING MAGNIFICATION USING AN OPTICAL MICROSCOPE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. DEVICE EVALUATION: LENS WAS RETURNED IN SMALL VIAL. VISUAL INSPECTION FOUND CLEAR SURGICAL RESIDUE ON THE LENS SURFACE, MISSING PIECE OF HAPTIC. (B)(4). | -11.0/+1.0/080 | OD | Physician | 03/01/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6589757&pc=MTA | 2023826-2017-00875 | 2017-02-22 | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.0/+1.5/081 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF CLEAR RESIDUE ON LENS. CONCLUSION: PATIENT'S ACD IS BELOW 3.0 MM AND AS PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRADICTED FOR PATIENTS WITH ACD BELOW 3.0 MM. CLAIM# (B)(4). Manufacturer Narrative: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF CLEAR RESIDUE ON LENS. CONCLUSION: PATIENT'S ACD IS BELOW 3.0 MM AND AS PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRADICTED FOR PATIENTS WITH ACD BELOW 3.0 MM. CLAIM# (B)(4). | -15.0/+1.5/081 | OS | Physician | 04/28/2017 | 2017 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6491230&pc=MTA | 2023826-2017-00638 | 2017-03-06 | Injury | STAAR SURGICAL COMPANY | 2017-04-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM12.0, DIOPTER -15.0/1.5/072 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCING LOW VAULT AND LENS ROTATION. AS OF THE DATE OF MDR REPORTING THE LENS HAS NOT YET BEEN RETURNED. Manufacturer Narrative: THE LENS WAS EXCHANGED FOR ANOTHER LENS, DIFFERENT MODEL AND DIOPTER, ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))]. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). NOT RETURNED TO MANUFACTURER. | -15.0/1.5/072 | OS | Physician | 03/14/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6491432&pc=MTA | 2023826-2017-00637 | 2017-03-06 | Injury | STAAR SURGICAL COMPANY AG | 2017-04-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM12.0, DIOPTER -13.5/1.5/097 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCING LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR REPORTING THE LENS HAS NOT YET BEEN RETURNED. Manufacturer Narrative: THE LENS WAS EXCHANGED FOR ANOTHER LENS, DIFFERENT MODEL AND DIOPTER, ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: PATIENT WEIGHT: UNK. THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. CONCLUSION: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))]. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/1.5/097 | OD | Physician | 03/14/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6624611&pc=MTA | 2023826-2017-00950 | 2017-03-21 | Injury | STAAR SURGICAL COMPANY | 2017-06-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+3.0/091 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.5/+3.0/091 | OD | Physician | 05/14/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6589766&pc=MTA | 2023826-2017-00876 | 2017-04-14 | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+1.50/X119 DIOPTER, IN THE PATIENT'S RIGHT +EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A PIECE OF HAPTIC MISSING FROM THE LENS AND CLEAR RESIDUE ON THE LENS SURFACE. CONCLUSION: OFF_LABEL USE- AS PER DFU, THE IMPLANTATION OF LENS MODEL VTICMO13.2 IS CONTRAINDICATED FOR PATIENTS OF AGE LESS THAN 21 YEARS. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+1.50/119 | OD | Physician | 04/27/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6624727&pc=MTA | 2023826-2017-00953 | 2017-05-02 | Injury | STAAR SURGICAL COMPANY | 2017-06-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.5/+3.5/081 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC DEFORMED, THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE), AND THE LENS HAVING A CURVED SHAPE. CONCLUSION: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+3.5/081 | OD | Physician | 05/16/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6723380&pc=MTA | 2023826-2017-01138 | 2017-05-25 | Injury | STAAR SURGICAL COMPANY | 2017-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.0/+3.0/102 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION AND WAS EXCHANGED FOR A LONGER LENS OF SIMILAR DIOPTER. THE EXCHANGE RESOLVED THE PROBLEM. ON (B)(6) 2017, THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS AND CLEAR RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: LENS EXCHANGE. (B)(4). | -11.0/+3.0/102 | OD | Physician | 06/30/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6791947&pc=MTA | 2023826-2017-01219 | 2017-06-18 | Injury | STAAR SURGICAL COMPANY | 2017-08-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 VTICMO13.2, -8.50/1.5/57 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. LOW VAULT AND LENS ROTATION WAS NOTICED. THE LENS WAS REPOSITIONED ON (B)(6) 2017. THE VAULT WAS STILL TOO LOW AND THE LENS ROTATED AGAIN AFTER REPOSITIONING. THE LENS IS PLANNED TO BE EXCHANGED FOR A LONGER LENS. Manufacturer Narrative: DATE RECEIVED BY MANUFACTURER" IN FOLLOW-UP # 1 SHOULD BE (B)(6) 2017. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THE LENS WAS EXCHANGED ON (B)(6) 2017 FOR A LONGER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID INSIDE A LENS CASE/ VIAL AND HAD CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF FIBERS ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4). | -8.50/1.5/57 | OD | Physician | 07/17/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6993880&pc=MTA | 2023826-2017-01662 | 2017-06-27 | Injury | STAAR SURGICAL COMPANY | 2017-11-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -10.5/+2.5/079 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS REPOSITIONED. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A LONGER LENS, THE SURGEON NOTING LOW VAULT, LENS ROTATION, AND LENS DISLOCATION/SUBLUXATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS RED AND CLEAR RESIDUE AND DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+2.5/079 | OS | Physician | 10/06/2017 | 2017 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6796062&pc=MTA | 2023826-2017-01307 | 2017-07-07 | Injury | STAAR SURGICAL COMPANY | 2017-08-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.50/+3.5/101 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD A LOW VAULT AND WAS REPOSITIONED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS -16.5/+3.5/101 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD A LOW VAULT AND ROTATION AND WAS REPOSITIONED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.50/+3.5/101 | OD | Physician | 07/21/2017 | 2017 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6796078&pc=MTA | 2023826-2017-01306 | 2017-07-07 | Injury | STAAR SURGICAL COMPANY | 2017-08-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.50/+3.00/087 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD A LOW VAULT AND WAS REPOSITIONED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS HAD LOW VAULT, LENS ROTATION BY 90 DEGREES AND WAS REPOSITIONED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.50/+3.00/087 | OS | Physician | 07/21/2017 | 2017 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8254846&pc=MTA | 2023826-2019-00063 | 2017-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0/+5.0/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE CAUSE OF THE EVENT IS LISTED AS "WRONG SIZE." THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. METHOD CODE ADDED CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -13.0/+5.0/097 | OD |  | 12/17/2018 | 2018 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7131207&pc=MTA | 2023826-2017-01902 | 2017-07-18 | Injury | STAAR SURGICAL COMPANY | 2017-12-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICM5_13.7, -2.0/+3.5/96 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL OBSERVATION FOUND THE LENS WITH NO VISIBLE DAMAGE AND CLEAR RESIDUE ON THE LENS SURFACE. CORRECTIONS: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICM5_13.7, -2.0/+3.5/96 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND LENS ROTATION. THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -2.0/+3.5/96 | OD | Physician | 11/21/2017 | 2017 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978096&pc=MTA | 2023826-2017-01616 | 2017-07-21 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4, DIOPTER -15.5/+3.0/081 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT, THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND CLEAR RESIDUE), AND THE LENS HAS A CURVED SHAPE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))] (B)(4). | -15.5/+3.0/081 | OS | Physician | 09/28/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7033554&pc=MTA | 2023826-2017-01727 | 2017-08-24 | Injury | STAAR SURGICAL COMPANY | 2017-11-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, DIOPTER -9.5/+2.5/090 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6)2017, THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE AND PARTICLES ON THE LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+2.5/090 | OS | Physician | 10/20/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7198057&pc=MTA | 2023826-2018-00017 | 2017-10-31 | Injury | STAAR SURGICAL COMPANY | 2018-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -9.50/+1.5/059 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR LONGER LENS THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). | -9.50/+1.5/059 | OS | Physician | 12/19/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7149243&pc=MTA | 2023826-2017-01957 | 2017-11-04 | Injury | STAAR SURGICAL COMPANY | 2017-12-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -14.0/+2.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION - PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: DATE RECEIVED BY MANUFACTURER IN INITIAL MDR CHANGED FROM 08-DEC-2017 TO 07-DEC-2017. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+2.0/092 | OS | Physician | 12/07/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7183434&pc=MTA | 2023826-2018-00055 | 2017-11-09 | Injury | STAAR SURGICAL COMPANY | 2018-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -11.0/+1.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND PRESENCE OF DEBRIS AND RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/+1.5/90 | OD | Physician | 12/15/2017 | 2017 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7183457&pc=MTA | 2023826-2018-00054 | 2017-11-20 | Injury | STAAR SURGICAL COMPANY | 2018-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -8.0/+2.0/87 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING AND THICK CLEAR RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.0/+2.0/87 | OD | Physician | 12/15/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7184025&pc=MTA | 2023826-2018-00056 | 2017-11-20 | Injury | STAAR SURGICAL COMPANY | 2018-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -8.50/+2.0/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND CLEAR AND WHITE RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/+2.0/92 | OS | Physician | 12/15/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7201661&pc=MTA | 2023826-2018-00092 | 2017-12-20 | Injury | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VTICH12.1 IMPLANTABLE COLLAMER LENS, -1.50/+4.00/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTES LOW VAULT, LENS ROTATION, LENS DISLOCATION/SUBLUXATION. REPORTEDLY, THE LENS WAS ALSO REPOSITIONED. ADDITIONAL COMMENTS STATE, "PRIMARY ICL IS TOO SHORT AND TOO MOBILE IN THE SULCUS. NEEDS TO BE REPLACED WITH A 12.6 MM MODEL." CURRENTLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: UPDATED INFORMATION: ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS. (B)(4) Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -1.50/+4.00/093 | OD | Physician | 12/21/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8182196&pc=MTA | 2023826-2018-01868 | 2017-12-28 | Injury | STAAR SURGICAL COMPANY | 2018-12-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/104 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING, WITH LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BENT AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.0/104 | OD | Physician | 11/22/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7639928&pc=MTA | 2023826-2018-00959 | 2018-01-13 | Injury | STAAR SURGICAL COMPANY | 2018-06-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.0/+2.5/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS, THE SURGEON NOTING LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT AND THE OPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) (B)(4). | -14.0/+2.5/086 | OS | Physician | 06/01/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7456981&pc=MTA | 2023826-2018-00589 | 2018-01-18 | Injury | STAAR SURGICAL COMPANY | 2018-04-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -13.0/+3.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED ON (B)(6) 2018 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND THE LENS HAS A CURVED SHAPE AND RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.0/+3.5/90 | OD | Physician | 03/27/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7456982&pc=MTA | 2023826-2018-00591 | 2018-02-21 | Injury | STAAR SURGICAL COMPANY | 2018-04-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -16.0/+3.0/87 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS REMOVED ON (B)(6) 2018 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF ITS HAPTIC AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.0/+3.0/87 | OD | Physician | 03/27/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7465301&pc=MTA | 2023826-2018-00606 | 2018-03-06 | Injury | STAAR SURGICAL COMPANY | 2018-04-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -12.5/+1.5/097 | OD | Physician | 04/04/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7481617&pc=MTA | 2023826-2018-00663 | 2018-03-16 | Injury | STAAR SURGICAL COMPANY | 2018-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -12.0/+2.5/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO LOW VAULT . THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: SHOULD HAVE BEEN MENTIONED IN INTIAL MDR. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO LOW VAULTING AND LENS ROTATION. (B)(4) Manufacturer Narrative: PATIENT CODE: (B)(4), SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE. (B)(4). | -12.0/+2.5/92 | OD | Physician | 04/04/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7495916&pc=MTA | 2023826-2018-00715 | 2018-04-03 | Injury | STAAR SURGICAL COMPANY | 2018-05-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+5.5/171 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A LARGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A LIQUID INSIDE A VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND THE TORN PIECE MISSING. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.0/+5.5/171 | OD | Physician | 04/17/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594259&pc=MTA | 2023826-2018-00889 | 2018-04-04 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.0/+2.0/067 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. REPORTEDLY, THE LENS WAS ALSO REPOSITIONED PRIOR TO LENS EXCHANGE. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: REPORT SOURCE: USER FACILITY IS INCORRECT, SHOULD BE DISTRIBUTOR. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) | -13.0/+2.0/067 | OS | Physician | 05/16/2018 | 2018 | 49
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7853006&pc=MTA | 2023826-2018-01331 | 2018-04-06 | Injury | STAAR SURGICAL COMPANY | 2018-09-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) (B)(4). | -9.5/+2.0/087 | OD | Physician | 08/09/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7765090&pc=MTA | 2023826-2018-01168 | 2018-04-30 | Injury | STAAR SURGICAL COMPANY | 2018-08-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+4.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE CAUSE OF THE EVENT IS LISTED AS A PATIENT-RELATED FACTOR. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -15.5/+4.0/091 | OD | Physician | 07/13/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7704644&pc=MTA | 2023826-2018-00987 | 2018-05-26 | Injury | STAAR SURGICAL COMPANY | 2018-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. (B)(4). | -13.0/+2.5/094 | OD | Physician | 06/19/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7704666&pc=MTA | 2023826-2018-01051 | 2018-05-26 | Injury | STAAR SURGICAL COMPANY | 2018-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. (B)(4). | -11.5/+1.5/095 | OD | Physician | 06/19/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7885165&pc=MTA | 2023826-2018-01398 | 2018-05-26 | Injury | STAAR SURGICAL COMPANY | 2018-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+2.5/093 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FIBERS ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.50/+2.5/093 | OS | Physician | 08/20/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7783456&pc=MTA | 2023826-2018-01211 | 2018-06-07 | Injury | STAAR SURGICAL COMPANY | 2018-08-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+4.5/107 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -13.0/4.5/107 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. IT WAS REPORTED THAT THERE WAS LENS ROTATION DUE TO LOW VAULT. LENS REPOSITIONING WAS PERFORMED BUT DID NOT REACH DESIRED OUTCOME. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -13.0/+4.5/107 | OD | Physician | 12/12/2017 | 2017 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7740028&pc=MTA | 2023826-2018-01142 | 2018-06-12 | Injury | STAAR SURGICAL COMPANY | 2018-08-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. H6-METHOD CODE ADDED. (B)(4) Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.0/+2.0/092 | OD | Physician | 07/06/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7765097&pc=MTA | 2023826-2018-01167 | 2018-06-27 | Injury | STAAR SURGICAL COMPANY | 2018-08-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.0/+1.0/086 | OS | Physician | 07/13/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824411&pc=MTA | 2023826-2018-01261 | 2018-07-10 | Injury | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -2.50/+3.5/136 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD LOW VAULTING. THE LENS HAD SHIFTED DOWN TO LOWER MARGIN OF PUPIL AND THE CONJUCTIVA WAS OVERHANGING FOR ALMOST 1MM ALL AROUND CORNEA. THE LENS REMAINS IMPLANTED. THE REPORTER INDICTED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LOW VAULT OF THE LENS WAS FIRST NOTICED ON THE 1ST POST-OP DAY AND THE DECENTRATION WAS FOUND ON THE 6TH POST-OP DAY. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -2.50/+3.5/136 | OD | Physician | 07/29/2018 | 2018 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7973764&pc=MTA | 2023826-2018-01474 | 2018-07-26 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.50/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS, EXCESSIVE VAULT, AND ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2018. THE LENS WAS REMOVED AND A SHORTER LENGTH LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -13.50/2.0/095 | OD | Physician | 09/18/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7870353&pc=MTA | 2023826-2018-01386 | 2018-08-02 | Injury | STAAR SURGICAL COMPANY | 2018-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.5/+2.0/091 | OS | Physician | 08/17/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7870364&pc=MTA | 2023826-2018-01383 | 2018-08-02 | Injury | STAAR SURGICAL COMPANY | 2018-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.5/+2.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.5/+2.5/094 | OD | Physician | 08/17/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7870361&pc=MTA | 2023826-2018-01381 | 2018-08-14 | Injury | STAAR SURGICAL COMPANY | 2018-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5/+3.0/106 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS REPOSITIONED. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE IS LISTED AS USER ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.5/+3.0/106 | OD | Physician | 08/17/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7972154&pc=MTA | 2023826-2018-01526 | 2018-08-14 | Injury | STAAR SURGICAL COMPANY | 2018-10-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.0/+3.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+3.0/093 | OD | Physician | 09/20/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7870358&pc=MTA | 2023826-2018-01385 | 2018-08-16 | Injury | STAAR SURGICAL COMPANY | 2018-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0/+3.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE IS LISTED AS USER ERROR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL (RESIDUE) ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -8.0/+3.0/092 | OS | Physician | 08/17/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8129212&pc=MTA | 2023826-2018-01729 | 2018-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.00/4.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2018 AND (B)(6) 2018, THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2018 THE LENS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: DEVICE CODE:(B)(4). ADDITIONAL DATA: DEVICE EVALUATED BY MFR-DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN, AND TORN PIECE IS MISSING. CLAIM #(B)(4). Manufacturer Narrative: UNK. (B)(4). WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/4.0/092 | OS |  | 11/06/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7890063&pc=MTA | 2023826-2018-01414 | 2018-08-20 | Injury | STAAR SURGICAL COMPANY | 2018-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON DECIDED TO IMPLANT AN 13.2 MM VTICMO13.2, -10.0/+2.5/062 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. LENS LENGTH WHICH WAS TOO SHORT CAUSING LOW VAULT. IT WAS HIS FAULT NOT TO SELECT THE 13.7MM LENS LENGTH WHICH IS NOW IN THE PATIENTS EYE (AFTER EXCHANGE ON ((B)(6) 2018) AND SHOWS A GOOD VAULTING OF ~400¿M. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL, CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC DEFORMED, LENS HAS A CURVED SHAPE, AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON DECIDED TO IMPLANT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -10/2.5/62 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON 08/07/2018. THE LENS WAS EXCHANGED ON 08/30/2018 DUE TO LENS LENGTH BEING TOO SHORT AND CAUSING LOW VAULT. IT WAS REPORTED "IT WAS HIS FAULT NOT TO SELECT THE 13.7MM LENS LENGTH WHICH IS NOW IN THE EYE (AFTER EXCHANGE) AND SHOWS A GOOD VAULTING OF ~400M." THIS EXCHANGE RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.0/+2.5/062 | OS | Physician | 08/21/2018 | 2018 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7906081&pc=MTA | 2023826-2018-01345 | 2018-08-22 | Injury | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -12.05/+2.5/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -12.5/+2.5/72 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -12.05/+2.5/072 | OS | Physician | 09/04/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8092652&pc=MTA | 2023826-2018-01714 | 2018-09-01 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.5/+2.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING AND LENS ROTATAION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WASRESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+2.5/088 | OS | Physician | 10/25/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8194497&pc=MTA | 2023826-2018-01902 | 2018-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/108 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -13.5/+1.5/108 | OD |  | 11/30/2018 | 2018 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8054766&pc=MTA | 2023826-2018-01625 | 2018-09-28 | Injury | STAAR SURGICAL COMPANY | 2018-11-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -8.50/2.0/106 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO OBSERVATION OF LOW VAULT WITH ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/2.0/106 | OD | Physician | 10/25/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8194048&pc=MTA | 2023826-2018-01900 | 2018-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -12.5/+1.5/091 | OD |  | 11/30/2018 | 2018 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8354487&pc=MTA | 2023826-2019-00293 | 2018-12-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATE THAT THE SURGEON IMPLANTED A 12.1MM, VTICM5_12.1, -11.5/1.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, INA MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: CORRECTED DATA EVENT: THE REPORTER INDICATE THAT THE SURGEON IMPLANTED A 12.1MM, VTICM5_12.1, -11.5/1.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON 06 JUL 2018. ON 29 DEC 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH LENS ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. DEVICE CODE:(B)(4) - ADDED TO INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/1.5/90 | OD |  | 01/23/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4741318&pc=MTA | 2023826-2015-00167 |  | Injury | STAAR SURGICAL COMPANY | 2015-04-30 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 OF -10.5/+1.6/100 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS ROTATION ASSOCIATED WITH A LOW VAULT. THE LENS WAS REPOSITIONED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). RESULTS: MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID AND IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. THIS IS A RE-SUBMISSION FOR SUPPLEMENTAL MDR #2 WITH CORRECT MFR NUMBER PER FDA REQUEST. SUBMITTED IN ERROR WITH INCORRECT MFR NUMBER USED WAS (2023826-2016-00167) THIS IS NOW CORRECTED TO MFR # 2023826-2015-00167. CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. WEIGHT: UNK. DATE OF EVENT: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). THIS IS A RE-SUBMISSION OF THE INITIAL MDR ELECTRONICALLY THROUGH WEB TRADER PER FDA REQUEST. CLAIM# (B)(4). | -10.5/+1.6/100 | OS |  | 01/19/2015 | 2015
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6519134&pc=MTA | 2023826-2017-00687 |  | Injury | STAAR SURGICAL COMPANY | 2017-04-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+3.0/092 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND LENS ROTATION. LARGE CYSTS WERE PRESENT IN THE EYE, INTERFERING WITH FOOTPLATES POSITIONING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). CONCLUSION: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS UNDER 21 YEARS OLD. (B)(4). | -12.50/+3.0/092 | OS | Physician | 12/01/2015 | 2015 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6791745&pc=MTA | 2023826-2017-01316 |  | Injury | STAAR SURGICAL COMPANY | 2017-08-14 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0/10.5/5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6)2017 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/10.5/5 | OD | Physician | 07/20/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7046770&pc=MTA | 2023826-2017-01746 |  | Injury | STAAR SURGICAL COMPANY | 2017-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+1.0/090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND THE LENS HAD TILTED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/19. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LENS. PATIENT CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. (B)(4). | -14.5/+1.0/090 | OS | Physician | 10/23/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5553052&pc=MTA | 2023826-2016-00388 | 2014-04-13 | Injury | STAAR SURGICAL COMPANY | 2016-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -12.00/+4.50/X 89 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2014. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED (B)(6) 2014. THERE WAS ENLARGEMENT OF OR ADDITION OF A NEW PERIPHERAL IRIDECTOMY. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). CONCLUSION: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL BELOW THE AGE OF 21 YEARS IS CONTRAINDICATED. THIS PATIENT IS (B)(6). (B)(4) Manufacturer Narrative: DIOPTER: -12.00/+4.50/X089. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -12.00/+4.50/ 89 | OS | Physician | 03/08/2016 | 2016 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5989111&pc=MTA | 2023826-2016-01352 | 2016-06-20 | Injury | STAAR SURGICAL COMPANY | 2016-09-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Unintended Movement | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.0/+2.0/088 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION FOUND HAPTIC TORN AND SPOTTY OPTIC. (B)(4): (UNINTENDED MOVEMENT) WAS INCORRECT; ROTATION DUE TO LOW VAULTING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+2.0/088 | OS | Physician | 08/30/2016 | 2016 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5923601&pc=MTA | 2023826-2016-01206 | 2016-03-25 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS, -7.5/+1.5/099 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2009. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION (THE LENS WAS RETURNED DRY IN THE CASE/VIAL; VISUAL INSPECTION FOUND UNSPECIFIED FOREIGN MATERIAL ON SURFACE AND NO VISIBLE DAMAGE). (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | -7.5/+1.5/099 | OD | Physician | 08/02/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6031741&pc=MTA | 2023826-2016-01133 | 2016-03-12 | Injury | STAAR SURGICAL COMPANY | 2016-10-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -10.0/+3.0/92 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT DEVELOPED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED IN THE PATIENTS EYE. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). MODEL# - CORRECTED MODEL# VTICMO13.2. (B)(4). Manufacturer Narrative: (B)(4). | -10.0/+3.0/92 | OS | Physician | 05/17/2016 | 2016 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6684956&pc=MTA | 2023826-2017-01053 | 2017-04-07 | Injury | STAAR SURGICAL COMPANY | 2017-07-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, DIOPTER -8.5/+0.5/072 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT BEGAN TO EXPERIENCE PIGMENT DISPERSION. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS WITH A SIMILAR DIOPTER. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: PREVIOUSLY STATED THAT THE PATIENT EXPERIENCED PIGMENT DISPERSION. CORRECTED INFORMATION RECEIVED FROM THE REPORTER STATES THAT THE PATIENT EXPERIENCED HIGH VAULT, WHICH THEN CAUSED THE PIGMENT DISPERSION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | -8.5/+0.5/072 | OD | Physician | 06/04/2017 | 2017 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5642753&pc=MTA | 2023826-2016-00594 | 2016-03-30 | Injury | STAAR SURGICAL COMPANY | 2016-05-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.00/+1.50/075 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS REPORTED AS HAVING A REFRACTIVE SURPRISE AND REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.00/+1.5/075 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, LENS ROTATION, AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). | -8.00/+1.50/075 | OS | Physician | 07/12/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5834807&pc=MTA | 2023826-2016-00998 | 2016-03-22 | Injury | STAAR SURGICAL COMPANY | 2016-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Torn Material; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/006 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND LENS TEAR. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC TORN/ SPLIT, HAPTIC TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL PRESENT ON THE LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+1.5/006 | OD | Physician | 07/01/2016 | 2016 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6332495&pc=MTA | 2023826-2017-00148 | 2016-12-20 | Injury | STAAR SURGICAL COMPANY | 2017-02-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Unintended Movement | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -7.50/+6/008 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS ROTATION. THERE WAS A PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION. AFTER THE CATARACT SURGERY, THE PATIENT GOT HIS BCVA PRE OP. ON PATIENT'S LAST VISIT ON (B)(6) 2017, THE PATIENT'S BEST CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: REVIEW OF THE FILE INDICATES THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS BECAUSE PATIENT HAS KERATOCONUS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. THIS LENS MODEL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE. (B)(4). LENS NOT RETURNED. | -7.50/+6/008 | OD | Physician | 01/20/2017 | 2017 | 49
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4816518&pc=MTA | 2023826-2015-00639 | 2015-04-25 | Injury | STAAR SURGICAL COMPANY | 2015-06-03 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Malposition of Device | No Code Available | N/A | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LOW VAULT AND ROTATION OF THE LENS WERE NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. REFERENCE MFR. REPORT # 2023826-2015-00638 FOR PATIENT'S RIGHT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGERY. NEW INCISION. EXPLANTED. VAULTING. (B)(4). Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH & MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED DRY WITH NO VISIBLE DAMAGE OBSERVED. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -15.0/2.5/95 | OS | Other | 05/11/2015 | 2015 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5652082&pc=MTA | 2023826-2016-00652 | 2016-03-28 | Injury | STAAR SURGICAL COMPANY | 2016-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Positioning Problem; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM VTICMO12, -17.00/+2.00/X080 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2015. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS ROTATED 70 DEGREES AND WAS REPOSITIONED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENGTH LENS WITH A SIMILAR DIOPTER ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL ON LENS SURFACE (DRY DISCOLORED MATERIAL). (B)(4). CONCLUSION: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS OVER THE AGE 45 YEARS OLD. CORRECTED DATA: "THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 -17.00/+2.00/X080 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE ON (B)(6) 2015. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS ROTATED 70 DEGREES AND WAS REPOSITIONED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS WITH A SIMILAR DIOPTER ON (B)(6) 2016. THIS RESOLVED THE PROBLEM." (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.00/+2.00/080 | OS | Physician | 04/08/2016 | 2016 | 54
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5668415&pc=MTA | 2023826-2016-00713 | 2014-08-25 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Positioning Problem; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4, -19.50/1.50/X009 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2014. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED. A REPOSITIONING OF THE LENS WAS PERFORMED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2016, UCVA WAS 20/33. SEE MFR. #2023826-2016-00714 FOR LEFT EYE. Manufacturer Narrative: (B)(4). CHANGE "INDO-CORNEAL ANGLES" TO "IRIDO-CORNEAL ANGLES". (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED. | -19.50/1.50/009 | OD | Physician | 04/22/2016 | 2016 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5668391&pc=MTA | 2023826-2016-00714 | 2014-08-26 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Positioning Problem; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4, -17.00/+1.50/X005 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2014. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED. A REPOSITIONING OF THE LENS WAS PERFORMED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2016, UCVA WAS 20/20. SEE MFR. #2023826-2016-00713 FOR RIGHT EYE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED. | -17.00/+1.50/005 | OS | Physician | 04/22/2016 | 2016 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5706645&pc=MTA | 2023826-2016-00254 | 2015-11-04 | Injury | STAAR SURGICAL COMPANY | 2016-06-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Torn Material; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.0/+2.5/115 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -16.0/+2.5/115 | OD | Physician | 01/29/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4600891&pc=MTA | 2023826-2015-00314 | 2014-10-21 | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LENS ROTATION, ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED. THE ICL WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DATE OF BIRTH - UNK. PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -13.0/+3.0/089 | OS | Other | 02/13/2015 | 2015 | 47
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4744632&pc=MTA | 2023826-2015-00494 | 2015-03-17 | Injury | STAAR SURGICAL COMPANY | 2015-05-02 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LOW VAULT WAS NOTED, AS WELL AS LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW AND DEVICE HISTORY REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED DRY WITH PEN MARKS. NO VISIBLE DAMAGE WAS OBSERVED. MEDICAL REVIEW - THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST. SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS AND HAPTIC POSITION, ECT). CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -16.0/1.0/118 | OS | Physician | 04/02/2015 | 2015 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4806956&pc=MTA | 2023826-2015-00585 | 2014-04-11 | Injury | STAAR SURGICAL COMPANY | 2015-05-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION CODES: METHOD - WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). CONCLUSIONS - BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH AND MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED IN LIQUID AND NO VISIBLE DAMAGE WAS OBSERVED. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND WITHIN THE WORKORDER. MEDICAL REVIEW: ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -12.5/+3.5/090 | OS | Physician | 02/17/2015 | 2015 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5291346&pc=MTA | 2023826-2015-01612 | 2014-08-25 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS, -22.5/+6.0/93 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/17. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | -22.5/+6.0/93 | OD | Physician | 10/27/2015 | 2015 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5012305&pc=MTA | 2023826-2015-01020 | 2014-12-01 | Injury | STAAR SURGICAL COMPANY AG | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/089 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO INADEQUATE VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4). | -12.5/+2.0/089 | OS | Physician | 07/20/2015 | 2015 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4816519&pc=MTA | 2023826-2015-00638 | 2015-04-25 | Injury | STAAR SURGICAL COMPANY | 2015-06-03 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LOW VAULT AND ROTATION OF THE LENS WERE NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. REFERENCE MFR. REPORT # 2023826-2015-00639 FOR PATIENT'S LEFT EYE. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH & MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED DRY WITH A BROKEN HAPTIC, HOWEVER, THERE WAS NO VISIBLE DAMAGE OBSERVED. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITON, ETC.). CONCLUSION : BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -15.0/2.0/95 | OD | Other | 05/11/2015 | 2015 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5225078&pc=MTA | 2023826-2015-01514 | 2015-07-30 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/X102 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL WITH PERFECT VAULTING OF THE LENS. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | -13.0/+2.0/102 | OS | Physician | 10/16/2015 | 2015 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5225082&pc=MTA | 2023826-2015-01516 | 2015-08-05 | Injury | STAAR SURGICAL COMPANY | 2015-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/X071 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -11.5/+1.0/071 | OS | Physician | 10/16/2015 | 2015 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5225083&pc=MTA | 2023826-2015-01515 | 2015-08-19 | Injury | STAAR SURGICAL COMPANY | 2015-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/X131 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -10.5/+1.0/131 | OD | Physician | 10/16/2015 | 2015 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5972077&pc=MTA | 2023826-2016-01314 | 2015-10-14 | Injury | STAAR SURGICAL COMPANY | 2016-09-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+2.0/095 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS REPORTED AS HAVING A LOW VAULT, WITH LENS ROTATION AND A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2017 AND EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. THE PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/25. DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). | -9.50/+2.0/095 | OS | Physician | 08/25/2016 | 2016 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5453090&pc=MTA | 2023826-2016-00174 | 2015-11-09 | Injury | STAAR SURGICAL COMPANY | 2016-02-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+2.0/088 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED DUE TO LOW VAULTING, ROTATION AND REFRACTIVE CHANGE OVERTIME . THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -9.50/+2.0/088 | OS | Physician | 02/05/2016 | 2016 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5668354&pc=MTA | 2023826-2016-00699 | 2016-02-01 | Injury | STAAR SURGICAL COMPANY | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/+3.50/X141 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2016. LOW VAULT WITH LENS ROTATION WAS NOTED. THE PHYSICIAN TRIED TO REPOSITION THE LENS, BUT THE LENS MOVED AGAIN ON (B)(6) 2016. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2016, UCVA WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BROKEN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). | -10.00/+3.50/141 | OD | Physician | 04/22/2016 | 2016 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5665324&pc=MTA | 2023826-2016-00740 | 2016-03-10 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TICM125V4 IMPLANTABLE COLLAMER LENS, -17.5/+4.5/094 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS HAD DISLOCATED/SUBLUXED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | -17.5/+4.5/094 | OD | Physician | 04/18/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5694863&pc=MTA | 2023826-2016-00755 | 2016-04-25 | Injury | STAAR SURGICAL COMPANY | 2016-06-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.5/+4.5/102 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT, ROTATION, AND LENS DISLOCATION/SUBLUXATION. Manufacturer Narrative: CORRECTED INFORMATION: (B)(4). ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -10.5/+4.5/102 | OD | Physician | 05/14/2016 | 2016 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5850883&pc=MTA | 2023826-2016-01028 | 2016-05-04 | Injury | STAAR SURGICAL COMPANY | 2016-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/+4.5/084 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT AND ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: PREVIOUSLY STATED THAT THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE. CORRECTED STATEMENT-THE PROBLEM WAS NOT RESOLVED WITH THE LENS EXCHANGE. SEE MFR REPORT# 2023826-2018-01213 FOR CASE CONTINUATION. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | -14.50/+4.5/084 | OS | Physician | 07/20/2016 | 2016 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5850885&pc=MTA | 2023826-2016-01027 | 2016-05-04 | Injury | STAAR SURGICAL COMPANY | 2016-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.50/+4.5/102 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT AND ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND OPTIC TORN/SPLIT AND HAPTIC TORN/BROKEN/BENT/DEFORMED. CLAIM #(B)(4). Manufacturer Narrative: PREVIOUSLY STATED THAT THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE. CORRECTED INFORMATION-THE PROBLEM WAS NOT RESOLVED WITH THE LENS EXCHANGE. SEE MFR REPORT# 2023826-2018-01211 FOR CASE CONTINUATION. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WEIGHT: UNK. (B)(4). LENS NOT RETURNED. | -12.50/+4.5/102 | OD | Physician | 07/20/2016 | 2016 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873703&pc=MTA | 2023826-2016-01063 |  | Injury | STAAR SURGICAL COMPANY | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -1.5/+6.0./X080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS FOUND ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN AND BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -1.5/+6.0./080 | OD | Physician | 07/13/2016 | 2016 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5572539&pc=MTA | 2023826-2016-00411 |  | Injury | STAAR SURGICAL COMPANY AG | 2016-04-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -18.0/+6.0/X141 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: CORRECTED INFORMATION: THE LENS WAS EXPLANTED DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED ON (B)(6) 2015 FOR A LONGER LENS AND THE PROBLEMS WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. (B)(4). Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN AND BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | -18.0/+6.0/141 | OD | Physician | 02/19/2016 | 2016
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5648953&pc=MTA | 2023826-2016-00502 | 2016-02-15 | Injury | STAAR SURGICAL COMPANY | 2016-05-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Missing Value Reason | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6-8.0/+1.50/X171 IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2015. THE PATIENT EXPERIENCED LOW VAULT, LENS ROTATION AND GLARE AND HALOS. THE LENS WAS EXPLANTED ON (B)(6) 2016. A LONGER LENGTH LENS WAS IMPLANTED ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS OVER 45 YEARS OF AGE. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8/1.5/X171 | OD | Physician | 04/14/2016 | 2016 | 57
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6340150&pc=MTA | 2023826-2017-00240 | 2015-05-12 | Injury | STAAR SURGICAL COMPANY | 2017-02-17 | MTA | IMPLANTABLE COLLAMER LENS | Insufficient Information | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/+1.0/059 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD). THIS WAS PERFORMED ON (B)(6) 2015. ON (B)(6) 2015, THE LENS WAS THEN EXPLANTED. Manufacturer Narrative: LENS WAS EXCHANGED FOR THE SAME SIZE AND MODEL, SIMILAR DIOPTER LENS ON (B)(6) 2015. AS OF THE DATE OF SUPPLEMENTAL MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. (B)(4). METHOD CODE: WORK ORDER SEARCH PERFORMED. NO SIMILAR COMPLAINTS WERE FOUND FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. (B)(4). Manufacturer Narrative: PATIENT WEIGHT-UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -17.50/+1.0/059 | OD | Physician | 02/24/2016 | 2016 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5106749&pc=MTA | 2023826-2015-01183 | 2015-08-18 | Injury | STAAR SURGICAL COMPANY | 2015-09-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/X102 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE THE LENS REMAINS IMPLANTED. A SUPPLEMENTAL WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS CODE(S): (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, A CONCLUSION CAN DRAWN THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON THE INVESTIGATION PERFORMED, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. EVALUATION CONCLUSION CODE(S): - BASED ON THE COMPLAINT HISTORY, DEVICE HISTORY RECORD REVIEW AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED. | -18.0/+2.0/102 | OD | Physician | 08/28/2015 | 2015 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5908631&pc=MTA | 2023826-2016-00370 | 2015-12-03 | Injury | STAAR SURGICAL COMPANY | 2016-08-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, UNKNOWN DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT DEVELOPED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: DIOPTER OF THE LENS IS -10/+2,5/70. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). DEVICE NOT RETURNED. | -10/+2.5/70 | OS | Physician | 12/30/2015 | 2015 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7558981&pc=MTA | 2023826-2018-00835 | 2018-03-12 | Injury | STAAR SURGICAL COMPANY | 2018-05-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), DURING THE SAME SURGERY DUE TO A CAPSULE TEAR. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A PIECE OF LENS OPTIC AND ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS REVIEWED FOR SHARP/ROUGH EDGES AND NO EVIDENCE OF SHARP/ROUGH EDGES WERE FOUND. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). | -12.5/+1.5/095 | OD | Physician | 05/01/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7720893&pc=MTA | 2023826-2018-01099 |  | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | No Code Available | NA | Event Description: A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.0/+3.0/017 (SPHERE/CYLINDER/AXIS), WAS REPORTED AS BEING TAKEN OUT (REMOVED) FROM THE PATIENTS LEFT EYE (OS) AND REPLACED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, LENS REMOVED AND REPLACED. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -4.0/+3.0/017 | OS | Physician | 06/25/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8073825&pc=MTA | 2023826-2018-01676 | 2018-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.00/3.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE SURPRISE WAS OBSERVED. THE LENS WAS EXPLANTED ON (B)(6) 2018. IT WAS REPORTED THAT THIS EYE HAS LENS COLOBOMA AND THE PUPIL WAS MID-DILATED AT LAST VISIT. Manufacturer Narrative: CHANGED TO ADVERSE EVENT. EVENT - THE PATIENT DID NOT EXPERIENCE REFRACTIVE SURPRISE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/3.5/089 | OS |  | 10/16/2018 | 2018 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8149327&pc=MTA | 2023826-2018-01804 | 2018-11-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-12-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -3.5/+2.5/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, THE "PATIENT IS DISSATISFIED WITH THE MYOPIA REMAINING". THE LENS REMAINS IMPLANTED. SEE MFR REPORT# 2023826-2018-01801 FOR PREVIOUSLY IMPLANTED LENS. Manufacturer Narrative: DEVICE CODE 1401: PREVIOUS CODE IN INITIAL MDR NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -3.5/+2.5/100 | OS |  | 11/19/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372520&pc=MTA | 2023826-2019-00324 | 2018-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS DUE TO LENS ROTATION. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN AND RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.0/+1.0/090 | OS |  | 02/01/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372523&pc=MTA | 2023826-2019-00325 | 2018-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SAME SIZE, DIFFERENT MODEL LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION DEVICE EVALUATED BY MFR : DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -8.5/+1.5/088 | OD |  | 02/01/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8237923&pc=MTA | 2023826-2019-00080 | 2018-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.5/+1.0/121 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION DEVICE EVALUATION- THE LENS WAS RETURNED WITH MOISTURE AND RESIDUE ON LENS SURFACE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN. CORRECTED DATA DEVICE CODE- (B)(4). ADDED TO INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.5/+1.0/121 | OD |  | 12/11/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8237944&pc=MTA | 2023826-2019-00081 | 2018-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.5/+2.0/077 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION DEVICE EVALUATION- THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN. CORRECTED DATA DEVICE CODE- (B)(4). COMPLAINT-(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.5/+2.0/077 | OS |  | 12/11/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8296803&pc=MTA | 2023826-2019-00138 | 2018-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.00/1.0/086 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION- THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TEAR IN LENS HAPTIC. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/1.0/086 | OD |  | 01/10/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372536&pc=MTA | 2023826-2019-00331 | 2018-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: DATE RECEIVED BY MANUFACTURER- CORRECTED FROM "26-MAR-2019" TO "20-MAR-2019" FOR MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+1.0/090 | OS |  | 02/04/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8307842&pc=MTA | 2023826-2019-00224 | 2018-10-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+3.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/DEFORMED; WITH RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: YES;DEVICE EVALUATED BY MANUFACTURER VISUAL INSPECTION FOUND HAPTIC TORN STATEMENT IN PREVIOUS MDR IS NOT APPLICABLE. VISUAL INSPECTION FOUND HAPTIC DEFORMED. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -13.5/+3.0/092 | OS |  | 01/10/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8421242&pc=MTA | 2023826-2019-00348 | 2018-10-31 | Injury | STAAR SURGICAL COMPANY | 2019-03-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.00/1.5/094 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/1.5/094 | OD | Physician | 02/19/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8392895&pc=MTA | 2023826-2019-00360 | 2019-01-29 | Injury | STAAR SURGICAL COMPANY | 2019-03-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+1.5/077 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS A PATIENT-RELATED FACTOR. STATUS OF THE EYE IS REPORTED AS "ANTERIOR CHAMBER NARROWING." Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.5/+1.5/077 | OS | Physician | 02/08/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8421273&pc=MTA | 2023826-2019-00349 | 2019-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/1.0/097 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/1.0/097 | OD |  | 02/19/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8307849&pc=MTA | 2023826-2019-00215 | 2018-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+4.0/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+4.0/084 | OS |  | 01/10/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5092392&pc=MTA | 2023826-2015-01170 | 2015-08-24 | Injury | STAAR SURGICAL COMPANY | 2015-09-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Optical Decentration | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/+3.0/X086 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ON (B)(6) 2015 A REFRACTIVE SURPRISE WAS NOTED. THE LENS REMAINS IMPLANTED. A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED THEN ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -13.0/3.0/X086. DEVICE EVALUATED BY MANUFACTURER? NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINT WAS FOUND. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -13.00/+3.0/086 | OD | Physician | 08/24/2015 | 2015 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8254839&pc=MTA | 2023826-2019-00096 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-01-16 | MTA | IMPLANTABLE COLLAMER LENS | Optical Problem | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -3.0/+6.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. THE SURGEON REPORTS THAT, AT 2 MONTH POST-OP, THE PATIENT DOES NOT SEE WELL. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.; (B)(4). | -3.0/+6.0/086 | NA | Other | 12/20/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7197902&pc=MTA | 2023826-2018-00075 | 2017-12-18 | Injury | STAAR SURGICAL COMPANY | 2018-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -7.0/+2/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO THE LENS BEING IMPLANTED UPSIDE DOWN. THE LENS WAS EXCHANGED FOR A SIMILAR LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT, A TEAR THROUGH THE LENS AND PRESENCE OF RESIDUE ON THE LENS SURFACE. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A, 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, OF -7.0/+2/088 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A SIMILAR LENS AND THE PROBLEM WAS RESOLVED. MODEL #: IT IS CORRECTED FROM "VTICM5_12.6" TO "VTICM5_13.2". SERIAL #: IT IS CORRECTED FROM "(B)(4)" TO "(B)(4)". Manufacturer Narrative: METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | -7.0/+2/088 | OD | Physician | 12/19/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7051357&pc=MTA | 2023826-2017-01784 | 2017-10-26 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.50/+1.00/102 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/100. Manufacturer Narrative: THE REPORTER INDICATED THE ICL LENS (CRYSTAL) WAS IMPLANTED UPSIDE-DOWN UPON INJECTION. THE ICL LENS WAS TORN WHILE BEING REMOVED FROM THE PATIENT'S EYE ON (B)(4) 2017. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A PIECE OF ONE HAPTIC DEFORMED AND MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4). SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR SHORTER LENS. (B)(4). | -17.50/+1.00/102 | OS | Physician | 10/30/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5583323&pc=MTA | 2023826-2016-00450 |  | Injury | STAAR SURGICAL COMPANY AG | 2016-04-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Positioning Problem; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 -15.00/+05.00/X089 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2008. THERE WAS A LENS ROTATION AND LENS REPOSITIONING. THE PATIENT EXPERIENCED A SECOND LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. LENS REMAINS IMPLANTED. PLANNING TO EXCHANGE LENS. SEE MFR. #2023826-2016-00451 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE (B)(4). BUT NOT MARKETED IN THE U.S. DIOPTER: -15.00/+05.00/X089. DEVICE EVALUATED BY MANUFACTURER? NO. DEVICE REMAINS IMPLANTED. CONCLUSIONS: BASED ON THE COMPLAINT INFORMATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -15.00/+05.00/089 | OD | Physician | 03/17/2016 | 2016 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592591&pc=MTA | 2023826-2015-00283 | 2014-10-20 | Injury | STAAR SURGICAL COMPANY | 2015-03-11 | MTA | VISIAN TICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. ROTATION OF THE LENS WAS NOTED ON (B)(6)2014 AND THE LENS WAS REPOSITIONED. THE LENS WAS REMOVED AND REPLACED AND THIS RESOLVED THE PROBLEM. PATIENT'S LAST BCVA - 20/20. Manufacturer Narrative: (B)(4). EVALUATION CODES: METHOD - (OTHER): WORK ORDER SEARCH, DEVICE HISTORY RECORD RESULTS -(OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER (B)(4). Manufacturer Narrative: EVALUATION METHOD: MEDICAL REVIEW. RESULTS: - VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. MEDICAL REVIEW - CLINICAL STUDIES HAVE SHOWN THAT TORIC ICL ROTATION OCCURS IN 0.5% OF PATIENTS WHERE A TICL HAS BEEN IMPLANTED. SEVERAL FACTORS MAY CONTRIBUTE TO THIS PHENOMENON (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION ECT.). THERE IS NOT ENOUGH INFORMATION TO DETERMINE THE EXACT ROOT CAUSE, HOWEVER, IT SEEMS THE ICL WAS SHORT FOR THIS EYE. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -12/3/161 | OD | Physician | 02/09/2015 | 2015 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4903566&pc=MTA | 2023826-2015-00808 | 2015-05-20 | Injury | STAAR SURGICAL COMPANY | 2015-07-09 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015 AND FIVE DAYS POST-OP THE LENS HAD ROTATED 90 DEGREES. THE LENS WAS REPOSITIONED AND THREE DAYS LATER IT HAD ROTATED AGAIN. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND NOT REPLACED. Manufacturer Narrative: (B)(4). EVALUATION CODES: METHOD -(OTHER): WORK ORDER SEARCH RESULTS -(OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSIONS -(NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -9.50/1.0/101 | OS | Physician | 06/09/2015 | 2015 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8149323&pc=MTA | 2023826-2018-01808 | 2018-10-30 | Injury | STAAR SURGICAL COMPANY | 2018-12-10 | MTA | IMPLANTABLE COLLAMER LENS | Unintended Movement | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6)2018. ON (B)(6) 2018 THE LENS WAS SUCCESSFULLY EXCHANGED WITH AN ALTERNATE LENS DUE TO THE LENS WAS IMPLANTED "REVERSE SIDE." THE LENS SURFACE THAT SHOULD HAVE BEEN ON THE IRIS SIDE WAS INSERTED SO THAT IT WAS FACING THE CAPSULE; "FRONT AND BACK SIDE ARE REVERSED." THE CAUSE IS REPORTED AS UNKNOWN. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+1.0/093 | OD | Physician | 11/19/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5741984&pc=MTA | 2023826-2016-00842 | 2015-06-18 | Injury | STAAR SURGICAL COMPANY | 2016-06-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.50/+5.5/071 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD A REFRACTIVE SURPRISE AND ROTATION, AND WAS REPOSITIONED. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. PERFORMED VISUAL INSPECTION AND FOUND THE HAPTIC TORN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -5.50/+5.5/071 | OS | Physician | 05/30/2016 | 2016 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4625042&pc=MTA | 2023826-2015-00363 |  | Injury | STAAR SURGICAL COMPANY | 2015-03-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE ICL WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4): LENS NOT RETURNED. Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -13.0/+1.5/019 | OD | Other | 02/27/2015 | 2015 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4759154&pc=MTA | 2023826-2015-00521 |  | Injury | STAAR SURGICAL COMPANY | 2015-05-08 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) -(SECONDARY SURGERY); (NEW INCISION); UNINTENDED MOVEMENT (ROTATION); (EXPLANTED). EVALUATION METHOD: (OTHER) WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -4.5/+3.0/95 | OD | Physician | 04/09/2015 | 2015
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7356128&pc=MTA | 2023826-2018-00433 | 2018-01-18 | Injury | STAAR SURGICAL COMPANY | 2018-03-20 | MTA | IMPLANTABLE COLLAMER LENS | Use of Device Problem | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.0/+5.0/131 (SPHERE/ CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED AS THE SURGEON IMPLANTED THE LENS INTO THE WRONG EYE-IT SHOULD HAVE BEEN IMPLANTED IN THE OTHER EYE. THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | -8.0/+5.0/131 | OD | Physician | 02/23/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7356134&pc=MTA | 2023826-2018-00436 | 2018-01-31 | Injury | STAAR SURGICAL COMPANY | 2018-03-20 | MTA | IMPLANTABLE COLLAMER LENS | Use of Device Problem | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.0/+1.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A DIFFERENT LENS DUE TO THE SURGEON IMPLANTING THE INCORRECT LENS INTO THE EYE. THE SURGEON MIXED UP THE LENSES FOR THE OS AND OD EYES. THE PROBLEM WAS RESOLVED. REFERENCE MFR REPORT# 2023826-2018-00283 (CLAIM# (B)(4)) FOR PREVIOUS PART OF THIS CASE. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | -6.0/+1.5/095 | OD | Physician | 02/23/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7660094&pc=MTA | 2023826-2018-00999 | 2018-05-17 | Injury | STAAR SURGICAL COMPANY | 2018-07-03 | MTA | IMPLANTABLE COLLAMER LENS | Use of Device Problem | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/059 (SPHERE/CYLINDAR/AXIS) INTO THE PATIENT'S LEFT EYE (OS). ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH AN ALTERNATE LENS DUE TO THE INCORRECT LENS BEING IMPLANTED INTO THE EYE (OD LENS IMPLANTED INTO OS EYE). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS MARKED AS USER ERROR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE AND DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+2.0/059 | OS | Physician | 06/07/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8041444&pc=MTA | 2023826-2018-01590 | 2018-07-21 | Injury | STAAR SURGICAL COMPANY | 2018-11-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available; Missing Value Reason | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.00/+2.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2018 THE LENS WAS REPOSITIONED BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/+2.5/094 | OS | Physician | 10/10/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5549244&pc=MTA | 2023826-2016-00393 | 2016-01-13 | Injury | STAAR SURGICAL COMPANY | 2016-04-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available; Missing Value Reason | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -11.50/2.00/X113 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2015. EXCESSIVE VAULTING WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED (B)(6) 2016. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS. THIS RESOLVED THE PROBLEM. AC IS QUIET AND DEEP NOW. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: DIOPTER: -011.50/+2.00/X113. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -11.50/2.00/113 | OD | Physician | 03/11/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7591522&pc=MTA | 2023826-2018-00683 | 2018-02-19 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICM5_12.6, -11.5/+1.5/98 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED DURING THE SAME SURGERY BECAUSE THE PATIENT MOVED DURING THE PROCEDURE AND THE LENS TURNED UPSIDE DOWN. THERE WAS NO PATIENT INJURY. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TEAR IN A HAPTIC AND RESIDUE ON THE LENS SURFACE. (B)(4). | -11.5/+1.5/98 | OD | Physician | 04/11/2018 | 2018 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7770618&pc=MTA | 2023826-2018-01179 | 2018-06-18 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON WAS IMPLANTING A VTICMO12.6, -13.5/2.00/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS AND THE PLUNGER OVERSHOT THE LENS AND IT BECAME STUCK. IT WAS REPORTED THAT THE LENS TORE DURING REMOVAL AND THERE WAS NO PATIENT INJURY. DELIVERY SYSTEM MODEL AND LOT NUMBERS ARE NOT AVAILABLE. ON (B)(6) 2018 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED.Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE. THE REPORTER STATES THAT DURING IMPLANT, THE PLUNGER OVERSHOT THE LENS AND GOT STUCK. IT TORE DURING REMOVAL. THE LENS WAS IMPLANTED AND REMOVED ON (B)(6) 2018. ON (B)(6) 2018 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). | -13.5/2.00/95 | OD | Physician | 07/10/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4975055&pc=MTA | 2023826-2015-00971 | 2015-06-05 | Malfunction | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS AND THE LENS WAS DAMAGED. THERE WAS NO REPORTED PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING VERY WELL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EVENT PROBLEM CODES: PATIENT CODE: (B)(4)- NO CONSEQUENCE OR IMPACT TO PATIENT, LABELED. DEVICE CODE: (B)(4)- NO CODE AVAILABLE (LENS DAMAGE), LABELED DEVICE EVALUATED BY MANUFACTURER? NO - (OTHER): LENS NOT RETURNED. EVALUATION CODES: METHOD - (OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (OTHER): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). OTHER TEXT : LENS NOT RETURNED. | -13.0/+2.5/063 | NA | Physician | 07/14/2015 | 2015 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5068907&pc=MTA | 2023826-2015-01102 | 2015-08-10 | Malfunction | STAAR SURGICAL COMPANY | 2015-09-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VTICMO12.6, -13.00/+2.50/X121 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD). LENS GOT DAMAGED WHILE INJECTING INTO THE EYE. THE LENS WAS REMOVED AND NO ADVERSE EVENT WAS REPORTED. A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. (B)(4). | -13.00/+2.50/121 | OD | Physician | 08/12/2015 | 2015 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5102508&pc=MTA | 2023826-2015-01178 | 2015-08-26 | Malfunction | STAAR SURGICAL COMPANY | 2015-09-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/X092 DIOPTER IN THE PATIENT'S RIGHT EYE (OD). THE LENS FLIPPED UPSIDE DOWN DURING INSERTION INTO THE EYE. THE LENS WAS ACCIDENTALLY TORN WHILE IT WAS BEING REMOVED FROM THE EYE. NO ADVERSE EVENT WAS REPORTED. NO OTHER LENS WAS IMPLANTED AT THE TIME. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -15.5/+1.5/X092. (B)(4). METHOD CODE: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODE: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION : CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -15.5/+1.5/092 | OD | Physician | 08/26/2015 | 2015 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5381805&pc=MTA | 2023826-2016-00069 | 2015-12-31 | Malfunction | STAAR SURGICAL COMPANY | 2016-01-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/+3.0/092 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2015. THE SURGERY WAS UNEVENTFUL. AFTER THE SURGERY, THEY NOTED MICROBIAL GROWTH/DEPOSITS ON THE EXTERIOR OF THE GLASS BOTTLE, ESPECIALLY ON THE RIM AND NECK. THE LENS REMAINS IMPLANTED AND THE PATIENT IS DOING WELL. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: REVIEW OF THE COMPLAINT BY THE MANUFACTURING TEAM CONCLUDED THE "MICROBIAL GROWTH-LIKE DEPOSIT" ON THE EXTERIOR OF THE VIAL, IS MOST LIKELY THE RESIDUAL BALANCED SALT SOLUTION (BSS) ON THE VIAL. AS PART OF THE PRODUCT PROCESSING STEPS, IT IS POSSIBLE THAT BSS COULD BE DEPOSITED ON THE VIAL(S) AND IF THE BSS DRIES UP AND THE VIAL SURFACE IS NOT THOROUGHLY WIPED DOWN, AS PER THE PROCEDURE, THIS COULD CAUSE "MICROBIAL GROWTH-LIKE DEPOSIT" ON THE VIAL SURFACE. HENCE, RE-TRAINING OF THE CONCERNED OPERATORS WAS COMPLETED BY MANUFACTURING AS A CORRECTIVE ACTION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -16.0/+3.0/092 | NA | Physician | 12/31/2015 | 2015
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4937354&pc=MTA | 2023826-2015-00926 | 2015-06-18 | Malfunction | STAAR SURGICAL COMPANY | 2015-07-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) BUT THE LENS BROKE. THE LENS REMAINS IMPLANTED BUT WILL BE REMOVED. THE LENS WAS THE EXCHANGE LENS, SEE MFR REPORT (B)(4) FOR THE FIRST LENS.Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED ON (B)(6) 2015 AND WAS EXCHANGED FOR ANOTHER SAME MODEL LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). EVALUATION CODES: METHOD: WORK ORDER SEARCH RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM #: (B)(4). LENS NOT RETURNED. Manufacturer Narrative: (B)(4). | -14.0/+4.5/105 | OD | Other | 07/07/2015 | 2015 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5244753&pc=MTA | 2023826-2015-01538 | 2015-10-23 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED 13.2MM VTICMO13.2, -13.50/+2.50/X055 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS TORE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED AND NO OTHER LENS WAS IMPLANTED. NO ADVERSE EVENTS REPORTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -13.50/+2.50/X055. (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S) (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -13.50/+2.50/055 | OS | Physician | 10/29/2015 | 2015 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5440036&pc=MTA | 2023826-2016-00151 | 2014-12-30 | Malfunction | STAAR SURGICAL COMPANY | 2016-02-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -17.0/+3.5/087 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) AND THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED WITH NO REPORT OF ANY PATIENT INJURY. Manufacturer Narrative: CORRECTED DATA: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH . RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -17.0/+3.5/087 | OS | Physician | 12/11/2015 | 2015 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218815&pc=MTA | 2023826-2015-01502 | 2015-08-15 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+0.5/099 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS TORE/BROKE. THE LENS WAS REMOVED AND THERE WAS NO REPORTED PATIENT INJURY. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: THIS CASE IS AN ADVERSE EVENT NOT A PRODUCT PROBLEM. DEVICE EVALUATION - VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN/BENT DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EVENT PROBLEM AND EVALUATION CODES: METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -12.0/+0.5/099 | OD | Physician | 10/13/2015 | 2015 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5481740&pc=MTA | 2023826-2016-00167 | 2014-12-14 | Malfunction | STAAR SURGICAL COMPANY AG | 2016-03-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -9.5/2.5/086 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 AND THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED ON (B)(6) 2014 AND WAS EXCHANGED FOR ANOTHER SAME MODEL/DIFFERENT DIOPTER LENS. THE EXCHANGED LENS RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC TORN. CORRECTED DATA: (B)(4). Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID AND IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. CORRECTIONAL DATA: (B)(4). Manufacturer Narrative: CORRECTED DATA: PLEASE DISREGARD PREVIOUS MDR SUPPLEMENTAL #2. SUBMITTED IN ERROR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | -9.5/2.5/086 | OS | Physician | 12/11/2015 | 2015 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4754639&pc=MTA | 2023826-2015-00507 | 2015-04-02 | Malfunction | STAAR SURGICAL COMPANY | 2015-05-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) AND THE LENS TORE/BROKE DURING INJECTION. THE LENS WAS REMOVED WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER ICL LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -9.5/+1.0/093 | OS | Physician | 04/09/2015 | 2015 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4937372&pc=MTA | 2023826-2015-00907 | 2015-06-11 | Malfunction | STAAR SURGICAL COMPANY | 2015-07-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) BUT THE LENS TORE/BROKE. THE LENS WAS REMOVED WITH NO REPORTED INJURY. A SHORTER LENS WAS IMPLANTED ON (B)(6) 2015 BUT IT ALSO TORE - SEE MFR REPORT (B)(4) FOR THE EXCHANGED LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). EVALUATION CODES: METHOD: WORK ORDER SEARCH RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM#: (B)(4). | -13.5/-4.5/95 | OD | Physician | 07/07/2015 | 2015 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4949683&pc=MTA | 2023826-2015-00896 | 2015-07-01 | Malfunction | STAAR SURGICAL COMPANY | 2015-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) BUT THE LENS TORE/BROKE. THE LENS WAS REMOVED WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED FOR A LONGER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -18.0/+2.5/098 | OD | Physician | 07/03/2015 | 2015 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5518776&pc=MTA | 2023826-2016-00294 | 2015-07-16 | Malfunction | STAAR SURGICAL COMPANY | 2016-03-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+4.0/117 DIOPTER, AND THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO REPORTED INJURY AND THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -13.00/+4.0/117 | NA | Physician | 01/25/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119124&pc=MTA | 2023826-2015-01304 | 2015-08-20 | Malfunction | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/X084 DIOPTER, AND THE LENS TORE/BROKE. THE LENS WAS REMOVED WITH NO REPORTED INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. (B)(4). | -9.0/+1.0/084 | NA | Physician | 09/03/2015 | 2015 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218084&pc=MTA | 2023826-2015-01495 | 2015-08-27 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+1.5/083 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) AND THE LENS TORE/BROKE. THE LENS WAS NOT IMPLANTED AND THERE WAS NO REPORTED PATIENT INJURY. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -14.5/+1.5/083 | OS | Physician | 10/13/2015 | 2015 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5185745&pc=MTA | 2023826-2015-01403 | 2015-09-19 | Malfunction | STAAR SURGICAL COMPANY | 2015-10-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+3.5/X140 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS TORE/BROKE. THE LENS WAS REMOVED WITH NO REPORTED PATIENT INJURY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -14.0/+3.5/140 | OD | Physician | 09/28/2015 | 2015 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5166980&pc=MTA | 2023826-2015-01376 | 2015-05-20 | Malfunction | STAAR SURGICAL COMPANY | 2015-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Crack | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.5/+2.0/096 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS WAS NOTED TO HAVE A CRACK UPON INSERTION. THERE WAS NO REPORTED PATIENT INJURY AND THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -8.5/+2.0/096 | OD | Physician | 09/23/2015 | 2015 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5335403&pc=MTA | 2023826-2015-01695 | 2015-11-17 | Malfunction | STAAR SURGICAL COMPANY | 2015-12-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Contamination with Chemical or Other Material | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.00/+3.00/X093 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD). THERE WAS A FIBRE ON THE ICL AFTER INSERTING THE LENS INTO THE EYE. THE LENS WAS EXPLANTED AND EXCHANGED FOR ANOTHER SIMILAR LENS AND SIZE. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -17.00/+3.00/X093. DEVICE EVALUATED BY MANUFACTURER? NO, DEVICE NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -17.00/+3.00/093 | OD | Physician | 12/02/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7571622&pc=MTA | 2023826-2018-00841 | 2018-03-26 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Contamination with Chemical or Other Material | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT BEFORE THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.5/+2.0/087 (SPHERE/CYLINDR/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD), HE NOTICED A FOREIGN BODY ON THE LENS. REPORTEDLY, A THREAD-SHAPED FOREIGN BODY WAS FOUND ON THE LENS DURING LOADING. THE LENS DID NOT CONTACT THE EYE. THIS OCCURRED ON (B)(6) 2018. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN WITH TORN PIECE MISSING AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. IT SHOULD BE NOTED THAT NO FOREIGN BODY WAS NOTED ON LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+2.0/087 | OD | Physician | 05/09/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7197906&pc=MTA | 2023826-2018-00069 | 2017-02-13 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Damaged by Another Device | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1, -17.0/+1.5/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. DURING IMPLANTATION, THE CARTRIDGE CRACKED DURING INJECTION, CAUSING THE LENS TO TEAR. THE LENS WAS EXPLANTED DURING THE SAME SURGERY AND THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED WITH A SIMILAR LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17.0/+1.5/93 | OD | Physician | 12/19/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6500835&pc=MTA | 2023826-2017-00663 | 2015-09-14 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18/+4.5/112 DIOPTER, WAS NOTED TO BE TORN/ BROKEN BEFORE INJECTION INTO THE PATIENT'S EYE. THE LENS WAS NOT IMPLANTED AND THERE WAS NO PATIENT CONTACT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY. VISUAL INSPECTION FOUND THAT HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED. FOREIGN MATERIAL (HAIR) WAS FOUND ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18/+4.5/112 | NA | Physician | 03/01/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6499917&pc=MTA | 2023826-2017-00656 | 2016-03-14 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+0.5/142 DIOPTER, WAS NOTED TO BE TORN/ BROKEN BEFORE INJECTION INTO THE PATIENT'S EYE. THE LENS WAS NOT IMPLANTED AND THERE WAS NO PATIENT CONTACT. ANOTHER TORRIC LENS OF THE SAME MODEL AND SIMILAR DIOPTER WAS USED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID. VISUAL INSPECTION FOUND THAT HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17.5/+0.5/142 | NA | Physician | 03/23/2016 | 2016
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7001416&pc=MTA | 2023826-2017-01678 | 2017-08-28 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/095 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017 AND THE LENS TORE/BROKE. THE LENS WAS REMOVED DURING THE SAME SURGERY WITH NO PATIENT INJURY. ANOTHER LENS WAS IMPLANTED 2 WEEKS LATER. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS EXCHANGE RESOLVED THE PROBLEM AND THE EXACT DATE OF THE LENS EXCHANGE WAS UNKNOWN. CLAIM #: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC WAS MISSING, HAPTIC TORN, AND CLEAR RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CONCLUSION CODE:OFF LABEL USE (PATIENT BELOW 21 YEARS OF AGE AT TIME OF IMPLANT). (B)(4). | -10.0/+1.0/095 | OD | Physician | 10/15/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7232833&pc=MTA | 2023826-2018-00155 | 2017-10-24 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/077 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), AND THE LENS TORE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED DURING THE SAME SURGERY WITH NO PATIENT INJURY. ANOTHER SAME MODEL, DIFFERENT DIOPTER LENS WAS IMPLANTED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN BUT THE LENS MAY HAVE BEEN SQUEEZED IN THE INJECTOR. Manufacturer Narrative: CORRECTED DATA: DATE OF THIS REPORT: IN THE INITIAL MDR, 05-JAN-2018 IS CORRECTED TO 04-JAN-2018. DATE RECEIVED BY MANUFACTURER: IN THE INITIAL MDR, 05-JAN-2018 IS CORRECTED TO 04-JAN-2018. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/+1.0/077 | OS | Physician | 01/04/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7495137&pc=MTA | 2023826-2018-00671 | 2018-02-18 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+2.5/081 (SPHERE/CYLINDER/AXIS), AND THE LENS TORE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO APPARENT INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID INSIDE A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN WITH PRESENCE OF DEBRIS ON THE LENS SURFACE. CORRECTED DATA: DATE OF THIS REPORT: IT IS CORRECTED TO "(B)(6) 2018" IN THE INITIAL MDR. DATE RECEIVED BY MANUFACTURER: IT IS CORRECTED TO "14-APR-2018" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/+2.5/081 | NA | Physician | 04/14/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7434154&pc=MTA | 2023826-2018-00553 | 2018-03-11 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.00/+1.5/108 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD) AND THE SURGEON NOTED THE LENS WAS TORN. THE LENS WAS CUT TO REMOVE FROM THE EYE DURING THE SAME SURGERY. THERE WAS NO PATIENT INJURY. THE TORN PIECE OF THE LENS REMAINED IN THE CARTRIDGE. THE LENS WAS NOT REPLACED WITH ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A LIQUID INSIDE A LENS CASE/ VIAL AND HAD RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC, HAPTIC TORN AND THE LENS RETURNED IN THREE PIECES. THERE WAS PRESENCE OF RESIDUE ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.00/+1.5/108 | OD | Physician | 03/26/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7434227&pc=MTA | 2023826-2018-00554 | 2018-03-11 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+2.5/088 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS) AND THE SURGEON NOTED THE LENS WAS TORN. THE ANTERIOR CHAMBER WAS IRRIGATED/EVACUATED OF VISCO/FLUIDS. THE LENS WAS REMOVED FROM THE EYE DURING THE SAME SURGERY. THERE WAS NO PATIENT INJURY. THE TORN PIECE OF THE LENS REMAINED IN THE CARTRIDGE. THE LENS WAS NOT REPLACED WITH ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID INSIDE A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS TORN IN THREE PIECES AND A PIECE OF HAPTIC MISSING. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.50/+2.5/088 | OS | Physician | 03/26/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7607748&pc=MTA | 2023826-2018-00914 | 2018-05-07 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/082 (SPHERE/CYLINDER/AXIS), FROM THE LENS VIAL AND NOTICED, WHILE LOADING, THE LENS WAS TORN. THERE WAS NO PATIENT CONTACT. ANOTHER SAME MODEL LENS WAS IMPLANTED. THE SURGEON FELT THE LENS WAS DEFECTIVE. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL AND THERE WAS SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+2.5/082 | NA | Physician | 05/21/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7629848&pc=MTA | 2023826-2018-00947 | 2018-05-10 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -15.5/+4.5/099 (SPHERE/CYLINDER/AXIS), AND THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED DURING THE SAME SURGERY , WITH NO PATIENT INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL AND THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN; HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.5/+4.5/099 | NA | Physician | 05/28/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8795473&pc=MTA | 2023826-2019-01231 | 2019-06-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Failure to Unfold or Unwrap | No Consequences Or Impact To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -7.0/+3.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE ON (B)(6) 2019. THE SURGEON REPORTS THE LENS "WOULDN'T LAY FLAT-LOOKED LIKE THERE WAS A FOLD IN IT." INTRAOPERATIVELY, THE SURGEON REMOVED THE LENS AND SUCCESSFULLY IMPLANTED A BACK-UP LENS. AFTER THE INITIAL LENS WAS REMOVED, NOTHING OBVIOUS WAS NOTED ON THE LENS. Manufacturer Narrative: "PHAKIC INTRAOCULAR LENS, PRODUCT CODE: MTA" SHOULD BE CORRECTED TO "PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB" IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: SURGERY OCCURRED IN THE LEFT (OS) EYE. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNKNOWN. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -7.0/+3.0/088 | OS |  | 06/19/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5154928&pc=MTA | 2023826-2015-01266 | 2015-09-16 | Malfunction | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Fracture | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM013.2 IMPLANTABLE COLLAMER LENS, -11.0/+5.0/X145 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS FOUND TO HAVE A LATERAL CRACK DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED. NO OTHER LENS WAS IMPLANTED. NO INJURY TO THE PATIENT REPORTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT. IS NOT. MARKETED IN THE U.S. DIOPTER: -11.0/+5.0/X145. (B)(4). METHOD CODE(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODE(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -11.0/+5.0/145 | OD | Physician | 09/16/2015 | 2015 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7169879&pc=MTA | 2023826-2018-00052 | 2017-11-01 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS REMOVED DURING THE SAME SURGERY DUE TO EXCESSIVE VAULTING WITH NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.(B)(4). | -9.0/+1.5/082 | OD | Physician | 12/10/2017 | 2017 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8254849&pc=MTA | 2023826-2019-00095 | 2018-03-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE. THIS OCCURRED ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED WITH AN ALTERNATE LENS INTRA-OPERATIVELY. THE PROBLEM WAS RESOLVED. NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. METHOD CODE ADDED, RESULT CODE UPDATED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/+1.0/091 | OD |  | 12/20/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721255&pc=MTA | 2023826-2018-01082 | 2018-04-16 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT, WHILE REMOVING A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/104 (SPHERE/CYLINDER/AXIS) FROM THE VIAL, THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2018. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE LENS WAS RETURNED IN PIECES AND WAS UNABLE TO BE EVALUATED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.5/+2.0/104 | NA | Physician | 07/02/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7749557&pc=MTA | 2023826-2018-01158 | 2018-06-17 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -14.5/1.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY AND THE PATIENT WAS IN GOOD CONDITION AFTER SURGERY. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.5/1.5/90 | NA | Physician | 07/04/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7704799&pc=MTA | 2023826-2018-01053 | 2018-06-18 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.5/+1.0/092 (SPHERE/CYLINDER/AXIS), AND THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. THERE WAS EVIDENCE OF RESIDUE ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.5/+1.0/092 | NA | Physician | 06/20/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824050&pc=MTA | 2023826-2018-01283 | 2018-07-06 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/089 (SPHERE/CYLINDAR/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED INTRAOPERATIVELY, THE SURGEON NOTING LENS TEAR/BREAK DURING INJECTION INTO THE EYE. THIS OCCURRED ON (B)(6) 2018. REPORTEDLY, THE LENS WAS TORN DURING LOADING INTO THE CARTRIDGE. 'PATIENT INJURY' IS INDICATED AS NO. ON (B)(6) 2018 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+2.0/089 | OS | Physician | 08/03/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8408110&pc=MTA | 2023826-2019-00413 | 2018-08-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -15.0/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. PATIENT CONTACT BUT NO PATIENT INJURY WAS REPORTED. ON (B)(6) 2018 A REPLACEMENT LENS WAS IMPLANTED AND THE REPORTER STATED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.0/1.0/093 | OS |  | 02/14/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8171724&pc=MTA | 2023826-2018-01872 | 2018-08-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-12-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.0/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT WITH THE DEVICE BUT NO PATIENT INJURY. ON (B)(6) 2018 A REPLACEMENT LENS WAS IMPLANTED AND IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.0/1.0/092 | OS |  | 11/20/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372285&pc=MTA | 2023826-2019-00303 | 2018-09-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, - 12.0/+2.0/093 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: IN PREVIOUS MDR FOLLOW-UP TYPE WAS FOR DEVICE EVALUATION. DEVICE EVALUATED BY MANUFACTURER: YES; DEVICE EVALUATED BY MANUFACTURER. RESULT CODE 114: OPERATIONAL PROBLEM IDENTIFIED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | - 12.0/+2.0/093 | OS |  | 01/28/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7947554&pc=MTA | 2023826-2018-01487 | 2018-09-12 | Malfunction | STAAR SURGICAL COMPANY | 2018-10-09 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A VTICM5_12.1, -13.50/1.0/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING INJECTION. THE LENS WAS REMOVED AND A REPLACEMENT LENS WAS LATER IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FOREIGN MATERIAL (RESIDUE) ON THE LENS SURFACE. CORRECTED DATA: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 -13.5/+1.0/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED INTRAOPERATIVELY; REPORTER STATES NO PATIENT INJURY. ON (B)(6) 2018 A REPLACEMENT LENS WAS SUCCESSFULLY IMPLANTED, RESOLVING THE PROBLEM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.50/1.0/097 | OD | Physician | 10/02/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8065575&pc=MTA | 2023826-2018-01623 | 2018-09-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -18.00/1.5/017 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE DURING LOADING IN THE CARTRIDGE. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. THE SURGERY WAS ABORTED AND A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2018. IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED, LOADING AND SURGERY WENT WELL AND THE PATIENT IS DOING WELL. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN A ND THE LENS WAS CURVED IN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.00/1.5/017 | NA |  | 10/25/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8299125&pc=MTA | 2023826-2019-00211 | 2018-09-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -6.50/2.0/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING INJECTION/DELIVERY INTO THE PATIENT'S LEFT EYE (OS). THE TORN LENS WAS REMOVED ON ORIGINAL SURGERY DATE OF (B)(6) 2018 WITH NO INJURY TO THE PATIENT. A REPLACEMENT LENS WAS ORDERED ON THE SAME DAY OF THE INCIDENT AND IMPLANTED ON (B)(6) 2018 WITH NO INJURY TO THE PATIENT AND THE PROBLEM WAS RESOLVED. THE REPORTER STATED THAT THE CAUSE OF THIS EVENT WAS USER ERROR " LENS OUT OF AXIAL POSITION TURNED LENS AND LENS TORE". Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.50/2.0/104 | OS |  | 01/04/2019 | 2019 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8113536&pc=MTA | 2023826-2018-01757 | 2018-10-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/073 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) AND THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE LENS WAS EXCHANGED WITH ANOTHER SAME MODEL/DIOPTER LENS ON (B)(6) 2018. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT; PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+1.5/073 | OS |  | 10/29/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8255229&pc=MTA | 2023826-2019-00104 | 2018-11-20 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/092 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THEY FELT THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. POST-OP PATIENT CONDITION WAS RESIDUAL ASTIGMATISM OF 1.25 D. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.0/+2.5/092 | NA |  | 12/17/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8298708&pc=MTA | 2023826-2019-00189 | 2018-12-19 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, - 12.0/+1.5/085 (SPHERE/CYLINDER/AXIS) DURING THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE OPIC AND HAPTIC TORN. PRESENCE OF RESIDUE WAS FOUND ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | - 12.0/+1.5/085 | NA |  | 01/07/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8382689&pc=MTA | 2023826-2019-00342 | 2019-01-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT WHILE INJECTING A 12.6MM, VTICMO12.6, -16/2.5/085 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(60 2019 THE LENS TORE. THE LENS WAS REMOVED IMMEDIATELY, NO PATIENT INJURY. A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. PATIENT IS DOING WELL AND HAPPY. Manufacturer Narrative: PRODUCT EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN AND BENT AND RESIDUE ON LENS SURFACE. IMPLANT/EXPLANT DATE: CORRECTED FROM "(B)(6) 2019" TO "(B)(6) 2019" FOR INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16/2.5/085 | OS |  | 01/30/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7197900&pc=MTA | 2023826-2018-00070 | 2017-11-15 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -10.5/+5.0/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED DURING THE SAME SURGERY DUE TO THE LENS BEING UPSIDE DOWN. THE LENS TORN DURING REMOVAL. THERE WAS NO PATIENT INJURY. ANOTHER LENS WAS IMPLANTED ON (B)(6) 2017 AND IT WAS SUCCESSFUL. IT IS REPORTED THAT POOR LOADING OF THE LENS IN THE CARTRIDGE MIGHT BE THE CONTRIBUTING FACTOR THAT LED TO THE LENS BEING IMPLANTED UPSIDE DOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL AND HAD CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF ITS HAPTIC AND PRESENCE OF RESIDUE ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/+5.0/90 | OS | Physician | 12/19/2017 | 2017 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7364242&pc=MTA | 2023826-2018-00450 | 2018-01-31 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+5.0/092 (SPHERE/CYLINDER/AXIS), AND THE LENS WENT INTO THE EYE UPSIDE DOWN. THE LENS WAS DAMAGED WHEN REMOVED, WITHIN THE SAME SURGERY. THERE WAS NO PATIENT INJURY. THE LENS WAS NOT REPLACED WITH ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (DEBRIS ON LENS) AND THE LENS MISSING A PIECE OF HAPTIC. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). | -10.0/+5.0/092 | NA | Physician | 02/25/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7481807&pc=MTA | 2023826-2018-00625 | 2018-03-20 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -8.5/3.0/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS IMPLANTED UPSIDE-DOWN AND WAS EXPLANTED IN THE SAME SURGERY. THE LENS WAS DAMAGED DURING EXPLANTATION. THERE WAS NO PATIENT INJURY. A REPLACEMENT LENS WAS IMPLANTED A WEEK LATER AND IT IS ALL FINE. IT IS REPORTED THAT THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS VIAL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TORN WITH PRESENCE OF RESIDUE ON THE LENS SURFACE. THE LENS WAS MISSING THE TORN PIECE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/3.0/89 | OD | Physician | 04/03/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119417&pc=MTA | 2023826-2015-01296 | 2015-02-13 | Malfunction | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON NOTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+6.0/X089 DIOPTER, WAS TORN BEFORE LOADING. THE LENS WAS NOT IMPLANTED AND THERE WAS NO REPORTED INJURY. THE LENS WAS EXCHANGED FOR ANOTHER ICL LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. IMPLANT AND EXPLANT DATES: NA. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. : EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | -11.5/+6.0/089 | NA | Physician | 09/01/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5155148&pc=MTA | 2023826-2015-01352 | 2015-06-24 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0/+2.0/X098 DIOPTER, AND NOTED THE LENS HAD TORN. THE LENS WAS REMOVED WITH NO REPORTED INJURY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EVENT PROBLEM AND EVALUATION CODES: METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) LENS NOT RETURNED. | -16.0/+2.0/098 | NA | Physician | 09/15/2015 | 2015 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4655095&pc=MTA | 2023826-2015-01120 | 2015-08-04 | Malfunction | STAAR SURGICAL COMPANY | 2015-04-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material; Unintended Movement | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.0/+3.0/099 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE LENS FLIPPED IN THE EYE AND UPON REMOVAL THE LENS TORE. THERE WAS NO REPORTED PATIENT INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. THE LENS WAS REHYDRATED AND BOTH THE LENGTH AND WIDTH WERE RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP, ETC.) TO PREVENT ANY OF THESE COMPLICATIONS FORM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -18.0/+3.0/099 | OD | Physician | 08/11/2015 | 2015 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5047786&pc=MTA | 2023826-2015-01117 | 2015-03-13 | Injury | STAAR SURGICAL COMPANY | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/3.0/94 DIOPTER, AND THE LENS WENT INTO THE PATIENT'S EYE UPSIDE DOWN. THE LENS WAS REMOVED WITH NO APPARENT INJURY. ANOTHER SAME MODEL LENS LENS WAS IMPLANTED ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -12.0/3.0/94 | NA | Physician | 08/04/2015 | 2015 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8073837&pc=MTA | 2023826-2018-01603 | 2018-05-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 A VTICMO 13.2, -13.00/3.0/115 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS FLIPPED. THE LENS WAS REMOVED AND EXCHANGED DURING INITIAL SURGERY WITH NO PATIENT INJURY. REPORTER STATES UPSIDEDOWN WHEN IMPLANTING THE LENS. THE LENS TORE WHEN REMOVING. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BENT AND RESIDUE ON THE LENS. THE REPORTER INDICATED THAT ON (B)(6) 2018 A VTICMO13.2, -13.00/3.0/115 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS FLIPPED UPSIDE DOWN DURING IMPLANTATION. THE LENS WAS REMOVED IN THE SAME SURGERY AND TORE DURING REMOVAL WITH NO PATIENT INJURY. A DIFFERENT LENS WAS LATER IMPLANTED. THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.00/3.0/115 | NA |  | 10/16/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8203029&pc=MTA | 2023826-2018-01826 | 2018-11-19 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-12-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 A 13.2MM VTICM5_13.2, -12.00/1.5/098 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENTS RIGHT EYE (OD). THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO BEING IMPLANTED UPSIDE DOWN. IT WAS REPORTED THE LENS WAS DAMAGED WHEN THE SURGEON REMOVED THE LENS. AND THAT SURGEON WILL IMPLANT A NEW LENS AT A LATER DATE TO RESOLVE THE PROBLEM. "NO PATIENT INJURY IS REPORTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/1.5/098 | OD |  | 12/04/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7834998&pc=MTA | 2023826-2018-01316 | 2018-07-16 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Consequences Or Impact To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_13.2, -14.50/+3.5/173 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE SO THE SURGEON REMOVED THE LENS WITHIN THE SAME SURGERY, THERE WAS NO PATIENT INJURY. ON (B)(6) 2018 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DIOPTER CORRECTED TO: "-14.00/+3.5/173. LENS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND THE LENS TORN INTO TWO PIECES. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_13.2 -14.00/+3.5/173 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE SO THE SURGEON REMOVED THE LENS WITHIN THE SAME SURGERY. THERE WAS NO PATIENT INJURY. ON (B)(6) 2018 A NEW LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IIS NOT MARKETED IN THE US. (B)(4). | -14.50/+3.5/173 | OS | Physician | 07/31/2018 | 2018 | 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5495686&pc=MTA | 2023826-2016-00260 | 2015-12-24 | Malfunction | STAAR SURGICAL COMPANY | 2016-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material; Unintended Movement | No Consequences Or Impact To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/095 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS WENT INTO THE EYE UPSIDE DOWN. THE LENS WAS DAMAGED WHILE REMOVING FROM THE EYE WITH NO REPORTED INJURY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXCHANGED ON (B)(6) 2016 FOR A SAME MODEL LENS. THIS RESOLVED THE PROBLEM. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -12.0/+1.5/095 | OD | Physician | 12/28/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6801531&pc=MTA | 2023826-2017-01346 | 2017-07-14 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.00/1.00/093 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT IS EXPERIENCING REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. ACCORDING TO THE REPORTER THE LENS IS UPSIDE DOWN AND AN EXCHANGE IS PLANNED FOR (B)(6) 2017. Manufacturer Narrative: EXPLANTED DATE: (B)(6) 2017. THE LENS WAS REPOSITIONED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS FOUND TO BE 20/20. (B)(4). CORRECTED DATA: THE STATEMENT "THE PATIENT IS EXPERIENCING REFRACTIVE SURPRISE" IS CORRECTED TO "THE SURGEON NOTED REFRACTIVE SURPRISE". THE CLAIM NUMBER IN THE INITIAL SHOULD BE "(B)(4)". Manufacturer Narrative: ADDITIONAL DATA: THIS PRODUCT IS MANUFACTURED IN THE US. RESULT: WORK ORDER SEARCH WAS PERFORMED- NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4). | -5.00/1.00/093 | OS | Physician | 07/19/2017 | 2017 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5662052&pc=MTA | 2023826-2016-00698 | 2016-02-17 | Injury | STAAR SURGICAL COMPANY | 2016-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Consequences Or Impact To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.00/+1.00/X094 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS TORE DURING DELIVERY INTO THE EYE. DURING THE SAME PROCEDURE THE LENS WAS REMOVED FROM THE EYE AND EXCHANGED FOR ANOTHER SAME MODEL WITH A SIMILAR DIOPTER. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2016; UCVA WAS 20/20. NO ADVERSE EVENT REPORTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE. (B)(4). THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14/+1.5/94 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON THE SAME DAY DUE TO A LENS TEAR/BREAK DURING INJECTION. THE LENS WAS EXCHANGED ON (B)(6) 2016 FOR ANOTHER SAME MODEL AND SIMILAR DIOPTER. THIS RESOLVED THE PROBLEM. ON (B)(6) 2016, THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.00/+1.00/094 | OD | Physician | 04/22/2016 | 2016 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5662051&pc=MTA | 2023826-2016-00697 | 2016-03-09 | Injury | STAAR SURGICAL COMPANY | 2016-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Consequences Or Impact To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.00/+2.00/X064 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS TORE DURING DELIVERY INTO THE EYE. THE LENS GOT STUCK IN THE INJECTOR TIP AND PLUNGER AND TORE. DURING THE SAME PROCEDURE THE LENS WAS REMOVED FROM THE EYE AND EXCHANGED FOR ANOTHER SAME MODEL WITH A SIMILAR DIOPTER. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2016; UCVA WAS 20/20. NO ADVERSE EVENT WAS REPORTED. Manufacturer Narrative: CORRECTED DATA: THIS CASE IS AN ADVERSE EVENT. THE LENS WAS EXCHANGED FOR A SIMILAR LENS WITH A -15/+2,5/96 DIOPTER ON (B)(6) 2016. (B)(4). Manufacturer Narrative: (B)(4). CLAIM #(B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. A VISUAL INSPECTION WAS PERFORMED OBSERVING THE OPTIC WAS TORN/SPLIT, HAPTIC WAS TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL ON THE LENS SURFACE (UNIDENTIFIED). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THE CORRECT LENS POWER IS -15.00/+2.50/X064 DIOPTER. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.00/+2.00/064 | OS | Physician | 04/22/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6058822&pc=MTA | 2023826-2016-01478 | 2016-06-22 | Injury | STAAR SURGICAL COMPANY | 2016-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Consequences Or Impact To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/90 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016 AND THE LENS WAS NOTED TO BE TORN. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A LENS THE SAME DIOPTER AND LENGTH AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN AND HAPTIC BROKEN. WORK ORDER SEARCH - NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THIS CASE IS AN ADVERSE EVENT, NOT A PRODUCT PROBLEM. THE CORRECT IMPLANT DATE PER MEDICAL REVIEW IS (B)(6) 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.5/90 | OS | Physician | 09/27/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6926706&pc=MTA | 2023826-2017-01555 | 2017-08-02 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+2.0/083 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE AND HAD CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS IMPLANTED AND REMOVED DURING THE SAME SURGERY AND THERE WAS NO PATIENT INJURY. CORRECT DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR SHORTER LENS. CONCLUSION CODE: OFF LABEL USE (PATIENT BELOW 21 YEARS OF AGE AT TIME OF IMPLANT). (B)(4). | -14.0/+2.0/083 | OD | Physician | 09/11/2017 | 2017 | 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5216032&pc=MTA | 2023826-2015-01450 | 2015-10-12 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50/+2.50/X009 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LENS TORE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED. NO ADVERSE EVENT WAS REPORTED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; VISUAL INSPECTION FOUND HAPTIC TORN AND PEN MARKS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -15.50/+2.50/X009. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -15.50/+2.50/009 | OD | Physician | 10/12/2015 | 2015 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7855722&pc=MTA | 2023826-2018-01354 | 2018-04-06 | Injury | STAAR SURGICAL COMPANY | 2018-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient; No Information | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -7.5/1.5/79 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO OBSERVATION OF EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -7.50/1.5/079 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. LATER, A NEW SHORTER LENS WAS IMPLANTED ON (B)(6) 2018. THE PROBLEM WAS RESOLVED. (B)(4). "LENS WAS RETURNED DRY" NOT INCLUDED IN THE SUPPLEMENTAL MDR#1. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4) | -7.5/1.5/79 | OD | Physician | 08/08/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5549343&pc=MTA | 2023826-2016-00368 | 2016-03-05 | Malfunction | STAAR SURGICAL COMPANY | 2016-04-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected | No Consequences Or Impact To Patient; No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 --06.00/+3.50/X122 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS TORN WHILE IMPLANTING INTO THE EYE. THE ICL WAS REMOVED DURING THE SAME PROCEDURE. NO INJURY TO THE PATIENT. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -06.00/+3.50/X122 (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). DEVICE NOT RETURNED. | -06.00/+3.50/122 | OD | Physician | 03/08/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5248742&pc=MTA | 2023826-2015-01547 | 2015-10-15 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material; Unintended Movement | No Consequences Or Impact To Patient; No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.0/+4.5/061 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS REMOVED ON (B)(6) 2015 DUE TO THE LENS UNFOLDED UPSIDE DOWN. THE LENS TORE WHEN REMOVED . THERE WAS NO REPORTED PATIENT INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 AND THE STATUS OF THE EYE IS EXCELLENT. Manufacturer Narrative: CORRECTED DATA: USER FACILITY IS INCORRECT, SHOULD BE DISTRIBUTOR. (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -13.0/+4.5/061 | OD | Physician | 11/03/2015 | 2015 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5632998&pc=MTA | 2023826-2016-00541 | 2016-02-29 | Injury | STAAR SURGICAL COMPANY | 2016-05-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Failure to Unfold or Unwrap; Activation Failure | No Consequences Or Impact To Patient; No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -3.00/+1.50/X173 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE. THE LENS DID NOT FULLY UNFOLD IN THE EYE. LENS WAS REMOVED AND EXCHANGED WITH ANOTHER SAME LENS MODEL BUT SMALLER DIOPTER SIZE. NO INJURY TO THE PATIENT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. UNIDENTIFIED FOREIGN MATERIAL WAS PRESENT ON THE LENS SURFACE. AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.96MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: (B)(4). DEVICE NOT RETURNED. | -3.00/+1.50/173 | OS | Physician | 04/04/2016 | 2016 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7136376&pc=MTA | 2023826-2017-01935 | 2017-11-08 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient; No Patient Involvement | NA | Event Description: THE REPORTER INDICATED THE SURGEON WAS ATTEMPTING TO LOAD A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/080 (SPHERE/CYLINDER/AXIS), AND THE LENS BROKE. THERE WAS NO PATIENT CONTACT. THE LENS WAS EXCHANGED FOR ANOTHER ICL LENS ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: IN THE INITIAL MDR, (B)(6) 2017 CORRECTED TO (B)(6) 2017. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS), LENS MISSING A PIECE OF HAPTIC, AND LENS HAVING A CURVED SHAPE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+2.0/080 | NA | Physician | 11/27/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8073812&pc=MTA | 2023826-2018-01675 | 2018-08-02 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient; No Patient Involvement; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A VTICMO13.7, -8.00/.05/115 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING DELIVERY. IT WAS REPORTED THAT THERE WAS NO PATIENT CONTACT AND NO PATIENT INJURY. BACK UP LENS WAS USED AND DID NOT RESOLVE THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.00/.05/115 | NA | Physician | 10/16/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5055150&pc=MTA | 2023826-2015-01099 | 2015-08-02 | Injury | STAAR SURGICAL COMPANY | 2015-09-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Information | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.00/05/122 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. SURGERY IS PENDING TO EXPLANT AND EXCHANGE THE LENS DUE EXCESSIVE VAULTING. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S BUT NOT MARKETED IN THE U.S. DIOPTER: -14.00/05/122. NO CODE AVAILABLE (VAULTING). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WAS NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED. | -14.00/05/122 | OD | Physician | 08/06/2015 | 2015 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5066131&pc=MTA | 2023826-2015-01176 | 2015-08-05 | Injury | STAAR SURGICAL COMPANY | 2015-09-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Information | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -5.50/+1.00/X081 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD). EXCESSIVE VAULT WAS NOTED ONE WEEK POST-OP. A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION IS AVAILABLE. SEE MFR # 2023826-2015-01177 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -5.50/+1.00/X081. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED. | -5.50/+1.00/081 | OD | Physician | 08/11/2015 | 2015 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5991808&pc=MTA | 2023826-2016-01359 | 2016-07-04 | Injury | STAAR SURGICAL COMPANY | 2016-09-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Information | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -6.0/+2.5/150 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND SPOT ON SURFACE; MEASUREMENT WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.0/+2.5/150 | OS | Physician | 09/06/2016 | 2016 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5991827&pc=MTA | 2023826-2016-01361 | 2016-09-04 | Injury | STAAR SURGICAL COMPANY | 2016-09-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Information | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0/+3.0/122 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS REMAINS IMPLANTED BUT THE PLAN IS TO REMOVE AND EXCHANGE FOR A SMALLER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED ON (B)(6) 2016 AND WAS EXCHANGED FOR A SHORTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.0/+3.0/122 | OD | Physician | 09/09/2016 | 2016 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5991829&pc=MTA | 2023826-2016-01362 | 2016-09-04 | Injury | STAAR SURGICAL COMPANY | 2016-09-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Information | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+3.5/075 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS REMAINS IMPLANTED BUT THE PLAN IS TO REMOVE AND EXCHANGE FOR A SMALLER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016 AND WAS EXCHANGED FOR A SHORTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: PATIENT CODE (B)(4) - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CONCLUSION CODE AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) UNDER 3.0MM. CORRECTED DATA: THE CORRECT REPORT SOURCE COUNTRY IS (B)(6). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/+3.5/075 | OS | Physician | 09/09/2016 | 2016 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7855444&pc=MTA | 2023826-2018-01353 | 2018-03-22 | Injury | STAAR SURGICAL COMPANY | 2018-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Information | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -8.5/1.5/85 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO OBSERVATION OF EXCESSIVE VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/1.5/85 | OD | Physician | 08/08/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4922994&pc=MTA | 2023826-2015-00822 | 2015-05-18 | Injury | STAAR SURGICAL COMPANY | 2015-07-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD AN EXCESSIVE VAULT AND THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS REMAINS IMPLANTED BUT THE PLAN IS TO EXPLANT THE LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -3.0/+6.0/110 | OS | Physician | 06/17/2015 | 2015 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5375633&pc=MTA | 2023826-2016-00067 | 2015-11-04 | Injury | STAAR SURGICAL COMPANY | 2016-01-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50/+3.5/093 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD AN EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). DEVICE EVALUATION: THE PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL ON THE LENS SURFACE. DATE OF EVENT WAS INCORRECT. THE DATE OF EVENT IS (B)(6) 2015. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | -13.50/+3.5/093 | OD | Physician | 12/01/2015 | 2015
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4629107&pc=MTA | 2023826-2015-00383 |  | Injury | STAAR SURGICAL COMPANY | 2015-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAS AN EXCESSIVE VAULT AND REMAINS IMPLANTED. THE REPORTER INDICATED THE SURGEON PLANS TO EXPLANT THE LENS AND EXCHANGE FOR A SHORTER LENS. Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CAE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: (B)(4) - NO INFORMATION; SIZE INCORRECT FOR PATIENT; (VAULTING); (IMPLANTED). DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. METHOD: MEDICAL REVIEW. RESULTS: MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -16.00/+4.0/094 | OD | Physician | 03/06/2015 | 2015 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4516764&pc=MTA | 2023826-2015-00152 | 2014-11-05 | Injury | STAAR SURGICAL COMPANY | 2015-02-13 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Information | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND LOW VAULT WAS NOTED. THE LENS WAS ROTATED AND REPOSITIONED. THE SURGEON IS PLANNING ON EXPLANTING THE LENS FOR A LONGER LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PT WEIGHT: UNK. EXPLANT DATE: NA. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS IMPLANTED. Manufacturer Narrative: MEDICAL REVIEW - CLINICAL STUDIES HAVE SHOWN THAT TORIC ICL ROTATION OCCURS IN (B)(4) OF PATIENTS WHERE A TICL HAS BEEN IMPLANTED. SEVERAL FACTORS MAY CONTRIBUTE TO THIS PHENOMENON (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION ECT.). ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | -6.51+11.0/12.5 | OD | Physician | 01/19/2015 | 2015 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5240537&pc=MTA | 2023826-2015-01520 | 2015-10-20 | Injury | STAAR SURGICAL COMPANY | 2015-11-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Information | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -5.0/+4.5/100 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD A LOW VAULT AND LENS ROTATION. THE SURGEON IS PLANNING TO EXPLANT THE LENS AND EXCHANGE FOR A LONGER LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4). | -5.0/+4.5/100 | OS | Physician | 10/21/2015 | 2015 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4613564&pc=MTA | 2023826-2015-00327 | 2014-12-29 | Injury | STAAR SURGICAL COMPANY | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Information | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS ROTATED AND REMAINS IMPLANTED. THE REPORTER INDICATED THE ROTATION WAS NOT ASSOCIATED WITH A LOW VAULT. Event Description: ADDITIONAL INFORMATION RECEIVED - THE FACILITY REPORTED THE SURGEON REPOSITIONED THE LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S UCVA IS OPTIMAL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) -DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). METHOD: MEDICAL REVIEW. RESULTS: MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -12.5/+4.0/087 | OD | Physician | 02/17/2015 | 2015 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5478181&pc=MTA | 2023826-2016-00091 | 2015-11-17 | Injury | STAAR SURGICAL COMPANY | 2016-03-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Information; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, - 18.00/+2.50/X079 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. AFTER IMPLANTATION, REFRACTIVE SURPRISE WAS NOTED. THE ICL REMAINS IMPLANTED. LAST OFFICE VISIT ON (B)(6) 2015, BCVA: 20/20, UCVA: 20/25. SEE MFR. REPORT #2023826-2016-00092 FOR RIGHT EYE. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -18.00/+2.50/X079. DEVICE EVALUATED BY MANUFACTURER?: NO. ICL REMAINS IMPLANTED. (B)(4). EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). THE ICL REMAINS IMPLANTED. | - 18.00/+2.50/079 | OS | Physician | 12/24/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5478184&pc=MTA | 2023826-2016-00092 | 2015-11-17 | Injury | STAAR SURGICAL COMPANY | 2016-03-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Information; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.00/+3.00/X094 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. AFTER IMPLANTATION, REFRACTIVE SURPRISE WAS NOTED. THE ICL REMAINS IMPLANTED. LAST OFFICE VISIT ON (B)(6) 2015, BCVA: 20/20, UCVA: 20/30. SEE MFR. REPORT #2023826-2016-00091 FOR LEFT EYE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -17.00/+3.00/X094. DEVICE EVALUATED BY MANUFACTURER?: NO - ICL REMAINS IMPLANTED.(B)(4). CONCLUSIONS CODE(S): CONCLUSION- CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -17.00/+3.00/094 | OD | Physician | 12/24/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5991812&pc=MTA | 2023826-2016-01358 | 2016-07-04 | Injury | STAAR SURGICAL COMPANY | 2016-09-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Information; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -5.50/+4.5/150 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT OPTIC WAS TORN/ SPLIT, HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL (SPOT) ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.50/+4.5/150 | OD | Physician | 09/06/2016 | 2016 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7855987&pc=MTA | 2023826-2018-01355 | 2018-06-26 | Injury | STAAR SURGICAL COMPANY | 2018-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Information; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -15/1.5/79 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO OBSERVATION OF EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DEFORMED, WITH FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. LENS HAPTIC WAS CAUGHT IN VIAL CAP. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO_12.6, -15/1.5/79 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. A SHORTER REPLACEMENT LENS WAS IMPLANTED POST-OPERATIVELY ON (B)(6) 2018 TO RESOLVE THE ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4) | -15/1.5/79 | OS | Physician | 08/08/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4574376&pc=MTA | 2023826-2015-00264 | 2014-10-30 | Injury | STAAR SURGICAL COMPANY | 2015-03-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 AND REMOVED IT DUE TO LOW VAULT. THE LENS WAS NOT REPLACED. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVALUATION: EVALUATION METHOD: LENS WORK ORDER SEARCH & MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS WAS WITHIN SPECIFICATION. (B)(4). | -11.5+1.0/080 | OS | Other | 02/04/2015 | 2015 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5593300&pc=MTA | 2023826-2016-00486 | 2016-01-27 | Injury | STAAR SURGICAL COMPANY | 2016-04-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.5/078 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD AN EXCESSIVE VAULT AND REMAINS IMPLANTED. Manufacturer Narrative: THE PATIENT HAD A LENS EXCHANGE WITH A MULTIFOCAL IOL. NO IMPACT ON PATIENT'S VISION REPORTED. PREVIOUS REPORT SOURCE, USER FACILITY, IS INCORRECT. THE CORRECT REPORT SOURCE IS DISTRIBUTOR. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4). | -11.5/+2.5/078 | OS | Physician | 01/28/2016 | 2016 | 50
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5593301&pc=MTA | 2023826-2016-00485 | 2016-01-27 | Injury | STAAR SURGICAL COMPANY | 2016-04-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+0.5/080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD AN EXCESSIVE VAULT AND REMAINS IMPLANTED. Manufacturer Narrative: PATIENT HAD A LENS EXCHANGE WITH MULTIFOCAL IOL. THERE IS NO IMPACT ON PATIENT'S VISION REPORTED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). THE PATIENT IS OVER THE AGE OF (B)(6) YEARS OLD. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4). | -10.0/+0.5/080 | OD | Physician | 01/28/2016 | 2016 | 50
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5659927&pc=MTA | 2023826-2016-00664 | 2016-03-18 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+0.5/058 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND REMAINS IMPLANTED. IT IS PLANNED TO EXCHANGE FOR A LONGER LENS. Manufacturer Narrative: CORRECTED DATA: (B)(4). ADDITIONAL INFORMATION: DESCRIBE EVENT OR PROBLEM - THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A LONGER LENS. THIS RESOLVED THE PROBLEM. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20 AS OF (B)(6) 2016. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. WEIGHT: UNK. OUTCOMES ATTRIBUTED TO ADVERSE EVENTS: UNK. EXPLANT DATE: N/A. (B)(4). EVENT PROBLEM AND EVALUATION CODES: (B)(4). LENS REMAINS IMPLANTED. | -8.5/+0.5/058 | OD | Physician | 03/21/2016 | 2016 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5937655&pc=MTA | 2023826-2016-01238 | 2016-05-04 | Malfunction | STAAR SURGICAL COMPANY | 2016-09-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/1.50/89 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. IT WAS INDICATED THE PATIENT EXPERIENCED EXCESSIVE VAULTING, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.91MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4). | -10.50/1.50/89 | OD | Physician | 08/10/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5710824&pc=MTA | 2023826-2016-00672 | 2016-05-06 | Injury | STAAR SURGICAL COMPANY | 2016-06-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/+3.0/X160 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: THE LENS WAS EXPLANTED ON (B)(6) 2016. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED IN THE PRESENCE OF KERATOCONUS. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. EXPLANT DATE: UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. (B)(4). | -16.0/+3.0/160 | OS | Physician | 05/10/2016 | 2016 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5816236&pc=MTA | 2023826-2016-00970 | 2016-06-02 | Injury | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+0.5/115 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND REMAINS IMPLANTED. THE LENS IS PLANNED TO EXCHANGE FOR A LONGER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, 12.5/+0.5/115 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED TO HAVE A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS, HOWEVER THE ICL VAULT DID NOT IMPROVE. PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) IS 20/20 AND UNCORRECTED VISUAL ACUITY WAS 20/20. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. AS PER THE DFU OF THIS LENS MODEL, IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRADICTED. THIS PATIENT HAD AN ACD OF 2.95 MM AT THE TIME OF IMPLANTATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4). | -12.5/+0.5/115 | OD | Physician | 06/22/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5816119&pc=MTA | 2023826-2016-00972 | 2016-06-08 | Injury | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/179 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND REMAINS IMPLANTED. THE LENS IS PLANNED TO EXCHANGE FOR A SHORTER LENS. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THIS RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20 AND UNCORRECTED VISUAL ACUITY WAS 20/30. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4). | -10.5/+1.0/179 | OS | Physician | 06/24/2016 | 2016 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5938278&pc=MTA | 2023826-2016-01253 | 2016-07-03 | Injury | STAAR SURGICAL COMPANY | 2016-09-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/095 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND WILL BE EXPLANTED. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE REPORTER CORRECTED THE INITIAL IMPLANT DATE TO (B)(6) 2016. THE SURGEON NOTICED THE LOW VAULT ON (B)(6) 2016. THE REPORTER STATED THAT THE INITIAL LENS WAS EXCHANGED WITH A LONGER LENS (SN: (B)(4) ON (B)(6) 2016. THIS EXCHANGE RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED.. (B)(4). | -11.0/+2.0/095 | OD | Physician | 08/10/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5955103&pc=MTA | 2023826-2016-01262 |  | Injury | STAAR SURGICAL COMPANY | 2016-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT, WITH SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES.THE LENS REMAINS IMPLANTED BUT WILL BE REMOVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). LENS NOT RETURNED. | -9.5/+1.5/090 | OS | Physician | 08/17/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7856947&pc=MTA | 2023826-2018-01356 | 2018-03-02 | Injury | STAAR SURGICAL COMPANY | 2018-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Information; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -7.5/4/75 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO OBSERVATION OF LOW VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS SURFACE. CORRECTED DATA: ORIGINAL MDR STATED "ON 06/07/18 THE LENS WAS EXCHANGED" WHICH SHOULD BE CORRECTED TO "ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO OBSERVATION OF LOW VAULT WITH LENS ROTATION. ON (B)(6) 2018 A LONGER LENGTH LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED." (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.5/4/75 | OS | Physician | 08/08/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5132184&pc=MTA | 2023826-2015-01232 | 2016-05-04 | Injury | STAAR SURGICAL COMPANY | 2015-10-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Information; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+4.50/X077 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THE PATIENT EXPERIENCED EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, IRIS ATROPHY, PUPIL OVALIZATION, AND SYNECHIA. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED. | -18.00/+4.50/077 | OS | Physician | 09/11/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5596430&pc=MTA | 2023826-2016-00492 | 2016-02-10 | Injury | STAAR SURGICAL COMPANY | 2016-04-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Information; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/X139 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND LENS ROTATION. THE SURGEON PLANS TO EXPLANT AND EXCHANGE FOR A LONGER LENS. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION WAS RECEIVED: THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A LONGER LENS WHICH RESOLVED THE PROBLEM. ON PATIENT'S LAST VISIT (B)(6) 2016, THE PATIENT'S UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. CONCLUSION: REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS PATIENT'S AGE BELOW 21 OR OVER 45 YEARS. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM #(B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | -11.5/+1.5/139 | OD | Physician | 02/11/2016 | 2016 | 50
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5909069&pc=MTA | 2023826-2016-01187 | 2016-07-25 | Injury | STAAR SURGICAL COMPANY | 2016-08-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Information; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+0.5/165 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT, LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2016 FOR A SHORTER LENS OF THE SAME MODEL AND SIMILAR DIOPTER WHICH RESOLVED THE PROBLEM. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20 ON (B)(6) 2016. DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). LENS NOT RETURNED. | -11.0/+0.5/165 | OS | Physician | 07/27/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4789537&pc=MTA | 2023826-2015-00598 | 2015-04-11 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-05-21 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED THE 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IMPLANTED UPSIDE DOWN IN THE PATIENT'S RIGHT (OD) EYE. THE LENS WAS EXCHANGE FOR ANOTHER SAME MODEL LENS AND NO PATIENT INJURY WAS REPORTED. PATIENT'S LAST BCVA WAS 20/20 . Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). | -12.5/1.5/092 | OD | Physician | 04/22/2015 | 2015 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5575803&pc=MTA | 2023826-2016-00442 | 2015-12-24 | Malfunction | STAAR SURGICAL COMPANY | 2016-04-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -13.00/+2.50/X090 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED AND REPLACED WITH BACKUP LENS, SAME MODEL AND SIZE DURING THE SAM PROCEDURE. THIS RESOLVED THE PROBLEM. NO ADVERSE EVENT WAS REPORTED. LAST VISIT ON (B)(6) 2016, UCVA WAS 20/16. Manufacturer Narrative: (B)(4). DEVICE EVALUATION-LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED, AND LENS TEAR. CLAIM # (B)(4). Manufacturer Narrative: DIOPTER: -13.00/+2.50/X090. (B)(4). | -13.00/+2.50/090 | OS | Physician | 03/15/2016 | 2016 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5766315&pc=MTA | 2023826-2016-00872 | 2016-06-26 | Malfunction | STAAR SURGICAL COMPANY | 2016-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.00/-2.50/170 DIOPTER, AND THE LENS TORE/BROKE WHILE INJECTING. THE REPORTER INDICATED THE HAPTIC WAS CATCHING ON THE FOAM TIP PLUNGER. THERE WAS NO PATIENT CONTACT. THE PATIENT'S BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. "THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO INSERT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12/+2.00/080 DIOPTER, AND THE LENS TORE/BROKE WHILE INJECTING. THE REPORTER INDICATED THE HAPTIC WAS CATCHING ON THE FOAM TIP PLUNGER. THERE WAS NO PATIENT CONTACT. THE PATIENT'S BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/20." NO MATERIAL INTEGRITY ISSUE REPORTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. CLAIM #(B)(4). LENS NOT RETURNED. | -9.00/-2.50/170 | NA | Physician | 06/07/2016 | 2016 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5613015&pc=MTA | 2023826-2016-00527 | 2016-02-01 | Malfunction | STAAR SURGICAL COMPANY | 2016-04-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.0/+3.0/085 DIOPTER, AND THE LENS TORE/BROKE. THE LENS WAS REMOVED AND THERE WAS NO REPORT OF ANY PATIENT INJURY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND THE OPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4). | -5.0/+3.0/085 | NA | Physician | 02/09/2016 | 2016 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8048832&pc=MTA | 2023826-2018-01600 | 2018-10-09 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Markings/Labelling Problem | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -09.00/1.5/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. IT WAS REPORTED THAT THE INSIDE PAITENT LABEL HAD AN INCORRECT SERIAL NUMBER, THE LENS WAS IMPLANTED, THE AXIS IS ALIGNED ACCORDING TO THE LABEL ON THE BOX NOT THE INSIDE LABEL. THE PATIENT OUT COME IS GOOD. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -09.00/1.5/107 | OS | Physician | 10/09/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8122178&pc=MTA | 2023826-2018-01667 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-12-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Markings/Labelling Problem | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A VTICMO 12.6 -9.50/+1.50/102, SERIAL (B)(4), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). IT WAS REPORTED THAT THE PATIENT CHART LABEL PROVIDED INSIDE THE PACKAGING BOX SHOW DISCREPANT LENS INFORMATION AS TO VTICMO 12.6 -9.50/+1.50/100, SERIAL # (B)(4). THE REPORTER STATES "THE DOCTOR FOUND THE LENS POWER TO BE SIMILAR TO WHAT HE ORDERED, SO HE IMPLANTED ACCORDINGLY". THE LENS REMAINS IMPLANTED AND IT WAS REPORTED THAT THERE IS NO NEGATIVE FEEDBACK OR COMPLAINTS ABOUT THE PATIENT'S CONDITION. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.50/+1.50/102 | OS | Non-Healthcare Professional | 10/31/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8122361&pc=MTA | 2023826-2018-01668 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-12-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Markings/Labelling Problem | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A VTICMO 12.6 -12.00/+1.50/083, SERIAL (B)(4), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). IT WAS REPORTED THAT THE PATIENT CHART LABEL PROVIDED INSIDE THE PACKAGING BOX SHOW DISCREPANT LENS INFORMATION AS TO VTICMO 12.6 -12.00/+1.5/082, SERIAL # (B)(4). THE REPORTER STATES "THE DOCTOR FOUND THE LENS POWER TO BE SIMILAR TO WHAT HE ORDERED, SO HE IMPLANTED ACCORDINGLY". THE LENS REMAINS IMPLANTED AND IT WAS REPORTED THAT THERE IS NO NEGATIVE FEEDBACK OR COMPLAINTS ABOUT THE PATIENT'S CONDITION. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/+1.50/083 | OS | Non-Healthcare Professional | 10/31/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8122438&pc=MTA | 2023826-2018-01669 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-12-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Markings/Labelling Problem | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A VTICMO 12.6 -9.50/+1.50/100, SERIAL (B)(4), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD). THE REPORTER STATES THAT THE SAME MODEL, POWER AND SERIAL NUMBER OF THIS LENS (VTICMO 12.6 -9.50/+1.50/100, SERIAL # (B)(4)) WAS SHOWN ON THE PATIENT CHART LABEL INSIDE THE BOX FOR ANOTHER LENS WITH VTICMO 12.6 -9.50/+1.50/102, SERIAL # (B)(4) INDICATED ON THE EXTERNAL BOX LABEL. THE LENS REMAINS IMPLANTED AND IT WAS REPORTED THAT THERE IS NO NEGATIVE FEEDBACK OR COMPLAINTS ABOUT THE PATIENT'S CONDITION. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.50/+1.50/100 | OD | Non-Healthcare Professional | 10/31/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6812227&pc=MTA | 2023826-2017-01342 | 2014-09-18 | Malfunction | STAAR SURGICAL COMPANY | 2017-08-22 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2014, AS THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -18.0/+2.0/127 INTO THE PATIENT'S EYE, HE DISCOVERED THE LENS BROKE. THE SURGEON STATED THAT THE HAPTICS OF THE LENS WERE BROKEN ON ONE SIDE WHEN THE LENS CAME OUT OF THE "SHOOTER." THE LENS WAS NOT IMPLANTED AND A NEW LENS WAS ORDERED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.0/+2.0/127 | NA | Physician | 07/24/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6509194&pc=MTA | 2023826-2017-00605 | 2015-04-02 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-20 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON PREPARED TO IMPLANT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.50/+3.5/083 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS), IT BECAME "UNSTABLE." A DIFFERENT LENS OF THE SAME TYPE AND SIZE AND SIMILAR DIOPTER WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. IT IS UNKNOWN IF THERE WAS PATIENT CONTACT; TO DATE, ATTEMPTS AT CUSTOMER CONTACT HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN THE LENS CASE/VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. METHOD CODES ADDED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CONCLUSION CODE: OFF-LABEL USE [ACD < 3.0MM (2.9MM)]. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.50/+3.5/083 | OS | Physician | 09/12/2015 | 2015
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6500531&pc=MTA | 2023826-2017-00657 | 2015-10-31 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON TRIED TO IMPLANT A 13.2MM VTICMO13.2 -12.5/1.5/082 DIOPTER IMPLANTABLE COLLAMER LENS. THE LENS TORE/BROKE BEFORE INJECTION INTO THE EYE. LENS WAS NOT IMPLANTED. NO PATIENT INJURY REPORTED. Manufacturer Narrative: THIS IS A PRODUCT PROBLEM NOT AN ADVERSE EVENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)46) | -12.5/1.5/082 | NA | Physician | 12/11/2015 | 2015 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6522664&pc=MTA | 2023826-2017-00716 | 2016-03-20 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | Event Description: THE SURGEON ATTEMPTED TO INSERT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS WITH A -10.50/+4.5/075 DIOPTER IN THE PATIENT'S LEFT EYE (OS) AND THE LENS TORE/BROKE BEFORE INJECTION IN THE EYE. LENS WAS NOT IMPLANTED. NO PATIENT CONTACT. Manufacturer Narrative: MANUFACTURED IN THE U.S., BUT NOT MARKETED. PRODUCT EVALUATION FOUND LENS RETURNED IN LIQUID IN THE LENS CONTAINER. VISUAL INSPECTION FOUND OPTIC TORN/SPLIT AND HAPTIC TORN/BROKEN/BENT/DEFORMED, AND FOREIGN MATERIAL ON LENS SURFACE. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.50/+4.5/075 | OS | Physician | 04/26/2016 | 2016 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6327785&pc=MTA | 2023826-2017-00237 | 2016-04-06 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12/+1/84 DIOPTER, WAS TORN BEFORE INJECTION INTO THE EYE. THE LENS WAS NOT IMPLANTED AND THERE WAS NO PATIENT CONTACT. A BACK-UP LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. DURING THE PATIENT'S LAST VISIT ON (B)(6) 2016, THE PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN. (B)(4). | -12/+1/84 | NA | Physician | 04/14/2016 | 2016 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6330729&pc=MTA | 2023826-2017-00239 | 2016-04-06 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6/2/084 DIOPTER, TORE/BROKE BEFORE INJECTION/DELIVERY IN THE PATIENT'S LEFT EYE (OS). THE SURGEON STATED "EVERYTHING IS FINE, PATIENT IS HAPPY." ON (B)(6) 2016, PATIENT'S UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN. CONCLUSION: THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4). | -6/2/084 | OS | Physician | 04/21/2016 | 2016 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6566822&pc=MTA | 2023826-2017-00820 | 2016-05-11 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.5/2.0/84 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS TORE/BROKE BEFORE INJECTION INTO THE EYE. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN/BENT/DEFORMED. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17.5/2.0/84 | OD | Physician | 07/01/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6605330&pc=MTA | 2023826-2017-00918 | 2016-06-09 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/3.5/89 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS TORE/BROKE DURING LOADING INTO THE CARTRIDGE. THE LENS WAS NOT IMPLANTED AND THERE WAS NO PATIENT CONTACT. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT, THE HAPTIC TORN/BROKEN/BENT/DEFORMED, SCRATCHES ON THE LENS SURFACE, AND A BLUE SPOT. (B)(4), VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN PATIENTS UNDER THE AGE OF 21 YEARS OLD. (B)(4). | -12.00/3.5/89 | OS | Physician | 07/19/2016 | 2016 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6125978&pc=MTA | 2023826-2016-01674 | 2016-07-11 | Injury | STAAR SURGICAL COMPANY | 2016-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.5/+3.5/84 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016 AND THE LENS WAS NOTED TO BE TORN. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A LONGER LENS THE SAME MODEL BUT A DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: EVENT WAS INCORRECTLY CLASSIFIED AS A PRODUCT PROBLEM. EVENT WAS INCORRECTLY CLASSIFIED AS A MALFUNCTION. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+3.5/84 | OS | Physician | 10/25/2016 | 2016 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6365375&pc=MTA | 2023826-2017-00284 | 2016-08-29 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.50/2.5/90 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THE LENS BROKE DURING INJECTION AND WAS EXCHANGED FOR THE SAME MODEL AND DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS TORN INTO TWO PIECES, TORN THROUGH THE HAPTICS AND OPTIC . A PIECE OF ONE HAPTIC WAS TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4). | -8.50/2.5/90 | OD | Physician | 02/01/2017 | 2017 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6590441&pc=MTA | 2023826-2017-00821 | 2016-08-30 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.5/2.0/103 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS TORE/BROKE BEFORE INJECTION INTO THE EYE (BEFORE CORNEAL INCISION). THE SURGERY WAS DEFERRED, NO DAMAGE ON PATIENT SIDE. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT, HAPTIC TORN/BROKEN/BENT/DEFORMED, AND FOREIGN MATERIAL ON LENS SURFACE (NOT SPECIFIED). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17.5/2.0/103 | OD | Physician | 09/21/2016 | 2016 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6260240&pc=MTA | 2023826-2017-00102 | 2016-09-18 | Malfunction | STAAR SURGICAL COMPANY | 2017-01-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -9.50/+3.0/082 DIOPTER, TORE/BROKE BEFORE INJECTION INTO THE EYE. THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.50/+3.0/082 | NA | Physician | 12/20/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6214842&pc=MTA | 2023826-2016-01846 | 2016-11-16 | Malfunction | STAAR SURGICAL COMPANY | 2016-12-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON WAS ATTEMPTING TO IMPLANT A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12/1.5/103 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) AND IT WAS NOTED THAT THE LENS WAS TORN. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID AND IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -12/1.5/103 | OD | Physician | 11/29/2016 | 2016 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6246925&pc=MTA | 2023826-2017-00062 | 2016-11-21 | Malfunction | STAAR SURGICAL COMPANY | 2017-01-12 | MTA | INTRAOCULAR LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.5/+3.0/080 DIOPTER, HAD TORN/BROKEN BEFORE INJECTION INTO THE LEFT (OS) EYE. THIS OCCURRED ON (B)(6) 2016. THE LENS WAS NOT IMPLANTED. AS OF THE DATE OF MDR SUBMISSION, THIS LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.50/+3.0/080 DIOPTER, HAD TORN/BROKEN BEFORE INJECTION INTO THE LEFT (OS) EYE. THIS OCCURRED ON (B)(6) 2016. THERE WAS NO PATIENT CONTACT AND THE LENS WAS NOT IMPLANTED. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: DATE RECEIVED BY MANUFACTURER ON INITIAL REPORT SHOULD BE 12/02/2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | -13.5/+3.0/080 | OS | Physician | 12/02/2016 | 2016 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6432865&pc=MTA | 2023826-2017-00491 | 2016-11-22 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+3.0/078 DIOPTER, RUPTURED DURING THE LOADING PROCESS AND BECAME STUCK IN THE CARTRIDGE. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A PARTIAL PIECE OF LENS RETURNED. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. | -13.0/+3.0/078 | NA | Physician | 02/22/2017 | 2017 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6234309&pc=MTA | 2023826-2017-00038 | 2016-11-24 | Malfunction | STAAR SURGICAL COMPANY | 2017-01-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/080 DIOPTER, CRACKED DURING INJECTION/DELIVERY INTO THE EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON THE SAME DAY. THE LENS WAS EXCHANGED FOR ANOTHER LENS AND THE PROBLEM WAS RESOLVED. THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: DEVICE EVALUATION: THE EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN AND BENT. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.0/080 | NA | Physician | 12/16/2016 | 2016 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6447914&pc=MTA | 2023826-2017-00540 | 2016-12-07 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6/+4.5/093 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). ON (B)(6) 2016 THE LENS WAS TORN BEFORE INJECTION INTO THE EYE. THE LENS WAS NOT IMPLANTED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BE RETURNED YET. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS AND A CARTRIDGE WERE RETURNED IN CARTRIDGE CASE INSIDE A BAG. THE LENS WAS RETURNED DRY AND VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS AND RESIDUE ON LENS SURFACE. THE CARTRIDGE WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE ON PRODUCT WITH NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6/+4.5/093 | OD | Physician | 03/03/2017 | 2017 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6619016&pc=MTA | 2023826-2017-00942 | 2017-01-25 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -2.0/+6/86 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS GOT TORN BEFORE INJECTION INTO THE EYE, LENS WAS NOT IMPLANTED. ON THE SAME DAY, THE BACK-UP LENS WITH LONGER SIZE 13.2MM WAS USED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS RETURNED IN LENS VIAL WITH LIQUID. VISUAL INSPECTION FOUND MISSING PIECE OF HAPTIC. NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -2.0/+6/86 | OS | Physician | 05/12/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6619902&pc=MTA | 2023826-2017-00922 | 2017-03-30 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON ATTEMPTED TO IMPLANT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/030 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD), IT TORE WHILE LOADING. THIS OCCURRED ON (B)(6) 2017. THE LENS WAS NOT IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION-LENS WAS RETURNED DRY IN THE LENS CASE/VIAL, WITH DEBRIS AND CLEAR AND WHITE RESIDUE COVERING THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. VIAL WAS SEALED WITH TAPE INSTEAD OF A STOPPER. NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.5/+1.0/030 | OD | Physician | 05/08/2017 | 2017 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6557223&pc=MTA | 2023826-2017-00779 | 2017-04-06 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+0.5/095 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS TORN/BROKEN BEFORE INJECTION INTO THE EYE. THE LENS WAS NOT IMPLANTED. THE SURGEON SAID THE LEADING PART, HELD BY THE FORCEPS, OF THE LENS WAS TORN IN THE BEGINING OF SUPER FUNNEL CARTRIDGE DURING LOADING PROCESS. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THE CLAIM IS A PRODUCT PROBLEM, NOT AN ADVERSE EVENT. THE CLAIM CONCERNS A MALFUNCTION, NOT A SERIOUS INJURY. INITIAL MDR STATES (B)(6) 2017. (B)(6) 2017 IS THE CORRECT DATE.(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/+0.5/095 | OD | Physician | 04/11/2017 | 2017 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6637170&pc=MTA | 2023826-2017-00963 | 2017-04-11 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+3.0/84 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) "201". IN THE SAME SURGERY, THE LENS WAS EXCHANGED DUE TO LENS TEAR BREAK DURING INJECTION INTO THE EYE. THE PROBLEM WAS RESOLVED. THE REPORTER STATED THERE WAS A RESISTANCE WHILE PUSHING THE LENS WITH FOAM TIP PLUNGER BECAUSE OF "OVD'S SHORTAGE" (NOT ENOUGH VISCOELASTIC). AT THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION - PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A SMALL VIAL AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND MISSING PIECE OF HAPTIC AND CLEAR RESIDUE ON LENS SURFACE. THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -13/+3/84 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON "(B)(6) 201" IS INCORRECT DATE; CORRECTED DATE IS (B)(6) 2017. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CONCLUSION: BELOW 21 YEARS OF AGE AT THE TIME OF IMPLANT. CLAIM#(B)(4). | -13.0/+3.0/84 | OS | Physician | 05/18/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6721298&pc=MTA | 2023826-2017-01179 | 2017-04-11 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5/3.0/91 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. A LENS TEAR/BREAK OCCURED BEFORE INJECTION INTO THE EYE. A LENS CRACK OCCURED. AN ALTERNATE LENS WAS IMPLANTED ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/16 AND THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/16. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE) AND THE LENS WAS COMPLETELY ROLLED UP. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5/3.0/91 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. A LENS TEAR/BREAK OCCURED BEFORE INJECTION INTO THE EYE. AN ALTERNATE LENS WAS IMPLANTED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/16 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/16. (B)(4). Manufacturer Narrative: PMA 510(K): - THE PRODUCT IS MANUFACTURED IN THE U.S, BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/3.0/91 | OS | Physician | 07/02/2017 | 2017 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6670145&pc=MTA | 2023826-2017-01028 | 2017-05-25 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+3/94 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. IN THE SAME SURGERY, THE LENS WAS EXPLANTED DUE TO LENS TEAR/BEAK DURING INJECTION INTO THE EYE. THE CUSTOMER STATED THAT HE THOUGHT THE INADEQUATELY HYDRATED FOAM TIP IN BSS IS CONTRIBUTING FACTOR TO THE LENS TEAR. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIALS ON LENS SURFACE (CLEAR RESIDUE) AND LENS WAS RETURNED IN A CURVED SHAPE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/+3/94 | OD | Physician | 05/31/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6737269&pc=MTA | 2023826-2017-01235 | 2017-06-27 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-24 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE. THE SURGEON "HAD TO USE THE BACKUP LENS FOR THE PHAKIC PATIENT IN THEATRE YESTERDAY. THE FIRST LENS TORN ON APPLICATION." Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -15.5/2.5/069 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON "HAD TO USE THE BACKUP LENS FOR THE PHAKIC PATIENT IN THEATRE YESTERDAY." THE FIRST LENS TORE ON APPLICATION ON (B)(6) 2017 AND HAD TO BE EXCHANGED WITH THE BACK UP LENS OF THE SAME MODEL AND DIOPTER. THE PROBLEM WAS RESOLVED. THE POST-OP STATED THAT THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) AND UNCORRECTED VISUAL ACUITY (UCVA) WERE BOTH FOUND TO BE 20/20. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID INSIDE A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND BROKEN. (B)(4). Manufacturer Narrative: THE FIRST LENS TORE ON APPLICATION ON (B)(6) 2017 AND HAD TO BE EXCHANGED WITH THE BACK-UP LENS OF THE SAME MODEL AND DIOPTER INTRAOPERATIVELY. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.5/2.5/069 | OS | Physician | 06/28/2017 | 2017 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7335041&pc=MTA | 2023826-2018-00342 | 2017-07-28 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICM5_13.7, -9.50/+4.0/80 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. PATIENT WAS DISSATISFIED WITH VISUAL ACUITY AS IT IS 20/25. THE SURGEON IS CURRENTLY CHECKING ON THE PATIENT'S CONDITION. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.50/+4.0/80 | OD | Physician | 02/13/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7089422&pc=MTA | 2023826-2017-01828 | 2017-08-11 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANTED A 12.6 MM VTICMO12.6, -15.0/+1.5/168 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. WHILE LOADING THE LENS IN THE CARTRIDGE, THE LENS GOT TORN. THE SURGEON ABORTED THE SURGERY. ON (B)(6) 2017, ANOTHER LENS WAS IMPLANTED AND IT WAS SUCCESSFUL WITH THE PATIENT DOING WELL. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL AND HAD CLEAR SURGICAL RESIDUE/ DEBRIS ON THE VIAL. VISUAL INSPECTION FOUND THE LENS MISSING PIECES OF ITS HAPTIC AND PRESENCE OF CLEAR RESIDUE ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.0/+1.5/168 | OS | Physician | 11/07/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438478&pc=MTA | 2023826-2018-00539 | 2017-11-17 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -11.5/2.5/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS INTRAOPERATIVELY EXPLANTED DUE TO THE LENS BEING TORE DURING INJECTION. THE PATIENT IS FINE AND DOES NOT WANT TO HAVE ANOTHER OPERATION. THE REPORTER BELIEVES THAT LACK OF HYDRATION MIGHT CONTRIBUTE TO THE TEAR. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND THE TORN PIECE MISSING. CORRECTED INFORMATION: LENS WAS IMPLANTED ON (B)(6) 2017. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: LENS WAS IMPLANTED ON (B)(6) 2017. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/2.5/89 | OD | Physician | 03/20/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7571616&pc=MTA | 2023826-2018-00854 | 2018-01-23 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-05 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/093 (SPHERE/CYLINDAR/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED INTRAOPERATIVELY, THE SURGEON NOTING LENS TEAR/BREAK DURING INJECTION INTO THE EYE AND "BROKEN LENS." THIS OCCURRED ON (B)(6) 2018. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE/DEBRIS ON LENS). CORRECTED DATA: REPORT SOURCE: USER FACILITY IS INCORRECT, SHOULD BE DISTRIBUTOR. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/+1.5/093 | OS | Physician | 05/11/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7378406&pc=MTA | 2023826-2018-00471 | 2018-02-01 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-28 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2018. INTRAOPERATIVELY THE LENS WAS EXPLANTED. ON (B)(6) 2018 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF ITS HAPTIC AND PRESENCE OF RESIDUE ON THE LENS SURFACE. CORRECTED DATA: TYPE OF REPORTABLE EVENT: IT IS CORRECTED FROM "SERIOUS INJURY" TO "MALFUNCTION" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.0/095 | OD | Physician | 03/07/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7410050&pc=MTA | 2023826-2018-00516 | 2018-03-08 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO LOAD A 12.6 MM VTICMO12.6, -11.5/+2/85 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE CARTRIDGE ON (B)(6) 2018 WHEN THE LENS TORE. THE LENS WAS BEING SLID SOFTLY WITH THE FORCEPS THRU THE NARROWEST PART OF THE CARTRIDGE WHEN IT SUDDENLY TORE FROM ONE OF THE CENTRAL HOLES TO THE HOLE IN THE DISTAL HAPTIC AND TO THE DISTAL LEFT HAPTIC. A REPLACEMENT LENS IS PLANNED FOR THE PATIENT. THE REPORTED BELIEVED THAT THE LENS HAD A LITTLE SCRATCH FROM MANUFACTURER. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL OBSERVATION FOUND THE LENS WITH THE HAPTIC TORN AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). WORK ORDER SEARCH: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/+2/85 | NA | Physician | 03/13/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7487578&pc=MTA | 2023826-2018-00583 | 2018-03-26 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON FOUND A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/091 (SPHERE/CYLINDER/AXIS), WAS CRACKED AFTER OPENING THE LENS VIAL. THE LENS WAS NOT IMPLANTED AND THERE WAS NO PATIENT CONTACT. THERE WAS NO MEDICAL OR SURGICAL INTERVENTION NECESSARY AND THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND A PIECE OF ONE HAPTIC MISSING. (B)(4). | -10.50/+1.5/091 | NA | Physician | 04/09/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7484264&pc=MTA | 2023826-2018-00642 | 2018-03-27 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-03 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON PLANNED TO IMPLANT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/139 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2018. HOWEVER, PRIOR TO IMPLANT, THE HAPTIC TORE. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON (B)(6) 2018. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND THE TORN HAPTIC PIECE IS MISSING. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+1.5/139 | OD | Physician | 04/05/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7535059&pc=MTA | 2023826-2018-00775 | 2018-04-08 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-22 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2018. THE LENS WAS REMOVED INTRAOPERATIVELY. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND THE LENS MISSING THE PIECE OF TORN PIECE. THERE WAS ALSO PRESENCE OF RESIDUE ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.0/+2.0/090 | OD | Physician | 04/27/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7535110&pc=MTA | 2023826-2018-00777 | 2018-04-09 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-22 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2018. THE LENS WAS REMOVED INTRAOPERATIVELY. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/+2.5/095 | OD | Physician | 04/27/2018 | 2018 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7617159&pc=MTA | 2023826-2018-00924 | 2018-04-19 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-19 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+3.0/095 (SPHERE/CYLINDAR/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY ON (B)(6) 2018. REPORTEDLY, THE LENS GOT STUCK IN THE PLUNGER AND THE HAPTIC TORE. ON (B)(6) 2018 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17.0/+3.0/095 | OS | Physician | 05/23/2018 | 2018 | 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7639946&pc=MTA | 2023826-2018-00954 | 2018-05-18 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-26 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+1.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), HE DISCOVERED THAT THE LENS WAS UPSIDE DOWN. REPORTEDLY, THE SURGEON "OVERFILLED" VISCOELASTIC INTO THE ANTERIOR CHAMBER. WHEN THE SURGEON EXTRACTED THE LENS, IT TORE. THIS OCCURRED ON (B)(6) 2018. THE LENS WAS REMOVED INTRAOPERATIVELY. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS A SMALL PIECE OF THE LENS RETURNED. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+1.0/095 | OD | Physician | 05/29/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6226914&pc=MTA | 2023826-2017-00024 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-01-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+1.0/087 DIOPTER, TORE/BROKE BEFORE INJECTION INTO THE EYE. THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT OPTIC WAS TORN, HAPTIC WAS BROKEN AND FOREIGN MATERIAL ON LENS SURFACE. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.0/+1.0/087 | NA | Physician | 06/11/2016 | 2016
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6421840&pc=MTA | 2023826-2017-00440 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-03-21 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT, AS THE SURGEON PREPARED TO IMPLANT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+1.0/105 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD), THAT THE "LENS WAS RECEIVED DEFECTIVE." A BACK-UP LENS WAS USED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT YET BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN AND THE OTHER HAPTIC HAD A PIECE TORN OFF AND MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS WERE FOUND. (B)(4). DEVICE NOT RETURNED. | -17.5/+1.0/105 | OD | Physician | 02/20/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7378739&pc=MTA | 2023826-2018-00465 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICMO12.6, -12.00/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENT'S EYE. IT WAS NOTICED THAT THE LENS WAS TORN IN THE MIDDLE. THE SURGEON PLANS TO EXCHANGE THE LENS. ATTEMPTS AT OBTAINING ADDITIONAL INFORMATION HAVE NOT BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO [OR #] SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/2.5/090 | NA | Physician | 02/27/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7286002&pc=MTA | 2023826-2018-00187 | 2017-12-13 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -11.00/1.5/X088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. AFTER IMPLANTATION, LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS NOTICED AS WELL AS THE LENS BEING TORN DURING INJECTION. A LENS EXCHANGE SURGERY IS PLANNED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.00/1.5/088 | OS | Physician | 01/23/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7434208&pc=MTA | 2023826-2018-00556 | 2018-03-16 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Mechanical Jam | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/107 (SPHERE/CYLINDER/AXIS), AND THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THE SURGEON REPORTED THE LENS GOT JAMMED IN THE FOAM TIP PLUNGER WHEN IMPLANTING AND WAS TORN. THE LENS WAS REMOVED AND EXCHANGED DURING THE SAME SURGERY WITH NO APPARENT INJURY. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WAS REMOVED DURING INITIAL SURGERY. ANOTHER LENS WAS IMPLANTED AT A LATER DATE. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/+2.5/107 | NA | Physician | 04/02/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5667306&pc=MTA | 2023826-2016-00576 | 2015-02-10 | Malfunction | STAAR SURGICAL COMPANY | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.50/+2.0/094 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS TORE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO APPARENT INJURY. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED INFORMATION: (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED AND OPTIC TORN/SPLIT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -6.50/+2.0/094 | OD | Physician | 02/11/2016 | 2016 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5817530&pc=MTA | 2023826-2016-00921 | 2016-06-15 | Malfunction | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/133 DIOPTER, FROM THE VIAL, AND THE LENS WAS NOTED TO BE IN A TORN CONDITION. THE LENS WAS NOT IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS RETURNED DRY IN LENS/CASE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN/BENT. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+2.0/133 | NA | Physician | 06/24/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5996366&pc=MTA | 2023826-2016-01373 | 2016-08-25 | Malfunction | STAAR SURGICAL COMPANY | 2016-10-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS - -17.50/3.50/90 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS TEARING DURING INJECTION INTO THE EYE. THE SURGEON IMPLANTED THE BACK-UP LENS. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: CORRECTED DATA: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS ALSO CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TO BE BROKEN. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK . (B)(4). | -17.50/3.50/90 | OS | Physician | 09/06/2016 | 2016 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7542831&pc=MTA | 2023826-2018-00793 | 2017-12-29 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Entrapment of Device; No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+3.0/089 (SPHERE/CYLINDER/AXIS), AND THE LENS BECAME STUCK IN THE PLUNGER. THE LENS WAS REMOVED FROM THE EYE, WITHIN THE SAME SURGERY, WITH NO APPARENT INJURY. UPON REMOVING FROM THE CARTRIDGE, THE LENS TORE. THE SURGERY WAS ABORTED. ANOTHER SAME MODEL LENS WAS IMPLANTED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS SURFACE. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+3.0/089 | NA | Physician | 05/07/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7356140&pc=MTA | 2023826-2018-00430 | 2018-02-22 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-20 | MTA | IMPLANTABLE COLLAMER LENS | Improper or Incorrect Procedure or Method | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON DAMAGED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.0/+1.0/094 (SPHERE/CYLINDER/AXIS) DURING THE LOADING PROCESS. THE WAS DISCOVERED WHILE IMPLANTING THE LENS INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2018. THE LENS WAS REMOVED INTRAOPERATIVELY. REPORTEDLY, THE DAMAGE WAS DUE TO USER ERROR. THE SURGEON PLANS TO IMPLANT A REPLACEMENT LENS IN THE FUTURE. Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2018. THE PATIENT'S UCVA IS 20/20 1 DAY POST OP. THE PATIENT CONDITION IS GOOD. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/+1.0/094 | OD | Physician | 02/23/2018 | 2018 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4962172&pc=MTA | 2023826-2015-01007 | 2015-06-08 | Injury | STAAR SURGICAL COMPANY | 2015-08-03 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) AND LOW VAULT WAS OBSERVED. . THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S BUT NOT MARKETED IN THE U.S. (B)(4). LENS REMAINS IMPLANTED. (B)(4). | -10/+2.0/103 | OD | Physician | 07/09/2015 | 2015 | 47
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6721306&pc=MTA | 2023826-2017-01178 | 2016-11-24 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0/2.0/98 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED IMMEDIATELY ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. SEE CLAIM# (B)(4) FOR PROBLEM RESOLUTION. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED AND DEVICE EVALUATION IS EXPECTED, BUT NOT YET PERFORMED. Manufacturer Narrative: LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S, BUT NOT MARKETED IN THE U.S. CONCLUSION CODE: ACD < 3.0MM. (B)(4). | -10.0/2.0/98 | OS | Physician | 07/02/2017 | 2017 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6397791&pc=MTA | 2023826-2017-00378 | 2016-11-29 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/141 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. AT THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED, SURGEON PLANS TO EXCHANGE. Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.0/+2.0/141 | OD | Physician | 02/09/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6397794&pc=MTA | 2023826-2017-00379 | 2016-11-29 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/049 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. AT THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED, SURGEON PLANS TO EXCHANGE. Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.5/+1.5/049 | OS | Physician | 02/09/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6335875&pc=MTA | 2023826-2017-00185 | 2017-01-03 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -15.50/4.5/104 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. EXCESSIVE VAULTING OCCURED WITH A VAULT VALUE OF 1400 MICROMETERS. THE LENS IS PLANNED TO BE EXCHANGED. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -15.50/4.5/104 | OD | Physician | 01/24/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6335881&pc=MTA | 2023826-2017-00186 | 2017-01-03 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-16 | MTA | ICL (INTRAOCULAR COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.5/4.5/80 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. EXCESSIVE VAULTING OCCURRED WITH A VAULT VALUE OF (B)(4) MICROMETERS. THE LENS IS PLANNED TO BE EXCHANGED. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE'S WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). LENS NOT RETURNED. | -13.5/4.5/80 | OS | Physician | 01/24/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6447610&pc=MTA | 2023826-2017-00523 | 2017-01-18 | Injury | STAAR SURGICAL COMPANY | 2017-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+2.5/98 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. IN THE SAME SURGERY, THE LENS WAS EXCHANGED FOR SHORTER LENS DUE TO EXCSSIVE VAULTING. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISISON, THE LENS WAS LOST, UNABLE TO RETURNED. Manufacturer Narrative: ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SMALL VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE NO VISIBLE DAMAGE TO THE LENS AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/+2.5/98 | OD | Physician | 03/01/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6498936&pc=MTA | 2023826-2017-00651 | 2017-02-24 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+4.0/84 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. IN THE SAME SURGERY, THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.5/+4.0/84 | OS | Physician | 03/21/2017 | 2017 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6585683&pc=MTA | 2023826-2017-00869 | 2017-02-25 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+3.5/089 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON THE SAME DATE AND DURING THE SAME SURGERY, THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCING EXCESSIVE VAULT. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC BROKEN AND BENT. DIMENSIONAL INSPECTION FOUND LENS IS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NOT RETURNED TO MANUFACTURER. | -10.5/+3.5/089 | OD | Physician | 04/28/2017 | 2017 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6708264&pc=MTA | 2023826-2017-01170 | 2017-05-18 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+3.0/084 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON THE SAME DAY AND DURING THE SAME SURGERY, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE, AND CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+3.0/084 | OD | Physician | 06/14/2017 | 2017 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6694738&pc=MTA | 2023826-2017-01122 | 2017-06-02 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/92 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. IN THE SAME SURGERY, THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE DEVICE WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC SLIGHTLY BENT. THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/+2.0/92 | OD | Physician | 06/12/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7021464&pc=MTA | 2023826-2017-01673 | 2017-06-22 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -6.5/+1.0/94 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING DURING THE SAME SURGERY. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.5/+1.0/94 | OS | Physician | 10/18/2017 | 2017 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6899575&pc=MTA | 2023826-2017-01509 | 2017-07-24 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-27 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A VTICM5_13.2 DIOPTER -12.0/+2.0/081 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON THE SAME DATE, INTRAOPERATIVELY, THE LENS WAS EXCHANGED WITH A SMALLER, DIFFERENT DIOPTER LENS DUE TO EXCESSIVE VAULT. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC BENT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -12.0/+2.0/081 | OD | Physician | 09/01/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6899572&pc=MTA | 2023826-2017-01506 | 2017-08-18 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-27 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A VTICM5_13.2 DIOPTER -10.5/+2.0/091 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON THE SAME DATE, INTRAOPERATIVELY, THE LENS WAS EXCHANGED WITH A SMALLER, DIFFERENT DIOPTER LENS DUE TO EXCESSIVE VAULT. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WAS EXCHANGED FOR A SMALLER, SIMILAR DIOPTER LENS. DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BENT AND A PIECE OF THE HAPTIC MISSING. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT (20 YRS), ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.99MM)]. (B)(4). | -10.5/+2.0/091 | OD | Physician | 08/31/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7033907&pc=MTA | 2023826-2017-01737 | 2017-08-31 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -7.50/+1.0/089 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS REMOVED DURING THE SAME SURGERY DUE TO EXCESSIVE VAULTING, WITH NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20 AND THE PATIENT WAS SATISFIED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE. VISUAL INSPECTION FOUND A PIECE OF HAPTIC MISSING. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.50/+1.0/089 | OD | Physician | 10/16/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6918212&pc=MTA | 2023826-2017-01536 | 2017-09-07 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/X090 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. CLAIM REPORT STATES AN EXCESSIVE VAULT VALUE OF 1080 UM WAS NOTED, INCIDENT DATE OF (B)(6) 2017. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: IMPLANTED DATE: THE DATE OF IMPLANTATION WAS CORRECTED FROM "(B)(6) 2017" TO "(B)(6) 2017" IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: UPDATED INFORMATION: REPORTEDLY, THE PATIENT IS DOING WELL-NO HIGH IOP, NO ANGLE CLOSURE, PUPIL REACTION IS NORMAL. THE SURGEON HAS DECIDED TO NOT EXCHANGE THE LENS AND WILL MONITOR THE PATIENT'S STATUS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -12.0/+2.5/090 | OS | Physician | 09/08/2017 | 2017 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7165962&pc=MTA | 2023826-2018-00021 | 2017-12-05 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50/+0.5/049 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL BE MONITORED. Manufacturer Narrative: CORRECTED DATA: "ADVERSE EVENT" IS CORRECTED TO " PRODUCT PROBLEM". TYPE OF REPORTABLE EVENT: IT IS CORRECTED FROM "SERIOUS INJURY" TO "MALFUNCTION". CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.50/+0.5/049 | OD | Physician | 12/06/2017 | 2017 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7166019&pc=MTA | 2023826-2018-00022 | 2017-12-05 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.00/+0.5/152 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL BE MONITORED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.00/+0.5/152 | OS | Physician | 12/06/2017 | 2017 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438484&pc=MTA | 2023826-2018-00545 | 2017-12-12 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -9/2/117 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS INTRAOPERATIVELY REMOVED DUE TO EXCESSIVE VAULTING. ON (B)(6) 2017, A SHORTER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9/2/117 | OD | Physician | 03/20/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7300917&pc=MTA | 2023826-2018-00278 | 2018-01-02 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -9.50/+2.0/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. INTRA-OPERATIVELY THE LENS WAS EXCHANGED WITH A LONGER LENS, THE SURGEON NOTING LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF DEBRIS ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | -9.50/+2.0/081 | OS | Physician | 01/31/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7267522&pc=MTA | 2023826-2018-00211 | 2018-01-07 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+3.0/103 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON NOTES EXCESSIVE VAULT; INCIDENT DATE OF (B)(6) 2018. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: SHOULD HAVE BEEN RECORDED IN INITIAL MDR. SURGEON STATES THAT "PATIENT DEVELOPED A VERY HIGH VAULT POSTOPERATIVELY. WE NEED TO EXPLANT THE LENS AND EXCHANGE WITH A 0.5MM SMALLER OVERALL DIAMETER". CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/+3.0/103 | OD | Physician | 01/19/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438482&pc=MTA | 2023826-2018-00538 | 2018-02-22 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -17.50/+1.0/66 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS INTRAOPERATIVELY EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: REPORTER STATED THAT, IN HIS OPINION, THE CAUSE OF THE EVENT IS THAT STAAR RECOMMENDED AN INCORRECT LENS SIZE FOR THIS PATIENT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17.50/+1.0/66 | OD | Physician | 03/20/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8076044&pc=MTA | 2023826-2018-01652 | 2018-06-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+3.0/100 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE DOCTOR JUDGED THAT THE SIZE WAS UNSUITABLE. THERE WAS NO REPORT OF ANY APPARENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED ANOTHER SAME MODEL LENS BUT DIFFERENT AXIS WAS IMPLANTED AS EXCHANGE LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+3.0/100 | NA |  | 10/18/2018 | 2018 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7745144&pc=MTA | 2023826-2018-01162 | 2018-06-17 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -14.00/+2.5/139 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. EXCESSIVE VAULT WAS OBSERVED. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED, IT WAS REPORTED THAT THE SURGEON PLANS TO EXCHANGE THE LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.00/+2.5/139 | OS | Physician | 07/04/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7716093&pc=MTA | 2023826-2018-01106 | 2018-06-21 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/139 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS AN EXCHANGE LENS BUT THE PROBLEM WAS NOT RESOLVED, THE LENS STILL HAD EXCESSIVE VAULTING. THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT OR EXCHANGE THE LENS. SEE MFR. REPORT 2023826-2018-01098 FOR THE INITIAL LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -13.0/+2.5/139 | OD | Physician | 06/24/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824061&pc=MTA | 2023826-2018-01276 | 2018-07-20 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/079 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/+1.5/079 | OS | Physician | 08/02/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8404826&pc=MTA | 2023826-2019-00354 | 2018-07-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/+4.0/100 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.00/+4.0/100 | NA |  | 02/07/2019 | 2019 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7857657&pc=MTA | 2023826-2018-01366 | 2018-08-10 | Malfunction | STAAR SURGICAL COMPANY | 2018-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/078 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD A HIGH VAULT . THE LENS REMAINS IMPLANTED AND THE PATIENT IS BEING MONITORED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+1.0/078 | OS | Physician | 08/12/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8237904&pc=MTA | 2023826-2019-00079 | 2018-08-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/098 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY, DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+1.0/098 | NA |  | 12/11/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8002559&pc=MTA | 2023826-2018-01536 | 2018-09-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-10-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/171 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY, DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THERE WAS NO APPARENT INJURY. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND ONE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+1.0/171 | NA |  | 09/24/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8073867&pc=MTA | 2023826-2018-01687 | 2018-09-17 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS), REMOVED THE LENS AND EXCHANGED LENS FOR A LONGER LENGTH LENS WITHIN THE SAME SURGERY ON (B)(6) /2018, DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/1.0/092 | OS | Physician | 10/18/2018 | 2018 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8237900&pc=MTA | 2023826-2019-00078 | 2018-11-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/099 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY, DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED WITHIN THE SAME SURGERY FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+2.0/099 | NA |  | 12/11/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8227440&pc=MTA | 2023826-2019-00049 | 2018-11-26 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -2.0/+4.0/079 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. INTRAOPERATIVELY THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. THE SURGEON STATES THAT THE "SURGERY WAS COMPLEX AND [THE] DOCTOR FINALLY DECIDED [TO] NOT IMPLANT THE ICL. IN THEIR OPINION THE ICL PUSHED THE IRIS IN THE SURGERY BECAUSE THE ICL IS TOO BIG FOR THE EYE". SEE MFR REPORT 2023826-2019-00046 FOR INITIALLY IMPLANTED LENS. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. DIMENSIONAL INSPECTION WAS WITHIN LENS SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -2.0/+4.0/079 | OS | Physician | 12/13/2018 | 2018 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8272691&pc=MTA | 2023826-2019-00153 | 2018-12-24 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -3.00/+2.5/143 (SPHERE/CYLINDER/AXIS) WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING (THE LENS WAS PUSHING THE IRIS) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE SURGEON PLANS TO IMPLANT A SHORTER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO DEVICE (CALIPER MALFUNCTION). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION- THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION WAS WITHIN LENS SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -3.00/+2.5/143 | NA |  | 12/27/2018 | 2018 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7136389&pc=MTA | 2023826-2017-01939 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -3.0/+4.5/002 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THE LENS HAD A HIGH VAULT BUT THE OUTCOME WAS GOOD. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -3.0/+4.5/002 | OD | Physician | 11/28/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7346748&pc=MTA | 2023826-2018-00421 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-03-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND REMAINS IMPLANTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.0/+1.0/090 | OD | Physician | 02/19/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8071358&pc=MTA | 2023826-2018-01678 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-11-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: IT WAS REPORTED THAT A 13.7MM, VTICMO13.7, -17/+2.5/113 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT RIGHT EYE (OD) AS A REPLACEMENT LENS DUE TO LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM (VAULT WAS A BIT LOW, BUT IT WAS THE BIGGEST LENS). ON (B)(6) 2018 IT WAS REPORTED THAT THE PROBLEM RESOLVED, THE PATIENT IS OK AND NO ADDITIONAL TREATMENTS WILL BE PERFORMED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17/+2.5/113 | OD | Physician | 10/25/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8204316&pc=MTA | 2023826-2018-01946 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-12-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | Event Description: RIGHT EYE (OD): THE REPORTER STATED THAT WHEN THE VAULT OF THE LEFT EYE TURNED OUT TO BE TOO HIGH, THEY CHOSE THE LENGTH OF 12.6MM FOR THE RIGHT EYE (OD), THEN THE VAULT OF THE RIGHT EYE (OD) WAS STILL LOW ENOUGH (1000UM). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BENN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: RIGHT EYE (OD): THE REPORTER STATED THAT, "WHEN THE VAULT OF THE LEFT EYE TURNED OUT TO BE TOO HIGH, THEY CHOSE THE LENGTH OF 12.6MM FOR THE RIGHT EYE (OD), THEN THE VAULT OF THE RIGHT EYE (OD) WAS STILL LOW ENOUGH (1000UM)". THE DIOPTER FOR THE RIGHT EYE (OD) IS -11.00/+0.5/008 (SPHERE/CYLINDER/AXIS). NO OTHER INFORMATION HAS BEEN PROVIDED. CLAIM#: (B)(4). Manufacturer Narrative: NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/+0.5/008 | OD |  | 11/28/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6938692&pc=MTA | 2023826-2017-01547 | 2017-09-11 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -7.5/+4.0/60 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. LOW VAULTING AND LENS ROTATION WERE NOTICED. THE SURGEON DECIDED TO LEAVE THE LENS IN THE EYE AND WILL HAVE STRICT MONITORING OF THE VAULT. THE PATIENT'S VISION IS STABLE AND GOOD (20/20). Manufacturer Narrative: METHOD CODE: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CORRECTED DATA: DATE OF BIRTH IS CORRECTED FROM (B)(6) TO (B)(6). (B)(4). Manufacturer Narrative: CORRECTED DATA: THE TYPE OF EVENT IS CORRECTED FROM "ADVERSE EVENT" TO "PRODUCT PROBLEM". TYPE OF REPORTABLE EVENT: IT IS CORRECTED FROM "SERIOUS INJURY" TO "MALFUNCTION". CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.5/+4.0/60 | OS | Physician | 09/12/2017 | 2017 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7696908&pc=MTA | 2023826-2018-01045 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-07-17 | MTA | IMPLANTABLE COLLAMER LENS | Insufficient Information | No Known Impact Or Consequence To Patient | NA | Event Description: A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/100 (SPHERE/CYLINDER/AXIS) WAS RETURNED DAMAGED WITH NO ADDITIONAL INFORMATION PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE PTIC AND HAPTIC TORN AND THE HAPTIC BENT. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+1.5/100 | NA | Physician | 06/25/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7716197&pc=MTA | 2023826-2018-01103 | 2018-03-15 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/087 (SPHERE/CYLINDER/AXIS), AND THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED DURING THE SAME SURGERY, WITH NO APPARENT INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED ON (B)(6) 2018. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.0/+2.5/087 | NA | Physician | 06/25/2018 | 2018 | 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7720848&pc=MTA | 2023826-2018-01125 | 2018-03-22 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/097 (SPHERE/CYLINDER/AXIS), AND THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED DURING THE SAME SURGERY, WITH NO APPARENT INJURY. A DIFFERENT MODEL LENS WAS IMPLANTED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -18.0/+1.5/097 | NA | Physician | 06/25/2018 | 2018 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7720866&pc=MTA | 2023826-2018-01126 | 2018-04-20 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/096 (SPHERE/CYLINDER/AXIS), AND THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. IT IS UNKNOWN IF THERE WAS ANY PATIENT CONTACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.0/+1.0/096 | NA | Physician | 06/25/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7783504&pc=MTA | 2023826-2018-01205 | 2018-06-25 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-14 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE. THIS OCCURRED ON (B)(6) 2018. THE LENS WAS REMOVED INTRA-OPERATIVELY. ON (B)(6) 2018 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. THE REPORTER STATES THAT IT'S POSSIBLE THE "DEHYDRATION OF THE FOAM TIP" CAUSED THE TEAR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.5/+2.0/093 | OD | Physician | 07/20/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8081331&pc=MTA | 2023826-2018-01704 | 2018-09-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+3.5/149 (SPHERE/CYLINDER/AXIS), WAS TORN/BROKEN BEFORE INJECTION INTO THE EYE. THERE WAS NO REPORT OF PATIENT CONTACT OR INJURY. A SHORTER LENS WAS IMPLANTED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL OBSERVATION FOUND THE OPTIC TORN AND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.5/+3.5/149 | NA |  | 10/22/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8332643&pc=MTA | 2023826-2019-00258 | 2018-12-19 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -16/1.5/14 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME SURGERY DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS USER ERROR. ADDITIONAL INFORMATION HAS BEEN REQUESTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC AND OPTIC WERE TORN. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16/1.5/14 | OD |  | 01/17/2019 | 2019 | 53
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6441308&pc=MTA | 2023826-2017-00518 | 2017-03-07 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.50/082 DIOPTER, AND THE LENS WAS NOTED TO BE FRONT TO BACK UPON INSERTION. THE HAPTIC BROKE WHILE BEING REMOVED DURING THE SAME SURGERY. THERE WAS NO APPARENT PATIENT INJURY. THE SURGERY WAS CANCELLED AND ANOTHER LENS WILL BE IMPLANTED AT A LATER DATE. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/400. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF ONE LENS HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE PATIENTS EYE ONE DAY POST-OP WAS SORE, WITH REDNESS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+1.50/082 | NA | Physician | 03/08/2017 | 2017 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7161181&pc=MTA | 2023826-2018-00001 | 2017-11-28 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure; Entrapment of Device; Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 12.6 MM VTICMO12.6, -12.50/2.5/117 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS "HALF IMPLANTED" WHEN THE PLUNGER OVERRODE THE LENS AND GOT STUCK IN THE CARTRIDGE. WHEN THE LENS WAS REMOVED FOR RELOADING, A TEAR IN THE PERIPHERICAL HAPTIC WAS NOTED. THE LENS WAS NOT REIMPLANTED. ON (B)(6) 2017, A SIMILAR LENS WAS IMPLANTED AND IT WAS SUCCESSFUL. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID INSIDE A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.50/2.5/117 | OS | Physician | 12/05/2017 | 2017 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4727970&pc=MTA | 2023826-2015-00463 | 2015-03-23 | Malfunction | STAAR SURGICAL COMPANY | 2015-04-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015 AND THE LENS TORE UPON INSERTION. THE LENS WAS REMOVED AND ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS WAS PERFORMED. THE SURGEON PLANS ON REPLACING THE LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: CONCLUSION: BASED ON THE COMPLAINT INFORMATION AND LENS WORK ORDER SEARCH, THE MOST PROBABLE CAUSE OF THE LENS TEARING DURING INJECTION IS USER OR TECHNICAL ERROR. NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A CASE/ VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND PEN MARKS ON LENS. | -16.5/+1.5/74 | OS | Physician | 03/26/2015 | 2015 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5015483&pc=MTA | 2023826-2015-01060 | 2015-07-14 | Malfunction | STAAR SURGICAL COMPANY | 2015-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.50/+1.5/045 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS TORE DURING INJECTION INTO THE EYE. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -09.50/ +1.5/045. DEVICE EVALUATED BY MANUFACTURER? NO, DEVICE NOT RETURNED.(B)(4). DEVICE NOT RETURNED. | -09.50/+1.5/045 | OS | Physician | 07/20/2015 | 2015 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5671675&pc=MTA | 2023826-2016-00644 | 2016-03-29 | Malfunction | STAAR SURGICAL COMPANY | 2016-05-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/100 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS TORE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO APPARENT INJURY. Manufacturer Narrative: (B)(4). REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. THIS LENS MODEL IS INDICATED FOR USE IN ADULTS (B)(6) OF AGE. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND THE LENS OPTIC TORN/SPLIT, THE HAPTIC TORN/BROKEN/BENT/DEFORMED, AND ALSO FOREIGN MATERIALS ON LENS SURFACE SPECIFIED AS SPOTS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | -10.00/+2.0/100 | OD | Physician | 04/05/2016 | 2016 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5676350&pc=MTA | 2023826-2016-00722 | 2016-04-15 | Malfunction | STAAR SURGICAL COMPANY | 2016-05-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO INSERT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/080 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) BUT THE LENS TORE DURING INJECTION INTO THE EYE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS ON (B)(6) 2016 WITH NO APPARENT INJURY. Manufacturer Narrative: ADDITIONAL DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4). | -10.0/+2.0/080 | OS | Physician | 04/26/2016 | 2016
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5781003&pc=MTA | 2023826-2016-00894 | 2016-05-12 | Malfunction | STAAR SURGICAL COMPANY | 2016-07-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+2.5/107 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS TORE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED WITH ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CONCLUSION: REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, PATIENT AGE BELOW 21 YEARS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -17.0/+2.5/107 | OD | Physician | 05/31/2016 | 2016 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6004360&pc=MTA | 2023826-2016-01379 | 2016-07-18 | Malfunction | STAAR SURGICAL COMPANY | 2016-10-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -10/1/92 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016 WHEN THE LENS WAS NOTED TO BE TORN DURING INJECTION AS A PROXIMAL HAPTIC WAS BROKEN. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND THE SURGEON CANCELED THE SURGERY DUE TO NO BACK UP LENS. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND HAPTIC TORN AND SPOTS ON LENS SURFACE. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | -10/1/92 | OS | Physician | 09/06/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5923621&pc=MTA | 2023826-2016-01222 | 2016-08-01 | Malfunction | STAAR SURGICAL COMPANY | 2016-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE CUSTOMER NOTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/092 DIOPTER, WAS TORN PRIOR TO LOADING. THE REPORTER INDICATED SINCE THERE WAS NO BACKUP LENS AVAILABLE, THE SURGEON IMPLANTED THE LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THERE WAS NO REPORT OF ANY APPARENT INJURY AND THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE PATIENT WILL BE MONITORED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT CODE(S): (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH - NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). | -13.0/+2.0/092 | OS | Physician | 08/05/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8122335&pc=MTA | 2023826-2018-01670 |  | Injury | STAAR SURGICAL COMPANY | 2018-12-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Markings/Labelling Problem | No Known Impact Or Consequence To Patient; No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A VTICMO 12.6 -12.00/+1.50/082, SERIAL (B)(4), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). THE REPORTER STATES THAT THE SAME MODEL, POWER AND SERIAL NUMBER OF THIS LENS (VTICMO 12.6 -12.00/+1.50/082, SERIAL # (B)(4) WAS SHOWN ON THE PATIENT CHART LABEL INSIDE THE BOX FOR ANOTHER LENS WITH VTICMO 12.6 -12.00/+1.50/083, SERIAL # (B)(4) INDICATED ON THE EXTERNAL BOX LABEL. THE LENS REMAINS IMPLANTED AND IT WAS REPORTED THAT THERE IS NO NEGATIVE FEEDBACK OR COMPLAINTS ABOUT THE PATIENT'S CONDITION Manufacturer Narrative: B5: THE REPORTER STATED THAT THE SERIAL # (B)(6) (VTICMO 12.6 -12.00/+1 .50/082) WAS SHOWN ON THE PATIENT CHART LABEL INSIDE THE BOX FOR ANOTHER LENS WITH SERIAL# (B)(6) (VTICMO 12.6 -12.00/+1.50/083) INDICATED ON THE EXTERNAL BOX LABEL. REPORTEDLY, "THE LENS WAS NOT IMPLANTED IN 2018 DUE TO PATIENT'S REASONS." THE VTICMO 12.6 LENS WITH A DIOPTER OF -12.00/+1.50/082 (SPHERE/CYLINDER/AXIS), SERIAL (B)(6) WAS IMPLANTED TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULTING. THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. COMPLAINT DESCRIPTIONS PACKAGE LABEL DOES NOT MATCH SERIAL NUMBER ON STICKER. UNABLE TO VERIFY COMPLAINT STATEMENT DUE TO LENS NOT RETURNED IN ORIGINAL PACKAGING. CLAIM #(B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, THE PROBABLE CAUSE OF THE FAILURE IS LIKELY DUE TO FAILURE TO FOLLOW PROCEDURE BY PACKAGING MORE THAN ONE LENS AT A TIME AND THEREFORE INADVERTENTLY PACKAGING THE INCORRECT PATIENT CHART LABEL. A SECOND PROBABLE CAUSE IDENTIFIED IS DUE TO INADEQUATE FINAL VERIFICATION PROCESS OF THE PATIENT CHART LABEL. CORRECTIVE ACTIONS FOR A SIMILAR DISCREPANCY REGARDING THE PACKAGING OF AN INCORRECT PATIENT CHART LABEL HAVE BEEN IDENTIFIED AND HAVE BEEN IMPLEMENTED UNDER CAPA M18-12. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/+1.50/082 | OS | Non-Healthcare Professional | 10/31/2018 | 2018 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7853015&pc=MTA | 2023826-2018-01343 | 2018-05-28 | Injury | STAAR SURGICAL COMPANY | 2018-09-06 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT, WHILE PREPARING TO IMPLANT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.0/+1.0/070 (SPHERE/CYLINDER/AXIS) THE LENS TORE/BROKE BEFORE/DURING LOADING. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY FROM THE LEFT (OD) EYE ON (B)(6) 2018. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN; ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS IS ALSO REPORTED. Manufacturer Narrative: CORRECTED DATA: INITIAL REPORT STATES LENS WAS IMPLANTED INTO THE "LEFT (OD) EYE". CORRECTED TO "RIGHT (OD) EYE". (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.0/+1.0/070 | OD | Physician | 08/10/2018 | 2018 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5668559&pc=MTA | 2023826-2016-00632 | 2016-03-29 | Injury | STAAR SURGICAL COMPANY | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/060 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) AND THE LENS TORE/BROKE DURING INJECTION. THE LENS WAS REMOVED WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE. (B)(4): SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. CONCLUSION: PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS OVER THE AGE OF 45 YEARS OLD AND WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSURE, LESS THAN 3.0MM. THE PATIENT WAS (B)(6) AND HAD AN ACD OF 2.80MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | -11.50/+2.0/060 | OS | Physician | 04/19/2016 | 2016 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5817533&pc=MTA | 2023826-2016-00918 | 2016-04-29 | Malfunction | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/088 DIOPTER, AND THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED WITH NO APPARENT INJURY. ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE (NOT SPECIFIED). (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: UNCORRECTED VISUAL ACUITY WAS 20/20, NOT BEST CORRECTED VISUAL ACUITY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+2.5/088 | NA | Physician | 06/22/2016 | 2016 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4923021&pc=MTA | 2023826-2015-00826 | 2015-03-30 | Injury | STAAR SURGICAL COMPANY | 2015-07-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. CORRECTIONAL DATA: SUSPECT MEDICAL DEVICE: DIOPTER -14.0/+2.0/070 IS INCORRECT, CORRECT DATA IS DIOPTER -14.50/1.5/084. (B)(4). | -14.0/+2.0/070 | OD | Physician | 06/19/2015 | 2015 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7765105&pc=MTA | 2023826-2018-01172 | 2018-07-09 | Injury | STAAR SURGICAL COMPANY | 2018-08-08 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -17.5/+5.0/099 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ACCORDING TO THE PATIENT'S (B)(6) VIDEO LOG, THE ICL IS "SLIGHTLY TURNED AND NEEDS TO BE ADJUSTED BY 10 DEGREES." REPORTEDLY, THE LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION IS CURRENTLY AVAILABLE. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS REPOSITIONED ON (B)(6) 2018. THE ISSUE IS RESOLVED AND THE PATIENT NO LONGER WEARS GLASSES. (B)(4). Manufacturer Narrative: ADDRESS: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -17.5/+5.0/099 | OS | Other | 07/13/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6314690&pc=MTA | 2023826-2017-00140 |  | Injury | STAAR SURGICAL COMPANY AG | 2017-02-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS ROTATED ABOUT THE AXIS AT 84 DEGREES. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -17/+3.5/90 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS ROTATED ABOUT THE AXIS AT 84 DEGREES. LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2016. THE LENS REMAINS IMPLANTED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4)- PREVIOUS CODE NOT APPLICABLE. LENS REPOSITIONING. (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND CLEAR RESIDUE ON THE LENS SURFACE. (B)(4). DEVICE MANUFACTURE DATE IS CORRECTED TO 04/03/2013 FROM 03/27/2013. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED OR MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -17/+3.5/90 | OD | Physician | 01/17/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6125957&pc=MTA | 2023826-2016-01678 | 2016-09-20 | Injury | STAAR SURGICAL COMPANY | 2016-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Handling Problem | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/2.5/092 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE LENS WAS IMPLANTED INSIDE OUT. THE LENS WAS EXCHANGED FOR A LENS THE SAME MODEL, SIZE BUT A DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS EXPLANTED AND EXCHANGED FOR A LENS THE SAME SIZE, MODEL, BUT A DIFFERENT DIOPTER. (B)(4). | -7.0/2.5/092 | OD | Physician | 10/26/2016 | 2016 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6226582&pc=MTA | 2023826-2017-00023 | 2016-09-06 | Malfunction | STAAR SURGICAL COMPANY | 2017-01-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/093 DIOPTER, TORE/BROKE BEFORE INJECTION INTO THE EYE. THE LENS WAS FOUND CRACKED JUST BEFORE IMPLANTATION. THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS BROKEN. (B)(4). Manufacturer Narrative: CORRECTED INFORMATION: FOLLOW-UP #1 - 01/05/2017. (B)(4) PREVIOUS CODE NOT APPLICABLE, SECONDARY INTERVENTION WAS PERFORMED (EXCHANGE). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+1.0/093 | NA | Physician | 12/09/2016 | 2016 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6191942&pc=MTA | 2023826-2016-01800 | 2016-11-07 | Injury | STAAR SURGICAL COMPANY | 2016-12-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS, -10.5/2.5/126 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016 AND IT WAS NOTED THAT THE LENS WAS DAMAGED, THE LENS WAS EXPLANTED ON (B)(6) 2016. THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: THE TORN LENS WAS EXCHANGED. PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND HAPTIC BROKEN. (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/2.5/126 | OD | Physician | 11/21/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6254539&pc=MTA | 2023826-2017-00088 | 2016-12-03 | Injury | STAAR SURGICAL COMPANY | 2017-01-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+1.0/075 DIOPTER, TORE DURING INJECTION/DELIVERY INTO THE EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON THE SAME DAY. THE LENS WAS EXCHANGED FOR A LENS OF THE SAME MODEL AND SIZE AND THE PROBLEM WAS RESOLVED. THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: PREVIOUS INFORMATION INCORRECT, IT IS AN ADVERSE EVENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.5/+1.0/075 | NA | Physician | 01/06/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6361177&pc=MTA | 2023826-2017-00272 | 2016-12-09 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-27 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/1.5/90 DIOPTER, TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR A NEW LENS WITH THE SAME MODEL AND SIMILAR DIOPTER. THE PROBLEM WAS RESOLVED AND ON (B)(6) 2016, THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: CORRECTED FROM ADVERSE EVENT TO PRODUCT PROBLEM. (CORRECTED): THE DOCTOR STATED THAT THE ICL DIDN'T MOVE "SMOOTHLY" FORWARD IN THE CARTRIDGE TIP. HE NOTED THE "FOLDED SHAPE OF THE ICL IN THE TIP LOOKED DISTORTED." THE DOCTOR EJECTED THE LENS, PUSHING IT WITH THE PLUNGER, TEARING THE LENS. THERE WAS NO PATIENT CONTACT. THE CARTRIDGE WAS NOT EVER LOADED INTO THE INJECTOR. AS OF THE DATE OF SUPPLEMENTAL MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. (PREVIOUS): STATED THAT THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION-LENS RETURNED IN A CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN, HAPTIC PIECE MISSING. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.5/1.5/90 | OD | Physician | 01/31/2017 | 2017 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6332671&pc=MTA | 2023826-2017-00190 | 2017-01-16 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/-1.5/75 DIOPTER, WHICH TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS NOT IMPLANTED DUE TO THE ICL LENS BEING STUCK IN THE CARTRIDGE (TRAILING HAPTIC). THE SURGEON "HAD TO REMOVE THE INJECTED ICL AND NOTICED BREAKAGE OF THE ONE EDGE OF THE HAPTIC. THE RIGHT EYE (OD) LENS IMPLANT IS DOING WELL" WHILE THE LEFT EYE (OS) LENS IS BEING "SENT IN THE SAME CONTAINER TO STAAR SURGICAL REPRESENTATIVE" AS OF (B)(6) 2017. THE SURGEON REQUESTED FOR A NEW ICL LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | -10.5/-1.5/75 | OS | Physician | 01/23/2017 | 2017 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6429973&pc=MTA | 2023826-2017-00466 | 2017-01-24 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/1.5/093 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXCHANGED DUE TO TEAR/BREAK DURING INJECTION INTO THE EYE ON THE SAME DAY. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/1.5/093 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS TORE/BROKE DURING INJECTION INTO THE EYE AND WAS EXCHANGED DURING THE SAME SURGERY. THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AS WELL AS DEBRIS AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). RESULT CODE : WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4). | -12.50/1.5/093 | OD | Physician | 02/22/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320741&pc=MTA | 2023826-2018-00325 | 2017-07-12 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.5/+3.0/098 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2017. ON THE SAME DATE THE LENS WAS EXPLANTED. IT IS UNKNOWN WHETHER EXPLANT OCCURRED INTRA-OPERATIVELY OR POST-OPERATIVELY. THE SURGEON PLANS TO EXCHANGE THE LENS ONCE THE REPLACEMENT LENS HAS BEEN RECEIVED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE LENS WAS REMOVED INTRA-OPERATIVELY. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC HAD A TEAR ON IT AND WAS BENT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). TWENTY-FOUR-OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.92MM)]. (B)(4). | -8.5/+3.0/098 | OD | Physician | 02/08/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6817510&pc=MTA | 2023826-2017-01255 | 2017-07-26 | Injury | STAAR SURGICAL COMPANY | 2017-08-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50/+3.0/084 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) AND NOTED THE LENS WAS TORN. THE LENS WAS REMOVED WITH NO APPARENT INJURY . THE REPORTER INDICATED THE LENS WAS DAMAGED DURING THE LOADING PROCESS. THE LENS WAS NOT EXCHANGED WITH ANOTHER LENS. Manufacturer Narrative: ON (B)(6) 2017 A NEW LENS OF SAME MODEL AND SIMILAR DIOPTER WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC TORN, HAPTIC BROKEN AND FIBERS PRESENT ON THE LENS SURFACE. (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.50/+3.0/084 | OS | Physician | 07/27/2017 | 2017 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050895&pc=MTA | 2023826-2017-01705 | 2017-09-27 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1, -15.0/+2.5/100 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. DURING THE SAME SURGERY, THE SURGEON EXPLANTED THE LENS AS HE SAW THAT THE LENS TORE BEFORE IT WAS IMPLANTED. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A SIMILAR LENS AND THE PROBLEM WAS RESOLVED Manufacturer Narrative: A WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND. CORRECTED INFORMATION: IT WAS REPORTED AS PRODUCT PROBLEM WITH NO ADVERSE EVENTS. THE REPORTER INDICATED THAT AS THE SURGEON WAS ATTEMPTING TO IMPLANT A A 12.1MM VTICMO12.1, -15.0/+2.5/100 (SPHERE, CYLINDER, AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) THAT A LENS TEAR WAS NOTICED PRIOR TO IMPLANTATION. THIS OCCURRED ON (B)(6) 2017. ON (B)(6) 2017, A SIMILAR LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND FIBERS ON LENS SURFACE. (B)(4). | -15.0/+2.5/100 | OS | Physician | 10/25/2017 | 2017 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7245223&pc=MTA | 2023826-2018-00195 | 2017-12-25 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-06 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -15.0/+1.5/172 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED AND REMOVED ON (B)(6) 2017. IT IS UNKNOWN WHETHER THE LENS WAS EXPLANTED INTRA-OPERATIVELY OR IN A 2ND SURGERY. Manufacturer Narrative: THE LENS WAS EXPLANTED INTRAOPERATIVELY. B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.0/+1.5/172 | OS | Physician | 01/10/2018 | 2018 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7300925&pc=MTA | 2023826-2018-00289 | 2018-01-28 | Injury | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.00/+2.50/110 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2018. REPORTEDLY, THE LENS REMAINS IN THE EYE. THE SURGEON PLANS TO EXCHANGE THE LENS ONCE THE REPLACEMENT LENS HAS BEEN RECEIVED. Manufacturer Narrative: THE LENS WAS EXCHANGED AND THE PATIENT IS DOING FINE. (B)(4). IT SHOULD BE NOTED A SUTURE WAS USED DURING INITIAL IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | -17.00/+2.50/110 | OD | Physician | 02/07/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6850682&pc=MTA | 2023826-2017-01406 | 2017-07-28 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-07 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Activation Failure | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -9.0/+3.0/X086 INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THE SURGEON THEN EXCHANGED THE LENS FOR THE SAME SIZE, SIMILAR DIOPTER LENS. THIS OCCURRED ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: AFTER THE LENS WAS IMPLANTED, THE DOCTOR DISCOVERED IT WAS "FLIPPED INSIDE THE EYE." INTRAOPERATIVELY, THE SURGEON REMOVED THE LENS AND THE LENS BROKE DURING EXPLANT. AN ALTERNATE LENS WAS IMPLANTED. (B)(4). Manufacturer Narrative: PATIENT POST-OP BCVA WAS 20/20 ON (B)(6) 2017. DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND A MISSING PIECE OF HAPTIC. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.97MM)]. (B)(4). | -9.0/+3.0/086 | OD | Physician | 08/09/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5940511&pc=MTA | 2023826-2016-01249 | 2016-06-09 | Injury | STAAR SURGICAL COMPANY | 2016-09-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+3.0/090 DIOPTER, AND THE LENS TORE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED AND ANOTHER SAME MODEL/DIOPTER LENS WAS IMPLANTED. THERE WAS NO REPORT OF ANY APPARENT INJURY. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/28. Manufacturer Narrative: IMPLANT DATE (B)(6) 2016; EXPLANT DATE (B)(6) 2016; EXCHANGE DATE (B)(6) 2016. DEVICE EVALUATION: THE LENS WAS RETURNED IN CASE/VIAL; VISUAL INSPECTION FOUND HAPTIC TORN (B)(4). Manufacturer Narrative: EXCHANGE DATE (B)(6) 2016; PREVIOUS DATE INCORRECT. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4). | -9.5/+3.0/090 | NA | Physician | 08/10/2016 | 2016 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5826216&pc=MTA | 2023826-2016-00977 | 2016-06-28 | Injury | STAAR SURGICAL COMPANY | 2016-07-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON WAS INSERTING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+3.5/081 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016 AND THE LENS TORE. THE LENS WAS NOT IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN/ SPLIT, HAPTIC BROKEN AND LENS BROKEN IN TWO PIECES. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTIONS. THE REPORTER STATED INDICATED THE SURGEON INSERTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+3.5/081 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE SURGEON NOTICED A LENS TEAR DURING DELIVERY OF ICL INTO THE ANTERIOR CHAMBER. NEXT DAY, THE LENS WAS REMOVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | -18.0/+3.5/081 | OS | Physician | 07/08/2016 | 2016 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7508619&pc=MTA | 2023826-2018-00722 | 2010-06-22 | Injury | STAAR SURGICAL COMPANY AG | 2018-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2010. THE REPORTER INDICATED THE LENS HAD A LOW VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2018. THE LENS WAS EXCHANGED FOR A LONGER LENS (13.7) BUT DIFFERENT MODEL AND DIOPTER - VTICM5_13.7, -8.00/+1.5/X089 (SPHERE/CYLINDER/AXIS). THE EXCHANGE RESOLVED THE PROBLEM. (B)(4). CORRECTED DATA: PRODUCT PROBLEM TO ADVERSE EVENT. MALFUNCTION TO SERIOUS INJURY. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. | -10.0/+1.5/089 | OD | Physician | 04/16/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6441301&pc=MTA | 2023826-2017-00517 | 2017-02-08 | Injury | STAAR SURGICAL COMPANY | 2017-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/X126 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING AND WILL BE REMOVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF ONE LENS HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.5/+2.5/126 | OD | Physician | 03/02/2017 | 2017 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6770823&pc=MTA | 2023826-2017-01288 | 2017-07-13 | Injury | STAAR SURGICAL COMPANY AG | 2017-08-07 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/2.5/75 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE PATIENT EXPERIENCED LOW VAULTING. THE LENS IS PLANNED FOR EXCHANGE. Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2018 FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A CENTRIFUGE VIAL WITH PRESENCE OF RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). | -11.50/2.5/75 | OD | Physician | 07/13/2017 | 2017 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6811835&pc=MTA | 2023826-2017-01364 | 2017-07-25 | Injury | STAAR SURGICAL COMPANY | 2017-08-22 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/169 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND THE LENS IS PLANNED TO BE EXCHANGED AT THE END OF (B)(6). Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED ON (B)(6) 2017 FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR AND RED RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/+2.0/169 | OD | Physician | 07/26/2017 | 2017 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6836437&pc=MTA | 2023826-2017-01378 | 2017-08-02 | Injury | STAAR SURGICAL COMPANY | 2017-08-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.0/032 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND THE PLAN IS TO EXPLANT AND EXCHANGE THE LENS WITHIN THE NEXT COUPLE OF WEEKS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2017 FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. AT ONE WEEK POST-OP, THE PATIENT'S CONDITION WAS GOOD, WITH UCVA -20/20. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.00/+1.0/032 | OD | Physician | 08/03/2017 | 2017 | 47
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6850658&pc=MTA | 2023826-2017-01400 | 2017-08-02 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+2/114 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. INTRAOPERATIVELY, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS FOUND ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: CORRECTED DATA: ON (B)(6) 2017, POST-REMOVAL OBSERVATION OF ANISOCORIA AND UNCORRECTED VISUAL ACUITY (UCVA) OF 20/400 WAS OBSERVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/+2/114 | OD | Physician | 08/08/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7177066&pc=MTA | 2023826-2018-00007 | 2017-12-01 | Injury | STAAR SURGICAL COMPANY | 2018-01-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.5/+1.5/176 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS. Manufacturer Narrative: THE LENS WAS EXCHANGED FOR A SHORTER LENS ON 13 MAR 2018 AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.5/+1.5/176 | OS | Physician | 12/11/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7177114&pc=MTA | 2023826-2018-00006 | 2017-12-01 | Injury | STAAR SURGICAL COMPANY | 2018-01-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.0/+3.0/168 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -8.0/3.0/168 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE REPORTER STATED THAT "ALL IS FINE" AND THAT THE PATIENT IS HAPPY. EXPLANTED DATE: (B)(6) 2018. (B)(4) Manufacturer Narrative: CORRECTED DATA: IT IS CORRECTED FROM "ADVERSE EVENT" TO " PRODUCT PROBLEM". TYPE OF REPORTABLE EVENT: IT IS CORRECTED FROM "SERIOUS INJURY" TO " MALFUNCTION". CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.0/+3.0/168 | OD | Physician | 12/11/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7403688&pc=MTA | 2023826-2018-00513 | 2018-01-09 | Injury | STAAR SURGICAL COMPANY | 2018-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+1.5/070 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON NOTES EXCESSIVE VAULTING. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4) Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.5/+1.5/070 | OD | Physician | 03/09/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7403692&pc=MTA | 2023826-2018-00515 | 2018-01-09 | Injury | STAAR SURGICAL COMPANY | 2018-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+3.0/097 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES EXCESSIVE VAULTING. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED WITH A SHORTER LENS ON (B)(6) 2018. THE PROBLEM WAS RESOLVED AND "PATIENT IS HAPPY - ALL IS FINE". (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL AND THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | -17.5/+3.0/097 | OS | Physician | 03/09/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7300924&pc=MTA | 2023826-2018-00284 | 2018-01-31 | Injury | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+1.0/010 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON REPORTS EXCESSIVE VAULT. CURRENTLY, THE LENS REMAINS IMPLANTED, HOWEVER IT IS PLANNED TO BE EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: PREVIOUSLY STATED THE LENS REMAINS IMPLANTED. UPDATED/CORRECTED INFORMATION-THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.5/+1.0/010 | OS | Physician | 02/01/2018 | 2018 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7749504&pc=MTA | 2023826-2018-01155 | 2018-02-12 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6 VTICMO12.6, -6/1/87 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REMOVED WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS, THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THE REPORTER INDICATED THAT A 12.6 VTICMO_12.6 -6/1/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AND REMOVED INTRAOPERATIVELY DUE TO EXCESSIVE VAULT ON (B)(6) 2018. A SHORTER REPLACEMENT LENS WAS IMPLANTED POST-OPERATIVELY ON (B)(6) 2018, RESOLVING THE PROBLEM. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC DEFORMED. THERE WAS RESIDUE ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, EXCHANGE. (B)(4). | -6/1/87 | OD | Physician | 07/04/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7747114&pc=MTA | 2023826-2018-01161 | 2018-06-17 | Injury | STAAR SURGICAL COMPANY | 2018-08-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.50/+3.0/016 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. EXCESSIVE VAULT WAS OBSERVED. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED, IT WAS REPORTED THAT THE SURGEON PLANS TO EXCHANGE THE LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 -13.50/+3.0/016 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.50/+3.0/016 | OD | Physician | 07/04/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7839690&pc=MTA | 2023826-2018-01324 | 2018-06-19 | Injury | STAAR SURGICAL COMPANY | 2018-08-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -15.5/+2.0/84 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. LOW VAULT WAS OBSERVED, THE LENS IS PLANNED TO BE EXCHANGED. IT WAS REPORTED THAT THE CAUSE OF THE EVENT WAS USER ERROR . ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTION DATA: DEVICE MANUFACTURE DATE CORRECTED TO "07-FEB-2018" IN THE INITIAL MDR. ADDITIONAL DATA: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -15.50/+2.0/84 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2018 DUE TO LOW VAULT. THE PROBLEM WAS NOT RESOLVED. THE REPORTER STATED, "VAULT IS A BIT LOW, BUT IT'S THE BIGGEST LENS". IT WAS REPORTED THAT THE CAUSE OF THE EVENT WAS USER ERROR. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.5/+2.0/84 | OD | Physician | 08/01/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824068&pc=MTA | 2023826-2018-01274 | 2018-07-20 | Injury | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUSLY STATED THE LENS REMAINS IMPLANTED UPDATED STATEMENT-ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS. THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/+1.5/086 | OD | Physician | 08/02/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7995981&pc=MTA | 2023826-2018-01553 | 2018-09-28 | Injury | STAAR SURGICAL COMPANY | 2018-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/178 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED INFORMATION: ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS AND THE PROBLEM IS RESOLVED. "AFTER EXCHANGE ALL IS FINE/PATIENT IS HAPPY." THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.0/+1.0/178 | OD | Physician | 09/28/2018 | 2018 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8021154&pc=MTA | 2023826-2018-01580 | 2018-09-28 | Injury | STAAR SURGICAL COMPANY | 2018-10-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/055 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUSLY STATED THE LENS REMAINS IMPLANTED. UPDATED INFORMATION: ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. "AFTER EXCHANGE ALL IS FINE/PATIENT IS HAPPY." (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+1.0/055 | OS | Physician | 10/05/2018 | 2018 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8073836&pc=MTA | 2023826-2018-01628 | 2018-10-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -11.50/5.0/063 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS WAS OBSERVED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY AND WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF -11.5/5.0/63 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT WAS OBSERVED. ON 03/01/2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. THIS EXCHANGE RESOLVED THE PROBLEM. PATIENT CODE 3191: SECONDARY SURGERY, LENS EXCHANGE, SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/5.0/063 | OS |  | 10/16/2018 | 2018 | 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8113346&pc=MTA | 2023826-2018-01754 | 2018-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50/+2.0/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING. THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS FOR A SHORTER LENS. THE LENS TO DATE REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: AGE OR DATE OF BIRTH-CORRECTED TO :(B)(6)1981. ADDITIONAL DATA: EVENT-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -12.50/+2.0/083 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS) ON (B)(6)2018. ON (B)(6)2018, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE REPORTER NOTED, "PATIENT IS HAPPY. NO FURTHER TREATMENT IS NECESSARY. LAST VISIT: (B)(6)2018; BCVA: 20/20; UCVA: 20/20; VAULT: 600UM". PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXCHANGED CLAIM #(B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.50/+2.0/083 | OS |  | 10/29/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8298907&pc=MTA | 2023826-2019-00190 | 2018-12-23 | Injury | STAAR SURGICAL COMPANY | 2019-02-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.00/+2.0/079 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD A LOW VAULT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2019. THE PATIENT DECIDED NOT TO HAVE ANOTHER LENS IMPLANTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.00/+2.0/079 | OD | Physician | 01/07/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7562898&pc=MTA | 2023826-2018-00802 | 2018-05-05 | Injury | STAAR SURGICAL COMPANY | 2018-06-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.50/+2.5/174 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD A LOW VAULT, WITH ROTATION. THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS BUT TO DATE THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2018. A LONGER LENS WAS IMPLANTED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THE REPORTER CLARIFIED STATING THAT LENS WAS EXCHANGED IN THE SAME SURGERY AS LENS REMOVAL ON (B)(6) 2018 FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH PRESENCE OF RESIDUE/ DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE LENS TO HAVE A CURVED SHAPE AND PRESENCE OF RESIDUE/ DEBRIS ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. | -7.50/+2.5/174 | OD | Physician | 05/07/2018 | 2018 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824433&pc=MTA | 2023826-2018-01255 | 2018-05-15 | Injury | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.50/+3.5/X077 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS). THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS FOR A LONGER LENS. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) THE REPORTER INDICATED THE LENS WAS IMPLANTED ON (B)(6) 2018. THE LENS WAS REPOSITIONED ON (B)(6) 2018 BUT THE ISSUE WAS NOT RESOLVED. THE LENS ROTATED A SECOND TIME ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE LENS IS NOW STABLE. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. (B)(4). CORRECTED DATA: PRODUCT PROBLEM TO ADVERSE EVENT MALFUNCTION TO SERIOUS INJURY CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.50/+3.5/077 | OS | Physician | 07/30/2018 | 2018 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4958187&pc=MTA | 2023826-2015-00981 | 2015-02-19 | Injury | STAAR SURGICAL COMPANY | 2015-07-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015 AND LOW VAULT, ROTATION AND LENS DISLOCATION/SUBLUXATION WAS NOTED. THE LENS REMAINS IMPLANTED AND THE SURGEON PLANS ON EXCHANGING IT FOR A LONGER LENGTH. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: LENS WAS EXPLANTED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND BENT. CORRECTED DATA:(B)(4) PREVIOUS CODE NOT APPLICABLE, LENS EXPLANTED. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN PATIENTS WHO ARE AMBLYOPIC OR BLIND IN FELLOW EYE. (B)(4). | -11.5/4.0/70 | OS | Physician | 07/08/2015 | 2015 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7361884&pc=MTA | 2023826-2018-00463 | 2018-01-31 | Injury | STAAR SURGICAL COMPANY AG | 2018-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS, -6.5/2.5/89 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2011. THE REPORTER INDICATED THE LENS HAD AN LOW VAULT. THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS. Manufacturer Narrative: UPDATED TO ADVERSE EVENT. UPDATED DESCRIPTION OF EVENT: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS -6.5/2.5/89 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE REPORTER OBSERVED A LOW VAULT SO THE SURGEON EXPLANTED THE LENS. NO INFORMATION REGARDING THE DETAILS OF THE SECONDARY SURGERY HAS BEEN PROVIDED DESPITE THE REQUESTS FOR DETAILS. PRODUCT EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT WITH RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.5/2.5/89 | OS | Physician | 02/21/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6021626&pc=MTA | 2023826-2016-01440 | 2016-09-06 | Injury | STAAR SURGICAL COMPANY | 2016-10-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+1.5/040 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS TORE/BROKE DURING THE IMPLANTATION AND WAS EXPLANTED ON THE SAME DAY. ANOTHER LENS OF THE SAME SIZE WAS IMPLANTED ON (B)(6) 2016. THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016. ANOTHER LENS WAS IMPLANTED ON THE SAME DAY. (B)(4). Manufacturer Narrative: (B)(4). PRODUCT EVALUATION: THE LENS WAS RETURNED IN LIQUID IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAPTIC TORN (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). LENS NOT RETURNED. | -15.0/+1.5/040 | OS | Physician | 10/03/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5937931&pc=MTA | 2023826-2016-01239 | 2016-05-04 | Malfunction | STAAR SURGICAL COMPANY | 2016-09-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient; No Information | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50+1.0/092 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. IT WAS INDICATED THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND THE LENS REMAINS IMPLANTED. Manufacturer Narrative: METHOD: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CONCLUSION: OFF-LABEL USE: PATIENT ACD BELOW 3.0 MM. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: PRODUCT PROBLEM (MALFUNCTIONS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4). DEVICE NOT RETURNED. | -11.50+1.0/092 | OS | Physician | 08/10/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5816262&pc=MTA | 2023826-2016-00971 | 2016-06-10 | Injury | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient; No Information | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.0/+3.0/094 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND REMAINS IMPLANTED. THE LENS IS PLANNED TO EXCHANGE WITH ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: CORRECTED DATA: THIS IS A PRODUCT PROBLEM, NOT AN ADVERSE EVENT. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/40. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WEIGHT: UNK. OUTCOMES ATTRIBUTED TO ADVERSE EVENTS: UNK. EXPLANT DATE: N/A. (B)(4). LENS REMAINS IMPLANTED. | -14.0/+3.0/094 | OD | Physician | 06/22/2016 | 2016 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5506284&pc=MTA | 2023826-2016-00273 | 2015-11-24 | Injury | STAAR SURGICAL COMPANY | 2016-03-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient; No Information; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+3.0/104 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD A LOW VAULT. THE REPORTER INDICATED THE SURGEON PLANS TO EXPLANT THE LENS BUT THE LENS IS CURRENTLY IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: METHOD: WORK ORDER SEARCH PERFORMED, NO SIMILAR COMPLAINTS FOUND. CONCLUSION: PATIENT ACD WAS BELOW 3.0 MM. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN LIQUID IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. PROCESS EVALUATION: WORK ORDER SEARCH. EVALUATION RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -14.5/+3.0/104 | OS | Physician | 01/21/2016 | 2016 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6993821&pc=MTA | 2023826-2017-01658 | 2017-08-28 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-01 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Patient Involvement | NA | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2017, AS THE SURGEON ATTEMPTED TO LOAD AND IMPLANT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -17.0/+0.50/X134 INTO THE PATIENT'S LEFT (OS) EYE, HE DISCOVERED THE LENS BROKE. REPORTEDLY, THIS WAS DISCOVERED PRIOR TO IMPLANT AND THERE WAS NO PATIENT CONTACT WITH THIS LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))]. (B)(4). | -17.0/+0.50/134 | OS | Physician | 10/06/2017 | 2017 | 47
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7261732&pc=MTA | 2023826-2018-00169 | 2017-11-22 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -11.5/1.5/90 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2017. THE SURGEON EXPERIENCED LENS DAMAGE BEFORE IMPLANTATION OF THE LENS AND A BACK-UP LENS WAS USED INSTEAD ON (B)(6) 2017 TO COMPLETE THE SURGERY. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND PORTION OF THE LENS MISSING, A TEAR ON THE HAPTIC, AND RESIDUE ON THE LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. (B)(4). | -11.5/1.5/90 | OD | Physician | 01/10/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7262034&pc=MTA | 2023826-2018-00170 | 2017-12-20 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | N/A | Event Description: THE REPORTER INDICATED THAT THE FOOT PLATE OF 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/2.5/105( SPHERE/ CYLINDER/ AXIS), TORE SHARPLY WHEN PUSHING THE LENS BY USING FTP (FOAM TIP PLUNGER . THE SURGEON EXPERIENCED LENS DAMAGE BEFORE INJECTION INTO THE EYE. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND PORTION OF THE LENS MISSING AND RESIDUE ON THE LENS SURFACE. SFC-45 CARTRIDGE WAS RETURNED IN A BIOHAZARD BAG. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE DEVICE AND CLEAR RESIDUE. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: IT SHOULD BE NOTED THAT BEFORE LENS TORE, "LENS MOVEMENT WAS NOT SMOOTH WHEN PULLING LENS TOWARDS NOZZLE". CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. (B)(4). | -13.0/2.5/105 | NA | Physician | 01/10/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7508767&pc=MTA | 2023826-2018-00684 | 2018-03-21 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | NA | Event Description: THE REPORTER INDICATED THAT A 13.2 MM VTICMO13.2, -14.5/+4.0/85 (SPHERE/CYLINDER/AXIS). IMPLANTABLE COLLAMER LENS TORE WHILE ADJUSTING THE LENS POSITION WITHOUT EXCESSIVE OPERATION. THE LENS WAS NOT IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS SURFACE. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.5/+4.0/85 | NA | Physician | 04/11/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8358682&pc=MTA | 2023826-2019-00292 | 2019-01-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Material Discolored | No Patient Involvement | N/A | Event Description: THE REPORTER INDICATED THAT AFTER OPENING THE VIAL OF A 12.1MM, VTICM5_12.1, -12.5/2.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, THE DOCTOR FOUND A FOREIGN BODY,IT LOOKED LIKE A BLACK LINE. THE LENS HAD NO PATIENT CONTACT, NOT IMPLANTED. ON (B)(6) 2019 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION : LENS WAS RETURNED IN DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A TEAR ON THE OPTIC AND PRESENCE OF RESIDUE ON LENS SURFACE. THE LENS WAS SOAKED IN BSS SOLUTION IN ORDER TO NOTE ANY FOREIGN BODIES, NONE WHERE NOTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.5/2.5/90 | NA |  | 01/23/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8254836&pc=MTA | 2023826-2019-00060 | 2017-10-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Patient Involvement | NA | Event Description: THE REPORTER INDICATED THAT, WHILE PREPARING TO LOAD A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+4.0/084 (SPHERE/CYLINDAR/AXIS) THEN INJECT INTO THE PATIENT'S LEFT EYE (OS), HE NOTICED THE LENS WAS TORE/BROKE. THIS OCCURRED ON (B)(6) 2017. REPORTEDLY, THIS WAS DISCOVERED PRIOR TO OPENING THE VIAL, HOWEVER IT IS ALSO NOTED THAT THIS COULD HAVE BEEN USER ERROR. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN AND A PORTION OF THE LENS MISSING. METHOD CODE ADDED, RESULT CODE UPDATED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.; (B)(4). | -10.0/+4.0/084 | OS |  | 12/17/2018 | 2018 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7898026&pc=MTA | 2023826-2018-01419 | 2018-02-22 | Malfunction | STAAR SURGICAL COMPANY | 2018-09-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | Event Description: THE REPORTER INDICATED THAT A VTICM5_12.6, -9.50/+2/028 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS RECEIVED DAMAGE AND BROKEN ON ONE SIDE. THERE WAS NO PATIENT CONTACT. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS WAS RETURNED DRY WITH RESIDUE IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND TEARS IN HAPTICS. (B)(4). | -9.50/+2/028 | NA | Physician | 08/22/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8073847&pc=MTA | 2023826-2018-01677 | 2018-08-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | Event Description: THE REPORTER INDICATED THAT A VTICM5_13.7, -8.00/0.5/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS DID NOT SLIDE SMOOTHLY INTO THE INJECTOR. LENS WAS INJECTED OUR OF INJECTOR ONTO TABLE WHERE IT APPERED TORN. NO PAITEN CONTACT AND NO PATIENT INJURY WAS REPORTED. BACK UP LENS WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: RETURNED TO MANUFACTURER CORRECTED TO: "25-OCT-2018" IN THE INITIAL MDR. RESULT CODE CORRECTED TO: "114" IN THE INITIAL MDR, PREVIOUS CODE NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). | -8.00/0.5/112 | NA |  | 10/16/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7906907&pc=MTA | 2023826-2018-01387 | 2018-08-21 | Malfunction | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | Event Description: THE REPORTER INDICATED THAT VTICMO13.2, -18.00/2.0/083 (SPHERE/CYLINDER/AXIS), INTRAOCULAR LENS TORE DURING LOADING. IT WAS REPORTED THAT THE LENS STUCK TO THE CARTRIDGE MORE THAN USUAL AND THAT THERE WAS NO PATIENT CONTACT. REPLACEMENT LENS WAS IMPLANTED ONE WEEK LATER. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: CORRECTED DATA: DIOPTER CORRECTED TO " -18.00/ 2.0/ 093 " IN THE INITIAL MDR. RETURNED TO MANUFACTURER DATE: 14-SEP-2018 IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.00/2.0/083 | NA | Physician | 08/28/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8371356&pc=MTA | 2023826-2019-00337 | 2018-11-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.00/2.00/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING LOADING, WHEN POSITIONING IN THE INJECTOR. THERE WAS NO PATIENT CONTACT WITH THE DEVICE. A NEW LENS WAS ORDERED AND IMPLANTED ON (B)(6) 2018, PROBLEM RESOLVED. IT WAS REPORTED THAT THE FORCEPS WERE CHECKED FOR BURRS AND THEY HAD NO DEVIATIONS. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS THE DEVICE. Manufacturer Narrative: DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND THAT A HAPTIC WAS TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/2.00/095 | NA |  | 01/28/2019 | 2019 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372268&pc=MTA | 2023826-2019-00305 | 2018-12-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | Event Description: THE REPORTER INDICATED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, - 11.0/+1.0/087 (SPHERE/CYLINDER/AXIS) TORE WHILE BEING LOADED AND THERE WAS NO PATIENT CONTACT. ANOTHER SAME MODEL LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS THE DEVICE BECAUSE THE LENS HOLDER FORCEP DID NOT GRAB THE LENS PROPERLY SO THERE WAS A LENS BITE. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | - 11.0/+1.0/087 | NA |  | 01/29/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8302662&pc=MTA | 2023826-2019-00227 | 2019-01-03 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | Event Description: THE REPORTER INDICATED THAT THE HAPTICS OF A 13.2MM, VTICM5_13.2, -13.00/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WERE BROKEN WHEN THE DOCTOR OPENED THE LENS. THERE WAS NO PATIENT CONTACT WITH THE LENS. A REPLACEMENT LENS WAS IMPLANTED AND THE PATIENT IS DOING WELL. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/1.5/090 | NA |  | 01/11/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372540&pc=MTA | 2023826-2019-00326 | 2019-01-25 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Patient Involvement | NA | Event Description: THE REPORTER INDICATED THAT, AFTER OPENING THE VIAL AND BEFORE LOADING A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/094 (SPHERE/CYLINDAR/AXIS), IT TORE/BROKE. THIS WAS MEANT TO BE IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THERE WAS NO PATIENT CONTACT WITH THE LENS AND THIS OCCURRED ON (B)(6) 2019. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. ON THE SAME DATE AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.5/+2.0/094 | OS |  | 02/01/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7206051&pc=MTA | 2023826-2018-00102 | 2017-11-16 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Cut In Material; Device Operates Differently Than Expected | No Patient Involvement; No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER STATED THAT THE SURGEON NOTICED A TEAR ON A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS,-8.5/+3.0/94 (SPHERE/CYLINDER/AXIS), WHEN THE LENS WAS REMOVED FROM THE VIAL. THE LENS WAS NOT USED, NO PATIENT CONTACT. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN A LIQUID INSIDE A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND THE LENS MISSING A PIECE OF ITS HAPTIC. THERE WAS PRESENCE OF FIBER- LIKE DEBRIS ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.5/+3.0/94 | NA | Physician | 12/20/2017 | 2017 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6200475&pc=MTA | 2023826-2016-01808 | 2016-05-26 | Malfunction | STAAR SURGICAL COMPANY | 2016-12-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Patient Involvement; No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON WAS ATTEMPTING TO IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/179 DIOPTER IN THE PATIENT'S LEFT EYE (OS) AND IT WAS NOTED THAT THE LENS WAS TORN DURING LOADING. BACKUP LENS OF THE SAME SIZE, MODEL AND DIFFERENT DIOPTER WAS USED. THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: (B)(4). BASED ON INFORMATION FROM REPORTER, IT CANNOT BE CONFIRMED THAT DEVICE MALFUNCTION OR USER ERROR CONTRIBUTED TO THIS EVENT. PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD CODE(S): CLAIM EVALUATED BASED ON REPORTER RESPONSE ON COMPLAINT QUESTIONNAIRE. (B)(4). LENS NOT RETURNED. | -13.5/+2.0/179 | OS | Physician | 11/22/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7072216&pc=MTA | 2023826-2017-01809 | 2017-10-26 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement; No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/145 DIOPTER, AND THE LENS TORE/BROKE BEFORE INJECTION INTO THE EYE. THE REPORTER INDICATED THE LENS WAS TORN WHILE LOADING INTO THE CARTRIDGE AND THERE WAS NO PATIENT CONTACT. THE SURGERY WAS ABORTED. ANOTHER SAME MODEL LENS WAS IMPLANTED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL WITH VISON OF 20/20. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: PATIENT IDENTIFIER: IT IS CORRECTED FROM (B)(4). (B)(4). ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND DEBRIS ON LENS, THE LENS MISSING A PIECE OF HAPTIC AND THE LENS SHAPE DISTORTED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+1.0/145 | NA | Physician | 11/11/2017 | 2017 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7339933&pc=MTA | 2023826-2018-00415 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-03-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; Device Operates Differently Than Expected | No Patient Involvement; No Known Impact Or Consequence To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON WAS UNABLE TO INJECT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, - 10.0/+2.0/091 (SPHERE/CYLINDER/AXIS), DUE TO WRONG VISCOELASTIC WAS USED. THE LENS WAS DAMAGED DURING THE LOADING PROCESS. THE LENS WAS NOT IMPLANTED AND THERE WAS NO REPORT OF PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY ON A MICRO CENTRIFUGE VIAL WITH PRESENCE OF CLEAR RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE LENS WAS MISSING A PART AND THE PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | - 10.0/+2.0/091 | NA | Physician | 02/12/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721248&pc=MTA | 2023826-2018-01078 | 2017-12-07 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Pain; Blurred Vision; Halo; Discomfort; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+3.0/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, GLARE/HALOES, BLURRED VISION, PAIN, AND "NO COMFORT FOR NEAR AND MIDDLE VISION." THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND THE HAPTIC BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.0/+3.0/100 | OD | Physician | 06/29/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8403798&pc=MTA | 2023826-2019-00367 | 2019-02-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated; No Apparent Adverse Event | Pain; Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -9.50/1.5/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN (B)(6) 2017. THE REPORTER INDICATED THAT AN ACCIDENT TO THE PATIENTS EYE HAS RESULTED IN BLUR VISION WHEN READING OR CONCENTRATING, AND PAIN WITH LIGHT CHANGES AND WHILE LOOKING AT A MOVING OBJECT. REPORTEDLY, THE PATIENT HAD A VISUAL CHECK WITH THE SURGEON AND THE VISUAL CHECK WAS "PERFECT." THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: (B)(4). Manufacturer Narrative: IMPLANT DATE: (B)(6) 2017. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO [OR#] SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.50/1.5/100 | OD |  | 02/05/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5248831&pc=MTA | 2023826-2015-01539 | 2015-05-30 | Injury | STAAR SURGICAL COMPANY | 2015-11-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Insufficient Information | Pain; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -04.00/+4.50/X114 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT STARTED TO HAVE SEVERE PAIN 15 MINUTES POST-OP. THE LENS WAS EXPLANTED THAT SAME DAY. NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PEN MARKS ON LENS SURFACE. (B)(4). CONCLUSION: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. PATIENT'S ACD IS 2.97 MM. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -04.00/+4.50/X114. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). METHOD CODE(S): EVALUATION METHOD: LENS WORK ORDER SEARCH RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | -04.00/+4.50/114 | OD | Physician | 10/30/2015 | 2015 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7261493&pc=MTA | 2023826-2018-00173 | 2018-01-10 | Injury | STAAR SURGICAL COMPANY | 2018-02-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Pain; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -0.11/+0.33/094 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE PATIENT BEGAN TO COMPLAIN OF OCULAR PAIN. REPORTEDLY, ALSO NOTED IS EXCESSIVE VAULT AND THAT THE LENS IS "COMPROMISING THE ANGLES." THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: LENS WAS EXCHANGED AS THE VAULT WAS EXCESSIVE. (B)(4) Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -0.11/+0.33/094 | OD | Physician | 01/10/2018 | 2018 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6527426&pc=MTA | 2023826-2017-00698 | 2015-03-14 | Injury | STAAR SURGICAL COMPANY | 2017-04-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Retinal Detachment; Blurred Vision; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -21.50/+4.5/100 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT EXPERIENCED RETINAL DETACHMENT, AND LOST OF BCVA. THE PATIENT WAS TREATED WITH "EXTERNAL CERCLAGE" AND RETINAL REPAIR WAS PERFORMED ON (B)(6) 2015. THE CUSTOMER STATED HE THOUGH THE RETINAL DETACHMENT IS INDEPENDENT OF THE LENS. ON (B)(6) 2015 THE LENS WAS EXPLANTED. AT PRESENT, THE PATIENT REFUSES TO IMPLANT ANOTHER LENS. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -21.50/+4.5/100 | OS | Physician | 03/29/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4923027&pc=MTA | 2023826-2015-00828 | 2015-06-19 | Injury | STAAR SURGICAL COMPANY | 2015-07-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Visual Disturbances | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD A REFRACTIVE SURPRISE AND WILL BE EXCHANGED FOR ANOTHER LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO -(OTHER): LENS REMAINS IMPLANTED. EVALUATION CODES: METHOD (OTHER): WORK ORDER SEARCH RESULTS -(OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS -(OTHER): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS IMPLANTED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.5/1/092 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCING REFRACTIVE SURPRISE. EXPLANT DATE: - UNK. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CONCLUSION CODE: THE PATIENT IS UNDER 21 YEARS OLD AND HAS AN ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS UNDER 21 YEARS OLD AND ACD BELOW 3.0MM. (B)(4). | -7.5/+1.0/092 | OD | Other | 06/22/2015 | 2015 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6158460&pc=MTA | 2023826-2016-01736 | 2016-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2016-12-09 05:00 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Visual Disturbances | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -7.00/40.5/090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT IS EXPERIENCING ISSUE WITH VISION WHEN IT IS DARK. THE LENS REMAINS IMPLANTED AT THE TIME OF THIS REPORT. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS -7.00/0.5/090 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS) ON (B)(4) 2015. THE PATIENT IS EXPERIENCING ISSUE WITH VISION WHEN IT IS DARK. ON (B)(4) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL, DIFFERENT LENGTH LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY, "PATIENT IS HAPPY". CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.00/40.5/090 | OS | Physician | 11/11/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7287284&pc=MTA | 2023826-2018-00243 | 2017-06-21 | Injury | STAAR SURGICAL COMPANY | 2018-02-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances; Halo | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+1.5/102 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS TORN IN THE HAPTIC AREA, DURING DELIVERY INTO THE EYE AND REMAINS IMPLANTED. THE PATIENT COMPLAINS OF GLARE AND HALOS, ESPECIALLY AT NIGHT. THE SURGEON PLANS TO EXPLANT THE LENS. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.5/+1.5/102 | OS | Physician | 01/23/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7954143&pc=MTA | 2023826-2018-01500 | 2018-04-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-10-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND THE PATIENT EXPERIENCED GLARE/HALOS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.5/+1.5/093 | OS |  | 09/10/2018 | 2018 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6574193&pc=MTA | 2023826-2017-00831 |  | Injury | STAAR SURGICAL COMPANY | 2017-05-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Visual Disturbances; Halo | NA | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS; -16.00/2.5/085 DIOPTER; IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THIS WAS A 2ND LENS IMPLANTED IN THIS PATIENT (EXCHANGED TO A SHORTER LENS) AND THE PATIENT REPORTED THE EXCHANGE DID NOT RESOLVE THE PROBLEM. THE PATIENT REPORTED VISION WAS POOR IN BAD LIGHTING AND HALOS AT NIGHT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.00/2.5/085 | OS | Patient | 03/20/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6574194&pc=MTA | 2023826-2017-00830 |  | Injury | STAAR SURGICAL COMPANY | 2017-05-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Visual Disturbances; Halo | NA | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS; -16.50/2.5/097 DIOPTER; IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THIS WAS A 2ND LENS IMPLANTED IN THIS PATIENT (EXCHANGED FOR A SHORTER LENS) AND THE PATIENT REPORTED THE EXCHANGE DID NOT RESOLVE THE PROBLEM. THE PATIENT REPORTED VISION WAS POOR IN BAD LIGHTING AND HALOS AT NIGHT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.50/2.5/097 | OD | Patient | 03/20/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6218810&pc=MTA | 2023826-2017-00006 | 2015-09-22 | Injury | STAAR SURGICAL COMPANY | 2017-01-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Visual Disturbances; Halo; Discomfort | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.50/+5.0/090 DIOPTER IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. THE PATIENT EXPERIENCED VISION DISCOMFORT, GLARES/HALOES, AND BLURRED VISION. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THE CORRECT DATE OF IMPLANTATION IS (B)(6) 2015. (B)(4). Manufacturer Narrative: HIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). | -07.50/+5.0/090 | OS | Physician | 12/06/2016 | 2016 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6442665&pc=MTA | 2023826-2017-00511 | 2016-11-18 | Injury | STAAR SURGICAL COMPANY | 2017-03-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Visual Disturbances; Halo; Discomfort; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.5/+0.5/80 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCED GLARE/HALOES. THE ISSUE WAS RESOLVED, HOWEVER, THE PATIENT KEEPS EXPRESSING SUBJECTIVE DISCOMFORT. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION: OFF LABEL USE (ACD<3.00MM). (B)(4). | -9.5/+0.5/80 | OD | Physician | 02/28/2017 | 2017 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6442671&pc=MTA | 2023826-2017-00510 | 2016-11-18 | Injury | STAAR SURGICAL COMPANY | 2017-03-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Visual Disturbances; Halo; Discomfort; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/96 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCED GLARE/HALOES. THE ISSUE WAS RESOLVED, HOWEVER THE PATIENT KEEPS EXPRESSING SUBJECTIVE DISCOMFORT. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION-LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/+1.5/96 | OS | Physician | 02/28/2017 | 2017 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7740001&pc=MTA | 2023826-2018-01147 | 2016-07-13 | Injury | STAAR SURGICAL COMPANY | 2018-08-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -1.5/+5.5/060 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SAME SIZE, SIMILAR DIOPTER LENS DUE TO GLARE/HALOES. THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE, THE SURGEON NOTES "HAPPY PATIENT". Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -1.5/+5.5/060 | OD | Physician | 07/06/2018 | 2018 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824090&pc=MTA | 2023826-2018-01288 | 2018-07-30 | Injury | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE PROBLEM WAS RESOLVED WITH THE LENS EXPLANT. THE CAUSE OF THE EVENT IS LISTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.0/+1.5/088 | OS | Physician | 08/03/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050569&pc=MTA | 2023826-2017-01751 |  | Injury | STAAR SURGICAL COMPANY AG | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/120 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) IN 2014. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO GLARE AND HALOS, THE PATIENT HAD LARGE PUPIL SIZE. THE LENS WAS EXCHANGED FOR ANOTHER SAME SIZE (13.2 MM) BUT DIFFERENT DIOPTER (11.0/+2.0/109) LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.50/+2.0/120 | OD | Physician | 10/23/2017 | 2017
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6454953&pc=MTA | 2023826-2017-00563 | 2016-12-21 | Injury | STAAR SURGICAL COMPANY | 2017-04-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17/3.0/79 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. SURGEON STATED THAT THIS WAS DUE TO "BIG PUPIL," WHICH HE TOLD THE PATIENT WILL GET BETTER ON ITS OWN. AFTER EXPLANT, THE EYE IS FINE AND THE SURGEON DOES NOT PLAN TO EXCHANGE THE LENS. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND SCRATCHES ON A HAPTIC. THERE WAS EVIDENCE OF CLEAR AND DARK BROWN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | -17/3.0/79 | OD | Physician | 03/07/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6454975&pc=MTA | 2023826-2017-00564 | 2016-12-21 | Injury | STAAR SURGICAL COMPANY | 2017-04-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+2.5/096 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. SURGEON STATED THAT THIS WAS DUE TO "BIG PUPIL," WHICH HE TOLD PATIENT WILL GET BETTER ON ITS OWN. AFTER EXPLANT, THE EYE IS FINE AND THE SURGEON DOES NOT PLAN TO EXCHANGE THE LENS. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF DEBRIS ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.50/+2.5/096 | OS | Physician | 03/07/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7201664&pc=MTA | 2023826-2018-00094 | 2017-06-20 | Injury | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -3.5/+1.5/001 (SPHERE/ CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED. REPORTEDLY, THE PATIENT EXPERIENCED "READING PROBLEMS" AND LENS WAS EXPLANTED "DESPITE VERY GOOD VA." Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND PRESENCE OF FOREIGN MATERIAL (RESIDUE) ON LENS. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | -3.5/+1.5/001 | OD | Physician | 12/22/2017 | 2017 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7201666&pc=MTA | 2023826-2018-00096 | 2017-07-06 | Injury | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -3.5/+1.5/160 (SPHERE/ CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED. REPORTEDLY, THE PATIENT EXPERIENCED "READING PROBLEMS" AND LENS WAS EXPLANTED "DESPITE VERY GOOD VA." Manufacturer Narrative: SHOULD HAVE BEEN RECORDED IN INITIAL MDR. THE PATIENT EXPERIENCED GLARE/HALOES ALONG WITH "READING PROBLEMS." (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). CLAIM # (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | -3.5/+1.5/160 | OS | Physician | 12/22/2017 | 2017 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7456987&pc=MTA | 2023826-2018-00595 | 2017-09-26 | Injury | STAAR SURGICAL COMPANY | 2018-04-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -9.50/+1.00/85 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018, THE LENS WAS EXPLANTED DUE TO GLARES AND HALOES. IT WAS REPORTED THAT THE PATIENT HAS RECOVERED AND THERE IS NO MORE PROBLEMS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.50/+1.00/85 | OS | Physician | 03/27/2018 | 2018 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6108232&pc=MTA | 2023826-2016-01649 | 2016-10-17 | Injury | STAAR SURGICAL COMPANY | 2016-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+2.00/092 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE REPORTER INDICATED THE PATIENT HAS SIGNIFICANT PROBLEMS WITH GLARE/HALOS. THE LENS REMAINS IMPLANTED BUT THE PLAN IS TO EXPLANT/EXCHANGE FOR A DIFFERENT MODEL LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR ANOTHER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. FOREIGN MATERIAL (FIBERS) WAS PRESENT ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). CORRECTED FROM USER FACILITY TO DISTRIBUTOR. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.50/+2.00/092 | OD | Physician | 10/17/2016 | 2016 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6108233&pc=MTA | 2023826-2016-01650 | 2016-10-17 | Injury | STAAR SURGICAL COMPANY | 2016-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+2.00/095 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER INDICATED THE PATIENT HAD SIGNIFICANT PROBLEMS WITH GLARE/HALOS. THE LENS REMAINS IMPLANTED BUT THE PLAN IS TO EXPLANT/EXCHANGE FOR A DIFFERENT MODEL LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR ANOTHER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE AND FOREIGN MATERIAL - FIBERS PRESENT ON LENS SURFACE. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTIONAL DATA: THE LENS WAS EXCHANGED FOR ANOTHER LENS BUT THE PROBLEM WAS NOT RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.00/+2.00/095 | OS | Physician | 10/17/2016 | 2016 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6701687&pc=MTA | 2023826-2017-01152 | 2015-09-01 | Injury | STAAR SURGICAL COMPANY | 2017-07-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.0/+1.0/63 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT EXPERIENCED EXCESSIVE VAULT, GLARE/HALOES, SLOW PUPILLARY REACTION, OVERCORRECTION, AND SEES "CENTERFLOW" REFLECTIVES EVEN ON DAYTIME. THE CUSTOMER STATED THAT THE SURGEON PLANNED TO EXCHANGE THE LENS FOR THE PATIENT. AS OF THE DAY OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (FIBERS). DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.0/+1.0/63 | OD | Physician | 06/13/2017 | 2017 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6801631&pc=MTA | 2023826-2017-01301 | 2016-07-11 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+3.0/104 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PATIENT EXPERIENCED GLARE AND HALOS. THE PROBLEM WAS RESOLVED AFTER THE EXCHANGE. THE PATIENT'S POST-OP SHOWED UNCORRECTED VISUAL ACUITY (UCVA) TO BE 20/25. Manufacturer Narrative: (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS WERE FOUND. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTICS SLIGHTLY CURVED. (B)(4). | -14.5/+3.0/104 | OD | Physician | 07/19/2017 | 2017 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6440720&pc=MTA | 2023826-2017-00503 | 2017-02-24 | Injury | STAAR SURGICAL COMPANY | 2017-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.5/005 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND GLARE/HALOES. ALTHOUGH THE PATIENT STILL COMPLAINED OF GLARE AND "READING PROBLEMS," THE SURGEON BELIEVED AFTER "ADAPTION," TIME, AND HEALING THE PROBLEM WOULD BE RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | -10.0/+2.5/005 | OD | Physician | 02/27/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6440729&pc=MTA | 2023826-2017-00504 | 2017-02-24 | Injury | STAAR SURGICAL COMPANY | 2017-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.5/154 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF ICA, AND GLARE/HALOES. ALTHOUGH THE PATIENT STILL COMPLAINED OF "READING PROBLEMS" AND "STRANGE" VISUAL EFFECTS AT DISTANCE, THE SURGEON STATED THAT OBJECTIVELY EVERYTHING WAS FINE AND HE BELIEVED THAT AFTER TIME AND HEALING THE PROBLEM WOULD BE RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON LENS. (B)(4). Manufacturer Narrative: ADDITIONAL DATA:WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | -10.0/+2.5/154 | OS | Physician | 02/27/2017 | 2017 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6657591&pc=MTA | 2023826-2017-01020 | 2017-03-01 | Injury | STAAR SURGICAL COMPANY | 2017-06-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/1.5/74 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON 01 MAR 2017 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION, AND GLARE/HALOES. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. (B)(4). RESULT CODE: PREVIOUS CODE NOT APPLICABLE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND RESIDUE ON LENS), THE LENS MISSING A PIECE OF HAPTIC, AND THE LENS HAS A CURVED SHAPE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.0/1.5/74 | OD | Physician | 05/26/2017 | 2017 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320737&pc=MTA | 2023826-2018-00328 | 2018-02-02 | Injury | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.00/+3.00/119 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON NOTES THE FOLLOWING: EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND GLARE/HALOES. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.85MM)]. (B)(4). | -6.00/+3.00/119 | OD | Physician | 02/08/2018 | 2018 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7516771&pc=MTA | 2023826-2018-00745 | 2018-04-03 | Injury | STAAR SURGICAL COMPANY | 2018-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/+2.0/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND GLARE/HALOS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION - LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.0/+2.0/095 | OD | Physician | 04/27/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7740014&pc=MTA | 2023826-2018-01141 | 2018-06-09 | Injury | STAAR SURGICAL COMPANY | 2018-08-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.5/+1.5/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, GLARE/HALOES, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE AND DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | -6.5/+1.5/080 | OS | Physician | 07/06/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5596417&pc=MTA | 2023826-2016-00500 | 2015-12-02 | Injury | STAAR SURGICAL COMPANY | 2016-04-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/+2.00/X019 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES, GLARE/HALOES WAS OBSERVED ON (B)(6) 2015 THE ICL WAS EXPLANTED ON (B)(6) 2016, THE INCISION WAS ENLARGED TO REMOVE THE ICL. ANTERIOR CHAMBER IRRIGATION/ELEVATION OF THE REMAINING VISCO/FLUIDS AND INTRAOCULAR INJECTION (MEDICATION) WAS PERFORMED. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. A SUTURE WAS REQUIRED. LAST VISIT ON (B)(6) 2016, BCVA: 20/20. Manufacturer Narrative: (B)(4). PATIENT IS ABOVE THE AGE OF 45 AND PER DFU FOR THIS MODEL, THIS LENS MODEL IS CONTRADICTED FOR THE PATIENTS ABOVE THE AGE OF 45 YEARS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID AND IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4) SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. NEW INCISION. SECONDARY SURGERY. ADDITIONAL SUTURE. LENS EXPLANTED. VAULTING. (B)(4). | -11.00/+2.00/019 | OS | Physician | 03/22/2016 | 2016 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6908149&pc=MTA | 2023826-2017-01499 | 2017-09-01 | Injury | STAAR SURGICAL COMPANY | 2017-10-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1, -13/+5/96 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO LOW VAULTING, LENS ROTATION, AND GLARES/HALOES. A NEW LENS IS PLANNED TO BE IMPLANTED NEXT MONTH. THE PATIENT IS CURRENTLY FINE. Manufacturer Narrative: THE NEW LENS WAS NOT IMPLANTED YET DUE TO PATIENT WORKING ISSUES. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS) AND THE LENS HAS A CURVED SHAPE. (B)(4) Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2018 WITH A VTICM5 13.2 IMPLANTABLE COLLAMER LENS -12.5/+5 /93 DIOPTER. THE SURGEON NOTES THAT THE VISUAL ACUITY AND VAULT ARE OKAY, AND NO ROTATION. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4) SECONDARY SURGERY, LENS EXPLANT. (B)(4). | -13/+5/96 | OS | Physician | 09/06/2017 | 2017 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7740026&pc=MTA | 2023826-2018-01150 | 2018-07-09 | Injury | STAAR SURGICAL COMPANY | 2018-08-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -3.5/+1.5/174 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), GLARE/HALOES, AND A "SLIGHT TILT." THE CAUSE IS LISTED AS "SIZING." THE LENS REMAINS IMPLANTED AND PATIENT IS UNDER EVALUATION. Manufacturer Narrative: STATES THE LENS REMAINS IMPLANTED. CORRECTED INFORMATION ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. THERE WAS CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -3.5/+1.5/174 | OS | Physician | 07/09/2018 | 2018 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8191674&pc=MTA | 2023826-2018-01877 | 2018-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO 13.2, -5.0/+1.50/5 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT AND GLARE/HALOES. A REPLACEMENT LENS, SHORTER LENGTH WAS LATER IMPLANTED IN A SECOND SURGERY. IT WAS REPORTED THAT THE REPLACEMENT RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTION PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXCHANGE. PATIENT CODE 2140 - 2227 IN THE PREVIOUS MDR SHOULD BE INCLUDED IN THE ORIGINAL MDR. DEVICE CODE 3001 IN THE PREVIOUS MDR SHOULD BE INCLUDED IN THE ORIGINAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.0/+1.50/5 | OD |  | 11/21/2018 | 2018 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7853013&pc=MTA | 2023826-2018-01336 | 2018-08-07 | Injury | STAAR SURGICAL COMPANY | 2018-09-06 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -3.0/+4.0/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, THE PATIENT IS EXPERIENCING GLARE AND "HALOS DUE TO THE AQUAPORT". THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -3.00/+4.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY THE PATIENT EXPERIENCED GLARES AND "HALOS DUE TO AQUAPORT". THE LENS WAS EXCHANGED ON (B)(6) 2018. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S., LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -3.0/+4.0/095 | OS | Physician | 08/09/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7853023&pc=MTA | 2023826-2018-01333 | 2018-08-07 | Injury | STAAR SURGICAL COMPANY | 2018-09-06 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -2.5/+2.5/098 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. REPORTEDLY, THE PATIENT IS EXPERIENCING GLARES AND "HALOS DUE TO THE AQUAPORT." THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -2.50/+2.5/098 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. REPORTEDLY THE PATIENT EXPERIENCED GLARES AND "HALOS DUE TO AQUAPORT". THE LENS WAS EXCHANGED ON (B)(6) 2018. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION OF THE LENS FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -2.5/+2.5/098 | OD | Physician | 08/09/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8365953&pc=MTA | 2023826-2019-00268 | 2019-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-25 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6, -13.50/3.0/098 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT, PIGMENT DISPERSION, UNREACTIVE(FIXED) PUPIL, AND GLARE/HALOES WAS OBSERVED. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.50/3.0/098 | OD |  | 02/13/2019 | 2019 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6110075&pc=MTA | 2023826-2016-01634 | 2016-04-06 | Injury | STAAR SURGICAL COMPANY | 2016-11-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Optical Decentration; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -13.00/+3.5/085 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARES/HALOES, BLURRED VISION, AND THE LENS BEING DECENTERED. THE LENS WAS EXCHANGED FOR A LENS OF THE SAME MODEL AND SIZE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN CASE/VIAL; SURGICAL RESIDUE WAS CLEAR; VISUAL INSPECTION FOUND NO VISIBLE DAMAGE PREVIOUS DEVICE CODE (B)(4) WAS INCORRECT; REPLACED BY (B)(4) METHOD CODE WORK ORDER SEARCH PERFORMED AND NO SIMILAR COMPLAINTS FOUND CONCLUSION CODE PATIENT ACD WAS BELOW 3.0 MM. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.00/+3.5/085 | OS | Physician | 06/01/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8176968&pc=MTA | 2023826-2018-01855 | 2016-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.0/+2.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO GLARE/HALOS. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL LENS (VTICM). THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR AND DEVICE. Manufacturer Narrative: EVENT: THE REPORTER INDICATED THE LENS EXCHANGE RESOLVED THE PROBLEM. DEVICE EVALUATED BY MFR : DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE/RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM #(B)(4). Manufacturer Narrative: DATE OF EVENT CORRECTED FROM (B)(6) 2016 TO UNK. DEVICE EVALUATION- LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND RESIDUE ON LENS SURFACE AND NO VISIBLE DAMAGE TO THE LENS. PATIENT CODE- 2227 ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.0/+2.5/092 | OD |  | 11/19/2018 | 2018 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8255318&pc=MTA | 2023826-2019-00107 | 2018-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+2.0/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT EXPERIENCED GLARE/HALOS. THE LENS WAS EXCHANGED FOR A PHAKIC INTRAOCULAR LENS (PIOL). THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED ON 12-FEB-2019 STATING THE REPLACEMENT LENS WAS AN "ARTIFLEX". DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.; (B)(4). | -18.00/+2.0/098 | OD |  | 12/17/2018 | 2018 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8255338&pc=MTA | 2023826-2019-00109 | 2018-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -3.5/+6.0/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT EXPERIENCED GLARE/HALOS DUE TO BIG PUPIL SIZE. THE LENS WAS NOT REPLACED WITH ANOTHER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: ADDITIONAL DATA: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -3.5/+6.0/098 | OS |  | 12/17/2018 | 2018 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320749&pc=MTA | 2023826-2018-00324 | 2016-12-01 | Injury | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.0/+1.0/041 (SPHERE/ CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED. REPORTEDLY, THE PATIENT EXPERIENCED "PROBLEMS READING AND DOING COMPUTER WORK." THE SURGEON STATES THE EYE IS "CALM AND QUIET" AFTER THE EXCHANGE AND THAT THE PROBLEM IS RESOLVED. Manufacturer Narrative: - DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN/BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. ALTHOUGH LENS SN: (B)(4) WAS FOUND OUT OF TOLERANCE FOR AXIS, THERE IS NO REASONABLE PROBABLE CAUSE IN MANUFACTURING THAT WOULD HAVE RESULTED IN A REFRACTIVE CHANGE. WORK ORDER SEARCH FOUND NO OTHER SIMILAR DEFECT WITHIN THE SAME LOT. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND "PROBLEMS READING AND DOING COMPUTER WORK." (B)(4) Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.0/+1.0/041 | OS | Physician | 02/08/2018 | 2018 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6598926&pc=MTA | 2023826-2017-00901 | 2017-03-04 | Injury | STAAR SURGICAL COMPANY | 2017-05-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Visual Disturbances; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.50/+1.5/84 IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED GLARE AND UNREACTIVE PUPIL (FIXED). AS OF THE DATE OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. THIS CASE IS UNDER MEDICAL CONSULT. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -13.50/+1.5/84 | OD | Physician | 05/03/2017 | 2017 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320720&pc=MTA | 2023826-2018-00321 | 2017-08-22 | Injury | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.0/+1.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXPLANTED, REPORTEDLY DUE TO OCCASIONAL "FLASHING LIGHTS." ATTEMPTS TO OBTAIN PATIENT'S CURRENT STATUS HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE PATIENT COMPLAINS OF CHROMATIC ABERRATION AND COLOR DISTORTION. REPORTEDLY THIS DID NOT OCCUR PRIOR TO ICL IMPLANT. PATIENT STATES THESE SYMPTOMS, INCLUDING "FLASHES OF LIGHT," PERSIST AFTER ICL EXPLANT. THE PATIENT STATES THESE DISTURB HER DAILY LIFE. OCT AND FUNDUS CAMERA TESTS WERE TAKEN TO ENSURE THE BACK OF THE EYE IS IN GOOD SHAPE. APPARENTLY, THE PATIENT'S FEEDBACK KEEPS CHANGING AND INVOLVED SURGEONS BELIEVE THE PATIENT MAY BE "MENTALLY UNSTABLE." THE SURGEON REPORTS THAT HE HAS DISCHARGED THE PATIENT. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC AND OPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4) Manufacturer Narrative: (B)(4). CONCLUSION: IT IS ADDED TO THE PREVIOUS SUPPLEMENTAL MDR. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+1.5/090 | OD | Physician | 02/07/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320758&pc=MTA | 2023826-2018-00323 | 2017-08-22 | Injury | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.5/+1.5/082 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXPLANTED, REPORTEDLY DUE TO OCCASIONAL "FLASHING LIGHTS." ATTEMPTS TO OBTAIN PATIENT'S CURRENT STATUS HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE PATIENT COMPLAINS OF CHROMATIC ABERRATION AND COLOR DISTORTION. REPORTEDLY THIS DID NOT OCCUR PRIOR TO ICL IMPLANT. PATIENT STATES THESE SYMPTOMS, INCLUDING "FLASHES OF LIGHT," PERSIST AFTER ICL EXPLANT. THE PATIENT STATES THESE DISTURB HER DAILY LIFE. OCT AND FUNDUS CAMERA TESTS WERE TAKEN TO ENSURE THE BACK OF THE EYE IS IN GOOD SHAPE. APPARENTLY, THE PATIENT'S FEEDBACK KEEPS CHANGING AND INVOLVED SURGEONS BELIEVE THE PATIENT MAY BE "MENTALLY UNSTABLE." THE SURGEON REPORTS THAT HE HAS DISCHARGED THE PATIENT. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). CONCLUSION: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.94MM)]. (B)(4). | -12.5/+1.5/082 | OS | Physician | 02/07/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6917932&pc=MTA | 2023826-2017-01501 | 2017-08-31 | Injury | STAAR SURGICAL COMPANY | 2017-10-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE. THE LENS WAS EXPLANTED A YEAR AFTER THE IMPLANTATION AS THE PATIENT WAS BOTHERED BY THE CENTRAL HOLE OF THE LENS AS THE LIGHT DISRUPTION FROM THE HOLE DIDN'T GET BETTER. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -14.0/+1.0/105 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXPLANTED AS THE PATIENT WAS BOTHERED BY THE CENTRAL HOLE OF THE LENS AS THE LIGHT DISRUPTION FROM THE HOLE DIDN'T GET BETTER. THERE IS NO PLAN TO IMPLANT ANOTHER LENS. PATIENT POST-OP BCVA WAS 20/20 ON (B)(6) 2017. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS FOUND. (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (MANUFACTURING SITE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM #: (B)(4). | -14.0/+1.0/105 | OD | Physician | 09/06/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7152459&pc=MTA | 2023826-2017-01966 | 2017-11-28 | Injury | STAAR SURGICAL COMPANY | 2017-12-28 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/072 (SPHERE/ CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED. REPORTEDLY, THE PATIENT EXPERIENCED "VERY STRONG GLARES," AND GHOST IMAGES. AFTER EXPLANT, IT IS NOTED THAT "AS BEFORE SURGERY CRYSTALLINE LENS IS CLEAR." Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THE LENS EXPLANTATION RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+2.0/072 | OS | Physician | 12/01/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7152539&pc=MTA | 2023826-2017-01965 | 2017-11-28 | Injury | STAAR SURGICAL COMPANY | 2017-12-28 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/093 (SPHERE/ CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED. REPORTEDLY, THE PATIENT EXPERIENCED "VERY STRONG GLARES," AND GHOST IMAGES. AFTER EXPLANT, IT IS NOTED THAT "AS BEFORE SURGERY...LENS IS CLEAR." Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMAGE. (B)(4) Manufacturer Narrative: THE LENS EXPLANT RESOLVED THE PROBLEM. (B)(4) Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+2.0/093 | OD | Physician | 12/01/2017 | 2017 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7245216&pc=MTA | 2023826-2018-00198 | 2017-12-29 | Injury | STAAR SURGICAL COMPANY | 2018-02-06 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.50/+2.50/105 (SPHERE/ CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH THE SAME SIZE, DIFFERENT DIOPTER LENS. REPORTEDLY, THE PATIENT WAS UNHAPPY POSTOPERATIVELY AND THE SURGEON DECIDED TO CHANGE THE LENS "BASED ON EITHER WRONG ASIGMATS AND/OR WRONG AXIS" OF THE LENS. THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL AND HAD CLEAR SURGICAL RESIDUE OR DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND PRESENCE OF CLEAR AND WHITE RESIDUE AND DEBRIS ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | -6.50/+2.50/105 | OD | Physician | 01/10/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7245212&pc=MTA | 2023826-2018-00197 | 2018-01-02 | Injury | STAAR SURGICAL COMPANY | 2018-02-06 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/091 (SPHERE/ CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH THE SAME SIZE, DIFFERENT DIOPTER LENS. REPORTEDLY, THE PATIENT WAS UNHAPPY POSTOPERATIVELY AND THE SURGEON DECIDED TO CHANGE THE LENS "BASED ON EITHER WRONG ASIGMATS AND/OR WRONG AXIS" OF THE LENS. THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH FOREIGN MATERIAL ON LENS SURFACE (CLEAR AND WHITE RESIDUE AND DEBRIS). (B)(4) Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.95MM)] (B)(4). | -11.0/+2.0/091 | OS | Physician | 01/10/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7395672&pc=MTA | 2023826-2018-00502 | 2018-01-17 | Injury | STAAR SURGICAL COMPANY | 2018-04-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.0/+3.5/89 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXPLANTED AND EXCHANGED DUE TO GLARE/HALOES. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS FOUND. (B)(4). | -12.0/+3.5/89 | OS | Physician | 03/07/2018 | 2018 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7897512&pc=MTA | 2023826-2018-01435 | 2018-04-16 | Injury | STAAR SURGICAL COMPANY | 2018-09-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.00/+3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND GLARE/HALOES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. (B)(4). | -12.00/+3.0/093 | OD | Physician | 08/22/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6562054&pc=MTA | 2023826-2017-00799 | 2017-03-07 | Injury | STAAR SURGICAL COMPANY | 2017-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Migration or Expulsion of Device; Inadequacy of Device Shape and/or Size | Visual Disturbances; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -805/+1.0/070 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SAME SIZE, SIMILAR DIOPTER LENS DUE TO THE PATIENT EXPERIENCING LOW VAULT AND GHOSTING (DOUBLE VISION). THE SURGEON DECIDED TO IMPLANT THE LENS HORIZONTALLY, STATING THE LENS OPTICAL ZONE HAD MIGRATED DOWNWARD. ALTHOUGH THERE IS STILL LOW VAULT, THE PATIENT AND SURGEON ARE HAPPY WITH THE RESULTS. AS OF THE DATE OF MDR REPORTING THE LENS HAS NOT YET BEEN RETURNED. Manufacturer Narrative: PREVIOUS DIOPTER: -805/+1.0/070; CORRECT DIOPTER: -8.5/+1.0/070. THE LENS HAS BEEN RETURNED AND EVALUATED. DEVICE EVALUATION: LENS WAS RETURNED DRY IN A SMALL VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC MISSING A PIECE. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: PREVIOUS DIOPTER: -805/+1.0/070; CORRECT DIOPTER: -8.5/+1.0/070. THE LENS HAS BEEN RETURNED AND EVALUATED. DEVICE EVALUATION: LENS WAS RETURNED DRY IN A SMALL VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC MISSING A PIECE. CLAIM# (B)(4). | -8.5/+1.0/070 | OS | Physician | 04/13/2017 | 2017 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7006956&pc=MTA | 2023826-2017-01685 | 2017-08-31 | Injury | STAAR SURGICAL COMPANY | 2017-11-07 | MTA | IMPLANTABLE COLLAMER LENS | Patient-Device Incompatibility; Device Operates Differently Than Expected | Visual Disturbances; No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.0/+1.0/067 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO DYSPHOTOPSIA; THE PATIENT WAS "BOTHERED BY THE CENTRAL HOLE CONSTANTLY, WHICH HE COULD NOT TOLERATE." THE EXPLANT RESOLVED THE PROBLEM. NO LENS EXCHANGE IS PLANNED. Manufacturer Narrative: (B)(4). RESULT: NO SIMILAR COMPLAINT TYPE EVENTS FOUND. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | -13.0/+1.0/067 | OS | Physician | 10/12/2017 | 2017 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119133&pc=MTA | 2023826-2015-01309 | 2015-09-01 | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Visual Impairment | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/X145 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THERE WAS LOSS OF BEST-CORRECTED VISUAL ACUITY POST-OP. THE LENS REMAINS IMPLANTED. THERE WERE POST-OP DEPOSITS ON THE SURFACE OF THE CRYSTALLINE LENS (VISCO) WHICH DECREASED THE VISUAL ACUITY . Manufacturer Narrative: METHOD (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: INVESTIGATION SUMMARY: THE MANUFACTURER'S DEVICE ANALYSIS RESULTS - THE EVENTS REPORTED BY PROF. (B)(6) ARE LIMITED TO A SINGLE SITE BY A SINGLE SURGEON IMPLANTING IMPLANTABLE COLLAMER LENSES (LCLS) BETWEEN THE DATES OF (B)(6) 2015. TEN PATIENTS WERE REPORTED TO HAVE MILKY WHITE PRECIPITATES FORM ON THE SURFACE OF THE ICL AND ON THE ANTERIOR LENS CAPSULE OF THE PHYSIOLOGIC CRYSTALLINE LENS. NO ADDITIONAL REPORTS OF PRECIPITATES HAVE BEEN RECEIVED FROM ANY OTHER SURGEON OR SITE WITH THE SAME LOT NUMBERS AS THE REPORTED EVENTS FROM PROF. (B)(6). ADDITIONALLY, IT IS IMPORTANT TO NOTE THAT IN THE DIRECTIONS FOR USE (DFU) OF THE LCL, THE COMPLICATIONS AND ADVERSE REACTIONS SECTION IDENTIFIES THE POTENTIAL RISK FOR "SECONDARY SURGICAL INTERVENTION TO REMOVE/REPLACE THE LENS". THE PRECAUTIONS SECTION ALSO SUGGESTS "THE LENS SHOULD NOT BE EXPOSED TO ANY SOLUTIONS OTHER THAN THE NORMALLY USED INTRAOCULAR IRRIGATING SOLUTIONS (E.G. ISOTONIC SALINE, BSS, VISCOELASTIC. ETC.)." DURING THE IN-DEPTH INVESTIGATION WITH PROF. (B)(6) TO IDENTIFY ROOT CAUSE, THE INTRAOPERATIVE REGIMEN REVEALED THAT INTRACAMERAL 0.1% CEFUROXIME ANTIBIOTIC FOR PROPHYLAXIS WAS USED IN THE 10 REPORTED CASES, ALONG WITH RINGER'S SOLUTION. LN AN ARTICLE PUBLISHED BY VASAVADA ET AL. (COMPARISON BETWEEN RINGER'S LACTATE AND BALANCED SALT SOLUTION ON POSTOPERATIVE OUTCOMES AFTER PHACOEMULSIFICATION: A RANDOMIZED CLINICAL TRIAL, INDIAN J OPHTHALMOL. 2009 MAY-JUN; 57(3): 191-195), THE AUTHORS CONCLUDE THAT RINGER'S LACTATE IS HYPOTONIC AND SLIGHTLY ACIDIC (OSMOLALITY 280MMOL, PH 6.0) AS COMPARED TO BSS (OSMOLALITY 302 MMOL, PH 7.4) AND AQUEOUS (OSMOLALITY 302 MMOL, PH 7.4). THEY FURTHER CONCLUDED THAT A STATISTICALLY SIGNIFICANT LARGER NUMBER OF PATIENTS WITH FLARE AND CELLS MAY BE DUE TO THE SLIGHTLY ACIDIC FLUCTUATING PH AND LACK OF BUFFER SYSTEM IN RINGER'S LACTATE, WHICH COULD AFFECT THE BLOOD AQUEOUS BARRIER STABILITY. ON THE OTHER HAND, 855 HAS AN ACETATE CITRATE BUFFER SYSTEM AND THE PH, THOUGH SLIGHTLY ALKALINE, IS STABLE. SUBSEQUENT TO THE SERIES OF REPORTED EVENTS, ADDITIONAL PATIENTS HAVE BEEN TREATED BY PROF. (B)(6) WITH AN ICL USING ONLY COMMERCIALLY AVAILABLE BALANCED SALINE SOLUTION (BSS, ALCON) AND OCUCOAT (BAUSCH & LOMB) INTRAOPERATIVE, WITHOUT THE USE OF 0.1% INTRACAMERAL CEFUROXIME ANTIBIOTIC. THERE WERE NO OBSERVED INCIDENTS OF DEPOSITS FOLLOWING THIS PROTOCOL REGIMEN. UPON REVIEW OF THE DEVICE HISTORY RECORDS (DHR), ALL PRODUCTS REPORTED IN THE INCIDENCES, BASED UPON LOT AND SERIAL NUMBER REVIEW, INDICATED SPECIFICATIONS WERE MET DURING MANUFACTURING PROCESS. ALL PRODUCTS PASSED QUALITY INSPECTIONS AND WERE RELEASED PER ACCEPTANCE CRITERIA. FINAL COMMENTS FROM THE MANUFACTURER: BASED UPON A THOROUGH REVIEW OF ALL AVAILABLE DATA AND ON CLINICAL PERSPECTIVE AS A PRACTICING OPHTHALMIC SURGEON WITH EXPERTISE IN THE LCL, DR. (B)(6) (MEDICAL CONSULTANT) AND STAAR CONCLUDED THAT THE OBSERVED PRECIPITATES WERE MOST LIKELY THE RESULT OF INTERACTION BETWEEN THE INTRAOCULAR MEDICATIONS USED AT THE TIME OF SURGERY. CLAIM # (B)(4). Manufacturer Narrative: THE CUSTOMER HAS REPORTED SINCE USING THE NEW SURGICAL PROTOCOL (USING BSS AND OCUCOAT ONLY) NO NEW COMPLAINT CASES OF LOSS OF BCVA, WITH PRECIPITATE ON THE CRYSTALLINE LENS, HAVE BEEN RECEIVED SINCE THE BEGINNING OF 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO -LENS IMPLANTED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | -11.0/+1.5/145 | OS | Physician | 09/03/2015 | 2015 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6024707&pc=MTA | 2023826-2016-01453 | 2016-09-22 | Injury | STAAR SURGICAL COMPANY | 2016-10-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | Visual Impairment | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+1.5/092 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016, AND THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THE PATIENT EXPERIENCED LOSS OF BEST CORRECTED VISUAL ACUITY (BCVA). THE LENS WILL BE TAKEN OUT AT A LATER TIME.Event Description: THE LENS WAS EXCHANGED FOR A SAME SIZE LENS ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND OPTIC TORN, HAPTIC TORN, AND DISCOLORED MATERIAL ON SURFACE. PREVIOUS PATIENT CODE (B)(4) (VISION, IMPAIRED) WAS INCORRECT; ONLY LOSS OF BCVA WAS REPORTED. METHOD CODE (DEVICE FROM SAME LOT EVALUATED); A WORK ORDER SEARCH WAS PERFORMED. (B)(4). | -17.0/+1.5/092 | OS | Physician | 10/03/2016 | 2016 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119420&pc=MTA | 2023826-2015-01311 | 2015-07-22 | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Visual Impairment; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/X132 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SIMILAR LENS WITH A DIFFERENT DIOPTER. THE EXCHANGE RESOLVED THE PROBLEM. THE PATIENT'S POST-OP ((B)(6) 2015) UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. SECONDARY INTERVENTION, EXCHANGE AND PROBLEM RESOLVED. AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS OVER 45 YEARS OLD. (B)(4). Manufacturer Narrative: EVALUATION RESULTS: UPON REVIEW OF THE DEVICE HISTORY RECORD, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON THE ABOVE INVESTIGATION, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. Manufacturer Narrative: (B)(4): METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. LENS IMPLANTED. | -9.5/+1.0/132 | OS | Physician | 09/09/2015 | 2015 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5365600&pc=MTA | 2023826-2016-00009 | 2015-11-11 | Malfunction | STAAR SURGICAL COMPANY | 2016-01-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Fracture; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Visual Impairment; Visual Disturbances; Halo | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.50/+2.5/X099 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS TORE DURING INSERTION INTO THE EYE. A SIGNIFICANT TEAR WAS NOTICED AFTER IMPLANT. LENS WAS LEFT IN THE EYE WITH SUBSEQUENT SYMPTOMS (HALOS AND OCCASIONAL DIPLOPIA) AND DISPLACEMENT (LENS ROTATION). PATIENT LAST OFFICE VISIT ON (B)(6) 2015, UCVA, 20/50. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -9.50/+2.50/X099. DEVICE EVALUATED BY MANUFACTURER? NO, DEVICE REMAINS IMPLANTED. (B)(4). METHOD CODE(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODE(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). DEVICE REMAINS IMPLANTED. | -9.50/+2.5/099 | OD | Physician | 12/15/2015 | 2015 | 23