## IOP Issue-1
Web Address | Report Number | Event Date | Event Type | Manufacturer | Date Received | Product Code | Brand Name | Device Problem | Patient Problem | PMA/PMN Number | Exemption Number | Number of Events | Event Text | Management | Outcomes
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12755755&pc=MTA | 2023826-2021-04210 | 2021-07-04 | Injury | STAAR SURGICAL COMPANY | 2021-11-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; Eye Pain | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -6.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "THE PATIENT HAS A VAGUE OF THE LEFT EYE BRINGING A VERY SERIOUS IMPACT ON THEIR LIVES TO MAKE IT VERY PAINFUL AND ANXIOUS, COMMUNICATED, AND DECIDED TO REMOVE CRYSTALS." Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISIBLE INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC BROKEN. UNABLE TO PERFORM DIMENSIONAL AND FUNCTIONAL INSPECTION DUE TO SIGNIFICANT DAMAGE OF THE LENS OPTIC PART. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6054105&pc=MTA | 2023826-2016-01523 | 2016-10-10 | Injury | STAAR SURGICAL COMPANY | 2016-10-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract, Induced; Corneal Edema; Hyphema; Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM MICL12.1 LENS, -3.50 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON THE SAME DAY DUE TO EXCESSIVE VAULTING, RESULTING IN ELEVATED INTRAOCULAR PRESSURE (IOP), SHALLOW ANTERIOR CHAMBER, CORNEAL EDEMA, HYPHEMA, AND DEVELOPMENT OF A CATARACT. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS VIAL. VISUAL INSPECTION FOUND A SMALL TEAR ON THE HAPTIC. (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED. | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218705&pc=MTA | 2023826-2015-01478 | 2015-08-10 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract, Induced; Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE DEVELOPMENT OF AN ANTERIOR SUBCAPSULAR OPACITY, PUPIL BLOCK, ELEVATED INTRAOCULAR PRESSURE AND ANGLE CLOSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: ANTERIOR CHAMBER IRRIGATION/ EVACUATION OF REMAINING VISCO/ FLUIDS OCCURRED WITH EXPLANT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE FOREIGN MATERIAL ON LENS SURFACE: FIBERS. (B)(4): ANTERIOR CHAMBER IRRIGATION/ EVACUATION OF REMAINING VISCO/ FLUIDS. CONCLUSION: THE PATIENT IS UNDER THE AGE OF 21 YEARS. THA PATIENT'S EYE HAS ANTERIOR CHAMBER DEPTH (ACD), LESS THAN 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5816274&pc=MTA | 2023826-2016-00974 | 2016-06-13 | Injury | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; Corneal Edema; Intraocular Pressure Increased; Blurred Vision; Loss of Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO AN EXCESSIVE VAULT, LENS OPACITY (CATARACT), AND LOSS OF VISION. THE PATIENT EXPERIENCED AN ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE, BLURRED VISION, IRIS ATROPHY, AND CORNEAL EDEMA. THE ANTERIOR CHAMBER IRRIGATION OF REMAINING VISCO FLUIDS WAS PERFORMED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/200. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND OPTIC TORN/SPLIT AND SPOT ON LENS SURFACE. DIMENSIONAL INSPECTION CONLUDED THAT THE LENS IS IN SPECIFICATION. (B)(4). Manufacturer Narrative: NO REPORTED LOSS OF VISION. DEVICE EVALUATION: VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11435324&pc=MTA | MW5099834 | 2019-06-03 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-05 | MTA | STAAR SURGICAL ICL IMPLANT | Insufficient Information | Cataract; Foreign Body Sensation in Eye; Itching Sensation; Blurred Vision; Excessive Tear Production; Eye Pain | NA | NA | 1 | Event Description: I HAD ICL SURGERY ON (B)(6) 2009. IN 2011, I STARTED TO HAVE ISSUES WITH MY VISION. I HAD TO START WEARING EYE GLASSES TO DRIVE AND THEN ENDED UP WITH -1.25 VISION TO THIS DATE. IN (B)(6) 2019 I STARTED TO HAVE AN ITCHING/POKING FEELING IN MY RIGHT EYE. MY EYE CONTINUED TO WATER AND ITCH AS IF SOMETHING WAS POKING ME. I VISITED SEVERAL EYE DOCTORS, INCLUDING MY SURGEON WHO HAD DONE MY ICL SURGERY. IT FEELS/FELT AS IF SOMETHING WAS WRONG WITH MY ICL. THE PAIN IS BEHIND MY EYE. THE SURGEON STATED THAT I HAVE A SMALL CATARACT NOW ON MY EYE AND THAT IS CAUSING MY BLURRY VISION AND THE ISSUE. I AM HAVING BLURRY VISION, BUT I AM STILL WITNESSING THE WATERY/ITCHING SENSATION. THE SURGEON DR. LIU ADVISED I NEED TO GET MY ICL IMPLANT TAKEN OUT AND INPUT A CATARACT LENSES. I STATED THAT I DO NOT FEEL THIS IS RIGHT. SOMETHING IS NOT RIGHT WITH THE CURRENT ILC LENSES INSIDE MY EYE. I HAVE SINCE DEALT WITH THIS CONTINUED EFFECT FOR THE PAST TWO YEARS. EYE DOCTORS CANNOT FIGURE IT OUT. IT IS NOT MY CATARACT CAUSING THIS SENSATION. IT'S ALMOST LIKE MY ICL IS TRYING TO POKE OUT OF MY EYE OR COME OUT. I AM FINALLY GETTING A SECOND OPINION THIS NEXT WEEK WITH ANOTHER EYE SURGEON TO CHECK MY ICL. TO SEE IF THAT IS IN FACT THE ISSUE, AS MY RIGHT EYE IS STILL ITCHING, WATERY. I CAN'T PUT A CONTACT IN MY RIGHT EYE AS IT IS TOO BOTHERSOME. I WANT TO REPORT THIS AS I FEEL AFTER THIS MANY YEARS PERHAPS THE LENSES HAS GONE BAD OR IT NEEDS REPLACED AFTER A CERTAIN TIME. WHEN I DID ICL BACK IN 2009 IT WAS NEW AND JUST BECOMING A THING. PERHAPS THE ICL'S BACK THEN ARE NOT AS UP TO DATE AS WHAT THEY HAVE CREATED NOW. STAAR SURGICAL ICL IMPLANTS. IMPLANT ICL: MICL12.1 RIGHT EYE. DIOPTER. -10.0/ OPTIC DIAMETER 4.9-5.8MM OVERALL DIAMETER: 12.1M. SN: (B)(4) LEFT EYE: DIOPTER.-10.5/ OPTIC DIAMETER 4.9- 5.8MM OVERALL DIAMETER: 12.2M SN: (B)(4). FDA SAFETY REPORT ID # (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11620596&pc=MTA | 2023826-2021-00861 | NA | Injury | STAAR SURGICAL COMPANY | 2021-04-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract; Foreign Body Sensation in Eye; Itching Sensation; Blurred Vision; Excessive Tear Production; Eye Pain | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2009. REPORTEDLY, THE PATIENT DEVELOPED LENS OPACITY CORTICAL, REFRACTIVE CHANGE OVER TIME, THE PATIENT HAD TO WEAR GLASSES TO DRIVE, A WATERY/ITCHING/POKING FEELING AND PAIN IN THE EYE. BLURRY VISION ALSO REPORTED. THE SURGEON STATES THE PATIENT'S ACD CHANGED FROM 3.93 IN 2009 TO 3.39 IN 2020. THE PATIENT'S AXIAL LENGTH INCREASED. "NO ONE HAS CONTROL OVER THAT!" Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9814249&pc=MTA | 2023826-2020-00475 | 2020-01-17 | Injury | STAAR SURGICAL COMPANY | 2020-03-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Headache; Blurred Vision; Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE, BLURRED VISION, LOSS OF BEST CORRECTED VISUAL ACUITY AND LENS OPACITY (POSTERIOR SUBCAPSULAR) WAS NOTED. THE PATIENT IS SENSITIVE TO BRIGHT LIGHT AND HEADACHE WHEN READING. THE LENS REMAINED IMPLANTED AND PATIENT WILL RETURN IN 3 MONTHS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE LENS WAS EXPLANTED AND REPLACED WITH AN INTRAOCULAR LENS ON (B)(6) 2020. CATARACT SURGERY WAS ALSO PERFORMED ON (B)(6) 2020, DURING LENS REPLACEMENT. THE REPORTER INDICATED THAT THE PATIENT WAS SEEN ON (B)(6) 2020, AND THAT SHE IS "STABLE FOLLOWING SURGERY, VISION IS BETTER, PRESENTLY THIS LOOKS CLEARER THAN RIGHT AFTER HER PROCEDURE." H6 - HEALTH EFFECT IMPACT CODE: 4625 - CATARACT SURGERY, CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9107373&pc=MTA | 2023826-2019-01752 | 2019-05-30 | Injury | STAAR SURGICAL COMPANY | 2019-09-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract; Inflammation; Pain; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS -18.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. A COUPLE OF HOURS AFTER THE INITIAL SURGERY, THE PATIENT EXPERIENCED PAIN AND A NEGATIVE EFFECT ON HER VISION. A CATARACT HAD DEVELOPED AND THE LENS WAS EXPLANTED. THE CATARACT WAS REMOVED AND A MONOFOCAL LENS WAS IMPLANTED ON (B)(6) 2019. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 1932: INFLAMMATION. DEVICE CODE 1494:OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12516827&pc=MTA | 2023826-2021-03729 | 2018-06-20 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract; Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM5_12.1, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP AND CATARACT. PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED. PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6- MEDICAL DEVICE PROBLEM CODE: 1494: OFF-LABEL USE: ACD<3.0MM CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098622&pc=MTA | 2023826-2020-03254 | 2020-12-04 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method | Cataract; Intraocular Pressure Increased; Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -3.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, ELEVATED IOP, AND PUNTIFORM OPACITY IN THE POSTERIOR CAPSULE. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR: NON RECOMMENDED OVD AND EXCESSIVE USE OF OVD DUE TO A NON TOTAL DILATION DURING SURGERY. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS WAS PERFORMED. REPORTEDLY, THE IOP WAS MANAGED AFTER THE EVACUATION OF OVD. PUNTIFORM OPACITY HAS DECREASED AFTER 48 HOURS OF SURGERY. Manufacturer Narrative: B5-UPDATED INFORMATION: REPORTEDLY, THE PROBLEM IS NOW RESOLVED. (B)(4). Manufacturer Narrative: B5. ELEVATED IOP IS AN EXPECTED OR FORESEEABLE SIDE EFFECT. (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10669737&pc=MTA | 2023826-2020-02329 | 2020-08-21 | Injury | STAAR SURGICAL COMPANY | 2020-10-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Cataract; Intraocular Pressure Increased; Iritis; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO INTRAOCULAR PRESSURE SPIKE (NOT CONTROLLABLE WITH DROPS) AND LENS WAS TOO LARGE, WHICH WAS CAUSING THE LENS TO BOW TOWARDS THE IRIS, INCREASING THE PRESSURE. THERE IS NO LENS IN THE EYE CURRENTLY, SURGEON WILL PROCEED WITH CATARACT EXTRACTION AS ANTERIOR CHANGES TO THE NATURAL LENS OCCURRED AFTER EXPLANTATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A CONTACT LENS CASE. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM #: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6: METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: B5 - "LENS WAS TOO LARGE" AMMENDED TO "LENS WAS TOO SMALL" IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT CLINICAL CODE 4581: NO CODE AVAILABLE (BOWING OF IRIS) SHOULD BE INCLUDED IN INITIAL MDR. H6 - HEALTH EFFECT IMPACT CODE 4625: ADDITIONAL SURGERY (PHACOEMULSIFICATION, IOL IMPLANTATION) SHOULD BE INCLUDED IN INITIAL MDR. ADDITIONAL INFORMATION: B5 - LOW VAULT, TRANSIENT LOSS OF BCVA, GLARE/HALOES, IRITIS, ICL-LENS CONTACT, AND ICL-IRIS CONTACT WERE ALSO REPORTED. THE PROBLEM IS RESOLVED AFTER PHACOEMULSIFICATION AND IOL IMPLANTATION. THE CAUSE OF THE EVENT IS REPORTED AS "DEVICE SIZE MAY HAVE BEEN TOO SMALL." H6 - HEALTH EFFECT IMPACT CODE 4581: NO CODE AVAILABLE (ICL-LENS TOUCH, ICL-IRIS CONTACT). CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10669743&pc=MTA | 2023826-2020-02328 | 2020-08-21 | Injury | STAAR SURGICAL COMPANY | 2020-10-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Cataract; Intraocular Pressure Increased; Iritis; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO INTRAOCULAR PRESSURE SPIKE (NOT CONTROLLABLE WITH DROPS) AND LENS WAS TOO LARGE, WHICH WAS CAUSING THE LENS TO BOW TOWARDS THE IRIS, INCREASING THE PRESSURE . THERE IS NO LENS IN THE EYE CURRENTLY, SURGEON WILL PROCEED WITH CATARACT EXTRACTION AS ANTERIOR CHANGES TO THE NATURAL LENS OCCURRED AFTER EXPLANTATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A CONTACT CASE. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: B5 - "LENS WAS TOO LARGE" AMMENDED TO "LENS WAS TOO SMALL" IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT CLINICAL CODE 4581: NO CODE AVAILABLE (BOWING OF IRIS) SHOULD BE INCLUDED IN INITIAL MDR. H6 - HEALTH EFFECT IMPACT CODE 4625: ADDITIONAL SURGERY (PHACOEMULSIFICATION, IOL IMPLANTATION) SHOULD BE INCLUDED IN INITIAL MDR. ADDITIONAL INFORMATION: B5 - TRANSIENT LOSS OF BCVA, GLARE/HALOES, IRITIS, ICL-LENS CONTACT, AND ICL-IRIS CONTACT WERE ALSO REPORTED. THE PROBLEM IS REPORTED AS RESOLVED AFTER PHACOEMULSIFICATION AND IOL IMPLANTATION. HOWEVER, MILD PUPIL CORECTOPIA IS REPORTED POST EXPLANT. THE CAUSE OF THE EVENT IS REPORTED AS "DEVICE SIZE MAY HAVE BEEN TOO SMALL." H6 - HEALTH EFFECT IMPACT CODE 4581: NO CODE AVAILABLE (ICL-LENS TOUCH, ICL-IRIS CONTACT). (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6- METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7155887&pc=MTA | 2023826-2017-01982 | 2017-09-12 | Injury | STAAR SURGICAL COMPANY | 2017-12-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Cataract; Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4, IMPLANTABLE COLLAMER LENS, -18.0 DIOPTER, ON (B)(6) 2015 IN THE PATIENT'S RIGHT EYE (OD). LENS OPACITY (ASCO) WAS OBSERVED ON (B)(6) 2017,ELEVATED INTRAOCULAR PRESSURE AND REFRACTIVE CHANGE OVERTIME DUE TO CATARACT WAS ALSO NOTICED. THE LENS WAS EXPLANTED ON (B)(6) 2017, PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS REPORTED. Manufacturer Narrative: FINAL PATIENT STATUS IS REPORTED AS "GOOD RESULT, PATIENT IS OK" WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: REPORTER STATED THAT "THE LENS IS NOT AVAILABLE (DISCARDED)". DATE OF THIS REPORT: IT IS CORRECTED FROM "30-NOV-2017" TO "29-NOV-2017" IN THE INITIAL MDR. DATE RECEIVED BY MANUFACTURER: IT IS CORRECTED FROM "30-NOV-2017" TO "29-NOV-2017" IN THE INITIAL MDR. DEVICE MANUFACTURE DATE: IT IS CORRECTED FROM "19-JUN-2017" TO "19-JUN-2014". (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8772247&pc=MTA | 2023826-2019-01176 | 2019-06-08 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Material Opacification; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Cataract; Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -8.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. "DUE TO FLOPPY IRIS AND HIGH PRESSURE RESULTING IN ANGLE CLOSURE", THE LENS WAS REMOVED ON (B)(6) 2019. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE CODE: 1426, SHOULD BE REMOVED FROM SUPPLEMENTAL #2 AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE: 1426 SHOULD HAVE BEEN ADDED IN SUPPLEMENTAL MEDWATCH REPORT. WORK ORDER SEARCH: "NO SIMILAR COMPLAINTS" SHOULD BE CORRECTED TO "THREE SIMILAR COMPLAINTS." RESULT CODE: 213 SHOULD BE CORRECTED TO 114. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION :THE REPORTER INDICATED THAT EXCESSIVE VAULT (INITIALLY 1.5, THEN CHANGED TO 5) WAS OBSERVED. PUPIL BLOCK WITH ELEVATED IOP (55 MMHG) AND ANGLE CLOSURE WITH ELEVATED IOP (55 MMHG) WERE OBSERVED ON (B)(6) 2019, LENS OPACITY (ASC) OBSERVED ON (B)(6) 2019. REPORTEDLY ENLARGEMENT OF PI AND ADDITION OF NEW PI WERE SURGICALLY PERFORMED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN AND "TOPOMAX RELATED SHIFT? SURGICALLY ENLARGED PI FAILED TO RESOLVE EDEMATUOUS CB WORSENED BY TOPOMAX?" PER REPORTER ON (B)(6) 2019, "THERE IS CATARACT FORMATION AND PATIENT IS SCHEDULED FOR CE WITH IOL IMPLANTATION." (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6859975&pc=MTA | 2023826-2017-01422 | 2017-08-10 | Injury | STAAR SURGICAL COMPANY | 2017-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Intraocular Pressure Increased; Pupillary Block; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.2, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, (PUPILLARY BLOCK) ELEVATED IOP, NARROWING OF THE ANGLE, ANGLE CLOSURE, SHALLOWING OF THE ACD, ANTERIOR LENS OPACITY, WATER VACUOLES, AND FIXED DILATED PUPIL. THE LENS WAS EXCHANGED. A POSSIBLE IRIDOPLASTY IS PLANNED. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. ON (B)(6) 2017, AN IRIDOTOMY ENLARGEMENT WAS DONE. THE REPORTER ALSO STATED THAT THE SOFTWARE DOES NOT ALWAYS PREDICT CORRECT SIZE IMPLANT. (B)(4). CORRECTION DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.2, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, (PUPILLARY BLOCK) ELEVATED IOP, NARROWING OF THE ANGLE, ANGLE CLOSURE, SHALLOWING OF THE ACD, ANTERIOR LEN OPACITY, WATER VACUOLES, AND FIXED DILATED PUPIL. ON (B)(6) 2017, AN IRIDOTOMY ENLARGEMENT WAS DONE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A SPECIMEN CUP. VISUAL INSPECTION FOUND CLEAR RESIDUE ON LENS AND A PIECE OF THE HAPTIC MISSING. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8233814&pc=MTA | 2023826-2019-00072 | 2018-07-05 | Injury | STAAR SURGICAL COMPANY | 2019-01-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem; Optical Problem | Cataract; Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PROBLEMS WITH PRESSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS WAS EXPLANTED DUE TO LOW VAULTING, ELEVATED INTRAOCULAR PRESSURE, LENS DISLOCATION/SUBLUXATION, GLARE/HALOS AND ANTERIOR SUBCAPSULAR CHANGES (LENS OPACITY). THE REPORTER INDICATED THE LENS HAPTICS WERE POSTERIOR TO SULCUS AND OPTIC WAS AGAINST THE NATURAL LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR (THE PATIENTS ANATOMY SEEMED TO BE THE CAUSE OF THE ICL FAILURE). THE INITIAL LENS WAS NOT EXCHANGED FOR ANOTHER LENS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SPECIMEN CUP, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8233815&pc=MTA | 2023826-2019-00071 | 2018-07-18 | Injury | STAAR SURGICAL COMPANY | 2019-01-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem; Optical Problem | Cataract; Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PROBLEMS WITH PRESSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS WAS EXPLANTED DUE TO LOW VAULTING, ELEVATED INTRAOCULAR PRESSURE, LENS DISLOCATION/SUBLUXATION, GLARE/HALOS AND ANTERIOR SUBCAPSULAR CHANGES (LENS OPACITY). THE REPORTER INDICATED THE LENS HAPTICS WERE POSTERIOR TO SULCUS AND OPTIC WAS AGAINST THE NATURAL LENS. A SECONDARY YAG WAS PERFORMED FOR ENLARGEMENT OF THE PERIPHERAL IRIDECTOMIES. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR (THE PATIENTS ANATOMY SEEMED TO BE THE CAUSE OF THE ICL FAILURE). THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. SEE MFR REPORT # 2023826-2019-00073 FOR INITIAL LENS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MFR: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SPECIMEN CUP, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6024717&pc=MTA | 2023826-2016-01448 | NA | Injury | STAAR SURGICAL COMPANY | 2016-10-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Intraocular Pressure, Delayed, Uncontrolled; Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO PRESSURE AND PAIN FOR PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SHALLOWING OF THE ANTERIOR CHAMBER DEPTH AND ELEVATED INTRAOCULAR PRESSURE. THE PATIENT EXPERIENCED BLURRY VISION. THE PATIENT WAS ADMINISTERED MEDICATION FOR THE PRESSURE. POST OP THE PATIENT HAD TRACE CORTICAL OPACITY AND BEST-CORRECTED VISUAL ACUITY WAS 20/25. THE VISUAL ACUITY WAS MORE STABLE BUT THE PUPIL WAS DILATED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. EVENT PROBLEM AND EVALUATION CODES: (B)(4). RESULT CODE(S): (OPERATIONAL PROBLEM) : VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSION CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6298841&pc=MTA | 2023826-2017-00208 | NA | Injury | STAAR SURGICAL COMPANY | 2017-02-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Conjunctivitis; Glaucoma; Inflammation; Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN HIS RIGHT EYE (OD). THE PATIENT REPORTED WAS PRESCRIBED ORAL MEDICATIONS/EYE DROPS FOR AT LEAST 3 CONSECUTIVE MONTHS DUE TO THE INSIDE OF HIS EYE WAS INFLAMED. ALSO WAS PRESCRIBED EYE DROPS FOR 3 CONSECUTIVE MONTHS FOR HIGH PRESSURE OR GLAUCOMA INSIDE THE EYE. THE LENS WAS IMPLANTED ON (B)(6) 2013. THE FACILITY REPORTED THEY WERE UNABLE TO CONFIRM THE PATIENT REPORT. THE DATE OF LAST VISIT WAS (B)(6) 2015. THE PATIENT HAD DEVELOPED VIRAL CONJUNCTIVITIS AT 4 MONTHS POST-OP AND THIS WAS NOT RELATED TO THE ICL. THE FACILITY REPORTED THE PATIENT HAD NO HISTORY OF ELEVATED IOP OR GLAUCOMA. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: PHYSICIAN'S REPORT STATED ALL PATIENT REPORTED ADVERSE EVENTS DID NOT OCCUR. CONJUNCTIVITIS DID OCCUR BUT NOT DEVICE RELATED. CORRECTED DATA: MFR DATE: THE CORRECTED DATE IS 02/28/2013. (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6298876&pc=MTA | 2023826-2017-00209 | NA | Injury | STAAR SURGICAL COMPANY | 2017-02-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Conjunctivitis; Glaucoma; Inflammation; Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN HIS LEFT EYE (OS). THE PATIENT REPORTED WAS PRESCRIBED ORAL MEDICATIONS/EYE DROPS FOR AT LEAST 3 CONSECUTIVE MONTHS DUE TO THE INSIDE OF HIS EYE WAS INFLAMED. ALSO WAS PRESCRIBED EYE DROPS FOR 3 CONSECUTIVE MONTHS FOR HIGH PRESSURE OR GLAUCOMA INSIDE THE EYE. THE LENS WAS IMPLANTED ON (B)(6) 2013. THE FACILITY REPORTED THEY WERE UNABLE TO CONFIRM THE PATIENT REPORT. THE DATE OF LAST VISIT WAS (B)(6) 2015. THE PATIENT HAD DEVELOPED VIRAL CONJUNCTIVITIS AT 4 MONTHS POST-OP AND THIS WAS NOT RELATED TO THE ICL. THE FACILITY REPORTED THE PATIENT HAD NO HISTORY OF ELEVATED IOP OR GLAUCOMA. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CONCLUSION: PHYSICIAN'S REPORT STATED ALL PATIENT REPORTED ADVERSE EVENTS DID NOT OCCUR. CONJUNCTIVITIS DID OCCUR BUT NOT DEVICE RELATED. MFG DATE: THE CORRECTED DATE IS 02/28/2013. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8159605&pc=MTA | 2023826-2018-01833 | 2018-11-07 | Injury | STAAR SURGICAL COMPANY | 2018-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Corneal Decompensation; Corneal Edema; Inflammation; Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL 13.2, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. EXCESSIVE VAULT, LEADING TO CORNEAL DECOMPENSATION, UNCONTROLLED IOP, NARROWING OF THE ANGLES, CORNEAL EDEMA AND SHALLOWING OF THE ACD WAS OBSERVED. PAIN AND BLURRY VISION WAS REPORTED AS ADDITIONAL SYMPTOMS. ADDITIONAL SURGICAL INTERVENTION ( AC WASHOUT, ICL "REPOSTIONING", REPEAT YAG PI) WAS PERFORMED. IT WAS REPORTED THAT IOP LOWERING DROPS AND PILLS WERE PRESCRIBED. THE LENS WAS EXPLANTED ON (B)(6) 2018. THE EXPLANT RESOLVED THE HIGH VAULT AND IOP. PATIENT IS DOING BETTER BUT DOES HAVE PIGMENT ON ENDOTHELIUM. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS OF -7.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. EXCESSIVE VAULT; CORNEAL DECOMPENSATION, UNCONTROLLED IOP, NARROWING OF THE ANGLES, CORNEAL EDEMA, INFLAMMATION, AND SHALLOWING OF ACD WAS OBSERVED. PAIN AND BLURRY VISION WERE REPORTED AS ADDITIONAL SYMPTOMS. ADDITIONAL SURGICAL INTERVENTION (AC WASHOUT, ICL REPOSITIONING, REPEAT YAG PI) WAS PERFORMED. IT WAS REPORTED THAT IOP LOWERING DROPS AND PILLS WERE PRESCRIBED. THE LENS WAS EXPLANTED ON (B)(6) 2018. THE EXPLANT RESOLVED THE HIGH VAULT AND IOP. PATIENT IS DOING BETTER BUT DOES HAVE PIGMENT ON ENDOTHELIUM. (B)(4). Manufacturer Narrative: (B)(4). WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8159607&pc=MTA | 2023826-2018-01834 | 2018-11-07 | Injury | STAAR SURGICAL COMPANY | 2018-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Corneal Decompensation; Corneal Edema; Inflammation; Intraocular Pressure Increased; Pain; Blurred Vision; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL 13.2, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. EXCESSIVE VAULT, LEADING TO CORNEAL DECOMPENSATION, UNCONTROLLED IOP, NARROWING OF THE ANGLES, CORNEAL EDEMA, AND SHALLOWING OF THE ACD WAS OBSERVED. PAIN AND BLURRY VISION WAS REPORTED AS ADDITIONAL SYMPTOMS. ADDITIONAL SURGICAL INTERVENTION ( AC WASHOUT, ICL "REPOSTIONING", REPEAT YAG PI) WAS PERFORMED. IT WAS REPORTED THAT IOP LOWERING DROPS AND PILLS WERE PRESCRIBED. THE LENS WAS EXPLANTED ON (B)(6) 2018. THE EXPLANT RESOLVED THE HIGH VAULT AND IOP. PATIENT IS DOING BETTER BUT DOES HAVE PIGMENT ON ENDOTHELIUM, AND MILD MYDRIASIS, LIKELY CONTRIBUTING TO GLARE COMPLAINT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS OF -7.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. EXCESSIVE VAULT, CORNEAL DECOMPENSATION, UNCONTROLLED IOP, NARROWING OF THE ANGLES, CORNEAL EDEMA, AND SHALLOWING OF THE ACD WAS OBSERVED. PAIN AND BLURRY VISION WERE REPORTED AS ADDITIONAL SYMPTOMS. ADDITIONAL SURGICAL INTERVENTION (AC WASHOUT, ICL REPOSITIONING, REPEAT YAG PI) WAS PERFORMED. IT WAS REPORTED THAT IOP LOWERING DROPS AND PILLS WERE PRESCRIBED. THE LENS WAS EXPLANTED ON (B)(6) 2018. THE EXPLANT RESOLVED THE HIGH VAULT AND IOP. PATIENT IS DOING BETTER BUT DOES HAVE PIGMENT ON ENDOTHELIUM AND MILD MYDRIASIS LIKELY CONTRIBUTING TO THE GLARE COMPLAINT. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, NARROWING OF ANGLES, SHALLOWING OF ACD, AC WASHOUT, ICL REPOSITIONING, YAG PI, IOP LOWERING DROPS AND PILLS, EXPLANT, MYDRIASIS (B)(4). Manufacturer Narrative: (B)(4). WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11231056&pc=MTA | 2023826-2021-00199 | 2020-11-20 | Injury | STAAR SURGICAL COMPANY | 2021-01-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Decompensation; Corneal Edema; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE AND CORNEAL DECOMPENSATION. POST-OP, THERE WAS SEMI-MYDRIASIS (DILATION OF PUPIL) AND CORNEAL EDEMA. A POSSIBLE CAUSE OF THE EVENT WAS OVERSIZE LENS (ICL). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: AT 2 MONTHS POST-OP, THERE IS NO MORE CORNEAL EDEMA OR HIGH IOP, ONLY SEMI-MYDRIASIS. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMEMSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13878560&pc=MTA | 2023826-2022-00910 | 2021-09-24 | Injury | STAAR SURGICAL COMPANY | 2022-03-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Corneal Decompensation; Corneal Edema; Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 13.2MM VICM5_13.2; -10.50 DIOPTER; INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS REPORTED NON-CONFORMITY - UNRESPONSIVE HYPETONUS DESPITE SURGICAL PROCEDURE AND COMPLICATIONS. THE PATIENT EXPERIENCED LOSS OF BCVA; ELEVATED INTRAOCULAR PRESSURE; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; CORNEAL DECOMPENSATION; BLURRED VISION. ON (B)(6) 2021 THE LENS WAS EXPLANTED. THE POSTOPERATIVE NOTES INDICATED MILD TRANSPARENT CRYSTALLINE; CORNEAL EDEMA "ABOUT PRESENT PUPIL REACTIVE WAITING TO PERFORM VISUAL FIELD". IT WAS ALSO NOTED THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#:(B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13529754&pc=MTA | 2023826-2022-00498 | 2022-02-05 | Injury | STAAR SURGICAL COMPANY | 2022-02-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Decompensation; Glaucoma; Intraocular Pressure Increased; Pain; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -2 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; PUPIL BLOCK; ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE; CORNEAL DECOMPENSATION; MEDICATION USED; PAIN; GLAUCOMA AND IT WAS NOTED THAT YAG SURGERY WAS PERFORMED AS WELL AS PARACENTESIS. ON (B)(6) 2022 THE LENS WAS EXPLANTED. POST OP RESULTS REPORTED THE PROBLEM WAS RESOLVED AND WITH THE LENS REMOVED THE PATIENT NOW HAS CORNEAL EDEMA; MID DILATED PUPIL; IRIS TRAUMA. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "ICL SEEMS TO BIG FOR THIS EYE DESPITE CORRECT MEASUREMENT AND CALCULATION PREOPERATIVELY". Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID, WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED CORRECTED VA IS 6/6 WITH SOME DEGREE WITH IRIS ATROPHY. H6 - HEALTH EFFECT CLINICAL CODE: 4581- IRIS ATROPHY. CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7899789&pc=MTA | 2023826-2018-01401 | 2017-06-24 | Injury | STAAR SURGICAL COMPANY | 2018-09-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Decompensation; Inflammation; Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. AT ONE DAY POST-OP, INFLAMMATION (TASS) PRESENTED IN THE EYE. CORTISONE WAS ADMINISTERED AND THE INFLAMMATION CLEARED UP. THE PATIENT ALSO EXPERIENCED ELEVATED INTRAOCULAR PRESSURE AND CORNEAL DECOMPENSATION. THE LENS WAS EXPLANTED ON (B)(6) 2018. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. THE POST-OP CONDITION OF THE PATIENT'S EYE IS AS FOLLOWS - SYNECHIAE OF CHAMBER ANGLE (PARTIAL), SOME SMALL SUBCAPSULAR VACUOLES IN PERIPHERY OF CRYSTALLINE LENS AND SLIGHT MYDRIASIS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS RESIDUE ON THE LENS. (B)(4). DEVICE HISTORY RECORD REVIEW - BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6540579&pc=MTA | 2023826-2017-00740 | 2017-01-11 | Injury | STAAR SURGICAL COMPANY | 2017-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Corneal Decompensation; Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, IRIS DISTORTION, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, CORNEAL DECOMPENSATION, UNREACTIVE PUPIL, ANGLE CLOSURE, GLARE/HALOES, BLURRED VISION. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE DR. SAID THE PROBLEM IS SLOWLY GETTING BETTER. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS EXCHANGED RESOLVED THE PROBLEM FOR THE PATIENT. LENS RETURNED IN LENS VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS, DEBRIS AND CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, IRIS DISTORTION, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, CORNEAL DECOMPENSATION, UNREACTIVE PUPIL, ANGLE CLOSURE, GLARE/HALOES, BLURRED VISION, AND LOSS OF BCVA. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8529658&pc=MTA | 2023826-2019-00632 | 2019-03-19 | Injury | STAAR SURGICAL COMPANY | 2019-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Decompensation; Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT, PUPIL BLOCK WITH IOP (INTRAOCULAR PRESSURE), ANGLE CLOSURE, CORNEAL DECOMPENSATION, LENS DISLOCATION/SUBLUXATION, AND UNREACTIVE (FIXED) PUPIL. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED. SURGEON IS NOT SURE YET IF THE EXCHANGE LENS RESOLVED THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED FOR THIS CASE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID , IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12125559&pc=MTA | 2023826-2021-02334 | 2020-09-20 | Injury | STAAR SURGICAL COMPANY | 2021-07-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Decompensation; Intraocular Pressure Increased; Retinal Detachment; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE (IOP), CORNEAL DECOMPENSATION AND RETINAL DETACHMENT. PHACO-EMULSIFICATION (CATARACT SURGERY), IOL IMPLANTATION AND RETINAL REPAIR WAS PERFORMED. THE PROBLEM WAS RESOLVED AND THE RETINA IS IN PLACE. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS -18.00 DIOPTER IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO RETINAL DETACHMENT. PHACO-EMULSIFICATION (CATARACT SURGERY); IOL IMPLANTATION AND RETINAL REPAIR WAS PERFORMED. THE PROBLEM IS RESOLVED. H6 - HEALTH EFFECT CLINICAL CODE 4581,1937,1790 IS NOT APPLICABLE IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5639098&pc=MTA | 2023826-2016-00630 | 2016-03-11 | Injury | STAAR SURGICAL COMPANY | 2016-05-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Corneal Edema; Endophthalmitis; Hypopyon; Intraocular Pressure Increased; Pupillary Block; Vitrectomy; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE PATIENT PRESENTED WITH POST-OP ENDOPHTHALMITIS. THE PATIENT UNDERWENT A VITRECTOMY WITH ANTIBIOTIC INJECTION. THE PATIENT EXPERIENCED HYPOPYON, PUPILLARY BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE AND CORNEAL EDEMA. THE LENS REMAINS IMPLANTED AND THE SURGEON INDICATED THE CAUSE OF THE EVENT WAS NOT DEVICE RELATED, MOST LIKELY DUE TO INTRODUCTION OF BACTERIA DURING OR AFTER SURGERY.Event Description: ADDITIONAL INFORMATION RECEIVED - THE CUSTOMER REPORTED THE LENS WAS EXPLANTED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: LENS IMPLANTED. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16850587&pc=MTA | 2023826-2023-01613 | 2023-03-15 | Injury | STAAR SURGICAL COMPANY | 2023-05-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICMO 13.7, -14.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. IT WAS ALSO REPORTED THAT CORNEAL EDEMA WAS NOTICED ON (B)(6) 2023, HOWEVER IT DID NOT REQUIRE TREATMENT BEYOND WAS IS STANDARD AFTER OCULAR SURGERY AND THE CORNEAL EDEMA RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12677856&pc=MTA | 2023826-2021-04046 | 2021-09-10 | Injury | STAAR SURGICAL COMPANY | 2021-10-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Unintended Movement | Corneal Edema; Intraocular Pressure Increased; Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS WAS INSERTED UPSIDE DOWN IN THE PATIENTS RIGHT EYE (OD). THE PATIENTS INTRAOCULAR PRESSURE (IOP) WAS ELEVATED DURING THE SURGERY AND THE PATIENT COMPLAINED OF BLURRED VISION. THE LENS WAS REPLACED DURING THE SAME SURGERY WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THERE WAS SOME CORNEAL EDEMA DUE TO THE INCREASED IOP AND BLURRED EYES. THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS WAS DISCARDED. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6711487&pc=MTA | 2023826-2017-01177 | 2017-06-16 | Injury | STAAR SURGICAL COMPANY | 2017-07-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Corneal Edema; Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, DIOPTER -18.00 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. FOLLOW UP ON (B)(6) 2017, THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULT, BLURRED AND DIMNESS OF VISION, ELEVATED IOP, IRIS PIGMENTATION, FIXED SEMI DILATED PUPIL, CORNEA OEDEMATUS. NEXT DAY, THE ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS, LENS REPOSITIONING, AND IRIDECTOMY WAS PERFORMED. FOLLOW UP ON (B)(6) 2017, THE PATIENT'S RIGHT EYE CORNEA GOT CLEARER WITH FIXED DILATED ATONIC PUPIL, PIGMENTATION IN ICL LENS, AND ANGLE CLOSURE. THE LENS WAS EXPLANTED ON (B)(6) 2017. THE IOP WAS CONTROLLED, AND ANGLES OPEN, PUPIL SEMI DILATED 4MM TO 5MM. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISISION THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). IRIS PIGMENTATION, FIXED SEMI DILATED PUPIL, ANGLE CLOSURE, DIMNESS OF VISION, SECONDARY SURGERY, YAG, LENS REPOSITIONING, IRIDECTOMY, ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS, LENS EXPLANT. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8903520&pc=MTA | 2023826-2019-01478 | 2019-07-16 | Injury | STAAR SURGICAL COMPANY | 2019-08-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTERD INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. EXCESSIVE VAULT, REFRACTIVE SURPRISE, ELEVATED IOP (INTRAOCULAR PRESSURE), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND BLURRED VISION WAS OBSERVED. ADDITION OF NEW PI (YAG) WAS REPORTED AS SECONDARY INTERVENTION. THE REPORTER STATED THAT THE PATIENT IS HAPPY WITH THE RESULTS AND DONT WANT TO CHANGE THE LENS . IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 3191: PHOTOPHOBIA, PATIENT CODE 1791: CORNEAL OEDEMA, METHOD CODE 4117: LENS REMAINS IMPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5688628&pc=MTA | 2023826-2016-00670 | NA | Injury | STAAR SURGICAL COMPANY | 2016-05-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Corneal Edema; Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED ELEVATED IOP, HAZY VISION, PIGMENT DISPERSION, AND MILD EDEMA. LASER PERIPHERAL IRIDOTOMY (LASER PI) WAS PERFORMED TO LOWER INTRAOCULAR PRESSURE. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5745383&pc=MTA | 2023826-2016-00851 | NA | Injury | STAAR SURGICAL COMPANY | 2016-06-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Difficult to Remove; Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Corneal Edema; Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -15.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) /2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO AN EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A 12.6MM ICL. THERE WAS NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER STATED THE PATIENT HAD CORNEAL EDEMA AS A RESULT OF DIFFICULT EXPLANT SURGERY. THE PATIENT EXPERIENCED ELEVATED IOP AND NARROWING OF THE ANGLE. THE REPORTER INDICATED THAT THE EVENT WAS PRODUCT-RELATED. Manufacturer Narrative: (B)(4). RESULTS: (OPERATIONAL PROBLEM) - VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND SMALL TEARS ON THE LENS HAPTIC. THE LENS WAS RETURNED IN VIAL IN LIQUID. (B)(4). Manufacturer Narrative: (B)(4) | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7070542&pc=MTA | 2023826-2017-01801 | 2017-08-17 | Injury | STAAR SURGICAL COMPANY | 2017-11-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. AT THE POST-OP VISIT ON (B)(6) 2017, CORNEAL EDEMA WAS OBSERVED, BUT THE PATIENT HAS FULLY RECOVERED AND IS NOW FINE. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/28. Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND A PIECE OF ONE LENS HAPTIC BROKEN. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5602196&pc=MTA | 2023826-2016-00514 | 2016-01-02 | Injury | STAAR SURGICAL COMPANY | 2016-04-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Corneal Edema; Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/50. THE PATIENT HAD CORNEAL EDEMA POST-OP. Manufacturer Narrative: DEVICE MFR DATE: 09/29/2015 TO 10/08/2015, PREVIOUS DATE IS INCORRECT. (B)(4). DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. WORK ORDER SEARCH: NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5602203&pc=MTA | 2023826-2016-00515 | 2016-01-02 | Injury | STAAR SURGICAL COMPANY | 2016-04-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Corneal Edema; Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/67. THE PATIENT HAD CORNEAL EDEMA POST-OP. Manufacturer Narrative: CORRECTED DATA: MFR DATE: 09/14/2015 TO 09/21/2015, PREVIOUS DATE IS INCORRECT. (B)(4). ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. WORK ORDER SEARCH: NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5652092&pc=MTA | 2023826-2016-00637 | NA | Injury | STAAR SURGICAL COMPANY | 2016-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Corneal Edema; Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE, ANGLE CLOSURE, SHALLOWING OF THE ANTERIOR CHAMBER AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT HAD CORNEAL EDEMA AND THE POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5652096&pc=MTA | 2023826-2016-00638 | NA | Injury | STAAR SURGICAL COMPANY | 2016-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Corneal Edema; Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE, ANGLE CLOSURE, SHALLOWING OF THE ANTERIOR CHAMBER AND ELEVATED INTRAOCULAR PRESSURE. THE REPORTER INDICATED THERE WAS ENDOTHELIAL TOUCH. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT HAD CORNEAL EDEMA AND THE POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND DEVICE HISTORY REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8903395&pc=MTA | 2023826-2019-01477 | 2019-07-16 | Injury | STAAR SURGICAL COMPANY | 2019-08-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTERD INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT, REFRACTIVE SURPRISE, ELEVATED IOP (INTRAOCULAR PRESSURE) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS OBSERVED. ADDITION OF NEW PI (YAG) WAS REPORTED AS SECONDARY INTERVENTION. THE REPORTER STATED THAT THE PATIENT IS SATISFIED WITH THE RESULTS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 OF -12.0 DIOPTER; IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT; REFRACTIVE SURPRISE; ELEVATED IOP; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; CORNEAL OEDEMA; BLURRED VISION; HERNIATED IRIS; IRIS ATROPHY; IRIS REPAIR; AND ADDITION OF SUTURE; ADDITION OF NEW PI (YAG) WAS REPORTED AS SECONDARY INTERVENTION. REPORTER STATED THAT PATIENT IS SATISFIED WITH THE RESULTS. PATIENT CODE 3191: HERNIATED IRIS; IRIS ATROPHY; IRIS REPAIR; ADDITION OF SUTURE; PATIENT CODE 1791:CORNEAL OEDEMA, DEVICE CODE 1401:REFRACTIVE SURPRISE, METHOD CODE 4117:LENS REMAINS IMPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957614&pc=MTA | 2023826-2020-00848 | 2019-07-08 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Corneal Edema; Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ELEVATED IOP (30 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED AND THE LENS WAS REMOVED ON (B)(6) 2020. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "THE PATIENT PRESENTED WITH HIGH INTRAOCULAR PRESSURE 4 MONTHS AFTER THE OPERATION, WHICH CONTINUED UNTIL NOW. THE LENS WAS REMOVED. BUT AFTER THE REMOVAL OF THE LENS, THE INTRAOCULAR PRESSURE WAS STILL HIGH" AND CORNEAL EDEMA WAS OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER STATED "CORNEAL EDEMA HAS NOT BEEN RESOLVED AT THIS MOMENT. THERE IS NO INFECTION. CORNEAL EDEMA ONLY BECAUSE OF HYPERTENSION. NO OTHER INTERVENTION PERFORMED BUT INTENDED TO HAVE SURGICAL INTERVENTION ON HYPERTENSION." CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: G3- (B)(6) 2021" SHOULD BE CORRECTED TO "07-APR-2020" IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957617&pc=MTA | 2023826-2020-00849 | 2019-07-10 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Corneal Edema; Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ELEVATED IOP (26MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED AND THE LENS WAS REMOVED ON (B)(6) 2020. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "THE PATIENT PRESENTED WITH HIGH INTRAOCULAR PRESSURE 4 MONTHS AFTER THE OPERATION, CANNOT BE CONTROLLED BY MEDICINE TREATMENT. THE LENS WAS REMOVED AT 8TH MONTH POST OPERATION. BUT, THE INTRAOCULAR PRESSURE WAS STILL HIGH AFTER REMOVAL OF LENS" AND CORNEAL EDEMA WAS OBSERVED. PATIENT NEED FURTHER SURGICAL INTERVENTION TO TREAT HIGH IOP. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. CLAIM#: (B)(4). Manufacturer Narrative: B5- THE REPORTER STATED "CORNEAL EDEMA HAS NOT BEEN RESOLVED AT THIS MOMENT. THERE ARE NO INFECTION. CORNEAL EDEMA ONLY BECAUSE OF HYPERTENSION. NO OTHER INTERVENTION PERFORMED BUT INTENDED TO HAVE SURGICAL INTERVENTION ON HYPERTENSION." CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: G3- (B)(6) 2021 SHOULD BE CORRECTED TO "07-APR-2021" IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H10- "(B)(6) 2021" SHOULD BE CORRECTED TO "(B)(6) 2020" IN PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11043753&pc=MTA | 2023826-2020-03159 | 2020-11-20 | Injury | STAAR SURGICAL COMPANY | 2020-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Optical Nerve Damage; Pupillary Block; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE AND THE PATIENT EXPERIENCED HIGH VAULT AND HIGH IOP (INTRAOCULAR PRESSURE). LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5:THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -5.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. EXCESSIVE VAULT, PUPIL BLOCK, ANGLE CLOSURE WITH ELEVATED IOP (47MMHG), PIGMENT DISPERSION, GLARE/HALOES AND BLURRED VISION WAS OBSERVED. ADDITIONAL SURGICAL INTERVENTION (PIS) WAS PERFORMED. THE REPORTER STATED "THE PATIENT WAS CYCLOPLEGED WITH ATROPINE FOR AN EXTENDED PERIOD OF TIME UNTIL THE IOP COULD BE STABILIZED AND THE CORNEAL EDEMA IMPROVE WITH STEROIDS, ORAL AND TOPICAL IOP LOWERING MEDICATIONS. THIS TOOK SEVERAL WEEKS TO RESOLVE". IT WAS REPORTED THAT THE PROBLEM RESOLVED, LAST VISIT DATE (B)(6) 2020 (BCVA:20/20, UCVA:20/20--, VAULT:1.5CT, IOP:12MMHG). FOR STATUS OF THE EYE (SIGN/SYMPTOMS) THE REPORTER STATED "HE SUBSEQUENTLY DEVELOPED UNEXPLAINED SYMPTOMATIC RIGHT SIDED OPTIC NERVE EDEMA AND WAS SENT TO THE ER FOR FULL NEUROLOGICAL AND OPHTHALMOLOGICAL EVALUATION. MRI WAS NEGATIVE EXCEPT FOR ENHANCEMENT OF THE RIGHT OPTIC NERVE POTENTIALLY CONSISTENT WITH OPTIC NEURITIS. THE HEAD NEUROLOGIST DID NOT FEEL THAT IT WAS CLASSIC NEURITIS BUT HE WAS TREATED FOR IT REGARDLESS". ADDITIONAL INFORMATION REPORTED "HE WAS GIVEN IV STEROIDS IN THE HOSPITAL AND SENT HOME WITH A FEW DAYS OF HIGH DOES ORAL STEROIDS." NO OTHER SYMPTOMS HAVE DEVELOPED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT MAY HAVE BEEN RETAINED OVD BEHIND THE IOL CAUSING ANGLE CLOSURE AND PUPILLARY BLOCK. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: CORRECTED DATA: B5 - SHOULD HAVE BEEN SPECIFIED THAT A YAG PI WAS PERFORMED. H6 - 4581: ANGLE CLOSURE, PIGMENT DISPERSION, ENLARGMENT OF PI (YAG), SHOULD HAVE BEEN SPECIFIED. CLAIM#: (B)(4). Manufacturer Narrative: AGE OR DOB: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. COMMON DEVICE NAME: PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNK. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. CLAIM # 431266 | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7097942&pc=MTA | 2023826-2017-01842 | 2017-10-29 | Injury | STAAR SURGICAL COMPANY | 2017-12-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.2, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ELEVATED IOP, ANGLE CLOSURE, NARROWING OF THE ANGLE, SHALLOWING OF THE ACD, BLURRY VISION, PAIN AND CORNEAL EDEMA. PARACENTESIS WAS DONE TO LOWER THE PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S POST-OP BCVA WAS 20/25. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7366387&pc=MTA | 2023826-2018-00456 | NA | Injury | STAAR SURGICAL COMPANY | 2018-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, NARROW ANGLES, ELEVATED INTRAOCULAR PRESSURE (IOP) AND CORNEAL EDEMA. THE PATIENT EXPERIENCED INTERMITTENT DECREASE OF VISUAL ACUITY, BLURRY VISION AND EYE PAIN. THE PATIENT RECEIVED MEDICAL TREATMENT FOR THE HIGH IOP. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE STATUS OF THE EYE POST-OP IS ASYMPTOMATIC. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE AND DEBRIS ON PRODUCT. THE LENS WAS FOUND WITH NO VISIBLE DAMAGE. DIMENSIONAL EVALUATION FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7366474&pc=MTA | 2023826-2018-00455 | NA | Injury | STAAR SURGICAL COMPANY | 2018-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, NARROW ANGLES, ELEVATED INTRAOCULAR PRESSURE (IOP) AND CORNEAL EDEMA. THE PATIENT EXPERIENCED INTERMITTENT DECREASE OF VISUAL ACUITY, BLURRY VISION AND EYE PAIN. THE PATIENT RECEIVED MEDICAL TREATMENT FOR THE HIGH IOP. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE STATUS OF THE EYE POST-OP IS ASYMPTOMATIC. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: TWO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6439246&pc=MTA | 2023826-2017-00486 | 2017-01-16 | Injury | STAAR SURGICAL COMPANY | 2017-03-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Corneal Edema; Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE PATIENT EXPERIENCED IRIS HERNIATION, AND REDUCTION SURGERY. ON THE NEXT DAY, THE PATIENT EXPERIENCED OCULAR PAIN AND BLURRY VISION, WITH SHALLOW ANTERIOR CHAMBER, MODERATE DIFFUSE CORNEAL EDEMA AND OCULAR HYPERTENSION. THE CENTRAL PORT OF THE LENS WAS BLOCKED BY IRIS PIGMENT. IRIDECTOMY SURGERY WAS PERFORMED, AND THE PROBLEM WAS RESOLVED. ENLARGEMENT OF PI OR ADDITION OF NEW PI SURGERY WAS ALSO PERFORMED. AS OF THE DAY OF MDR SUBMISSION THE LENS REMAINS IMPLANTED, WILL NOT BE RETURNED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH PERFORMED: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6872434&pc=MTA | 2023826-2017-01443 | 2017-08-20 | Injury | STAAR SURGICAL COMPANY | 2017-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL12.6, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, ELEVATED IOP, NARROWING OF THE ANGLE, BLURRY VISION, PAIN, PUPILLARY BLOCK, AND CORNEAL EDEMA. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND VIGAMOX AND PREDNISOLONE WERE USED AS STANDARD POST OPERATIVE PROTOCOL. THE PROBLEMS WERE RESOLVED. Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT TYPE EVENT WAS FOUND. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6658366&pc=MTA | 2023826-2017-00957 | 2017-05-02 | Injury | STAAR SURGICAL COMPANY | 2017-06-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, PUPILLARY BLOCK, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT EXPERIENCED BLURRY VISION AND CORNEAL EDEMA. THE REPORTER INDICATED MULTIPLE PARACENTISIS PERFORMED AND THE PERIPHERAL IRIDECTOMIES WERE REDONE. THE PATIENT WAS PRESCRIBED MEDICATION. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENTS POST-OP CONDITION IS - TRACE K EDEMA, 1+ CELLS WITH SHALLOW ANTERIOR CHAMBER, DILATED PUPIL WITH CORNEAL IRIS TOUCH PERIPHERALLY, AND BEST-CORRECTED VISUAL ACUITY 20/200. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE REPORTER INDICATED THE PROBLEM WAS NOT RESOLVED AFTER EXPLANTATION OF THE LENS. THE PATIENT CONTINUED TO HAVE INTRAOCULAR PRESSURE (IOP) INCREASE POST EXPLANT OF LENS. AT ONE WEEK POST-OP, THE PATIENT UNDERWENT AHMED VALVE SHUNT INSERTION FOR IOP, WITH GLAUCOMA SPECIALIST. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11234650&pc=MTA | 2023826-2021-00225 | 2020-11-20 | Injury | STAAR SURGICAL COMPANY | 2021-01-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. EXCESSIVE VAULT, PUPIL BLOCK, ANGLE CLOSURE WITH ELEVATED IOP (43MMHG), GLARE/HALOES AND BLURRED VISION WAS OBSERVED. ADDITIONAL SURGICAL INTERVENTION (PI) WAS PERFORMED. IT WAS REPORTED THAT THE PROBLEM RESOLVED, LAST VISIT DATE (B)(6) 2020 (BCVA:20/20, UCVA:20/25-, VAULT:1.0 CT, IOP:14MMHG). THE REPORTER STATED "HE DEVELOPED A STEROID RESPONSE WITH CORNEAL EDEMA (B)(6) 2020 AND WAS SWITCHED TO LOTEMAX FROM PREDNISOLONE ACETATE 1%. HAS SINCE BEEN TAPERED OFF ALL STEROIDS AND IOP LOWERING MEDICATIONS SUCCESSFULLY." IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT MAY HAVE BEEN RETAINED OVD BEHIND THE IOL. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - ANGLE CLOSURE, SHOULD HAVE BEEN SPECIFIED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6501830&pc=MTA | 2023826-2017-00562 | 2017-03-30 | Injury | STAAR SURGICAL COMPANY | 2017-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Pupillary Block; No Code Available | P990013 | NA | 1 | Event Description: THE REPORTER STATED THAT A MICL13.2, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO PUPILLARY BLOCK CAUSING ELEVATED IOP. THE REPORTER INDICATED THE PATIENT HAD EXCESSIVE VAULT, NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ACD. THE PATIENT'S REPORTED CONDITION AT THE TIME OF THE LAST VISIT ON (B)(6) 2017 WAS 20/70 BCVA, NORMAL IOP AND CORNEAL EDEMA. Manufacturer Narrative: AT THE TIME OF SURGERY PATIENT WAS (B)(6) AND PER DFU INDICATIONS VISIAN IMPLANTABLE COLLAMER LENSES (VICL) ARE INDICATED FOR USE IN PHAKIC EYE TREATMENT IN ADULTS (B)(6) OF AGE. THERE IS NO SIGNIFICANT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS. (B)(4). Manufacturer Narrative: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN GAUZE WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND MISSING PIECE OF THE HAPTIC. (B)(4). WORK ORDER SEARCH: NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9338922&pc=MTA | 2023826-2019-02167 | NA | Injury | STAAR SURGICAL COMPANY | 2019-11-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cyst(s); Intraocular Pressure Increased; Pain; No Code Available | NA | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED THAT CITED A CASE STUDY THAT FOUND PATIENT IMPLANTED WITH A 12.1MM IMPLANTABLE COLLAMER LENS, -6.00 DIOPTER IN THE PATIENTS RIGHT EYE (OD). AT ONE DAY POST-OPERATIVE, THE PATIENT SUFFERED PAINFUL EYE. THE PATIENT HAD SEVERE CORNEAL EDEMA. SLIT-LAMP EXAMINATION SHOWED IRIDO-CORNEAL CONTACT AT TEMPORAL MID-PERIPHERY, SHALLOW ANTERIOR CHAMBER WITH CENTRAL VAULTING OF ICL, PARTIAL OBSTRUCTION OF PERIPHERAL IRIDOTOMY, INTRAOCULAR PRESSURE WAS NOT MEASURED BUT THE EYE WAS VERY FIRM. ORAL AND SYSTEMATIC ANTIGLAUCOMA MEDICATIONS INCLUDING INTRAVENOUS (IV) MANNITOL 20% WERE ADMINISTERED. ADDITIONAL LASER PROCEDURE TO REOPEN PERIPHERAL IRIDOTOMY WAS PERFORMED. AN ULTRASOUND BIOMICROSCOPY (UBM) SHOWED A TEMPORAL POSTERIOR IRIS CYST AT 8 O'CLOCK, WHICH WAS PUSHING THE PIOL-IRIS COMPLEX FORWARD CAUSING IRIDO-CORNEAL ANGLE CLOSURE. PATIENT WAS READMITTED TO SURGERY FOR REPOSITIONING OF THE ICL. THE NEXT DAY, PATIENT DID NOT COMPLAIN OF SEVERE OCULAR PAIN. CORNEAL EDEMA IMPROVED MARKEDLY AND IOP WAS 20 MMHG. AT 1 MONTH FOLLOW-UP VISIT, UCVA WAS 0.9 BILATERALLY. CORNEAL EDEMA DISAPPEARED AND IOP WAS 20 MMHG. THE LENS REMAINED IMPLANTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. WEIGHT: UNK. DATE OF EVENT: UNK. EXPIRATION DATE UNK. IMPLANTED: UNK. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12749254&pc=MTA | 2023826-2021-04225 | 2021-07-04 | Injury | STAAR SURGICAL COMPANY | 2021-11-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Discomfort; Eye Pain | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO_12.1 -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO PATIENT EXPERIENCING DISCOMFORT AND PAIN, THE REPORTER STATED " PATIENTS IGNORANT STARRY EYES ARE UNCOMFORTABLE, PAIN". THIS EXPLANT RESOLVED THE PROBLEM. CAUSE WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14107540&pc=MTA | 2023826-2022-01096 | 2022-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Discomfort; Eye Pain | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 FOR RELIEF OF EYE PAIN AND DISCOMFORT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8774155&pc=MTA | 2023826-2019-01205 | 2015-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Dry Eye(s); Headache; Pain; Red Eye(s); Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER INTO THE PATIENT'S LEFTT EYE (OS) IN (B)(6) OF 2015. HYPERMETROPIA OF +1, ASTIGMATISM OF "AROUND" 0.5, PAIN, AND RED EYE IS REPORTED. LASIK SURGERY WAS PERFORMED (B)(6) 2017 WHICH RESOLVED THE PROBLEM. HOWEVER, RED EYE STILL OCCURS APPROXIMATELY EVERY 2 MONTHS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: "HEADACHE, DRY EYE AND NAUSEA IS REPORTED." SHOULD BE ADDED TO THE INITIAL MDR. PATIENT CODE 3191: NAUSEA. (B)(4). Manufacturer Narrative: PRODUCT IS MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8574636&pc=MTA | 2023826-2019-00729 | NA | Malfunction | STAAR SURGICAL COMPANY | 2019-05-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Dry Eye(s); Intraocular Pressure Increased; Eye Pain | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THAT HE HAD A 13.2MM, MICL13.2, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS IMPLANTED INTO HIS RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT STATES THAT HE IS EXPERIENCING CHRONIC EYE SYNDROME SINCE THE IMPLANT WAS PLACED. HE INDICATED THAT HE DID NOT HAVE THIS PRIOR TO THE ACTUAL SURGERY. HE STATES THAT WHEN HE TAPES HIS EYE SHUT THE CONDITION IS NOT AS SEVERE AS WHEN HE SLEEPS WITHOUT THE TAPE, WHEN HE SLEEPS HIS EYE IS SLIGHTLY SLIT. Manufacturer Narrative: "CHRONIC EYE SYNDROME" SHOULD BE CHANGED TO "CHRONIC DRY EYE SYNDROME." CLAIM#: (B)(4). Manufacturer Narrative: B5-ADDITIONAL SYMPTOMS REPORTED: PAINFUL DRY EYE AND ELEVATED IOP. H6-INVESTIGATION FINDINGS CODE CORRECTED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - "TREATMENT INCLUDES SERUM TEAR THERAPY; MOISTURE CHAMBER GOGGLES; FISH OIL; AND TAPING OF EYELIDS." SHOULD BE ADDED. H6 - 4650 - SERUM TEAR THERAPY. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8575042&pc=MTA | 2023826-2019-00730 | NA | Malfunction | STAAR SURGICAL COMPANY | 2019-05-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Dry Eye(s); Intraocular Pressure Increased; Eye Pain | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THAT HE HAD A 13.2MM, MICL13.2, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS IMPLANTED INTO HIS LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT STATES THAT HE IS EXPERIENCING CHRONIC EYE SYNDROME SINCE THE IMPLANT WAS PLACED. HE INDICATED THAT HE DID NOT HAVE THIS PRIOR TO THE ACTUAL SURGERY. HE STATES THAT WHEN HE TAPES HIS EYE SHUT THE CONDITION IS NOT AS SEVERE AS WHEN HE SLEEPS WITHOUT THE TAPE, WHEN HE SLEEPS HIS EYE IS SLIGHTLY SLIT. Manufacturer Narrative: EVENT DESCRIPTION: "CHRONIC EYE SYNDROME" SHOULD BE CHANGED TO "CHRONIC DRY EYE SYNDROME." CLAIM#: (B)(4). Manufacturer Narrative: B5-ADDITIONAL SYMPTOMS REPORTED: PAINFUL DRY EYE AND ELEVATED IOP. H6-INVESTIGATION FINDINGS CODE CORRECTED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5. "TREATMENT INCLUDES SERUM TEAR THERAPY; MOISTURE CHAMBER GOGGLES; FISH OIL; AND TAPING OF EYELIDS." SHOULD BE ADDED. H6. 4650 SERUM TEAR THERAPY. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9772182&pc=MTA | 2023826-2020-00374 | 2019-12-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Dry Eye(s); Intraocular Pressure Increased; Pain; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, - 11.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. AT ONE-DAY POST-OP, THE PATIENT REPORTED EYE PAIN, WHICH WAS NOT RESOLVED WITH ARTIFICIAL TEARS OR WARM COMPRESSES. THE PATIENT GOT A SECOND OPINION BUT WAS UNABLE TO SOLVE THE SOURCE OF THE PAIN. THE PATIENT REQUESTED THE LENS BE REMOVED AND THE LENS WAS EXPLANTED ON (B)(6) 2020. THERE ARE NO PLANS TO IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT REPORTED POST-OP, HER EYE PRESSURE HAD INCREASED AND WAS TREATED WITH COMBIGAN. WAS ALSO TREATED FOR DRY EYE FOR OVER A MONTH. THE PATIENT'S CURRENT CONDITION STILL HAS OCCASIONAL PAIN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A ZIPLOC BAG, IN A VIAL WITH LIQUID AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CORRECTED DATA: H6 - PATIENT CODE 3191: NO CODE AVAILABLE (MEDICATION: COMBIGAN) SHOULD HAVE BEEN INCLUDED IN INTIAL MEDWATCH REPORT. ADDITIONAL INFORMATION: H6 - METHOD CODE 3331, DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE: 1937 SHOULD BE REMOVED AS IT IS NOT APPLICABLE. B5 - IT WAS INDICATED FROM THE REPORTER THAT PRIOR TO SURGERY, IN (B)(6) 2019, THE PATIENT HAD A "HAIR CLIP INJURY" WHICH RESULTED IN A BRUISED EYE LID. DOCTOR FOUND NO EVIDENCE OF AN INJURY TO THE EYE ITSELF. PATIENT RECEIVED SEVERAL OTHER DOCTORS OPINIONS AND AT SOME POINT WAS PRESCRIBED CAMBIGAN TO TREAT ELEVATED IOP. COMBIGAN WAS NOTED TO HAVE BEEN FIRST USED ON (B)(6) 2020 AND WAS DISCONTINUED ON (B)(6) 2020 AND THEN AGAIN RESUMED ON (B)(6) 2020 WHICH REPORTED THE PATIENT HAD 13MMHG EYE PRESSURE. THE DAY BEFORE THEY FOUND HER EYE PRESSURE TO BE 23MMHG. PATIENT CONTINUED TO USE COMBIGAN AS OF (B)(6) 2020 AND HAS BEEN TO A CORNEAL SPECIALIST. WHO IS INVESTIGATING THE POSSIBILITY OF OPHTHALMIC NEURALGIA. CLAIM#: (B)(4). Manufacturer Narrative: REPORT NUMBER - MW5093020. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9772179&pc=MTA | 2023826-2020-00375 | 2020-02-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Dry Eye(s); Intraocular Pressure Increased; Pain; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, - 11.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE PATIENT REPORTED EYE PAIN, WHICH WAS NOT RESOLVED WITH ARTIFICIAL TEARS OR WARM COMPRESSES. THE PATIENT GOT A SECOND OPINION BUT WAS UNABLE TO SOLVE THE SOURCE OF THE PAIN. THE PATIENT REQUESTED THE LENS BE REMOVED AND THE LENS WAS EXPLANTED ON (B)(6) 2020. THERE ARE NO PLANS TO IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT REPORTED POST-OP, HER EYE PRESSURE HAD INCREASED AND WAS TREATED WITH COMBIGAN. WAS ALSO TREATED FOR DRY EYE FOR OVER A MONTH. THE PATIENT'S CURRENT CONDITION IS STILL HAS OCCASIONAL PAIN, ESPECIALLY AT NIGHT. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL IN A ZIPLOC BAG. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. H6- DEVICE HISTORY REVIEW: THE DEVICE HAS BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6- PATIENT CODE 1937 SHOULD BE REMOVED FROM INITIAL MDR AS IT IS NOT APPLICABLE. B5- IT IS ALSO NOTED THAT PRIOR TO SURGERY, IN (B)(6) OF 2019 SHE HAD A "HAIR CLIP INJURY" IN THE RIGHT EYE THAT RESULTED IN A BRUISED EYELID. THE DOCTOR FOUND NO EVIDENCE OF AN INJURY TO THE EYE ITSELF. PATIENT WENT TO SEVERAL OTHER DOCTORS FOR OPINIONS AND WAS AT SOME POINT PLACED ON COMBIGAN TO TREAT ELEVATED PRESSURE RELATED TO STEROID DROP USE. COMBIGAN WAS NOTED TO HAVE BEEN FIRST USED ON (B)(6) 2020 IN THE RIGHT EYE ONLY. COMBIGAN WAS USED IN BOTH EYES UNTIL (B)(6) 2020, AND WAS DISCONTINUED ON (B)(6) 2020 AND THEN AGAIN RESUMED ON (B)(6) 2020 WHICH REPORTED THE PATIENT HAD 34MMHG EYE PRESSURE IN THE LEFT EYE. THE DAY BEFORE THEY FOUND HER EYE PRESSURE TO BE 23MMHG. PATIENT CONTINUED USING COMBIGAN AS OF (B)(6) 2020 AND HAS BEEN TO A CORNEAL SPECIALIST WHO IS INVESTIGATING THE POSSIBILITY OF OPHTHALMIC NEURALGIA. CLAIM# (B)(4). Manufacturer Narrative: REPORT NUMBER - MW5093020. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9483580&pc=MTA | 2023826-2019-02411 | 2019-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Headache; Pain; Red Eye(s); No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2, -07.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REPORTEDLY, PATIENT "HAD RED LEFT EYE WITH PAIN AND HEADACHES. SLIT LAMP EXAMINATION SHOWED EXUDATION OF ANTERIOR CHAMBER, THE SURGEON WASHED THE ANTERIOR CHAMBER AND REMOVED THE LENS ON (B)(6) 2019. AN ANTI-INFLAMMATORY THERAPY WAS TAKEN. BACTERIAL CULTURE SHOWED POSITIVE (ANTERIOR CHAMBER AQUEOUS HUMOR) . NOW THE PATIENT IS IN GOOD CONDITION AND HIS VISUAL ACUITY HAS RECOVERED TO PRE-OPERATIVE UCVA AND BCVA IN THE LAST VISIT." CAUSE OF EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - 'ANTIBIOTICS (CERFUROXIME) WERE PRESCRIBED AS TREATMENT.' SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. H6 - PATIENT CODE 3191 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. H6 - PATIENT CODE 3191: MEDICAL INTERVENTION - ANTIBIOTIC TREATMENT (CEFUROXIME), ANTERIOR CHAMBER WASHOUT SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INPECTION FOUND HAPTIC TORN. METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7853704&pc=MTA | 2023826-2018-01352 | 2018-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-09-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Hypopyon; Inflammation; Intraocular Pressure Increased; Pain; Red Eye(s); Blurred Vision; Vitritis; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED ON (B)(6) 2018 AND OBSERVED TO HAVE INTRAOCULAR INFLAMMATION AND VITRITIS 1 DAY POST-OP. UPON EXAMINATION, THE SURGEON " FOUND REACTION AND MEMBRANES IN AC" AND "WAS OBLIGATED TO KEEP PATIENT INSIDE THE HOSPITAL". THE PATIENT "UNDERGOES INTRA-VITREAL ANTIBIOTIC INJECTION" AND " THE DOCTOR IS WORRIED ABOUT ENDOPHTHALMITIS". PER (B)(6) 2018 EMAIL FROM REPORTER, THE SURGEON HAS EXPLANTED THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). "NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT" NOT INCLUDED IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED ON 11 DEC 2018. THE CULTURE IS (B)(6). PATIENT IS STABLE NOW AND THE PROBLEM IS RESOLVED AFTER INTRAVITREAL INJECTION. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION EVENT- ADDITIONAL INFORMATION WAS RECEIVED ON (B)(6)2018: "FEMALE PATIENT SEEKING REFRACTIVE SURGERY RT UNEVENTFUL ICL IMPLANTATION WAS DONE ROUTINE POST OPERATIVE TTT VIGAMOX ED 1×5 ECONOPRED ED 1×5 2ND DAY POST OPERATIVE THE PATIENT PRESENTED WITH DIMINUTION OF VISION, PAIN AND RED EYE. BY EXAM CILIARY INJECTION, HYPOPYON IN AC, HIGH IOP 32, PUPILLARY MEMBRANE. FUNDUS NOT SEEN. ADMISSION OF THE PATIENT U/S WAS FREE. TREAT AS TASS FREQUENT TOPICAL STEROIDS, HOURLY TOPICAL ANTIBIOTICS, CYCLOPLIC, TOPICAL ANTIGLAUCOMA. BY THE END OF THE DAY NO IMPROVEMENT EXCEPT THE IOP DECREASED U/S SHOWED VITRITIS INTRAVITREAL VANCO AND CEFTAZIDIME 1MG WAS GIVEN 2ND DAY THE HYPOPYON INCREASED AND NO IMPROVMENT AC WASH WAS DONE BUT THE MEMBRANES WERE BEHIND ICL SO I DECIDED TO EXPLANT THE ICL AND PEELING THE MEMBRANES IN FRONT OF THE CRYSTALLINE LENS BUT THE YELLOW REFLEX WAS STILL PRESENT A VITRORETINAL COLLEAGUE STARTED TO PUT PPV SET EXAMINED THE POSTERIOR SEGMENT. WE FOUND DENSE INTRAVITREAL MEMBRANES; VITREOUS SAMPLE WITH THE AQUEOUS SAMPLE WAS TAKEN AND SENT TO MICROBIOLOGISTS. PPV WAS DONE AND INTRAVITREAL VANCO AND CEFTAZIDIME WAS GIVEN AT THE END OF THE SURGERY. THE PATIENT IMPROVED 2ND POST OPERATIVE AFTER 1 WEEK SHE REGAINED HER PREOP BCVA." EVENT- ADDITOINAL INFORMATION RECEIVED ON 30 NOV 2018 FROM REPORTER: A 13.2MM VICMO13.2, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2018. ON (B)(6)2018 THE LENS WAS EXPLANTED DUE TO OBSERVATION OF ELEVATED INTRAOCULAR PRESSURE (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), SUSPECTED ENDOPHTALMITIS AND INFLAMMATION. MEDICAL INTERVENTION (INTRAOCULAR INJECTION) WAS REPORTED. PATIENTS CONDITION AT LAST VISIT DATE: (B)(6)2018, RECOVERED POST OPERATIVE BCVA. THE FOLLOWING WAS REPORTED AS ADDITIONAL INFORMATION IN COMPLAINT FORM: " IT IS INDICATED THAT ITS A CASE OF TASS WITH HYPOPYON IN AC AND PUPILLARY MEMBRANES, INTRAVITREAL MEMBRANES WAS ALSO FOUND AND CASE IMPROVED AFTER PPV WITH INTRAVITREAL VANCOMYCIN AND CEFTAZIDIME INJECTION. PATIENT CODE:(B)(4). CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7356138&pc=MTA | 2023826-2018-00439 | 2018-02-21 | Injury | STAAR SURGICAL COMPANY | 2018-03-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Endophthalmitis; Intraocular Infection; Intraocular Pressure, Delayed, Uncontrolled; Iritis; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED. THE SURGEON NOTES ACUTE ENDOPHTHALMITIS WITH ACUTE IRITIS. REPORTEDLY, THE PATIENT HAS BEEN TREATED WITH AN INTRAOCULAR INJECTION AND TOPICAL DROPS INCLUDING ANTIBIOTICS. THE CLAIM REPORTER LISTS THE CAUSE OF THE ADVERSE EVENT AS HYPOTONY AND WOUND LEAKAGE. NO ADDITIONAL INFORMATION REGARDING PATIENT'S CURRENT STATUS IS KNOWN AT THIS TIME. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER STATED, "THE LEFT EYE HAS IMPROVED AND IS NOW 20/30-2 WITH CONTACT LENS. NO CELLS WERE SEEN IN EXAM WITH POSTERIOR SYNECHAIE WITH SLIGHT PIGMENT OVER ANTERIOR LENS SURFACE NOTED ON (B)(6) 218 EXAM." THE REPORTER STATED WHEN ASKED ABOUT ANY IDEA AS TO WHY WOUND LEAKAGE AND HYPOTONY HAD OCCURRED THAT, "THE ISSUE THE PATIENT HAD WAS DUE TO WOUND LEAKAGE AND THE PATIENT BEING A SMOKER AND EYE RUBBER." (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5826223&pc=MTA | 2023826-2016-00987 | 2016-06-24 | Injury | STAAR SURGICAL COMPANY | 2016-07-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Endophthalmitis; Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND REMAINS IMPLANTED. THE PATIENT EXPERIENCED AN ELEVATED IOP, SUSPECTED ENDOPHTHALMITIS, PIGMENT DISPERSION, AND BLURRED VISION. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED WITH A LONGER LENS. THE PROBLEM WAS RESOLVED. (B)(4). CONCLUSION: PER DFU, THIS LENS MODEL IS CONTRAINDICATED IN THE PATIENT HAD LOW CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY AND AN ANTERIOR CHAMBER DEPTH LOWER THAN 3.00MM. CORRECTED DATA: THERE WAS NO ENDOPHTHALMITIS OR ELEVATED INTRAOCULAR PRESSURE RISE OBSERVED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9278780&pc=MTA | 2023826-2019-02089 | 2019-10-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Intraocular Pressure Increased; Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM5_12.6, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ENDOPHTHALMITIS WAS OBSERVED ON (B)(6) 2019. ELEVATED IOP (29.5 MMHG) WAS ASSESSED ON (B)(6) 2019. TREATMENT OF STEROID, PREDONIN, DIAMOX, EYE DROPS, AND INTERNAL MEDICINE. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY, 'AFTER 3 DAYS FROM SURGERY, PATIENT CLAIMED CLOUDY VISION. FLARE VALUE HIGH. REVOFLO, AND SANBETA PRESCRIBED. AFTER 4 DAYS FROM THE SURGERY, CONFIRMED TOXIC ANTERIOR SEGMENT SYNDROME (TASS) SYMPTOMS AND IOP WAS HIGH (29.5). ENFORCED STEROID AND STARTED INTERNAL USE OF PREDONIN, AND DIAMOX. AS OF (B)(6), THE INFLAMMATION HAS GOTTEN BETTER." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5- "ENDOPHTHALMITIS WAS OBSERVED ON (B)(6) 2019." SHOULD BE REMOVED FROM THE INITIAL MDR. H6- PATIENT CODE 1835 IS NOT APPLICABLE IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- IT WAS REPORTED THAT AFTER 12-15 DAYS FROM SURGERY FIBRIN WAS RECOVERED, DEPOSIT WAS OBSERVED ON THE SURFACE OF ICL LENS, HYPEREMIA WAS RECOVERED AND AN IOP 14 MMHG WAS OBSERVED. AFTER A MONTH FROM SURGERY FIBRIN LIKE TURBIDITY APPEARED AGAIN AND BETAMETHASONE AND STEROID APPLIED ON PATIENT. AFTER 1.5 MONTH FROM SURGERY FIBRIN LIKE TURBIDITY GOT BETTER. THE PATIENT WAS CHANGED TO FLUOROMETHOLONE 0.1% X 4 TIMES PER DAY FROM BETAMETHASONE X4 TIMES PER DAY. AFTER 3 MONTH FROM SURGERY SLIGHT PIGMENT OBSERVED ON ICL BUT LESS INFLAMMATION. THE PATIENT WAS PRESCRIBED A DECREASED DOSE FREQUENCY TO FLUOROMETHOLONE 0.1% X 2 TIMES PER DAY. AFTER 4 MONTH FROM SURGERY THE PIGMENT OBSERVED ON ICL DECREASED AND THER WAS NO INFLAMMATION. THE PATIENT'S DOSE FREQUENCY WAS DECREASED TO FLUOROMETHOLONE 0.1% X 1 TIME PER DAY. THE PATIENT WENT THROUGH 4 MONTHS OF MEDICAL TREATMENT AND THEN FULLY RECOVERED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18048066&pc=MTA | 2023826-2023-04715 | 2023-07-11 | Injury | STAAR SURGICAL COMPANY | 2023-11-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility | Eye Pain | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO ORBITAL PAIN DUE TO PATIENT PSYCHOLOGICAL EFFECTS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR: PSYCHOLOGICAL EFFECTS. REPORTEDLY, PATIENT HAS REVERTED TO GLASSES. Manufacturer Narrative: H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18048064&pc=MTA | 2023826-2023-04714 | 2023-07-18 | Injury | STAAR SURGICAL COMPANY | 2023-11-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility | Eye Pain | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO ORBITAL PAIN DUE TO PATIENT PSYCHOLOGICAL EFFECTS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR: PSYCHOLOGICAL EFFECTS. REPORTEDLY, PATIENT HAS REVERTED TO GLASSES. Manufacturer Narrative: H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10529812&pc=MTA | 2023826-2020-02074 | 2020-08-17 | Injury | STAAR SURGICAL COMPANY | 2020-09-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND ANGLE CLOSURE. A NEW PERIPHERAL IRIDOTOMY WAS ADDED (LASER) BUT NOTHING CHANGED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS REMAINS IMPLANTED. THE PATIENT IS SEEING WELL BUT EYE HURTS. Manufacturer Narrative: H6 - METHOD CODE 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14046187&pc=MTA | 2023826-2022-01093 | 2022-03-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-04-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Fatigue; Headache; Inflammation; Intraocular Pressure Increased; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP) AND AC INFLAMMATION CELLS. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5- CORRECTED TO "THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLNATABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP), AC INFLAMMATION CELLS, PAIN, HEADACHES, NAUSEA, NUMBNESS OF LEG, GENERALIZED BODY WEAKNESS, INFLAMMATION, PIGMENT CELLS, MEDICATION WAS PRESCRIBED. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS UNKNOWN." CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH IMPACT: CLINICAL CODE: 4581- NAUSEA; NUMBNESS OF LEG. SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL #1 MED WATCH REPORT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6487472&pc=MTA | 2023826-2017-00620 | NA | Injury | STAAR SURGICAL COMPANY | 2017-04-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Foreign Body Sensation in Eye; Pain | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT REPORTED HAD SEVERE PAIN BEHIND THE EYE AND FOREIGN BODY SENSATION. THE LENS REMAINS IMPLANTED BUT THE PLAN IS TO EXPLANT AND EXCHANGE FOR A SHORTER LENS. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10590903&pc=MTA | 2023826-2020-02255 | 2020-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Foreign Body Sensation in Eye; Visual Disturbances; Eye Pain | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM, VICMO12.6, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED, IT WAS REPORTED THAT AFTER LENS IMPLANT THE PATIENT EXPERIENCE FOREIGN BODY SENSATION, PAIN AND DOUBLE IMAGES AND REQUIRED REMOVAL. LENS EXPLANT RESOLVED THE PROBLEM. PATIENT RELATED FACTOR WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: D10 - INJECTOR MODEL CORRECTED TO MSI-PF IN THE INITIAL MDR. H6 - WORK ORDER SEARCH: ONE ADDITIONAL SIMILAR COMPLAINT TYPE EVENT WITHIN ASSOCIATED LOTS WAS FOUND. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10590933&pc=MTA | 2023826-2020-02256 | 2020-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Foreign Body Sensation in Eye; Visual Disturbances; Eye Pain | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM, VICMO12.6, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED, IT WAS REPORTED THAT AFTER LENS IMPLANT THE PATIENT EXPERIENCE FOREIGN BODY SENSATION, PAIN AND DOUBLE IMAGES AND REQUIRED REMOVAL. LENS EXPALNT RESOLVED THE PROBLEM. PATIENT RELATED FACTOR WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND FOREIGN RESIDUE ON THE LENS SURFACE AND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: D10 - INJECTOR MODEL CORRECTED TO MSI-PF IN THE INITIAL MDR. H6 - WORK ORDER SEARCH: ONE ADDITIONAL SIMILAR COMPLAINT TYPE EVENT WITHIN ASSOCIATED LOTS WAS FOUND. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16230473&pc=MTA | 2023826-2023-00271 | NA | Injury | STAAR SURGICAL COMPANY | 2023-01-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Glaucoma | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -13 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND GLAUCOMA. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON (B)(6) 2022 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND IT'S UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED; ADDITIONALLY IT NOTED "PATIENT GOT GLAUCOMA BECAUSE OF HIGH VAULT". Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16230029&pc=MTA | 2023826-2023-00272 | NA | Injury | STAAR SURGICAL COMPANY | 2023-01-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Glaucoma | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -12.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 . THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND GLAUCOMA. ON (B)(6) 2022 THE LENS WAS EXPLANTED; ON THIS SAME DATE A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND IT'S UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED; ADDITIONALLY IT NOTED "PATIENT GOT GLAUCOMA BECAUSE OF HIGH VAULT". Manufacturer Narrative: B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE SHOULD BE ADDED TO THE INITIAL MDR. B4 - CORRECTED TO: 15-JAN-2023 IN THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16048692&pc=MTA | 2023826-2022-04678 | NA | Injury | STAAR SURGICAL COMPANY | 2022-12-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Glaucoma; Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS). THE PATIENT REPORTED EYE PROBLEMS AND POSSIBLE LOSS OF VISION. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF A -12.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REPORTEDLY, "AS SOON AS I WOKE UP FROM ANESTHESIA I HAD UNBEARABLE PAIN IN THE LEFT EYE WITH VOMITING AND NAUSEA. IT TURNED OUT THAT I HAD AN ATTACK OF OCULAR HYPERTONIA. THE OPHTHALMOLOGIST PUT ME ON A MANNITOL INFUSION AND GAVE ME DIAMOX TABLETS WITH DROPS TO LOWER THE OCULAR TENSION WITHOUT RESULTS." ON (B)(6)2022, AN OPTICAL COHERENCE TOMOGRAPHY FOUND, "IRRIDOCORNEAN ANGLE IS CLOSED (2 DEGREES). THE LENS WAS EXPLANTED AND THE REPORTER STATED, "I HAD CORNEAL EDEMA, A LOSS OF ENDOTHELIAL CELLS (I ONLY HAVE 1000 AT THE MOMENT) AN OPACITY OF THE LENS (GLAUCOMAFLECKEN BECAUSE OF THE OCULAR HYPERTENSION CRISIS) PUPILLARY MYDRIASIS AND GLAUCOMA WITH SIGNIFICANT LOSS OF OPTIC NERVE FIBERS IN THE EYE." THE RIGHT LENS REMAINS IMPLANTED AND IS CAUSING CHRONIC PAIN AND DOUBLE VISION, AND WITH THE LEFT EYE DAMAGE THE REPORTER STATED, "SUFFERING ANISOMETROPIA DUE TO THE DIFFERENCE OF VISION." REPORTEDLY, "WE HAVE TWO HYPOTHESES TO KNOW WHAT HAPPENED TO MY LEFT EYE EITHER THE IMPLANT WAS BIGGER THAN MY EYE IN THIS CASE THE OPHTHALMOLOGIST HAD TO CHANGE THE IMPLANT IMMEDIATELY OR IT'S DUE TO OBSTRUCTION OF THE CENTRAL AQUAPORT WITH VISCOELASTIC AND INFLAMMATORY DEBRIS." H6: WORK ORDER SEARCH FOUND ONE SIMILAR COMPLAINT - 727444 FOR FELLOW RIGHT EYE (OD). CLAIM#: (B)(4). Manufacturer Narrative: AGE: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. MANUFACTURE DATE: UNK. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13417556&pc=MTA | 2023826-2022-00338 | 2022-01-04 | Injury | STAAR SURGICAL COMPANY | 2022-02-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Glaucoma; Eye Pain | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO LOW VAULT, LENS ROTATION, PAIN, AND GLAUCOMA. THE PROBLEM IS RESOLVED. Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4743980&pc=MTA | 2023826-2015-00510 | NA | Injury | STAAR SURGICAL COMPANY | 2015-05-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Glaucoma; Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN HER RIGHT EYE (OD). THE PATIENT REPORTED WAS PRESCRIBED EYE DROPS FOR AT LEAST THREE CONSECUTIVE MONTHS FOR HIGH PRESSURE OR GLAUCOMA INSIDE THE EYE. THE LENS WAS IMPLANTED ON (B)(6) 2012. THE FACILITY INDICATED THEY AGREED WITH THE PATIENT REPORT. THE PATIENT HAD HIGH PRESSURE (11 - 36MM HG) AND WAS DIAGNOSED WITH GLAUCOMA (OCULAR HYPERTENSION - GLAUCOMA SUSPECT) FOR THE PERIOD OF (B)(6) 2012. THE PATIENT WAS PRESCRIBED MEDICATION FOR THE HIGH PRESSURE AND GLAUCOMA. THE FACILITY INDICATED THE CAUSE OF THE HIGH PRESSURE AND GLAUCOMA WAS FAMILIAL VERSUS OUTFLOW PROBLEM, THE EVENT WAS NOT RELATED TO THE DEVICE. THE PATIENT'S POST-OP VA WAS 20/20. THE LENS REMAINS IMPLANTED AND THE PATIENT'S PROGNOSIS IS GOOD. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - MANY PATHOPHYSIOLOGICAL CHANGES COULD CAUSE OCULAR HYPERTENSION SUCH AS ANTERIOR CHAMBER DISTORTION, VISCOELASTICS, INFLAMMATION, HEMORRHAGE, PIGMENT DISPERSION, PUPILLARY BLOCK AND VITREOUS IN THE ANTERIOR CHAMBER. EXCESSIVE VAULTING AS A FACTOR COULD CAUSE PUPILLARY BLOCK, NARROW ANTERIOR CHAMBER ANGLE OR CLOSURE LEADING TO OCULAR HYPERTENSION. MOST LIKELY VISCOELASTIC SUBSTANCE COULD CAUSE THE TEMPORARY IOP RISE POSTOPERATIVELY WHICH CAN BE CONTROLLED MEDICALLY OR RESUME TO NORMAL PRESSURE AUTOMATICALLY. ALSO MYOPIA IS HIGHLY ASSOCIATED WITH OCULAR HYPERTENSION ESPECIALLY HIGH DEGREE MYOPIA. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4743982&pc=MTA | 2023826-2015-00509 | NA | Injury | STAAR SURGICAL COMPANY | 2015-05-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Glaucoma; Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN HER LEFT EYE (OS). THE PATIENT REPORTED WAS PRESCRIBED EYE DROPS FOR AT LEAST THREE CONSECUTIVE MONTHS FOR HIGH PRESSURE OR GLAUCOMA INSIDE THE EYE. THE LENS WAS IMPLANTED ON (B)(6) 2012. THE FACILITY INDICATED THEY AGREED WITH THE PATIENT REPORT. THE PATIENT HAD HIGH PRESSURE (10 - 21MM HG) AND WAS DIAGNOSED WITH GLAUCOMA (OCULAR HYPERTENSION - GLAUCOMA SUSPECT) FOR THE PERIOD OF (B)(6) 2012. THE PATIENT WAS PRESCRIBED MEDICATION FOR THE HIGH PRESSURE AND GLAUCOMA. THE FACILITY INDICATED THE CAUSE OF THE HIGH PRESSURE AND GLAUCOMA WAS FAMILIAL VERSUS OUTFLOW PROBLEM, THE EVENT WAS NOT RELATED TO THE DEVICE. THE PATIENT'S POST-OP VA WAS 20/20-2. THE LENS REMAINS IMPLANTED AND THE PATIENT'S PROGNOSIS IS GOOD. Manufacturer Narrative: (B)(4) IMPLANTED. DEVICE EVALUATED BY MANUFACTURER? NO: LENS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - MANY PATHOPHYSIOLOGICAL CHANGES COULD CAUSE OCULAR HYPERTENSION SUCH AS ANTERIOR CHAMBER DISTORTION, VISCOELASTICS, INFLAMMATION, HEMORRHAGE, PIGMENT DISPERSION, PUPILLARY BLOCK AND VITREOUS IN THE ANTERIOR CHAMBER. EXCESSIVE VAULTING AS A FACTOR COULD CAUSE PUPILLARY BLOCK, NARROW ANTERIOR CHAMBER ANGLE OR CLOSURE LEADING TO OCULAR HYPERTENSION. MOST LIKELY VISCOELASTIC SUBSTANCE COULD CAUSE THE TEMPORARY IOP RISE POSTOPERATIVELY WHICH CAN BE CONTROLLED MEDICALLY OR RESUME TO NORMAL PRESSURE AUTOMATICALLY. ALSO MYOPIA IS HIGHLY ASSOCIATED WITH OCULAR HYPERTENSION ESPECIALLY HIGH DEGREE MYOPIA. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4513709&pc=MTA | 2023826-2015-00142 | NA | Injury | STAAR SURGICAL COMPANY | 2015-02-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | Glaucoma; Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS IN HIS LEFT EYE (OS). THE PATIENT REPORTED WAS PRESCRIBED EYE DROPS FOR AT LEAST THREE CONSECUTIVE MONTHS DUE TO HIGH PRESSURE OR GLAUCOMA INSIDE THE EYE. THE ICL WAS IMPLANTED ON (B)(6) 2011 AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS IMPLANTED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16326950&pc=MTA | 2023826-2023-00266 | 2022-11-02 | Injury | STAAR SURGICAL COMPANY | 2023-02-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Glaucoma; Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON(B)(6)2022. THE LENS WAS REPORTED AS LOW VAULTING AND REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-00265. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - ENLARGE PERIPHERAL IRIDOTOMY. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO ELEVATED INTRAOCULAR PRESSURE AND GLAUCOMA. AN IRIDOTOMY WAS PERFORMED BUT THE TENSION DID NOT GO DOWN. THE PATIENT HAD A CATARACT LENS IMPLANTED AND THE PROBLEM WAS RESOLVED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-00265. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. H6: MEDICAL DEVICE CODE: 2993. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6855474&pc=MTA | 2023826-2017-01419 | 2016-09-30 | Injury | STAAR SURGICAL COMPANY | 2017-09-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Glaucoma; Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VICMO13.7, -8.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2014. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO ELEVATED IOP, PIGMENT DISPERSION, AND MILD VISUAL FIELD LOSS. SMALL SIGNS OF PIGMENT DISPERSION WAS SEEN AT PREOP EXAM, BUT (CONSIDERING DEVELOPMENT) WAS MIS-ASSESSED. SURGICAL ENLARGEMENT OF PI WAS PERFORMED AND IOP LOWERING EYE DROPS WERE USED. GLAUCOMATOUS DISC WAS SEEN IN THE EYE, BUT OTHERWISE THE EYE IS NORMAL. PATIENT'S POST-OP BCVA ON (B)(6) 2017 WAS 20/30. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO 13.7 IMPLANTABLE COLLAMER LENS -8.5 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. SURGICAL ENLARGEMENT OF PI WAS NOT PERFORMED. (B)(4). Manufacturer Narrative: EXPLANTED (B)(6) 2017. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6054107&pc=MTA | 2023826-2016-01522 | 2016-10-10 | Injury | STAAR SURGICAL COMPANY | 2016-10-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Glaucoma; Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 LENS, -3.50 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. FOLLOWING THE SURGERY, THE PATIENT EXPERIENCED NARROWING OF THE ANGLE AND SHALLOWING OF THE ACD. ON (B)(6) 2016, THE PATIENT HAD ELEVATED INTRAOCULAR PRESSURE (IOP) AND CLOSED ANGLE GLAUCOMA DUE TO EXCESSIVE VAULTING OF THE LENS. THE PATIENT RECEIVED A YAG PERIPHERAL IRIDOTOMY, AND THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED. | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8547014&pc=MTA | 2023826-2019-00667 | 2018-06-20 | Injury | STAAR SURGICAL COMPANY | 2019-04-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Glaucoma; Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED 12.1MM, VTICM5_12.1,-12.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION, BLURRED VISION AND IRIS ASTROPHY. PHACO-EMULSIFICATION AND GLAUCOMA FILTRATION SURGERY WERE REPORTED AS SECONDARY SURGICAL INTERVENTIONS. THE REPORTER STATED "THE PATIENT DEVELOPED PIGMENT DISPERSION GLAUCOMA, BEYOND ANTI-GLAUCOMA CONTROLLED AND THE SURGEON PLANNED FOR GLAUCOMA FILTRATION SURGERY WITH PHACO". POST ICL EXPLANT THE IOP WAS STILL HIGH AND REQUIRED GLAUCOMA SURGERY. PHACO TRABECULECTOMY WAS PERFORMED. THE GLAUCOMA WAS REPORTED AS "CONTROLLED", PATIENTS BCVA 6/20-2. PHACO SURGERY WAS PERFORMED AND IOL WAS IMPLANTED TO REPLACE THE NATURAL CRYSTALLINE LENS. Manufacturer Narrative: BCVA AT LAST POST-OP VISIT WAS 20/20-2. ADDITIONALLY, IT WAS REPORTED THAT THERE WAS NO ASCO (ANTERIOR SUBCAPSULAR LENS OPACITY). DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE OPTIC OF THE LENS WAS BENT AND THE HAPTIC WAS BENT AND DEFORMED. H3- DIMENSIONAL INSPECTION: LENS WAS FOUND TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9703234&pc=MTA | 2023826-2020-00267 | 2020-01-13 | Injury | STAAR SURGICAL COMPANY | 2020-02-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility | Glaucoma; Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -01.75 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP), PIGMENT DISPERSION AND UNREACTIVE (FIXED) PUPIL. THE SURGEON DECIDED NOT TO IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. THE PATIENTS POST-OP CONDITION WAS GLAUCOMA ASSOCIATED WITH PIGMENTARY DISPERSION. Manufacturer Narrative: ADDITIONAL INFORMATION: H6-LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY, WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - DEVICE CODE 1494: OFF-LABEL USE (SAFETY AND EFFECTIVENESS OF THE LENS HAS NOT BEEN ESTABLISHED IN PATIENTS WITH HISTORY OF PREVIOUS OCULAR SURGERY INCLUDING REFRACTIVE CORNEAL SURGERY). CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - METHOD CODE: 3331 SHOULD BE CORRECTED TO 4110 IN SUPPLEMENTAL MEDWATCH REPORT #2. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6262587&pc=MTA | 2023826-2017-00092 | 2015-10-15 | Injury | STAAR SURGICAL COMPANY | 2017-01-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Glaucoma; Intraocular Pressure, Delayed, Uncontrolled; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO COMBINED MECHANISM GLAUCOMA. THE INCREASED INTRAOCULAR PRESSURE WAS NOT CAUSED BY THE ICL BUT WAS DUE TO AN UNKNOWN SICKLE CELL TRAIT AND YAG PI. THE PATIENT HAD REFRACTIVE LENS EXCHANGE AS RECOMMENDED BY A GLAUCOMA PROVIDER AND WAS PERFORMED AT THE TIME OF EXPLANTATION. THE PATIENTS POST-OP BSCVA WAS 20/25. Manufacturer Narrative: LENS WAS RETURNED DRY, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH A PIECE OF THE HAPTIC MISSING. THERE WAS DRY, CLEAR RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8799310&pc=MTA | 2023826-2019-01274 | NA | Injury | STAAR SURGICAL COMPANY | 2019-07-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Glaucoma; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS EXPLANTED FROM THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019 DUE TO GLAUCOMA. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: RETURN DATE SHOULD BE CORRECTED TO 21-JUN-2019 IN THE FOLLOW UP #1 MDR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8799452&pc=MTA | 2023826-2019-01273 | NA | Injury | STAAR SURGICAL COMPANY | 2019-07-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Glaucoma; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 MM, MICL13.2, -13.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS EXPLANTED FROM THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019 DUE TO GLAUCOMA. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: RETURN DATE SHOULD BE CORRECTED TO 21-JUN-2019 IN PREVIOUS MDR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5193790&pc=MTA | 2023826-2015-01396 | NA | Injury | STAAR SURGICAL COMPANY | 2015-11-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Glaucoma; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO ANGLE CLOSURE GLAUCOMA. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: RESULTS (EVALUATION RESULT): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A DENT ON A LENS HAPTIC. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, VISUAL INSPECTION OF THE RETURNED PRODUCT AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): MEDICAL REVIEW. RESULTS: (EVALUATION RESULT): PER MEDICAL REVIEW - REPORTEDLY ICL WAS EXPLANTED FOUR DAYS POSTOPERATIVELY TO ADDRESS ELEVATED IOP DUE TO ANGLE CLOSURE GLAUCOMA. NO POSTOPERATIVE SEQUELAE WERE REPORTED. ACCORDING TO THE REPORT, DATED ON (B)(6) 2015, THE SURGEON STATED THAT HE BELIEVED THAT THERE WAS RETAINED VISCOELASTIC (OCUCOAT) BEHIND IMPLANTED ICL. IT SHOULD BE NOTED THAT ACUTE RISE IN IOP COULD BE A CONSEQUENCE OF RETAINED VISCOELASTIC AND PER DFU PERCAUTIONS: "3. INADEQUATE FLUSHING OF THE VISCOELASTIC FROM THE EYE MAY LEAD TO INTRAOCULAR PRESSURE (IOP) SPIKES. IOP SHOULD BE CHECKED 24 HOURS POSTOPERATIVELY." GIVEN THIS INFORMATION IT IS VERY LIKELY THAT PROCEDURE RELATED FACTOR (LEFT VISCOELASTIC) HAS CAUSED THE REPORTED EVENT. CONCLUSION: (EVALUATION CONCLUSION): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION AND MEDICAL REVIEW, IT IS VERY LIKELY THAT PROCEDURE RELATED FACTOR (LEFT VISCOELASTIC) HAS CAUSED THE REPORTED EVENT. (B)(4). Manufacturer Narrative: METHOD (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372539&pc=MTA | 2023826-2019-00285 | 2018-07-16 | Injury | STAAR SURGICAL COMPANY | 2019-02-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Glaucoma; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER, INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2018. ON (B)(6) 2019 THE SURGEON EXCHANGED THE LENS WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ANGLE CLOSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND GLAUCOMA. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS "DIAMETER OF LENS TOO LONG." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A ZIP LOCK BAG. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9244518&pc=MTA | 2023826-2019-01974 | 2019-10-01 | Injury | STAAR SURGICAL COMPANY | 2019-10-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Glaucoma; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -02.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REPORTEDLY, CELL COUNT DECREASE AND ACUTE GLAUCOMA WERE OBSERVED. PER THE REPORTER ON (B)(6) 2019, LENS WAS REMOVED AND EXCHANGED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. DEVICE CODE 1494: OFF LABEL USE, ACD < 3.0MM. DEVICE CODE 2993 NOT APPLICABLE. EVENT: THE LENS WAS IMPLANTED ON (B)(6) 2019. A VAULT OF 2000 MICRONS WAS ALSO REPORTED. THE LENS WAS EXCHANGED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE WAS REPORTED AS DUE TO THE DEVICE. DEVICE EVALUATED BY MFR: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: ONE ADDITIONAL SIMILAR COMPLAINT TYPE EVENT WITHIN ASSOCIATED LOTS WAS FOUND. CLAIM# (B)(4). Manufacturer Narrative: A LASER PI WAS PERFORMED 1 WEEK PRIOR TO IMPLANTATION. EXCESSIVE VAULT WAS REPORTED. CLAIM# (B)(4). Manufacturer Narrative: COUNT DECREASE AND ACUTE GLAUCOMA WERE OBSERVED.' CORRECTED TO ' CELL COUNT DECREASE AND CAUTION OF ACUTE GLAUCOMA WERE OBSERVED.' CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. WEIGHT: UNK. DATE OF EVENT: UNK. EXPLANTED: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12168068&pc=MTA | 2023826-2021-02624 | 2021-03-05 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6; -12.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCED HEADACHES. IT WAS NOTED THERE IS NO PLAN TO REPLACE THE LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: TYPE OF INVESTIGATION : LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. G4-PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170268&pc=MTA | 2023826-2021-02625 | 2021-03-05 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6; -12.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCED HEADACHES. IT WAS NOTED THERE IS NO PLAN TO REPLACE THE LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6997961&pc=MTA | 2023826-2017-01639 | NA | Injury | STAAR SURGICAL COMPANY | 2017-11-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Headache | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED HAS EXPERIENCED "MIGRAINE AURA". THE LENS WAS IMPLANTED ON (B)(6) 2013 AND REMAINS IMPLANTED. THE PATIENT REPORTED THE EVENT STARTED IN (B)(6) 2017 AND OCCURS EVERY COUPLE OF WEEKS. THE DOCTOR'S OFFICE CONFIRMED THE EVENT AND INDICATED THE CAUSE WAS UNKNOWN. PATIENT HAD LASIK ON (B)(6) 2013. THE PATIENT HAS NOT BEEN PRESCRIBED ANY MEDICATION AND THE DATE OF LAST VISIT WAS (B)(6) 2017. THE PATIENT'S VISUAL ACUITY POST-OP WAS 20/25. Manufacturer Narrative: (B)(4) Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND. (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4957436&pc=MTA | 2023826-2015-00990 | 2015-04-13 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-31 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Anxiety; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED THE 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2014 AND ON (B)(6) 2015 THE PATIENT WAS EXPERIENCING HEADACHES, ANXIETY AND INSOMIA. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND WAS NOT REPLACED. SEE MFR REPORT# 2023826-2015-00991 FOR THE OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4). A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. BASED ON THE COMPLAINT INFORMATION AND WORK ORDER SEARCH, WE ARE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4957478&pc=MTA | 2023826-2015-00991 | 2015-04-13 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-31 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Anxiety; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED THE 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2014 AND ON (B)(6) 2015 THE PATIENT WAS EXPERIENCING HEADACHES, ANXIETY AND INSOMNIA. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND WAS NOT REPLACED. SEE MFR REPORT# 2023826-2015-00990 FOR THE OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). METHOD - WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - BASED ON THE COMPLAINT INFORMATION AND WORK ORDER SEARCH, WE ARE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6560720&pc=MTA | 2023826-2017-00791 | NA | Injury | STAAR SURGICAL COMPANY | 2017-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2016. THE REPORTER INDICATED THERE WAS OFF-TARGET REFRACTION (REFRACTIVE SURPRISE) AND LOW VAULTING WITH THE LENS. THE PATIENT WAS EXPERIENCING HEADACHES. THE LENS REMAINS IMPLANTED BUT WILL BE EXPLANTED AT A LATER DATE. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11510275&pc=MTA | 2023826-2021-00681 | 2021-01-11 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Headache; Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM5_12.1, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. PATIENT RELATED FACTOR WAS REPORTED FOR CAUSE OF EVENT, THE REPORTER STATED "PATIENT EXPERIENCED HEADACHE." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14440020&pc=MTA | 2023826-2022-01618 | 2022-04-18 | Injury | STAAR SURGICAL COMPANY | 2022-05-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -9.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. PIGMENT DISPERSION, FIXED PUPIL AND IRIS ATROPHY WAS OBSERVED. ON (B)(6) 2022 PI YAG WAS PERFORMED. LENS REMAINS IMPLANTED. FOR CAUSE DETAILS THE REPORTER STATED, "GOOD SURGICAL OUTCOME, HAPPY VISION, BUT HAS HAD EPISODES OF HEADACHE AND BLURRED VISION ABOUT 2-3 WEEKS POST-OP. IOP ON EXAM WAS NORMAL, VAULT NORMAL." FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED "PERSISTENT, MILD DILATED, FIXED PUPIL WITH DIFFUSE TID OD ONLY. ACUITY GOOD, BUT FLUCTUATING." FOR ADDITIONAL INFORMATION THE REPORTER STATED "THIS IS THE THIRD CASE OF FIXED PUPIL WITH DIFFUSE TIDS. ALL WITHIN LAST MONTH, ALL RIGHT EYES, NO COMMON THEMES. HAPPY WITH VISION, SO NOT LIKELY TO REMOVE LENS, BUT CONCERNED ABOUT LONG TERM ISSUES (IOP,GLARE,NIGHT VISION). AFTER DISCUSSION WITH COLLEAGUES, FEEL THIS MAY BE UNUSUAL CASE OF MOXIFLOXACIN TOXICITY." Manufacturer Narrative: (B)(4). TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 - SHOULD BE CORRECTED TO IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. CLAIM#: (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6587683&pc=MTA | 2023826-2017-00874 | 2016-09-22 | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Headache; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -4.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, BLURRED VISION, AND SEVERE HEADACHE. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: PREVIOUS: LENS HAS NOT BEEN RETURNED. UPDATED: LENS HAS BEEN RETURNED AND EVALUATED. DEVICE EVALUATION: LENS WAS RETURNED DRY IN A SMALL VIAL, WITH CLEAR DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. METHOD CODE: ADDITIONAL CODE ADDED. RESULT CODE: UPDATED BASED ON ADDITIONAL NFORMATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). (SECONDARY SURGERY, LENS EXCHANGE, SIGNIFICANT REDUCTION OF ICA). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12052958&pc=MTA | 2023826-2021-02041 | 2021-02-23 | Injury | STAAR SURGICAL COMPANY | 2021-06-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Headache; Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO TRANSIENT LOSS OF BEST CORRECTED VISUAL ACUITY (BCVA), GLARE/HALOS, HEADACHE AND BLURRED VISION. THE PIS WERE ENLARGED BY YAG. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE ICL IS IN A GOOD POSITION AND NO MORE HEADACHES OR VA LOSS. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12052969&pc=MTA | 2023826-2021-02042 | 2021-02-23 | Injury | STAAR SURGICAL COMPANY | 2021-06-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Headache; Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO TRANSIENT LOSS OF BEST CORRECTED VISUAL ACUITY (BCVA), GLARE/HALOS, HEADACHE AND BLURRED VISION. THE PIS WERE ENLARGED BY YAG. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE ICL IS IN A GOOD POSITION AND NO MORE HEADACHES. Manufacturer Narrative: ENLARGE PIS. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11441206&pc=MTA | 2023826-2021-00619 | 2021-01-02 | Injury | STAAR SURGICAL COMPANY | 2021-03-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Discomfort | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -02.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT HAD A HEADACHE AND DISCOMFORT. THE SYMPTOM WAS RESOLVED AFTER THE LENS WAS REMOVED. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE REPORTER INDICATED THERE WAS NO ABNORMAL ISSUE IN POST-OP, THE PATIENT FELT DISCOMFORT BECAUSE OF HER SUBJECTIVE SYMPTOMS AND PERSONALITY. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11441242&pc=MTA | 2023826-2021-00618 | 2021-01-02 | Injury | STAAR SURGICAL COMPANY | 2021-03-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Discomfort | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -02.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT HAD A HEADACHE AND DISCOMFORT. THE SYMPTOM WAS RESOLVED AFTER THE LENS WAS REMOVED. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE REPORTER INDICATED THERE WAS NO ABNORMAL ISSUE IN POST-OP, THE PATIENT FELT DISCOMFORT BECAUSE OF HER SUBJECTIVE SYMPTOMS AND PERSONALITY. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12073174&pc=MTA | 2023826-2021-02030 | 2021-05-11 | Injury | STAAR SURGICAL COMPANY | 2021-06-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Discomfort | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. IT WAS REPORTED THAT THE PATIENT COMPLAINED OF DISCOMFORT IN SEEING THINGS, SEVERE HEADACHE, AFFECTING NORMAL LIFE AND WORK. ON (B)(6) 2021 THE LENS WAS EXPLANTED PER PATIENTS REQUEST AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION PATIENT RELATED FACTOR WAS THE CAUSE OF THIS EVENT. FOR ADDITIONAL INFORMATION THE REPORTER STATED " THE PATIENT WAS (B)(6) YEARS OLD WITH PRESBYOPIA. IN ORDER TO DELAY THE TIME OF WEARING PRESBYOPIA AFTER OPERATION, IT IS SUGGESTED TO KEEP A CERTAIN DEGREE AFTER OPERATION. THE PRESBYOPIA PROBLEM WAS COMMUNICATED TO THE PATIENT MANY TIMES BEFORE OPERATION, BUT THE PATIENT WAS UNWILLING TO KEEP THE PRESBYOPIA DEGREE AND INSISTED ON COMPREHENSIVE CORRECTION." Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12073149&pc=MTA | 2023826-2021-02029 | 2021-05-11 | Injury | STAAR SURGICAL COMPANY | 2021-06-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Discomfort | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -5.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. IT WAS REPORTED THAT THE PATIENT COMPLAINED OF DISCOMFORT IN SEEING THINGS, SEVERE HEADACHE, AFFECTING NORMAL LIFE AND WORK. ON (B)(6) 2021 THE LENS WAS EXPLANTED PER PATIENTS REQUEST AND THE PROBLEM RESOLVED. IN THE REPORTER'S OPINION PATIENT RELATED FACTOR WAS THE CAUSE OF THIS EVENT. FOR ADDITIONAL INFORMATION THE REPORTER STATED "THE PATIENT WAS (B)(6) WITH PRESBYOPIA. IN ORDER TO DELAY THE TIME OF WEARING PRESBYOPIA AFTER OPERATION, IT IS SUGGESTED TO KEEP A CERTAIN DEGREE AFTER OPERATION. THE PRESBYOPIA PROBLEM WAS COMMUNICATED TO THE PATIENT MANY TIMES BEFORE OPERATION, BUT THE PATIENT WAS UNWILLING TO KEEP THE PRESBYOPIA DEGREE AND INSISTED ON COMPREHENSIVE CORRECTION." Manufacturer Narrative: B5: LENS WAS IMPALNTED ON (B)(6) 2021. CLAIM (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15442519&pc=MTA | 2023826-2022-03148 | 2022-03-05 | Injury | STAAR SURGICAL COMPANY | 2022-09-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Discomfort | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO DISCOMFORT AND HEADACHE. THE EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15442553&pc=MTA | 2023826-2022-03149 | 2022-03-05 | Injury | STAAR SURGICAL COMPANY | 2022-09-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Discomfort | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO DISCOMFORT AND HEADACHE. THE EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17477607&pc=MTA | 2023826-2023-03351 | 2023-05-01 | Injury | STAAR SURGICAL COMPANY | 2023-08-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Discomfort | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 LENS WAS EXPLANTED DUE TO PATIENT DISSATISFACTION- HEADACHE AND DISCOMFORT AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16766106&pc=MTA | 2023826-2023-01417 | 2023-03-20 | Injury | STAAR SURGICAL COMPANY | 2023-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Headache; Discomfort; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED. REPORTEDLY, "THE PATIENT UNDERWENT BILATERAL ICL LENS IMPLANTATION IN OUR HOSPITAL ON (B)(6) 202 AND (B)(6) 2022 RESPECTIVELY. AFTER THE OPERATION, HE COMPLAINED OF A HEADACHE AND DISCOMFORT ." THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- IMPACT CODE: "1880" AND "2330" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT:UNKNOWN, ETHNICITY:UNKNOWN, RACE:UNKNOWN. PATIENT RELATED FACTOR. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16766108&pc=MTA | 2023826-2023-01418 | 2023-03-20 | Injury | STAAR SURGICAL COMPANY | 2023-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Headache; Discomfort; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED. REPORTEDLY, "THE PATIENT UNDERWENT BILATERAL ICL LENS IMPLANTATION IN OUR HOSPITAL ON (B)(6) 2023. AFTER THE OPERATION, HE COMPLAINED OF A HEADACHE AND DISCOMFORT ." THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: A4-A6:UNK. H6: PATIENT RELATED FACTOR. CLAIM# (B)(4). Manufacturer Narrative: H6-HEALTH EFFECT- IMPACT CODE: "1880" AND "2330" SHOULD BE ADDED. D.6A SHOULD BE CORRECTED TO 02/18/2023. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0 INTO THE PATIENT'S LEFT EYE (OS) ON 18/FEB/2023. THE PATIENT EXPERIENCED HEADACHES AND DISCOMFORT. ON 04/MARCH/2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14030225&pc=MTA | 2023826-2022-01092 | 2022-03-12 | Injury | STAAR SURGICAL COMPANY | 2022-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Headache; Inflammation; Intraocular Pressure Increased; Blurred Vision; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP), PIGMENT DISPERSION, GLARE/HALOES, BLURRED VISION, AC INFLAMMATION CELLS AND IRREGULAR PUPIL. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2022. AT 3 DAY POST-OP, FIBRIN CONTRACTING AS PATIENT DEVELOPED SEVERE INFLAMMATION POST EXPLANT. EARLY CATARACT FORMATION NOTED. THE PATIENT IS BEING MONITORED. CLAIM # (B)(4). Manufacturer Narrative: B5- PATIENT ALSO EXPERIENCED HEADACHES; NAUSEA; WEAKNESS; MEDICATION REQUIRED. SHOULD HAVE BEEN INCLUDED IN INITIAL MED WATCH REPORT. H6- HEALTH IMPACT: CLINICAL CODE: 4581- NAUSEA; WEAKNESS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN, EXPLANT DATE: UNKNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15964027&pc=MTA | 2023826-2022-04484 | 2022-11-08 | Injury | STAAR SURGICAL COMPANY | 2022-12-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Headache; Intraocular Pressure Increased; Blurred Vision; Dizziness | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS WITH A -7.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED. REPORTEDLY, "AFTER THE PATIENT'S SURGERY, VISION, REFRACTIVE STATE, INTRAOCULAR PRESSURE, AND ARCH HIGH ARE NORMAL. THE PATIENT CONSCIOUSLY DIZZY AND HEADACHE." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: 1880 ADDED FOR HEADACHE. 2194 ADDED FOR DIZZINESS. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15965745&pc=MTA | 2023826-2022-04483 | 2022-11-08 | Injury | STAAR SURGICAL COMPANY | 2022-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Headache; Intraocular Pressure Increased; Blurred Vision; Dizziness | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS WITH A -8.5 DIOPTER WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED. REPORTEDLY, "AFTER THE PATIENT'S SURGERY, VISION, REFRACTIVE STATE, INTRAOCULAR PRESSURE, AND ARCH HIGH ARE NORMAL. THE PATIENT CONSCIOUSLY DIZZY AND HEADACHE." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: PATIENT INFORMATION: UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: 1880 ADDED FOR HEADACHE. 2194 ADDED FOR DIZZINESS. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7404820&pc=MTA | 2023826-2018-00508 | 2018-03-06 | Injury | STAAR SURGICAL COMPANY | 2018-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE SURGEON INDICATED THAT A 13.2MM MICL13.2, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. PATIENT REPORTED TO HAVE BLURRY VISION AND HEADACHE. ON (B)(6) 2018 THE SURGEON OBSERVED EXCESSIVE VAULT, ELEVATED IOP (INTRAOCULAR PRESSURE), NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ACD (ANTERIOR CHAMBER DEPTH). PATIENT WAS PUT ON GLAUCOMA MEDICATION AND AN ANTERIOR CHAMBER WASH OUT WAS PERFORMED. THE LENS WAS EXPLANTED ON (B)(6) 2018. GLAUCOMA DROPS WERE GIVEN AND IOP (INTRAOCULAR PRESSURE) RETURNED TO NORMAL. SURGEON REPORTED THAT THE CAUSE OF EVENT IS UNKNOWN- COMBINATION OF TRAUMA (PATIENT ACCIDENTLY HIT HER EYE WITH HAND ON (B)(6) 2018), LENS LENGTH, SMALL 2ND LPI (LASER PERIPHERAL IRIDOTOMY). Manufacturer Narrative: (B)(4) Manufacturer Narrative: PATIENT CODE: SECONDARY SURGICAL INTERVENTION, EXPLANT. NARROWING OF THE ANGLE, ANGLE CLOSURE, SHALLOWING OF THE ACD. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781720&pc=MTA | 2023826-2020-00391 | NA | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Headache; Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 LENS, -10.0 DIOPTER INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), ANGLE CLOSURE, BLURRED VISION, AND HEADACHES. THE PROBLEM IS RESOLVED. REPORTEDLY THE PATIENT IS "SEEING WELL, NO HEADACHES, PATIENT REPORTS HAVING SLIGHT TWITCH IN THIS EYE NO PAIN." THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6324404&pc=MTA | 2023826-2017-00233 | 2016-12-28 | Injury | STAAR SURGICAL COMPANY | 2017-02-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; Nausea; Pain; Pupillary Block; Blurred Vision; Vomiting; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXTREME VAULTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED: THE REPORTER INDICATED THE PATIENT CALLED POST-SURGERY WITH EXTREME PAIN, HEADACHE, NAUSEA AND VOMITING. THE PATIENTS PRESSURE WAS 50. THE WOUND WAS BURPED BUT WOULD NOT GO DOWN BELOW 37. A SECONDARY YAG PI WAS PERFORMED BUT WITH NO PRESSURE RELIEF. THERE WAS NARROWING OF THE ANGLE AND MECHANICAL ANGLE CLOSURE. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND A SHORTER LENS WAS IMPLANTED ON (B)(6) 2016. THE PATIENTS CURRENT VA IS 20/20 AND IS DOING WELL. THE REPORTER STATED THE EVENT WAS PRODUCT-RELATED. DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID AND THERE WAS CLEAR RESIDUE ON THE LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7065520&pc=MTA | 2023826-2017-01781 | 2017-10-13 | Injury | STAAR SURGICAL COMPANY | 2017-11-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; Nausea; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (PUPILLARY BLOCK) AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH (ACD). A REVISION OF THE PERIPHERAL IRIDECTOMIES WAS PERFORMED. THE PATIENT EXPERIENCED BLURRY VISUAL ACUITY, HEADACHE AND NAUSEA. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25 AND IOP WAS 21. Manufacturer Narrative: (B)(4) . DATE OF EVENT: IT IS CORRECTED FROM UNK TO (B)(6) 2017. (B)(4) Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7065648&pc=MTA | 2023826-2017-01782 | 2017-10-13 | Injury | STAAR SURGICAL COMPANY | 2017-11-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; Nausea; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (PUPILLARY BLOCK). THE PATIENT EXPERIENCED BLURRY VISION, HEADACHE AND NAUSEA. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25 AND IOP WAS 18. Manufacturer Narrative: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND PIECES OF LENS HAPTIC TORN OFF AND MISSING. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6853814&pc=MTA | 2023826-2017-01415 | NA | Injury | STAAR SURGICAL COMPANY | 2017-09-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO ELEVATED INTRAOCULAR PRESSURE AND PIGMENT DISPERSION. THE PATIENT EXPERIENCED ADDITIONAL SYMPTOMS - HEADACHE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S CURRENT CONDITION POST-OP - BSCVA - 20/30, INTRAOCULAR PRESSURE - 18MM HG. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PIGMENT DISPERSION/ANATOMICAL ANOMALY (ABNORMAL SULCUS ANATOMY). THE PATIENT HAS PLANNED THERAPY OF LASER VISION CORRECTION (LVC). Manufacturer Narrative: LENS WAS RETURNED IN LENS CASE. VISUAL INSPECTION FOUND HAPTIC DEFORMED. CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5728010&pc=MTA | 2023826-2016-00823 | 2015-10-02 | Injury | STAAR SURGICAL COMPANY | 2016-06-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE AND PIGMENT DISPERSION. THE PATIENT EXPERIENCED HEADACHES. THE REPORTER STATED THE CAUSE OF THE EVENT WAS DUE TO ABNORMAL POSTERIOR IRIS ANATOMY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS DISCARDED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5728011&pc=MTA | 2023826-2016-00822 | 2015-10-02 | Injury | STAAR SURGICAL COMPANY | 2016-06-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE AND PIGMENT DISPERSION. THE PATIENT EXPERIENCED HEADACHES. THE REPORTER STATED THE CAUSE OF THE EVENT WAS DUE TO ABNORMAL POSTERIOR IRIS ANATOMY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS DISCARDED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6765864&pc=MTA | 2023826-2017-01265 | 2017-07-06 | Injury | STAAR SURGICAL COMPANY | 2017-08-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR A SMALLER SIZE LENS ON (B)(6) 2017 DUE TO HIGH VAULT, ELEVATED IOP, AND NARROWING OF THE ANGLE. HEADACHE WAS ALSO REPORTED AS AN ADDITIONAL SYMPTOM. THE REPORTER STATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENTS POST-OP BCVA WAS 20/20. ADDITIONAL INFORMATION HAS BEEN REQUESTED, IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS RETURNED IN A LIQUID PRESENT IN A VIAL. VISUAL INSPECTION FOUND A PIECE OF HAPTIC MISSING FROM THE LENS. (B)(4). Manufacturer Narrative: THE PATIENT'S POST-OP BCVA WAS 20/20" IS INCORRECT. THE PATIENT'S POST-OP REFRACTION BCVA WAS FOUND TO BE 20/20 ON (B)(6) 2017 BEFORE THE EXCHANGE. CONCOMITANT MEDICAL PRODUCTS: INJECTOR, CARTRIDGE AND INJECTOR SYSTEM MODEL ARE KNOWN. (B)(4). Manufacturer Narrative: (B)(4). A WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320744&pc=MTA | 2023826-2018-00327 | 2018-02-03 | Injury | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -14.00, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE SURGEON NOTES THE FOLLOWING: EXCESSIVE VAULT, ELEVATED IOP, IRIS PIGMENTS ON ICL, "MASSIVE" HEADACHE, APPOSITIONAL ANGLE CLOSURE. ADDITIONAL INFORMATION PROVIDED: 1 WEEK POST-OP THE IOP BECAME ELEVATED, PATIENT WAS GIVE IOP DROPS (COSOPT+ALPHAGAN). ON (B)(6) 2018 IT WAS REPORTED THAT "THE CENTRAL HOLE IS OPEN." THERE IS "THE PRESENCE OF APPOSITIONAL ANGLE CLOSURE...ACETAZOLAMIDE WAS ADDED WHEN THE PRESSURE REACHED HIGH THIRTIES ON FULL TOPICAL MEDICATIONS. STEROIDS WERE STOPPED TO RULE OUT THE POSSIBILITY OF STEROID RESPONSE. NOW THE IOP IS 19MMHG ON COSOPT ONLY WITH A HIGH VAULT...AND LOTS OF IRIS PIGMENTS ON THE ICL..." THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: TWO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5668555&pc=MTA | 2023826-2016-00732 | NA | Injury | STAAR SURGICAL COMPANY | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Headache; Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO HIGH VAULT AND PRESSURE SPIKE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE PATIENT COMPLAINED OF A HEADACHE. RESULTS: (OPERATIONAL PROBLEM): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF BOTH HAPTICS TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL. CONCLUSIONS: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6453292&pc=MTA | 2023826-2017-00534 | 2017-01-30 | Injury | STAAR SURGICAL COMPANY | 2017-04-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Manufacturing, Packaging or Shipping Problem | Headache; Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED HEADACHE, ELEVATED IOP. THE CUSTOMER STATES THERE WAS NO CENTER HOLE ON THE LENS. ON (B)(6) 2017 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN SMALL VIAL, THERE WAS MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE, NO HOLE ON THE LENS, CLEAR SURGICAL RESIDUE ON LENS SURFACE, AND THERE WAS INK STAINING ON THE SURFACE OF THE LENS. THE LENS WAS REMEASURED, AND PASSED THE LENS VERIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6918173&pc=MTA | 2023826-2017-01503 | 2017-07-18 | Injury | STAAR SURGICAL COMPANY | 2017-10-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Shelf Life Exceeded; Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.7, -7.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO INTERMITTENT ELEVATED IOP, HEADACHES, EXCESSIVE VAULTING, NARROWING OF THE ANGLE, AND SHALLOWING OF THE ACD. THE LENS WAS EXCHANGED WITH A NEW LENS. Manufacturer Narrative: PATIENT CODE: (B)(4) NARROWING OF THE ANGLE, SHALLOWING OF THE ACD, SECONDARY SURGERY, LENS EXCHANGE. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138562&pc=MTA | 2023826-2020-01184 | 2019-11-05 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Optical Problem | Headache; Intraocular Pressure Increased; Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -04.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ELEVATED IOP (29MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS ASSESSED ON (B)(6) 2020, GLARE/HALOES WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR (ACCOMMODATION SPASM). REPORTEDLY, "GLARE/HALOES, HEAD PAIN DIMINISHED, PATIENT STILL CANNOT DRIVE BY NIGHT. PATIENT HAS HAD SOME STOMATOLOGY CAL PROBLEMS. AFTER RESOLVING THOSE, THE SYMPTOMATOLOGY AMELIORATED SIGNIFICANTLY." PER THE REPORTER ON (B)(6) 2020, "PATIENT WAS SCHEDULED FOR LENS REMOVAL BUT CANCELED BECAUSE OF THE SIGNIFICANT AMELIORATION OF THE SX AND NORMAL IOP WITHOUT TREATMENT." Manufacturer Narrative: B4 - CORRECTED TO (B)(6) 2020 IN INTIAL MDR. G3 - CORRECTED TO (B)(6) 2020 IN INTIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138567&pc=MTA | 2023826-2020-01185 | 2019-11-05 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL | Patient-Device Incompatibility; Optical Problem | Headache; Intraocular Pressure Increased; Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -04.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ELEVATED IOP (27MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS ASSESSED ON (B)(6) 2020, GLARE/HALOES WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR (ACCOMMODATION SPASM). REPORTEDLY, "GLARE/HALOES, HEAD PAIN DIMINISHED, PATIENT STILL CANNOT DRIVE BY NIGHT. PATIENT HAS HAD SOME STOMATOLOGY CAL PROBLEMS. AFTER RESOLVING THOSE, THE SYMPTOMATOLOGY AMELIORATED SIGNIFICANTLY." PER THE REPORTER ON (B)(6) 2020, "PATIENT WAS SCHEDULED FOR LENS REMOVAL BUT CANCELED BECAUSE OF THE SIGNIFICANT AMELIORATION OF THE SX AND NORMAL IOP WITHOUT TREATMENT." Manufacturer Narrative: B4 - CORRECTED TO (B)(6) 2020 IN INTIAL MDR. G3 - CORRECTED TO (B)(6) 2020 IN INTIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1 - CORRECTED TO ADVERSE EVENT B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT B5 - PHOTOPHOBIA WAS ALSO REPORTED. SECONDARY YAG WAS PERFORMED. H6 - HEALTH EFFECT CLINICAL CODE 1880 H6 - HEALTH EFFECT IMPACT CODE 4644 H6 - HEALTH EFFECT IMPACT CODE 4625: ADDITIONAL SURGERY (SECONDARY YAG) ADDITIONAL INFORMATION: H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510(K) - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13582873&pc=MTA | 2023826-2022-00595 | 2022-01-12 | Injury | STAAR SURGICAL COMPANY | 2022-02-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Intraocular Pressure Increased; Visual Disturbances; Halo; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM512.6, -12.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE PATIENT EXPERIENCED ELEVATED IOP (INTRAOCULAR PRESSURE), HEADACHE AND NAUSEA, MEDICAL INTERVENTION (PILOCARPINE/SAMPIRO 3-4/DAY) WAS PRESCRIBED. ON (B)(6) 2022 THE PATIENT VISITED AN EMERGENCY CLINIC DUE TO EYE PAIN, HEADACHE AND NAUSEA, IT WAS REPORTED THAT IOP (INTRAOCULAR PRESSURE) WAS NORMAL. PATIENT COMPLAINED OF FEELING SICK FROM HALO AND GLARE (LIGHTS ON THE STREET AT NIGHT) AND INCREASED EYE PAIN IN THE EVENING. THE PATIENT ALSO COMPLAINED OF FEELING SICK UNDER FLUORESCENT LIGHTS AND NOT BEING ABLE TO SLEEP DUE TO PAIN DURING SLEEP. LENS REMAINS IMPLANTED. IT WAS REPORTED THAT ELEVATED IOP (INTRAOCULAR PRESSURE) IS TRANSIENT, CAUSE OF EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION HALO, GLARE AND NIGHT BLINDNESS ARE BEING MONITORED. THERE IS NO PLAN TO REPLACE THE LENS NOW. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: D4 - SERIAL# (B)(6) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. (B)(4). CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13582877&pc=MTA | 2023826-2022-00594 | 2022-01-12 | Injury | STAAR SURGICAL COMPANY | 2022-02-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Intraocular Pressure Increased; Visual Disturbances; Halo; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM512.6, -11.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE PATIENT EXPERIENCED ELEVATED IOP (INTRAOCULAR PRESSURE), HEADACHE AND NAUSEA, MEDICAL INTERVENTION (PILOCARPINE/SAMPIRO 3-4/DAY) WAS PRESCRIBED. ON (B)(6) 2022 THE PATIENT VISITED AN EMERGENCY CLINIC DUE TO EYE PAIN, HEADACHE AND NAUSEA, IT WAS REPORTED THAT IOP (INTRAOCULAR PRESSURE) WAS NORMAL. PATIENT COMPLAINED OF FEELING SICK FROM HALO AND GLARE (LIGHTS ON THE STREET AT NIGHT) AND INCREASED EYE PAIN IN THE EVENING. THE PATIENT ALSO COMPLAINED OF FEELING SICK UNDER FLUORESCENT LIGHTS AND NOT BEING ABLE TO SLEEP DUE TO PAIN DURING SLEEP. LENS REMAINS IMPLANTED. IT WAS REPORTED THAT ELEVATED IOP (INTRAOCULAR PRESSURE) IS TRANSIENT, CAUSE OF EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION HALO, GLARE AND NIGHT BLINDNESS ARE BEING MONITORED. THERE IS NO PLAN TO REPLACE THE LENS NOW. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: D4 - SERIAL # (B)(6) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. DEVICE MANUFACTURING DATE: UNK, NO SERIAL NUMBER REPORTED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4767216&pc=MTA | 2023826-2015-00531 | 2015-03-09 | Injury | STAAR SURGICAL COMPANY | 2015-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Headache; Intraocular Pressure Increased; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. AT THE SAME DAY POST-OPERATIVE EXAMINATION, THE PATIENT HAD 20/60 UNCORRECTED VISUAL ACUITY, INTRAOCULAR PRESSURE BY APPLANATION WAS 41MM HG AND LENS HAD A 200% VAULT. THE PATIENT COMPLAINED OF SLIGHT HEADACHE AND SHE WAS EXPERIENCING "STARBURSTS SIMILAR TO MIGRAINE AURA". LPI ENHANCEMENT WAS NECESSARY AND "BURPING" OF THE WOUND WAS PERFORMED. THE INTRAOCULAR PRESSURE WAS LOWERED TO SATISFACTORY LEVELS AND MEDICATION WAS PRESCRIBED. AT ONE DAY POST-OPERATIVE VISIT, THE SURGEON DECIDED THE LENS WAS TOO LARGE AND THE LENS WAS EXPLANTED ON (B)(6) 2015. THE LENS WAS EXCHANGED FOR A SHORTER LENS. AT SAME DAY POST-OPERATIVE VISIT, THE INTRAOCULAR PRESSURE WAS 44MM HG, UNCORRECTED VISUAL ACUITY WAS 20/50, WITH 1+ CORNEAL EDEMA AND 75% VAULTING OF LENS. "BURPING" OF THE WOUND LOWERED THE INTRAOCULAR PRESSURE AND PATIENT WAS PRESCRIBED MEDICATION. AT POST-OPERATIVE VISIT ON (B)(6) 2015, PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/30, INTRAOCULAR PRESSURE WAS 18MM HG WITH CLEAR CORNEA AND LENS HAD 75% VAULT. Manufacturer Narrative: (B)(4). INCISION OPENED. SECONDARY SURGERY. NEW IRIDECTOMY. EXPLANTED. VAULTING. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - GIVEN THE REPORTED INFORMATION, MOST LIKELY THE EXCESSIVE VAULTING CAUSED SHALLOW ANTERIOR CHAMBER, ANGLE NARROWING OR CLOSURE, AND/OR PUPILLARY BLOCK WHICH WERE ASSOCIATED WITH INCREASED IOP. THE SECOND IOP INCREASE AFTER THE LENS EXCHANGE COULD BE CAUSED BY SURGICAL RELATED ANTERIOR CHAMBER ANGLE DISTORTION, PRIOR SUSPECT GLAUCOMA, PRIOR OPEN ANGLE GLAUCOMA OR VISCOELASTIC SUBSTANCE USED IN THE SURGERY. ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING WAS A CONSEQUENCE OF WRONG LENS USE WHICH INCLUDE THE FOLLOWING: IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENT BASED UPON DIFFERENT TECHNIQUES, POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN A INDIVIDUAL CASE, IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST AND IMPROPER LENS LABEL. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: EVALUATION: RESULTS - VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF DARK SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7794679&pc=MTA | 2023826-2018-01199 | 2018-01-25 | Injury | STAAR SURGICAL COMPANY | 2018-08-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Intraocular Pressure Increased; Vomiting; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. TWO DAYS AFTER THE SURGERY, THE PATIENT SUFFERED FROM SEVERE VOMITING, HEADACHE, AC LOSS AND TENSION OF MORE THAN 40. THE SURGEON ADMINISTERED MANNITOL BY IV AND THE SYMPTOMS DECREASED. AFTER 1 WEEK, THE LENS WAS EXPLANTED ON (B)(6) 2018. A SHORTER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4542564&pc=MTA | 2023826-2015-00176 | 2015-01-15 | Injury | STAAR SURGICAL COMPANY | 2015-02-24 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Intraocular Pressure Increased; Vomiting; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THAT EVENING AFTER SURGERY THE PATIENT HAD HEADACHE AND VOMITING, SUGGESTIVE OF INCREASED INTRAOCULAR PRESSURE. THE PATIENT HAS ALSO COMPLAINED OF GLARE. AT THE LAST POST-OP VISIT, THE PATIENT'S SYMPTOMS HAVE SUBSIDED BUT INTRAOCULAR PRESSURE WAS 24. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - INTRAOCULAR PRESSURE RISE; VOMITING; HEADACHE; (GLARE). DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THIS LENS THAT WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6373025&pc=MTA | 2023826-2017-00299 | 2017-01-05 | Injury | STAAR SURGICAL COMPANY | 2017-03-02 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Headache; Intraocular Pressure, Delayed, Uncontrolled; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS, -9.00 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017, THE PATIENT EXPERIENCED PIGMENT DISPERSION, HEADCHES, AND OCULAR PRESSURE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT AND THEREFORE, HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). DEVICE REMAINS IMPLANTED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6381321&pc=MTA | 2023826-2017-00325 | 2017-01-05 | Injury | STAAR SURGICAL COMPANY | 2017-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Headache; Intraocular Pressure, Delayed, Uncontrolled; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017 THE PATIENT EXPERIENCED PIGMENT DISPERSION, HEADACHES, AND OCULAR PRESSURE. ON (B)(6) 2017 IT WAS DOCUMENTED THAT THERE WAS NO FREE PIGMENT IN ANTERIOR CHAMBER OF THE PATIENT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). INTRAOCULAR PRESSURE, DELAYED UNCONTROLLED NOT REPORTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. RESULT: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WERE REPORTED FOR UNITS WITH IN THE SAME LOT. CLAIM#(B)(4). DEVICE REMAINS IMPLANTED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4614654&pc=MTA | 2023826-2015-00344 | 2014-05-30 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Nausea; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO THE PATIENT EXPERIENCED GLARE/HALOS, HEADACHE AND WAS NAUSEOUS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/24. Manufacturer Narrative: (B)(4): LENS NOT RETURNED. Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - GLARE/HALOS ARE A SUBJECTIVE PERCEPTION OF STARBURSTS OF LIGHT (LIGHT AROUND OR COMING OFF OBJECTS/LIGHT SOURCES). USUALLY MORE DETECTABLE AT NIGHT TIME, BUT IT CAN ALSO BE DETECTED DURING DAYTIME. POTENTIAL CAUSES ARE: RESIDUAL REFRACTIVE ERRORS, OPTICAL ZONE OF THE TREATMENT/IMPLANT SMALLER THAN PRE-OPERATIVE PUPIL SIZE IN MESOPIC CONDITIONS, DECENTERED TREATMENT/IMPLANT, LIGHT FROM THE IRIDOTOMY SITES, ETC. USUALLY THE PERCEPTION OF GLARE IS TRANSIENT, ONLY RELATED TO SOME LIGHTING CONDITIONS AND DIMINISHES/DISAPPEARS OVER A COURSE OF WEEKS/MONTHS. IN SOME CASES THE PERCEPTION MAY PERSIST AND BE SIGNIFICANT. IF THE PATIENT IS HAVING SIGNIFICANT DISTURBANCES AND IF THE IMPLANT IS THE SOURCE OF GLARE, IT CAN BE EXPLANTED THROUGH THE SAME INCISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4614667&pc=MTA | 2023826-2015-00345 | 2014-05-30 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Nausea; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO THE PATIENT EXPERIENCED GLARE/HALOS, HEADACHE AND WAS NAUSEOUS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/24. Manufacturer Narrative: (B)(4): LENS NOT RETURNED. Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - GLARE/HALOS ARE A SUBJECTIVE PERCEPTION OF STARBURSTS OF LIGHT (LIGHT AROUND OR COMING OFF OBJECTS/LIGHT SOURCES). USUALLY MORE DETECTABLE AT NIGHT TIME, BUT IT CAN ALSO BE DETECTED DURING DAYTIME. POTENTIAL CAUSES ARE: RESIDUAL REFRACTIVE ERRORS, OPTICAL ZONE OF THE TREATMENT/IMPLANT SMALLER THAN PRE-OPERATIVE PUPIL SIZE IN MESOPIC CONDITIONS, DECENTERED TREATMENT/IMPLANT, LIGHT FROM THE IRIDOTOMY SITES, ETC. USUALLY THE PERCEPTION OF GLARE IS TRANSIENT, ONLY RELATED TO SOME LIGHTING CONDITIONS AND DIMINISHES/DISAPPEARS OVER A COURSE OF WEEKS/MONTHS. IN SOME CASES THE PERCEPTION MAY PERSIST AND BE SIGNIFICANT. IF THE PATIENT IS HAVING SIGNIFICANT DISTURBANCES AND IF THE IMPLANT IS THE SOURCE OF GLARE, IT CAN BE EXPLANTED THROUGH THE SAME INCISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335509&pc=MTA | 2023826-2020-01564 | 2020-07-02 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Nausea; Pain; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS HAD EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF ICA AND PIGMENT DISPERSION. THE PATIENT EXPERIENCED EYE PAIN, BLURRING OF VISION AND NAUSEA. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE LENS WAS EXPLANTED ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT DOES NOT FEEL HEADACHE ANYMORE. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: UNK. EXPIRATION DATE UNK. NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8671658&pc=MTA | 2023826-2019-01019 | 2015-11-12 | Injury | STAAR SURGICAL COMPANY | 2019-06-05 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Headache; Nausea; Pain; Red Eye(s); Blurred Vision; Burning Sensation; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 2015. THE SURGEON REPORTS PIGMENT DISPERSION, RED EYE, AND PERMANENT EYE PAIN SINCE IMPLANTATION. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, HOWEVER THE SURGEON THINKS THE FOOTPLATES OF THE ICL SEEM TO BE LOW, PLACED ON/IN BETWEEN ZONULAS OR ON/IN THE CILIARY BODY-WHICH COULD CAUSE THE PAIN. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PATIENT CODE1880: HEADACHE. CLAIM#: (B)(4). Manufacturer Narrative: PER EMAIL ON 2019/07/25. OD IMPLANT SURGERY WAS ON "(B)(6)" OR (B)(6) 2015. CLAIM#: (B)(4). Manufacturer Narrative: DATE OF EVENT: DATE IS APPROXIMATE. INITIAL OPERATION OCCURRED EITHER (B)(6) 2015. AFTER IMPLANT PATIENT REPORTS HYPERMETROPIA OF +1 AND ASTIGMATISM OF 0.75. GLASSES WERE WORN AND PARTIALLY IMPROVED THE SITUATION. REPORTEDLY LASIK WAS PERFORMED IN (B)(6) 2017 WHICH IMPROVED THE PROBLEM. HOWEVER, IN (B)(6) 2018 THE RIGHT EYE BECAME WORSE AND THE PATIENT STATES PAIN IS CONSTANT ESPECIALLY WHEN READING. THERE IS ALSO A SLIGHT BURNING BEHIND THE EYE AND A FEELING OF NAUSEA ON THE RIGHT SIDE, EVEN AT NIGHT. IMPLANT DATE: DATE IS APPROXIMATE. ADDITIONAL PATIENT CODE ADDED, 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LASIK), LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: PRODUCT IS MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9257758&pc=MTA | 2023826-2019-02014 | 2018-09-03 | Injury | STAAR SURGICAL COMPANY | 2019-10-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PATIENT COMPLAINED OF A HEADACHE. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9266062&pc=MTA | 2023826-2019-02013 | 2018-09-03 | Injury | STAAR SURGICAL COMPANY | 2019-11-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PATIENT COMPLAINED OF A HEADACHE. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5112607&pc=MTA | 2023826-2015-01187 | 2015-04-09 | Injury | STAAR SURGICAL COMPANY | 2015-09-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO PATIENT SUFFERED FROM HEADACHE. NO OTHER LENS WAS IMPLANTED. SEE MFR #2023826-2015-01188 FOR LEFT EYE. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: BASED ON THE DHR REVIEW, THE MANUFACTURING AND PACKAGING PROCESSES WAS PERFORMED ACCORDING TO PROCEDURES AND NO FINDINGS OR DEVIATIONS WERE REPORTED. THE PRODUCT EVALUATION CONFIRMED THE LENS WAS IN SPECIFICATION. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINES, HOWEVER, THE COMPLAINT IS NOT DEVICE RELATED. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THAT PIECES OF THE LENS HAPTICS WERE TORN OFF. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -11.00 (B)(4). METHOD CODE: EVALUATION METHOD: WORK ORDER SEARCH. EVALUATION CODE: EVALUATION RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS CODE: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6048731&pc=MTA | 2023826-2016-01083 | 2016-01-26 | Injury | STAAR SURGICAL COMPANY | 2016-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Headache; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE PATIENT EXPERIENCING HEADACHE. Manufacturer Narrative: CORRECTED INFORMATION: (B)(4) SHOULD BE UNDER PATIENT CODE. ADDITIONAL INFORMATION: CONCLUSION CODE: PATIENT IS UNDER THE AGE OF 21 AND PER DFU FOR THIS MODEL, THIS LENS MODEL IS CONTRADICTED FOR THE PATIENT'S UNDER THE AGE OF 21 YEARS. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO DAMAGE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6052593&pc=MTA | 2023826-2016-01084 | 2016-01-26 | Injury | STAAR SURGICAL COMPANY | 2016-10-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Headache; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE PATIENT EXPERIENCING HEADACHE. Manufacturer Narrative: (B)(4). CONCLUSION CODE: PATIENT IS UNDER THE AGE OF (B)(6) AND PER DFU FOR THIS MODEL, THIS LENS MODEL IS CONTRADICTED FOR THE PATIENTS UNDER THE AGE OF (B)(6) YEARS. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE HAPTIC BENT AND PRESENCE OF FOREIGN MATERIAL (SPOT) ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10310562&pc=MTA | 2023826-2020-01501 | NA | Injury | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED INTERMITTENT HEADACHES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE HEADACHES HAVE BEEN RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A STERILE POUCH, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5788401&pc=MTA | 2023826-2016-00898 | 2016-06-07 | Injury | STAAR SURGICAL COMPANY | 2016-07-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Headache; Pain; Blurred Vision; Vomiting; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE LENS BEING TOO LARGE AND HYPOTONY (LOW IOP). THE PATIENT EXPERIENCED BLURRED VISUAL ACUITY, EYE PAIN, HEADACHE, AND VOMITING. THE REPORTER INDICATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/25. Manufacturer Narrative: (B)(4). RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN A SMALL CLEAR CONTAINER. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13057261&pc=MTA | 2023826-2021-04915 | 2021-02-08 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Pain; Dizziness; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED OF EYE DISTENSION, EYE PAIN, DIZZINESS AND HEADACHE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED OF EYE DISTENSION, EYE PAIN, DIZZINESS AND HEADACHE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13057173&pc=MTA | 2023826-2021-04916 | 2021-02-09 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Pain; Dizziness; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED OF EYE DISTENSION, EYE PAIN, DIZZINESS AND HEADACHE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. PATIENT IN GOOD CONDITION AFTER SURGERY.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED OF EYE DISTENSION, EYE PAIN, DIZZINESS AND HEADACHE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. PATIENT IN GOOD CONDITION AFTER SURGERY. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5047711&pc=MTA | 2023826-2015-01091 | 2015-04-25 | Injury | STAAR SURGICAL COMPANY | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Pain; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -09.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO PATIENT DISSATISFACTION WITH HEADACHE AND OPHTHALMODYNIA. NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. DIOPTER: -09.00. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS WAS NOT RETURNED. | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5047815&pc=MTA | 2023826-2015-01092 | 2015-04-25 | Injury | STAAR SURGICAL COMPANY | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Pain; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -09.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO PATIENT DISSATISFACTION WITH HEADACHE AND OPHTHALMODYNIA. NO OTHER LENS WAS IMPLANTED. SEE MFR. # 2023826-2015-01091 FOR RIGHT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -09.50 (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS WAS NOT RETURNED. | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17986867&pc=MTA | 2023826-2023-04600 | 2023-01-13 | Injury | STAAR SURGICAL COMPANY | 2023-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Headache; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS, HEADACHE, AND PATIENT DISSATISFACTION. PROBLEM RESOLVED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6694800&pc=MTA | 2023826-2017-01121 | 2017-05-10 | Injury | STAAR SURGICAL COMPANY | 2017-07-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Headache; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED HEADACHE, GLARE/HALOES. ON (B)(6) 2017 THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. THE CUSTOMER STATED THAT THE PATIENT DOES NOT HAVE PLAN TO EXCHANGE THE LENS. AS OF THE DAY OF MDR SUBMISSION THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A LENS HAPTIC SLIGHTLY BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119129&pc=MTA | 2023826-2015-01262 | 2015-09-01 | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Hyphema; Hypopyon; Inflammation; Irritation; Pain; Visual Impairment | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER STATED AT ONE DAY POST-OP ((B)(6) 2015), THE PATIENT'S EYE HAD HYPHEMA AND HYPOPHON. MEDICATIONS WERE ADMINISTERED FOR INFLAMMATION BUT THE REPORTER STATED IT LOOKED MORE LIKE AN IRRITATED IRIS. THE PATIENT WAS IN PAIN AND CULTURES WERE TAKEN. THE LENS VAULT IS GOOD. THE PATIENT RETURNED FOR A FOLLOW-UP VISIT ON (B)(6) 2015 AND THE VA WAS 20/20, THERE IS STILL SOME CELL IN THE EYE AND THE VISION IS NOT AS CLEAR. THE REPORTER STATED THE TEST ON THE CULTURE CAME BACK AS TOXIC ANTERIOR SEGMENT SYNDROME. THE REPORTER STATED THE PATIENT'S EYE HAS SETTLED DOWN AND IS RECOVERING. THE PATIENT'S VA WAS 20/20. Manufacturer Narrative: (B)(4). RESULTS: PER MEDICAL REVIEW - ENDOPHTHALMITIS IS A PAINFUL INFECTION OF OCULAR TISSUES THAT MAY OCCUR POSTOPERATIVELY. UNLESS IT IS VIRAL IN ORIGIN, A CULTURE WILL USUALLY BE POSITIVE. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) IS A PAINLESS INFLAMMATORY POSTOPERATIVE CONDITION THAT HAS BEEN ASSOCIATED WITH DRUGS, DEVICES, AND SURGICAL ISSUES. CULTURE IS USUALLY NEGATIVE. IT IS NOT CLEAR WHICH OF THESE CONDITIONS DEVELOPED IN THE PATIENT'S EYE, AND THE CAUSE OF EITHER CONDITION IN THIS CASE IS UNKNOWN. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, CARTRIDGE AND FOAM TIP PLUNGER LOT NUMBER SEARCHES, DEVICE HISTORY RECORD REVIEW AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). METHOD - (PROCESS EVALUATION): LENS WORK ORDER SEARCH, CARTRIDGE LOT NUMBER SEARCH, FOAM TIP PLUNGER LOT NUMBER SEARCH, DEVICE HISTORY RECORD REVIEW. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. A CARTRIDGE LOT NUMBER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. A FOAM TIP PLUNGER LOT NUMBER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS, INJECTOR AND CARTRIDGE THAT WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, LENS WORK ORDER SEARCH, CARTRIDGE AND FOAM TIP PLUNGER LOT NUMBER SEARCHES AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS IMPLANTED. | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8504763&pc=MTA | 2023826-2019-00588 | 2019-02-27 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Inflammation; Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM. MICL12.6, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS IS TOO LARGE FOR THE PATIENT, INCREASED EYE PRESSURE. WILL BE REPLACING WITH A SHORTER LENGTH LENS. NO PATIENT INJURY. HOWEVER, THE PATIENT LATER STATED " IF THERE WAS SOMETHING WRONG THE IMPLANT WOULD HAVE BEEN REMOVED". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: DOCTOR IS NOT REMOVING OR REPLACING THE CURRENT IMPLANT. THE DOCTOR FELT BECAUSE THE VIEW WAS HAZY DUE TO INFLAMMATION AND PRESSURE THE VAULT SIZE HAD TO BE TOO BIG, WHICH MUST BE LEADING TO THE INCREASE IN EYE PRESSURE. THE PATIENT WAS MONITORED FOR APPROXIMATELY 2-3 MONTHS. THE PATIENTS PRESSURE WENT DOWN TO NORMAL RANGE AND VISION IS 20/20. THE VAULT SIZE IS 100%. PT IS VERY HAPPY WITH OUTCOME. PATIENT CODE:1932 ADDED TO INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14028818&pc=MTA | 2023826-2022-01108 | NA | Injury | STAAR SURGICAL COMPANY | 2022-04-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Inflammation; Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -7.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ELEVATED IOP WAS OBSERVED. THE LENS REMAINS IMPLANTED. CHRONIC INFLAMMATION IS REPORTED AND DOCTOR STILL CONTINUES TO MONITOR PATIENT WITH COMBIGAN LE. CAUSE REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13389438&pc=MTA | 2023826-2022-00352 | 2021-12-30 | Injury | STAAR SURGICAL COMPANY | 2022-01-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1; -5.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (52MMHG) WHICH COULD NOT BE CONTROLLED WITH MAXIMAL TOPICAL ANTIGLAUCOMA +ORAL DIAMOX MEDICATION. PATIENT'S DRAINAGE ANGLE WAS HEAVILY PIGMENTED WITH PIGMENTS AND EXCESSIVE PIGMENTS WERE NOTED ANTERIOR CHAMBER. THE MEDICATION REDUCED THE IOP TO 9MMHG WITH ADDITIONAL ORAL GLYCEROL AT A LATER DATE. IT WAS ALSO INDICATED AT A LATER DATE INFLAMMATION OCCURRED AND ANTI-INFLAMMATORY EYEDROPS AND PREDNISOLONE WAS GIVEN. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND IT WAS REPORTED THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14030262&pc=MTA | 2023826-2022-01091 | 2022-02-24 | Injury | STAAR SURGICAL COMPANY | 2022-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP) AND PIGMENT DISPERSION. THE VAULT WAS NORMAL. POST-OP THERE WAS MILD INFLAMMATION IN ANTERIOR CHAMBER. IOP WAS 15MMHG AND MEDICATION WAS PRESCRIBED. THE PATIENT IS BEING MONITORED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15506509&pc=MTA | 2023826-2022-03375 | 2022-07-03 | Injury | STAAR SURGICAL COMPANY | 2022-09-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Inflammation; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -5.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED ELEVATED IOP WITHOUT PUPIL BLOCK; PIGMENT DISPERSION AND PRO-LONGED INFLAMMATION; "11 O'CLOCK IRIS TAGS". ON (B)(6) 2022 THE LENS WAS REPOSITIONED TO FREE UP THE LENS FROM THE IRIS TOUCH AND GONIO WASH WAS PERFORMED TO REDUCE THE PIGMENTS IN AC. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | Lens Repositioning | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873864&pc=MTA | 2023826-2016-01066 | 2016-06-18 | Injury | STAAR SURGICAL COMPANY | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Inflammation; Intraocular Pressure Increased; Blurred Vision; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -2.5 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT, GLARE, HALOES, BLURRED VISION, INFLAMMATION, ANGLE CLOSURE WITH ELEVATED IOP AND VERY SHALLOW AC. THE PATIENT HAD A LOSS OF BEST CORRECTED VISUAL ACUITY. THE LENS WAS EXPLANTED AND THERE WAS ENLARGEMENT OF INCISION. THE PATIENTS CONDITION ON THE LAST VISIT WAS HYPERAEMIA AND SMALL DISC HEMORRHAGE INFERIORLY. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: (B)(4). CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -2.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON 06/20/2016 DUE TO EXCESSIVE VAULTING, GLARE/HALOES, BLURRED VISION, INFLAMMATION, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE, AND VERY SHALLOW ANTERIOR CHAMBER. THE PATIENT HAD A LOSS OF BEST CORRECTED VISUAL ACUITY (BCVA). THE LENS WAS EXPLANTED ON (B)(6) 2016 AND THERE WAS ENLARGEMENT OF INCISION. THE SURGEON PERFORMED AN ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS AND A PARS PLANA VITRECTOMY TO DRAIN THE ACQUEOUS MISDIRECTION. THE PATIENT'S CONDITION ON THE LAST VISIT WAS HYPERAEMIA AND SMALL DISC HEMORRHAGE INFERIORLY. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/30. USER FACILITY IS THE INCORRECT REPORT SOURCE, SHOULD BE DISTRIBUTOR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873698&pc=MTA | 2023826-2016-01062 | 2016-06-18 | Injury | STAAR SURGICAL COMPANY | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Inflammation; Intraocular Pressure Increased; Blurred Vision; Loss of Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -02.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SHALLOW ANTERIOR CHAMBER, ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE AND INFLAMMATION. THE PATIENT EXPERIENCED LOSS OF BEST-CORRECTED VISUAL ACUITY, GLARE/HALOS AND BLURRED VISION. THE INCISION WAS ENLARGED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. THE PATIENT'S EYE IS IN GOOD CONDITION AFTER EXPLANTATION AND THE INFLAMMATION WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7714154&pc=MTA | 2023826-2018-01016 | 2018-06-12 | Injury | STAAR SURGICAL COMPANY | 2018-07-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Inflammation; Intraocular Pressure Increased; Iritis; Uveitis; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2, -13.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO OBSERVATION OF EXCESSIVE VAULT, ELEVATED IOP (INTRAOCULAR PRESSURE) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, IRIS ATROPHY, INCREASED CELL, AND FLARE/INFLAMMATION. PATIENT WAS ON ANTIBIOTICS AND ANTI-INFLAMMATORIES TO ADDRESS THE INFLAMMATION. IT WAS REPORTED THAT AT PATIENTS LAST FOLLOW UP THERE WAS NO PAIN, NO ACHE, DECREASED INFLAMMATION AND PATIENT IS DOING WELL. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A SMALL PLASTIC CONTAINER. VISUAL INSPECTION FOUND A HAPTIC TEAR. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4529095&pc=MTA | 2023826-2015-00185 | 2014-12-12 | Injury | STAAR SURGICAL COMPANY | 2015-02-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Inflammation; Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP), INFLAMMATION AND POSSIBLE IRIDOCYCLITIS. THE LENS WAS NOT REPLACED AND THE PATIENT'S LAST BCVA WAS 20/15. Manufacturer Narrative: DATE OF BIRTH - UNK. PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THIS LENS THAT WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: REFRACTIVE LENS EXCHANGE (PHACO-IOL) WAS PERFORMED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). CORRECTED DATA: (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7714241&pc=MTA | 2023826-2018-01017 | 2018-06-12 | Injury | STAAR SURGICAL COMPANY | 2018-07-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Inflammation; Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO OBSERVATION OF EXCESSIVE VAULT, ELEVATED IOP (INTRAOCULAR PRESSURE) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, INCREASED CELL, AND FLARE/INFLAMMATION. PATIENT WAS ON ANTIBIOTICS AND ANTI-INFLAMMATORIES TO ADDRESS THE INFLAMMATION. IT WAS REPORTED THAT AT PATIENTS LAST FOLLOW UP THERE WAS NO PAIN, NO ACHE, DECREASED INFLAMMATION AND PATIENT IS DOING WELL. Manufacturer Narrative: (B)(4) Manufacturer Narrative: DEVICE EVALUATION:LENS WAS RETURNED IN LIQUID, IN A SMALL PLASTIC CONTAINER. VISUAL INSPECTION FOUND A HAPTIC TEAR. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13830688&pc=MTA | 2023826-2022-00880 | NA | Injury | STAAR SURGICAL COMPANY | 2022-03-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Intraocular Pressure Increased; Pupillary Block; Red Eye(s); Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1; -10.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE; AC CELL++; CHRONIC INFLAMMATION; PUPIL BLOCK; DISCOMFORT AND RED EYE; OVER A PERIOD OF TIME AFTER IMPLANTATION; MULTIPLE MEDICATIONS WERE ALSO PRESCRIBED. THE SURGEON THEN PERFORMED A SECONDARY YAG ON (B)(6) 2021 WHICH THEN RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14710489&pc=MTA | 2023826-2022-01917 | 2022-03-14 | Injury | STAAR SURGICAL COMPANY | 2022-06-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient Device Interaction Problem | Inflammation; Intraocular Pressure Increased; Red Eye(s) | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -8.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO INFLAMMATION AND PIGMENT DISPERSION OCCURRED.; THEN STEROID EYE DROPS WERE PRESCRIBED AND INDUCED IOP. A TRABECULECTOMY WAS PERFORMED. CAUSE IS DUE TO A PATIENT RELATED FACTOR OF PIGMENT DISPERSION SYNDROME. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. (B)(4). CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14710500&pc=MTA | 2023826-2022-01918 | 2022-03-14 | Injury | STAAR SURGICAL COMPANY | 2022-06-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient Device Interaction Problem | Inflammation; Intraocular Pressure Increased; Red Eye(s) | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -7.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(4) 2022. ON (B)(4) 2022 THE LENS WAS EXPLANTED DUE TO INFLAMMATION AND PIGMENT DISPERSION OCCURRED.; THEN STEROID EYE DROPS WERE PRESCRIBED AND INDUCED IOP. A TRABECULECTOMY WAS PERFORMED. CAUSE IS DUE TO A PATIENT RELATED FACTOR OF PIGMENT DISPERSION SYNDROME. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. (B)(4). CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8742035&pc=MTA | 2023826-2019-01155 | 2019-05-31 | Injury | STAAR SURGICAL COMPANY | 2019-06-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Inflammation; Intraocular Pressure Increased; Uveitis; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, PIGMENT DISPERSION AND ANTERIOR CHAMBER INFLAMMATION. THE REPORTER INDICATED THE EXCESSIVE VAULTING CAUSED PIGMENT DISPERSION (INCREASING INTRAOCULAR PRESSURE CONTROLLED ON DROPS AND HEAVY PIGMENT ON ICL), AND ANTERIOR CHAMBER INFLAMMATION SPIKES AS STEROIDS WERE TAPERED. THE PATIENT HAD CONSTANT ACHING SENSATION BUT THE PATIENT IS DOING WELL AT ONE-DAY POST-OP EXPLANT EXAM. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: DATE OF EVENT: CORRECTED TO (B)(6) 2019 IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MFR: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND A HAPTIC TEAR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8742046&pc=MTA | 2023826-2019-01154 | 2019-05-31 | Injury | STAAR SURGICAL COMPANY | 2019-06-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Inflammation; Intraocular Pressure Increased; Uveitis; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, PIGMENT DISPERSION AND ANTERIOR CHAMBER INFLAMMATION. THE REPORTER INDICATED THE EXCESSIVE VAULTING CAUSED PIGMENT DISPERSION (INCREASING INTRAOCULAR PRESSURE CONTROLLED ON DROPS AND HEAVY PIGMENT ON ICL), AND ANTERIOR CHAMBER INFLAMMATION SPIKES AS STEROIDS WERE TAPERED. THE PATIENT HAD CONSTANT ACHING SENSATION BUT THE PATIENT IS DOING WELL AT ONE-DAY POST-OP EXPLANT EXAM. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: DAVE OF EVENT: CORRECTED TO (B)(6) 2019 IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9528000&pc=MTA | 2023826-2019-02499 | 2019-10-21 | Injury | STAAR SURGICAL COMPANY | 2019-12-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Pain; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED HAD AN MICL IMPLANTABLE COLLAMER LENS, IMPLANTED IN HER LEFT EYE (OS) ON (B)(6) 2019 WITH GOOD RESULTS. THE PATIENT REPORTED ON (B)(6) 2019, AFTER A 3 HOUR FLIGHT, THERE WAS A RED RING AROUND HER LEFT EYE AND HER PUPILS WERE DIFFERENT SIZES. TWO DAYS LATER, SHE DEVELOPED PAIN AND LIGHT SENSITIVITY IN THE LEFT EYE. THE PATIENT WENT TO HER DOCTOR AND THEY SAID THE PATIENT HAD REBOUND INFLAMMATION AND WAS PUT BACK ON STEROIDS. THE PATIENT HAS REPORTED CLOUDINESS AND HALF HER VISION WENT GRAY AND SHE COULD NOT SEE. TO THIS DAY, THE PATIENT HAS 2 DIFFERENT SIZE PUPILS AND VISION IN LEFT EYE IS CLOUDY. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED AN 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED "GRAY VISION TRANSIENT". THE CAUSE OF THE EVENT WAS UNKNOWN. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14187607&pc=MTA | 2023826-2022-01182 | NA | Injury | STAAR SURGICAL COMPANY | 2022-04-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Inflammation; Red Eye(s); Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5 12.1 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS HAD LOW VAULTING AND THERE WAS PIGMENT DISPERSION. THE IOP WAS CONTROLLED WITH TOPICAL ANTIGLAUCOMA EYEDROPS. THE PATIENT HAD PIGMENT DISPERSION SYNDROME 2 WEEKS POST ICL IMPLANTATION THAT COMPLETELY RESOLVED 2 MONTHS LATER. HOWEVER, SHE AGAIN HAD PIGMENTS IN A/C 1 MONTH LATER. PATIENT IS STILL UNDER CLOSE MONITORING. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- PATIENT EXPERIENCED INFLAMMATION, PAIN, AND REDNESS.- SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13317683&pc=MTA | 2023826-2022-00196 | 2020-04-27 | Injury | STAAR SURGICAL COMPANY | 2022-01-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(4) 2020. THE SURGEON REPORTS ELEVATED IOP; CONCOMITANT MEDICATIONS PRESCRIBED. REPORTEDLY THE PROBLEM IS RESOLVED.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS ELEVATED IOP; CONCOMITANT MEDICATIONS PRESCRIBED. REPORTEDLY THE PROBLEM IS RESOLVED. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17278995&pc=MTA | 2023826-2023-02608 | 2021-08-16 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO ELEVATED IOP(51MMHG)WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESSED ON (B)(6) 2021) ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, IOP WAS TREATED WITH TABLETS OVER TIME. PRESSURE KEPT INCREASING WITHOUT THERAPY(RX) SO SURGEON DECIDED TO EXPLANT. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-HEALTH IMPACT CODE: "4644" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.00 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED AN ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE AND MEDICATION WAS PRESCRIBED. REPORTEDLY "THE IOP WAS TREATED WITH TABLETS OVER TIME. BECAUSE THE PRESSURE KEPT INCREASING WITHOUT THERAPY THE SURGEON DECIDED TO FINALLY EXPLANT THE LENS". ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION PROVIDED: "PATIENT IS HAPPY". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17278997&pc=MTA | 2023826-2023-02609 | 2021-08-16 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO ELEVATED IOP(34MMHG)WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESSED ON (B)(6) 2021) ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, IOP WAS TREATED WITH TABLETS OVER TIME. PRESSURE KEPT INCREASING WITHOUT THERAPY(RX) SO SURGEON DECIDED TO EXPLANT. Manufacturer Narrative: CORRECTED DATA: H6-HEALTH IMPACT CODE: "4644" SHOULD BE ADDED. B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.00 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED AN ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE AND MEDICATION WAS PRESCRIBED. REPORTEDLY "THE IOP WAS TREATED WITH TABLETS OVER TIME. BECAUSE THE PRESSURE KEPT INCREASING WITHOUT THERAPY THE SURGEON DECIDED TO FINALLY EXPLANT THE LENS". ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION PROVIDED: "PATIENT IS HAPPY". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840569&pc=MTA | 2023826-2021-04398 | 2021-11-03 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS ELEVATED IOP. REPORTEDLY A WASHOUT WAS PERFORMED, DIAMOX, IOPIDINE, AND IATANOPROST WERE PRESCRIBED. STEROID TREATMENT WAS STOPPED AND NSAID (BROMEGENAC) WAS STARTED WHICH LOWERED THE IOP INITIALLY BUT IT WAS RAISED AGAIN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS ELEVATED IOP. REPORTEDLY A WASHOUT WAS PERFORMED, DIAMOX, IOPIDINE, AND IATANOPROST WERE PRESCRIBED. STEROID TREATMENT WAS STOPPED AND NSAID (BROMEGENAC) WAS STARTED WHICH LOWERED THE IOP INITIALLY BUT IT WAS RAISED AGAIN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. CLAIM# (B)(4). Manufacturer Narrative: REPORTEDLY THE LENS WAS EXPLANTED (B)(6) 2021. H6-INVESTIGATION CODE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840585&pc=MTA | 2023826-2021-04399 | 2021-11-03 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS ELEVATED IOP. REPORTEDLY A WASHOUT WAS PERFORMED, DIAMOX, IOPIDINE, AND IATANOPROST WERE PRESCRIBED. STEROID TREATMENT WAS STOPPED AND NSAID (BROMEGENAC) WAS STARTED WHICH LOWERED THE IOP INITIALLY BUT IT WAS RAISED AGAIN. Manufacturer Narrative: B5: UPDATED INFORMATION: REPORTEDLY THE LENS WAS EXPLANTED ON (B)(6) 2021. CLAIM#: (B)(4). Manufacturer Narrative: H6- INVESTIGATION TYPE: 4110- LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16200099&pc=MTA | 2023826-2023-00234 | 2022-08-23 | Injury | STAAR SURGICAL COMPANY | 2023-01-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -16.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE OF 39 MMHG. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16464622&pc=MTA | 2023826-2023-00764 | 2022-08-23 | Injury | STAAR SURGICAL COMPANY | 2023-03-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -16.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS ELEVATED INTRAOCULAR PRESSURE 35 MMHG; WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. THE SURGEON NOTES "THE PATIENT DID NOT WANT TO UNDERGO A SECONDARY LENS IMPLANTATION PROCEDURE". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17672709&pc=MTA | 2023826-2023-03650 | 2023-01-17 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP). THE PATIENT WAS TREATED WITH INTRAOCULAR PRESSURE LOWERING DRUGS (CAPTOPROLOL EYE SOLUTION+BROMONIDINE EYE SOLUTION). ONE MONTH AFTER SURGERY, THE IOP WAS 25.0MMHG. AFTER ADDING ORAL INTRAOCULAR PRESSURE LOWERING DRUGS, THE IOP REMAINED HIGH BY ABOUT 25-30MMHG. AT LAST POST-OP VISIT ON (B)(6) 2023, THE IOP WAS 20MMHG AND PROBLEM WAS RESOLVED. Manufacturer Narrative: H6: HEALTH EFFECT IMPACT CODE: 4644 - CAPTOPROLOL EYE SOLUTION; BROMONIDINE EYE SOLUTION; ORAL INTRAOCULAR PRESSURE LOWERING DRUGS. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17680175&pc=MTA | 2023826-2023-03651 | 2023-01-17 | Injury | STAAR SURGICAL COMPANY | 2023-09-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP) AND WAS TREATED WITH INTRAOCULAR PRESSURE LOWERING DRUGS (CAPTOPROLOL EYE SOLUTION+BROMONIDINE EYE SOLUTION). ONE MONTH AFTER SURGERY, THE IOP WAS 25.0MMHG. AFTER ADDING ORAL INTRAOCULAR PRESSURE LOWERING DRUGS, THE IOP DROPPED TO A NORMAL LEVEL AND ROUTINE FOLLOW-UP EXAMINATIONS WERE CARRIED OUT. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. D6B: NA. H6: HEALTH EFFECT IMPACT CODE: 4644 - CAPTOPROLOL EYE SOLUTION; BROMONIDINE EYE SOLUTION; ORAL INTRAOCULAR PRESSURE LOWERING DRUGS. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17581257&pc=MTA | 2023826-2023-03593 | 2023-06-23 | Injury | STAAR SURGICAL COMPANY | 2023-08-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -8.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) .ELEVATED IOP WAS OBSERVED. IT WAS REPORTED THAT ON (B)(6) 2023 THE LENS WAS ROTATED VERTICALLY AND THE PROBLEM RESOLVED. SURGEON STATES " DOING GREAT, OFF ALL DROPS". CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17581260&pc=MTA | 2023826-2023-03592 | 2023-06-23 | Injury | STAAR SURGICAL COMPANY | 2023-08-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF _7.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) .ELEVATED IOP WAS OBSERVED. IT WAS REPORTED THAT ON (B)(6) 2023 THE LENS WAS ROTATED VERTICALLY AND THE PROBLEM RESOLVED. SURGEON STATES " LOOKS GREAT, BELOW BASELINE, OFF ALL DROPS". CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17777978&pc=MTA | 2023826-2023-04108 | 2023-07-15 | Injury | STAAR SURGICAL COMPANY | 2023-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE. THE REPORTER STATED " PREOPERATIVE BASAL IOP WAS HIGHER THAN NORMAL, AND NO FURTHER EXAMINATION WAS PERFORMED TO EXCLUDE GLAUCOMA. AFTER THE CRYSTAL IS REMOVED, THE INTRAOCULAR PRESSURE IS TEMPORARILY STABILIZED TO NORMAL VALUE AFTER DRUG TREATMENT. THE LENS WAS REMOVED AND TRANSFERRED TO THE GLAUCOMA DEPARTMENT FOR FURTHER CONSULTATION." Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18346570&pc=MTA | 2023826-2023-05625 | 2023-10-26 | Injury | STAAR SURGICAL COMPANY | 2023-12-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -7.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXPLANTED DUE TO ELEVATED IOP (30MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESED ON 26/OCT/2023) ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18437835&pc=MTA | 2023826-2024-00032 | 2023-12-07 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF -4.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, ELEVATED IOP WITHOUT PUPIL BLOCK WAS OBSERVED. ON (B)(6) 2023, THE LENS WAS REMOVED. REPORTEDLY, "LEFT EYE ANTERIOR CHAMBER STILL HAS DENCE AC CELLS AND INV PIGMENTS DESPITE FREQUENT TOMZAL STERIOD." CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13488086&pc=MTA | 2023826-2022-00427 | NA | Injury | STAAR SURGICAL COMPANY | 2022-02-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM MICL12.6; -8.5 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO A PATIENT'S EYE. THE PATIENT HAD INTRAOCULAR PRESSURE SPIKES AND THE SURGEON HAD TO PERFORMED AN EMERGENCY PERIPHERAL IRIDOTOMY. IT WAS ALSO NOTED THEY BELIEVE THE LENS IS TOO LARGE FOR THE PATIENT'S EYE AND CALCULATIONS WILL BE PERFORMED TO ORDER A NEW LENS SIZE BEFORE IMPLANTING INTO THE PATIENT. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15221421&pc=MTA | 2023826-2022-02751 | 2022-07-15 | Injury | STAAR SURGICAL COMPANY | 2022-08-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF -6.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 25-MAY-2022. ON 07-JUL-2022: ELEVATED IOP 50MMHG, PIGMENT DISPERSION, AND PRO-LONGED INFLAMMATION WAS RECORDED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: B5 - LENS WAS EXPLANTED ON (B)(6) 2022. PROBLEM WAS RESOLVED. REPORTEDLY "MINIMAL PIGMENTS, MORE QUITE EYE." CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14092478&pc=MTA | 2023826-2022-01113 | NA | Injury | STAAR SURGICAL COMPANY | 2022-04-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM MICL 12.6 IMPLANTABLE COLLAMER LENS OF -12.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO PRESSURE. "PRESSURE IS CONTROLLED UNDER MAXIMUM MEDS IN ONE EYE AND THE OTHER EYE STILL HAS PRESSURE OF 34 DEGREES UNDER MAXIMUM MEDS." NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14092585&pc=MTA | 2023826-2022-01112 | NA | Injury | STAAR SURGICAL COMPANY | 2022-04-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM MICL 12.6 IMPLANTABLE COLLAMER LENS OF -9.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO PRESSURE. "PRESSURE IS CONTROLLED UNDER MAXIMUM MEDS IN ONE EYE AND THE OTHER EYE STILL HAS PRESSURE OF 34 DEGREES UNDER MAXIMUM MEDS." NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6505274&pc=MTA | 2023826-2017-00666 | 2017-03-08 | Injury | STAAR SURGICAL COMPANY | 2017-04-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE PATIENT BEGAN TO EXPERIENCE ELEVATED IOP. THE SURGEON STATED THAT HE BELIEVED THE HIGH IOP WAS CAUSED BY THE PATIENT'S STEROID (DEXAMETHASONE) TREATMENT. TO DATE, THE LENS REMAINS IMPLANTED. SURGEON STATED THAT "AFTER MEDICATION" THE IOP WENT DOWN TO NORMAL VALUE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6641024&pc=MTA | 2023826-2017-00975 | 2017-05-17 | Injury | STAAR SURGICAL COMPANY | 2017-06-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -14 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED ELEVATED IOP. THE REPORTER ALSO MENTIONED THAT THE SURGEON CHANGED THE MEDICATIONS EYEDROPS WITHOUT CORTISONE TO FIND OUT IF THE HIGH IOP IS THE OUTCOME OF GIVING CORTISONE EYEDROPS. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. REFERENCE MDR NUMBER 2023826-2017-00701 FOR THE SAME PATIENT'S EYE. Manufacturer Narrative: SINCE THE PHYSICIAN MODIFIED THE MEDICATION, THE IOP IS REPORTED AS NORMAL. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CONCLUSION: OFF LABEL USE (BELOW 21 YEARS OF AGE AT THE TIME OF IMPLANT, LENS IMPLANT IN A PSEUDO PHAKIC EYE). (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6632436&pc=MTA | 2023826-2017-00940 | NA | Injury | STAAR SURGICAL COMPANY | 2017-06-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER STATED THAT A 13.2MM MICL 13.2 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, WAS EXPLANTED DUE TO PATIENTS INTRAOCULAR PRESSURE BEING HIGH. THE REPORTER INDICATED THAT THE EVENT CAUSE COULD BE DUE TO THE LENS BEING BIG. ADDITIONAL INFORMATION WAS REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID , IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6952443&pc=MTA | 2023826-2017-01570 | NA | Injury | STAAR SURGICAL COMPANY | 2017-10-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGERY IMPLANTED A MICL12.6, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THREE WEEKS AFTER THE IMPLANTATION, THE IOP WAS NOTED TO BE 29. THE IOP WAS SLIGHTLY ASYMMETRIC A WEEK PRIOR BY 5 BUT WITHIN THE NORMAL LIMITS. GONIOSCOPY WAS COMPLETELY OPEN. THE VAULT WAS 400. THE PI'S WERE PATENT AND FUNCTIONING. DILATED EXAM WAS NORMAL AND NO DROP IN IOP FOLLOWING DILATION. THE PATIENT WAS ON BID PRED FORTE AND THIS WAS THOUGHT TO BE THE CAUSE ACCORDING TO THE SURGEON. AFTER TWO DAYS OF DISCONTINUING THE PRED FORTE, THE PATIENT'S IOP WAS 17. THE PROBLEM WAS RESOLVED. UCVA IS 20/20, IOP NORMALIZED, AND PATIENT IS HAPPY. Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5981810&pc=MTA | 2023826-2016-01321 | 2016-04-18 | Injury | STAAR SURGICAL COMPANY | 2016-09-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VICMO 13.2 IMPLANTABLE COLLAMER LENS - -13.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13361443&pc=MTA | 2023826-2022-00301 | 2020-02-12 | Injury | STAAR SURGICAL COMPANY | 2022-01-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1 VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -13.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM.Event Description: THE REPORTER INDICATED THAT A 12.1 VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -13.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10634642&pc=MTA | 2023826-2020-02293 | 2020-08-07 | Injury | STAAR SURGICAL COMPANY | 2020-10-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED LATER THAT SAME DAY DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DEFORMED WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13295627&pc=MTA | 2023826-2022-00213 | 2020-12-22 | Injury | STAAR SURGICAL COMPANY | 2022-01-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICMO 13.7 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON THE SAME DAY DIFFERENT SURGERY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT A 13.7MM VICMO 13.7 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON THE SAME DAY DIFFERENT SURGERY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13941313&pc=MTA | 2023826-2022-00951 | 2021-01-12 | Injury | STAAR SURGICAL COMPANY | 2022-03-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL 13.2 OF -8.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS DEVICE. "HIGH VAULT DESPITE APPROPRIATELY SIZED ICL BASED ON WTW. Manufacturer Narrative: CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11822759&pc=MTA | 2023826-2021-01436 | 2021-02-26 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM ,VICM5_13.7, -13.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. A SHORTER LENGTH LENS WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED TO BE OTHER: " WRONG LENGTH OF LENS. YES BECAUSE THE INITAL LENS WAS TOO LONG AND THE VAULT TOO HIGH". Manufacturer Narrative: (B)(6). PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119262&pc=MTA | 2023826-2021-02299 | 2021-03-14 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -8.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXPLANTED AND THEN REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12233847&pc=MTA | 2023826-2021-02882 | 2021-04-20 | Injury | STAAR SURGICAL COMPANY | 2021-07-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP (23MMHG) WITH OUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESSED ON (B)(6) 2021) AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR, LENS INCREASED IOP, MADE AC SHALLOWER. Manufacturer Narrative: B5: ADDITIONAL INFORMATION: REPORTEDLY THE LENS WAS EXCHANGED ON (B)(6) 2021. "SURGERY WENT WELL, PATIENT HAPPY!" VISION IS 20/25, IOP WITHIN NORMAL RANGE. H6: INVESTIGATION FINDING 114 WAS REPORTED PREVIOUSLY IN ERROR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16917178&pc=MTA | 2023826-2023-01761 | 2021-04-23 | Injury | STAAR SURGICAL COMPANY | 2023-05-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS -12.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11939026&pc=MTA | 2023826-2021-01671 | 2021-05-11 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO ELEVATED IOP (33MMHG, ASSESSED ON (B)(6) 2021) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY, PT SYMPTOMS: FOREIGN BODY SENSATION AND ITCHINESS. OBJECTIVE SLE: CORNEA: 2+ DMF. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN A SPECIMEN. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155775&pc=MTA | 2023826-2021-02533 | 2021-06-16 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING WITH ELEVATED INTRAOCULAR PRESSURE (IOP). THE SURGEON PERFORMED ADDITIONAL LASER SURGERY TO THE IRIDOTOMY (ENLARGED) AND THAT RESOLVED THE ISSUE. THE PATIENT IS FINE AND THERE WAS NO NEED TO EXPLANT THE LENS, THE LENS REMAINED IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK.. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: NA. ENLARGED IRIDOTOMY BY LASER. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12308409&pc=MTA | 2023826-2021-03077 | 2021-07-19 | Injury | STAAR SURGICAL COMPANY | 2021-08-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5 13.7, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP (50MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESED ON (B)(6) 2021) AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, ELEVATED IOP AS REDUCED TO 15MMHG BY TREATING WITH DIAMOX. Manufacturer Narrative: PT INFO: UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12399422&pc=MTA | 2023826-2021-03396 | 2021-07-30 | Injury | STAAR SURGICAL COMPANY | 2021-08-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS REMAINED IMPLANTED. THE PATIENT WILL BE MONITORED STRICTLY AND THE IOP HAS DECREASED SLIGHTLY DURING THE LAST WEEKS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13309724&pc=MTA | 2023826-2022-00221 | 2021-08-31 | Injury | STAAR SURGICAL COMPANY | 2022-01-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1 VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM.Event Description: THE REPORTER INDICATED THAT A 12.1 VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13309095&pc=MTA | 2023826-2022-00219 | 2021-11-01 | Injury | STAAR SURGICAL COMPANY | 2022-01-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -8.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: D6A- IMPLANT DATE CORRECTED TO (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13309644&pc=MTA | 2023826-2022-00220 | 2021-11-01 | Injury | STAAR SURGICAL COMPANY | 2022-01-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM.Event Description: THE REPORTER INDICATED THAT A 12.6 VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13267274&pc=MTA | 2023826-2022-00120 | 2021-11-04 | Injury | STAAR SURGICAL COMPANY | 2022-01-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -11.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13267305&pc=MTA | 2023826-2022-00121 | 2021-11-04 | Injury | STAAR SURGICAL COMPANY | 2022-01-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -11.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256669&pc=MTA | 2023826-2022-00118 | 2021-11-06 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 VICMO 12.6 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: RACE-ETHNICITY:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15501957&pc=MTA | 2023826-2022-03342 | 2022-08-18 | Injury | STAAR SURGICAL COMPANY | 2022-09-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -15.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE. THIS EXCHANGE RESOLVED THE PROBLEM. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED "BORDERLINE IOP IMPROVING- STILL ON GTTS." Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15503201&pc=MTA | 2023826-2022-03373 | 2022-08-18 | Injury | STAAR SURGICAL COMPANY | 2022-09-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -16.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE. THIS EXCHANGE RESOLVED THE PROBLEM. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED "BORDERLINE IOP IMPROVING- STILL ON GTTS." Manufacturer Narrative: CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16455859&pc=MTA | 2023826-2023-00761 | 2022-12-09 | Injury | STAAR SURGICAL COMPANY | 2023-02-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -14.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE 28 MMHG; WITHOUT PUPIL BLOCK. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: B5: ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR THE SAME MODEL/ SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE "ALL FINE!" AND "PATIENT IS HAPPY". CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16554736&pc=MTA | 2023826-2023-00869 | 2022-12-09 | Injury | STAAR SURGICAL COMPANY | 2023-03-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENS ON (B)(6) 2023. THE PROBLEM WAS RESOLVED. THE EYE IS FINE AND THE PATIENT IS HAPPY. CLAIM # (B)(4). Manufacturer Narrative: PATIENT WEIGHT: UNK. PATIENT ETHNICITY: UNK. PATIENT RACE: UNK. EXPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16433266&pc=MTA | 2023826-2023-00524 | 2022-12-26 | Injury | STAAR SURGICAL COMPANY | 2023-02-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IN A SUBSEQUENT SURGERY LATER THAT DAY, THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17741438&pc=MTA | 2023826-2023-03980 | 2023-07-07 | Injury | STAAR SURGICAL COMPANY | 2023-09-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -01.25 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC BROKEN WITH FIBRE ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17754907&pc=MTA | 2023826-2023-03981 | 2023-07-07 | Injury | STAAR SURGICAL COMPANY | 2023-09-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -01.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND OPTIC TORN/DEFORMED AND HAPTIC BROKEN/STRETCHED OUT WITH FIBRE ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17831137&pc=MTA | 2023826-2023-04215 | 2023-07-23 | Injury | STAAR SURGICAL COMPANY | 2023-09-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -04.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND HIGH IOP. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DOCTOR CHOSE A DIFFERENT ICL BY ERROR FROM HIS STOCK, SO USER HAD TO REPLACE IT SECOND DAY POST-OP. Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK RACE: UNK CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18360955&pc=MTA | 2023826-2023-05558 | 2023-11-23 | Injury | STAAR SURGICAL COMPANY | 2023-12-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611940&pc=MTA | 2023826-2019-00825 | NA | Injury | STAAR SURGICAL COMPANY | 2019-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, MICL 13.7, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. HIGH PRESSURE AND EXCESSIVE VAULT WAS OBSERVED, THE LENS WAS REMOVED ON (B)(6) 2018. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12281694&pc=MTA | 2023826-2021-03061 | NA | Injury | STAAR SURGICAL COMPANY | 2021-08-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL 13.2, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP AND EXCESSIVE VAULT ON (B)(6) 2021. REPORTEDLY, PATIENT HAS DONE BETTER WITH EXCHANGE AND HAS BETTER IOP. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12281698&pc=MTA | 2023826-2021-03060 | NA | Injury | STAAR SURGICAL COMPANY | 2021-08-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL 13.2, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP AND EXCESSIVE VAULT ON (B)(6) 2021. REPORTEDLY, PATIENT HAS DONE BETTER WITH EXCHANGE AND HAS BETTER IOP. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: SEX- UNK, WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, DATE OF EVENT- UNK. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16694223&pc=MTA | 2023826-2023-01293 | NA | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -9.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 . THE SURGEON REPORTS THERE WAS EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE; AND MEDICATION WAS USED. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5241755&pc=MTA | 2023826-2015-01526 | 2015-10-14 | Injury | STAAR SURGICAL COMPANY | 2015-11-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -12.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE WAS NOTED ON (B)(6) 2015. LENS REMAINS IMPLANTED. Manufacturer Narrative: THE SURGEON DID A REPOSITIONING OF THE NASAL HAPTIC WHICH WAS NOT RIGHTLY LOCATED IN THE SULCUS. THE VAULTING WAS RECORDED AS 400UM. AFTER REPOSITIONING, THE VAULTING INCREASED AGAIN, SO THE SURGEON DECIDED TO EXCHANGE THE LENS FOR A SHORTER LENS, SAME MODEL AND DIOPTER ON (B)(6) 2016. THE VAULTING WAS STABLE AT 400UM AND THE PATIENT WAS FINE. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. CONCLUSION: AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.89MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -12.50. DEVICE EVALUATED BY MANUFACTURER? NO, DEVICE REMAINS IMPLANTED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). LENS REMAINS IMPLANTED. | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13641098&pc=MTA | 2023826-2022-00644 | 2022-01-27 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. AT 4 HOUR POST-OP, THE INTRAOCULAR PRESSURE WAS 36MMHG BUT THE PRESSURE CAME DOWN TO 14MMHG. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. AT A POST-OP VISIT THE PATIENTS REFRACTION HAD IMPROVED TO 6/6 UNCORRECTED, +0.25 SCRIPT. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H1 - CORRECTED TO MALFUNCTION IN INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15117846&pc=MTA | 2023826-2022-02425 | 2022-06-30 | Injury | STAAR SURGICAL COMPANY | 2022-07-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -03.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THERE WAS ELEVATED INTRAOCULAR PRESSURE (IOP). THE IOP WAS RAISED ON DAY OF SURGERY. DIAMOX WAS USED TO BRING IOP DOWN BUT IOP WAS UP TO 50MMHG ON (B)(6) 2022, SO PATIENT WAS BROUGHT BACK TO THE THEATRE FOR AC WASHOUT. IOP 12 AFTER AND VISION IS 6/12 AFTER HAVING BEEN HM (HAND MOVEMENTS). STATUS OF EYE - PUPIL DILATED STILL, POTENTIAL SPHINCTER DAMAGE FROM HIGH IOP. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: POST-OP, THE PATIENT WAS 6/12, WITH FIXED DILATED PUPIL AND DAMAGED SPHINCTER. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10823655&pc=MTA | 2023826-2020-02586 | 2020-10-17 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS WAS OBSERVED AND ELEVATED IOP (37MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS ASSESSED ON (B)(6) 2020. USE OF STEROIDS AND IOP LOWERING MEDICATIONS PRESCRIBED. POST-OP MEDICATIONS STOPPED, AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS "STEROID RESPONSE". PER THE REPORTER ON (B)(6) 2020, "THE CASE IS RESOLVED; THE STEROIDS WERE STOPPED AND IOP LOWERING DROPS STOPPED AND THE EYE IS FINE. THE STEROIDS WERE STANDARD POST OP CARE." Manufacturer Narrative: B1: ADVERSE EVENT. B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. H1: SERIOUS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10823662&pc=MTA | 2023826-2020-02585 | 2020-10-17 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS WAS OBSERVED AND ELEVATED IOP (37MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS ASSESSED ON (B)(6) 2020 AND USE OF IOP LOWERING MEDICATIONS PRESCRIBED. POST-OP MEDICATIONS STOPPED, AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS "STEROID RESPONSE". PER THE REPORTER ON 15/OCT/ 2020, "THE CASE IS RESOLVED - THE STEROIDS WERE STOPPED AND IOP LOWERING DROPS STOPPED AND THE EYE IS FINE." Manufacturer Narrative: B1: ADVERSE EVENT. B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. H1: SERIOUS INJURY. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNKNOWN. IMPLANT DATE: N/A. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16773859&pc=MTA | 2023826-2023-01423 | 2022-02-23 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM512.1 IMPLANTABLE COLLAMER LENS OF -9.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023, ELEVATED IOP WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "BETTER CONTROL OF OCULAR INFLMMATION AND IOP," AND PATIENT CURRENT MEDICATION: AZARGA, ALPHAGEN, PRED FORTE, VIGAMOX. PATIENT DIAGNOSED ON BE MYOPIC CHOROIDAL NEOVASCULARIZATION WHICH REQUIRED INTRAVITREAL ANTI-VEGF INJECTION." Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 12.1MM VICM512.1 IMPLANTABLE COLLAMER LENS OF -9.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP AND PIGMENT DISPERSION WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "BETTER CONTROL OF OCULAR INFLAMMATION AND IOP," AND PATIENT CURRENT MEDICATION: AZARGA, ALPHAGEN, PRED FORTE, VIGAMOX. PATIENT DIAGNOSED ON BE MYOPIC CHOROIDAL NEOVASCULARIZATION WHICH REQUIRED INTRAVITREAL ANTI-VEGF INJECTION." D9: RETURN DATE: 31-MAY-2023. H6: 4627 SHOULD BE ADDED FOR CORRECTION. CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16773860&pc=MTA | 2023826-2023-01424 | 2022-02-23 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM512.1 IMPLANTABLE COLLAMER LENS OF -9.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. PIGMENT DISPERSION WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "BETTER CONTROL OF OCULAR INFLMMATION AND IOP," AND PATIENT'S CURRENT MEDICATION: AZARGA, ALPHAGEN, PRED FORTE, VIGAMOX. PATIENT DIAGNOSED ON BE MYOPIC CHOROIDAL NEOVASCULARIZATION WHICH REQUIRED INTRAVITREAL ANTI-VEGF INJECTION." Manufacturer Narrative: CORRECTED DATA: B5: SHOULD BE CORRECTED TO : THE REPORTER INDICATED THAT A 12.1MM VICM512.1 IMPLANTABLE COLLAMER LENS OF -9.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO PIGMENT DISPERSION. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "BETTER CONTROL OF OCULAR INFLAMMATION AND IOP," AND PATIENT'S CURRENT MEDICATION: AZARGA, ALPHAGEN, PRED FORTE, VIGAMOX. PATIENT DIAGNOSED ON BE MYOPIC CHOROIDAL NEOVASCULARIZATION WHICH REQUIRED INTRAVITREAL ANTI-VEGF INJECTION." D4: SERIAL NUMBER: SHOULD BE CORRECTED TO (B)(6). D4: UDI: SHOULD BE CORRECTED TO: (B)(4). D6B: (B)(6) 2023 SHOULD BE ADDED FOR CORRECTION. D9: RETURN DATE: 31-MAY-2023 SHOULD BE ADDED FOR CORRECTION. H6: 4627 SHOULD BE ADDED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14178946&pc=MTA | 2023826-2022-01224 | 2022-03-19 | Injury | STAAR SURGICAL COMPANY | 2022-04-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ELEVATED IOP WAS OBSERVED. POST-OP OBSERVATION AT 1DAY IOP 19; 1WEEK IOP 27; 1 MONTH IOP 46 AND THE DOCTOR GAVE A STEROID EYE DROP; 3 MONTH IOP 33. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(6). CLAIM #: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B1-ADVERSE EVENT B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) B5-THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ELEVATED IOP WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS OF -10.50 DIOPTER AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. D6B- (B)(6) 2022. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: CORRECTED TO " THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(3) 2021. PATIENT EXPERIENCED ELEVATED IOP AND MEDICATION WAS PRESCRIBED. ON (B)(3) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED UNKNOWN. H6: HEALTH IMPACT: CLINICAL CODE: (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12361974&pc=MTA | 2023826-2021-03282 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-08-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER INTO THE PATIENT'S EYE. THE SURGEON REPORTS HIGH PRESSURE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15502149&pc=MTA | 2023826-2022-03357 | 2022-08-04 | Injury | STAAR SURGICAL COMPANY | 2022-09-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5 DIOPTER, WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18145297&pc=MTA | 2023826-2023-05047 | 2023-09-09 | Injury | STAAR SURGICAL COMPANY | 2023-11-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18152256&pc=MTA | 2023826-2023-05048 | 2023-09-09 | Injury | STAAR SURGICAL COMPANY | 2023-11-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON 02-SEP-2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8699719&pc=MTA | 2023826-2019-01050 | 2019-05-15 | Injury | STAAR SURGICAL COMPANY | 2019-06-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED IOP (21MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, EXCESSIVE VAULT WAS OBSERVED. PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS DONE, THE LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS DEVICE, "ICL TOO LARGE". Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# : (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9186556&pc=MTA | 2023826-2019-01905 | 2019-09-11 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM MICL13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. THE SURGEON REPORTS VAULTING AT 300%, ELEVATED IOP, AND HIGH RISK OF ANGLE CLOSURE. LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION HAS BEEN RECEIVED. Manufacturer Narrative: TYPE OF REPORTABLE EVENT: MALFUNCTION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11813620&pc=MTA | 2023826-2021-01419 | 2021-01-27 | Injury | STAAR SURGICAL COMPANY | 2021-05-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11815697&pc=MTA | 2023826-2021-01420 | 2021-01-27 | Injury | STAAR SURGICAL COMPANY | 2021-05-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13252941&pc=MTA | 2023826-2022-00114 | 2021-09-28 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 VICMO 12.6 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13065390&pc=MTA | 2023826-2021-04742 | 2021-10-17 | Injury | STAAR SURGICAL COMPANY | 2021-12-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -14.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13253002&pc=MTA | 2023826-2022-00115 | 2021-10-22 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 VICMO 12.6 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13253038&pc=MTA | 2023826-2022-00116 | 2021-10-22 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 VICMO 12.6 IMPLANTABLE COLLAMER LENS OF -12.5 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13184529&pc=MTA | 2023826-2022-00103 | 2021-11-24 | Injury | STAAR SURGICAL COMPANY | 2022-01-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 VICMO 12.6 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13172928&pc=MTA | 2023826-2022-00098 | 2021-11-26 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -5.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON THE SAME DAY SEPARATE SURGERY THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT WITH ELEVATED IOP. A REPLACEMENT LENS WAS LATER IMPLANTED OF SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13172930&pc=MTA | 2023826-2022-00097 | 2021-11-26 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -4.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON THE SAME DAY SEPARATE SURGERY THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT WITH ELEVATED IOP. A REPLACEMENT LENS WAS LATER IMPLANTED OF SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16728674&pc=MTA | 2023826-2023-01333 | 2022-12-24 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED, THAT A 12.6MM, VICMO12.6, IMPLANTABLE COLLAMER LENS OF A -14.0 DIOPTER, WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS INTRAOPERATIVELY IMPANTED AND REMOVED. DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK, AND ANGLE CLOSURE, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17704133&pc=MTA | 2023826-2023-03950 | 1998-11-05 | Injury | STAAR SURGICAL COMPANY | 2023-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT PRESENTED WITH REDNESS IN THE EYE, PIGMENT DISPERSION AND ELEVATED IOP IN THE EARLY POST-OP PERIOD. IT WAS REPORTED THAT THE PROBLEM RESOLVED WITH MEDICAL MANAGEMENT. THE REPORTER STATES THAT THE DISPERSED PIGMENTS DECREASED WITH NORMALIZATION OF IOP AFTER CONSERVATIVE MANAGEMENT WITH NO BLURRING OF VISION. HOWEVER, THE PUPIL REMAINED SLIGHTLY DILATED AND IRREGULAR. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. THE REPORTER ALSO STATES THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: REDNESS, PIGMENT DISPERSION. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE - 4581 - PUPIL ANOMALY. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B3: 05-NOV-1998 SHOULD BE REMOVED FROM DATE OF EVENT. CLAIM#: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138789&pc=MTA | 2023826-2020-01186 | 2019-07-11 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -08.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. PUPIL BLOCK, ANGLE CLOSURE WITH ELEVATED IOP (38MMHG) WAS ASSESSED ON (B)(6) 2019. UNREACTIVE (FIXED) PUPIL, MID-DILATED IRIS WITH PERIPHERAL IRIDO-CORNEAL TOUCH AND FLAT AC WAS OBSERVED. AN ADDITION OF NEW PI AND YAG WAS PERFORMED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS "PROBABLY VISCO ELASTIC TRAP." REPORTEDLY, "HAD TO DO SECOND PI YAG. VERY DIFFICULT BUT MANAGED TO GET PUPIL BLOCK REVERSED AND SETTLED DOWN." Manufacturer Narrative: CORRECTED DATA: B4 - CORRECTED DATE TO (B)(6) 2020. G4 - DATE SHOULD BE CORRECTED TO (B)(6) 2020 IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO [OR#] SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10706946&pc=MTA | 2023826-2020-02421 | 2020-09-07 | Injury | STAAR SURGICAL COMPANY | 2020-10-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING ELEVATED INTRAOCULAR PRESSURE (IOP), UNREACTIVE (FIXED) PUPIL, ANISOCORIA AND IRIS ATROPHY. THE LENS WAS REPOSITIONED (ROTATED) BUT THE PROBLEM WAS NOT RESOLVED. MEDICATION WAS PRESCRIBED BUT WITH NO EFFECT ON PUPIL CONSTRICTION. THE LENS REMAINED IMPLANTED. AT THE LAST POST-OP VISIT ON (B)(6) 2020, THE PATIENT IS RECOVERING SLOWLY TO NORMAL PUPIL APPEARANCE. THIS CASE SEEMS TO BE AN URRETS-ZAVALIA SYNDROME. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507466&pc=MTA | 2023826-2021-00660 | 2021-02-12 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -3.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXPLANTED DUE TO ELEVATED IOP AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA) THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12277915&pc=MTA | 2023826-2021-03074 | 2021-07-09 | Injury | STAAR SURGICAL COMPANY | 2021-08-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2021 A 12.1MM MICL12.1; -11.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) AS A REPLACEMENT LENS. THE SURGEON INDICATED THE PATIENT EXPERIENCED SHALLOW AC, ELEVATED INTRAOCULAR PRESSURE AND WAS TREATED WITH MEDICATIONS LIDOCAINE, MOISTAT AND MIOCHOL ALONG WITH TOPICAL PILOCARPINE. PATIENT'S ANGLES WERE STILL NARROW AND WENT ON ALPHAGAN. IT WAS ALSO NOTED THAT "SHE FEELS GOOD WITH SOME IMPROVEMENT IN HER VISION". POSTOPERATIVE REPORT INDICATED PATIENT'S IOP DECREASED TO AN ACCEPTABLE LEVEL. LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: H10 - INITIAL CLAIM IS NOT APPLICABLE DUE TO WAS FOUND TO BE A DUPLICATE OF MFR# 2023826-2021-03057. REFERENCE MFR# 2023826-2021-03057 FOR CORRECT CLAIM DETAILS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12278129&pc=MTA | 2023826-2021-03075 | 2021-07-09 | Injury | STAAR SURGICAL COMPANY | 2021-08-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2021 A 12.1MM MICL12.1; -11.5 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AS A REPLACEMENT LENS. THE SURGEON INDICATED THE PATIENT EXPERIENCED SHALLOW AC, ELEVATED INTRAOCULAR PRESSURE AND WAS TREATED WITH MEDICATIONS LIDOCAINE, MOISTAT AND MIOCHOL ALONG WITH TOPICAL PILOCARPINE. PATIENT'S ANGLES WERE STILL NARROW AND WENT ON ALPHAGAN. IT WAS ALSO NOTED THAT "SHE FEELS GOOD WITH SOME IMPROVEMENT IN HER VISION". POSTOPERATIVE REPORT INDICATED PATIENT'S IOP DECREASED TO AN ACCEPTABLE LEVEL. LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: H10: INITIAL CLAIM IS NOT APPLICABLE DUE TO CLAIM WAS FOUND TO BE A DUPLICATE OF MFR# 2023826-2021-03058. SEE MENTIONED MFR# FOR CORRECT CLAIM DETAILS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14215323&pc=MTA | 2023826-2022-01346 | 2022-03-04 | Injury | STAAR SURGICAL COMPANY | 2022-04-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6; -11.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: (B)(4). TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14215325&pc=MTA | 2023826-2022-01347 | 2022-03-04 | Injury | STAAR SURGICAL COMPANY | 2022-04-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6; -11.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: (B)(4). TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15506324&pc=MTA | 2023826-2022-03358 | 2022-08-05 | Injury | STAAR SURGICAL COMPANY | 2022-09-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF -4.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, PUPIL BLOCK WITH ELEVATED IOP, ANGLE CLOSURE WITH ELEVATED IOP, PIGMENT DISPERSION, AND UNREACTIVE (FIXED) PUPIL. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18509050&pc=MTA | 2023826-2024-00217 | 2023-12-20 | Injury | STAAR SURGICAL COMPANY | 2024-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS OF -8.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ANGLE CLOSURE WITH ELEVATED IOP WAS OBSERVED DURING POST-OP EXAMINATION. THE REPORTER STATED THAT THE SURGEON ROTATED THE LENS TO VERTICAL POSITION ON (B)(6) 2023 SINCE IT WAS FOUND THE VAULT WAS HIGH DURING RE-EXAMINATION ON (B)(6) 2024, BUT THE PROBLEM DID NOT RESOLVE. THE SURGEON FOUND THE CENTRAL HOLE ON THE LENS WAS NOT PERFORATED DURING THE RE-EXAMINATION ON (B)(6) 2023. PER REPORTER THE LENS WAS PERFORATED BY THE DOCTOR ON (B)(6) 2023 AND THE PATIENTS VISION, IOP RETURNED TO NORMAL. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: ANGLE CLOSURE. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15571958&pc=MTA | 2023826-2022-03586 | NA | Injury | STAAR SURGICAL COMPANY | 2022-10-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -15.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS TO ADDRESS EXCESSIVE VAULT AND ELEVATED INTRAOCULARE PRESSURE. IT WAS REPORTED THAT IT RESOLVED PROBLEM BUT FOR STATUS OF THE EYE (SIGN/SYMPTOMS) THE REPORTER STATED "MILD PIGMENT DISPERSION AND BORDERLINE IOP IMPROVING-STILL ON GTTS". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15577782&pc=MTA | 2023826-2022-03587 | NA | Injury | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -16.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022 AS A REPLACEMENT LENS TO ADDRESS EXCESSIVE VAULT AND ELEVATED INTRAOCULARE PRESSURE. IT WAS REPORTED THAT IT RESOLVED THE PROBLEM BUT FOR STATUS OF THE EYE (SIGN/SYMPTOMS) THE REPORTER STATED "MILD PIGMENT DISPERSION AND BORDERLINE IOP IMPROVING-STILL ON GTTS". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18561095&pc=MTA | 2023826-2024-00367 | NA | Injury | STAAR SURGICAL COMPANY | 2024-01-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). THE PATIENT EXPERIENCED NARROW ANTERIOR CHAMBER, ELEVATED IOP WITH PIGMENT DISPERSION. IT WAS REPORTED THAT ON (B)(6) 2023 THE LENS WAS EXCHANGED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: D6A:UNK. H6- HEALTH EFFECT- CLINICAL CODE 4581:NARROW ANTERIOR CHAMBER, PIGMENT DISPERSION H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15713915&pc=MTA | 2023826-2022-03878 | 2022-08-26 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Patient Device Interaction Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS REPORTED AS HAVING ROTATED AND THE PATIENTS PRESSURE HAD SPIKED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -15.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE PATIENT WAS "BURPED" SEVERAL DAYS IN A ROW AND THE IOP WAS NORMAL. THE LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS REMOVED. THE EYE IS FOGGY BUT IS SLIGHTLY IMPROVED. THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR AND THE DEVICE DID NOT FAIL TO PERFORM. THE PATIENT IS SEEING A GLAUCOMA SPECIALIST WHICH WAS NO DISCLOSED TO THE FACILITY. THE PATIENT SUFFERED A CILIARY BODY INJURY WHICH CAUSED THE VISION TO CHANGE AND HER IOP TO BECOME UNCONTROLLABLE. A IOL ROTATION WAS PERFORMED BUT DUE TO THE ROTATION HER PRESSURES BEGAN TO SPIKE. IT WAS DECIDED TO REMOVE THE LENS FOR PATIENT SAFETY. THE PATIENT IS CURRENTLY SEEKING GLAUCOMA TREATMENT AND THERE ARE NO PLANS TO REPLACE THE LENS. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - "BURPING THE PRESSURE." CORRECTED DATA: D2: COMMON DEVICE NAME: PHAKIC INTRAOCULAR LENS, PRODUCT CODE: MTA. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11168648&pc=MTA | 2023826-2021-00133 | 2019-04-04 | Injury | STAAR SURGICAL COMPANY | 2021-01-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -7.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10535260&pc=MTA | 2023826-2020-01976 | 2020-08-06 | Injury | STAAR SURGICAL COMPANY | 2020-09-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, IN THE PATIENTS EYE. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND ANGLE CLOSURE. THE LENS WILL BE EXPLANTED AND EXCHANGED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS IMPLANTED IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. DATE OF EVENT: (B)(6) 2020. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT CLINICAL CODE: 4581- ANGLE CLOSURE. B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS; -05.50 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON 06/AUG/2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING, ELEVATED IOP WITH ANGLE CLOSURE. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE "SUGGESTED LENS WAS TOO LARGE". CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10963437&pc=MTA | 2023826-2020-02978 | 2020-08-07 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT, ELEVATED IOP AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2020 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11722535&pc=MTA | 2023826-2021-01208 | 2021-03-10 | Injury | STAAR SURGICAL COMPANY | 2021-04-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP AND UNREACTIVE PUPIL (PUPIL IS REACTIVE TO LIGHT BUT VERY DILATED). FOR STATUS OF THE EYE (SIGNS/SYMPTOMS): "AWAITING FOLLOW UP" WAS REPORTED." DEVICE" WAS REPORTED TO BE THE CAUSE OF THIS EVENT. FOR CAUSE DETAILS THE REPORTER STATED " SURGEON INTERPRETATION IS THAT THE ANGLE IS OPEN BUT THAT THE CONFIGURATION AND VAULT OF ICL IS SUCH THAT THE PUPIL CANNOT CONSTICT OVER THE LIP OF THE OPTIC RESULTING IN THE PUPIL BEING MECHANICALLY PREVENTED FROM RETURNING TO NORMAL SIZE." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- IT WAS LATER REPORTED THAT IT WAS AN IRIS CONFIGURATION ISSUE AND NOT AN ISSUE WITH VAULT OR THE ICL. CLAIM # (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721772&pc=MTA | 2023826-2021-01127 | 2021-03-16 | Injury | STAAR SURGICAL COMPANY | 2021-04-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P990013 | NA | 1 | Event Description: IT WAS REPORTED THAT A 13.2MM, MICL13.2, -8.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT (1.42MM VAULT), MILD DILATED PUPIL AND ELEVATED IOP (REQUIRING DROPS). THIS EXCHANGE RESOLVED THE PROBLEM. DEVICE, PATIENT- RELATED FACTOR, UNKNOWN, ICL TOO LARGE WAS REPORTED FOR CAUSE OF THIS EVENT. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED " IMPROVED-COMFORTABLE". Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. WEIGHT: UNK. RACE: UNK. (B)(4). DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11925869&pc=MTA | 2023826-2021-01746 | 2021-04-28 | Injury | STAAR SURGICAL COMPANY | 2021-06-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM ,VICMO13.2, -17.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. HIGH VAULT AND ANGLE CLOSURE WITH ELEVATED IOP WAS OBSERVED. IT WAS REPORTED THAT MEDICAL INTERVENTION (ANTIGLAUCOMATEOUS TREATMENT) WAS PERFORMED WHILE WAITING FOR SURGERY. IT WAS RECENTLY REPORTED THAT LENS EXCHANGE WAS PERFORMED AND THAT THE PATIENTS CURRENT CONDITION IS VERY WELL. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS " OTHER: WTW VALUE OF THE PATIENT MAY CAUSE AN ERROR IN CALCULATION OF LENS DIAMETER, THE LENS MUST BE EXCHANGED BY A 12.6MM LENS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12308408&pc=MTA | 2023826-2021-03089 | 2021-04-30 | Injury | STAAR SURGICAL COMPANY | 2021-08-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -15.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ELEVATED IOP (50MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VICMO 12.6; -15.00 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. CLAIM # (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: ONE SIMILAR COMPLAINT EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (CLAIM# (B)(4) EXCESSIVE VAULT; LENS EXPLANTED; PROBLEM RESOLVED). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178435&pc=MTA | 2023826-2021-02702 | 2021-06-11 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS -16.0 DIOPTER INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ANGLE CLOSURE, AND ELEVATED IOP. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "DEVICE BEING TOO LARGE CAUSED ANTERIOR VAULT, RESULTING IN NARROWING ANTERIOR CHAMBER AND INCREASED IOP." THE PROBLEM IS RESOLVED WITH THE EXPLANT. Manufacturer Narrative: B5 - PER EMAIL ON 15/OCT/2021, THE REPORTER STATED THAT "REPLACEMENT LENS WERE IMPLANTED ON (B)(6) 2021. RIGHT EYE: (B)(6) & LEFT EYE: (B)(6). ALL PROBLEMS HAVE BEEN RESOLVED AT THIS TIME." CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12271650&pc=MTA | 2023826-2021-03011 | 2021-07-01 | Injury | STAAR SURGICAL COMPANY | 2021-08-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS OTHER (WHITE TO WHITE MEASUREMENT). Manufacturer Narrative: PT INFO: UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12341302&pc=MTA | 2023826-2021-03173 | 2021-07-16 | Injury | STAAR SURGICAL COMPANY | 2021-08-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL 13.2, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP(32MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESSED ON (B)(6) 2021) AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE, RECOMMENDED LENGTH WAS TOO LONG. REPORTEDLY, EYE IS CLEAR, GOOD VAULT, NORMAL. PATIENT REPORTED "RAINBOW EFFECT IN VISION". Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12628814&pc=MTA | 2023826-2021-03987 | 2021-09-01 | Injury | STAAR SURGICAL COMPANY | 2021-10-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED LATER THAT SAME DAY DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE INCISION WAS ENLARGED AND ADDITIONAL SUTURE WAS REQUIRED. THE PROBLEM WAS RESOLVED WITH EXPLANT OF LENS. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12606079&pc=MTA | 2023826-2021-03851 | 2021-09-10 | Injury | STAAR SURGICAL COMPANY | 2021-10-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP(23MMHG) WITHOUT PUPIL BLOCK(ASSESSED ON (B)(6) 2021) AND ANGLE CLOSURE (ASSESED ON (B)(6) 2021), AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, ANGLE CLOSURE.REPORTEDLY, CORNEA AND LENS OK. PIO NORMAL. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO [OR#] SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256529&pc=MTA | 2023826-2022-00157 | 2021-09-16 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -17.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "HIGH VAULT; ELEVATED IOP; NARROW ANGLE". Manufacturer Narrative: B5 - IT WAS LATER CLARIFIED THAT THE LENS WAS EXPLANTED ON (B)(6) 2021. ON (B)(6) 2021 THE REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: B5: IT WAS LATER RE-REPORTED THAT ON 15-OCT-2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS, AND NOT REPLACED THE NEXT DAY IN ANOTHER SURGERY. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13206192&pc=MTA | 2023826-2022-00079 | 2021-10-25 | Injury | STAAR SURGICAL COMPANY | 2022-01-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -08.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP (24MMHG ASSESSED ON (B)(6) 2021) ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13362155&pc=MTA | 2023826-2022-00349 | 2021-12-09 | Injury | STAAR SURGICAL COMPANY | 2022-01-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7; -11.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE; AND ON (B)(6) 2021 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7; -11.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE; AND ON (B)(6) 2021 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC STRETCHED OUT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#:(B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#:(B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15179534&pc=MTA | 2023826-2022-02576 | 2022-04-12 | Injury | STAAR SURGICAL COMPANY | 2022-08-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT; ANGLE CLOSURE WITH ELEVATED IOP. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14929821&pc=MTA | 2023826-2022-02102 | 2022-05-20 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS OF DIFFERENT POWER DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, ANGLE CLOSURE WITH ELEVATED INTROCULAR PRESSURE AND UNREACTIVE (FIXED) PUPIL. THE EXCHANGE RESOLVED THE PROBLEM. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED " STABLE AN IOP RETURNED TO NORMAL VALUES. REFRACTION IS ALMOST EMMETROPE." IN THE REPORTER OPINION PATIENT- RELATED FACTOR WAS THE CAUSE OF THIS EVENT. FOR CAUSE DETAILS THE REPORTER STATED " OVER CORRECTION. HIGH VAULT AND ANGLE CLOSURE ACCOMPANIED BY DILATED PUPIL." Manufacturer Narrative: PATIENT INFORMATION: UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16916212&pc=MTA | 2023826-2023-01798 | 2023-02-23 | Injury | STAAR SURGICAL COMPANY | 2023-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -6.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; ELEVATED INTRAOCULAR PRESSURE; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; THE LENS WAS EXPLANTED ON THIS SAME DATE. ON (B)(6) 2023 A REPLACEMENT LENS OF SHORTER LENGTH AND DIFFERENT MODEL WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "WRONG WTW". Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17216375&pc=MTA | 2023826-2023-02434 | 2023-04-25 | Injury | STAAR SURGICAL COMPANY | 2023-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -14.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ELEVATED IOP(22MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESSED ON (B)(6)2023), UNREACTIVE(FIXED) PUPIL, AND EXCESSIVE VAULT WERE OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, LENS EXCHANGE REQUESTED. Manufacturer Narrative: B5 - LENS WAS EXPLANTED AND REPLACED ON (B)(6) 2023 WITH A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. REPORTEDLY, "SOLVED THE VAULT PROBLEM, IS PERFECT." CLAIM# (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: UNREACTIVE(FIXED) PUPIL. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17279000&pc=MTA | 2023826-2023-02611 | 2023-05-31 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM512.6, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, ANGLE CLOSURE WITH ELEVATED IOP(41MMHG, ASSESSED ON (B)(6) 2023), AND EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE, DIAMETER WAS TOO LARGE, REPLACED SMALLER LENS FIXING PROBLEM. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGLES, ANGLE CLOSURE. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18567731&pc=MTA | 2023826-2024-00371 | 2023-11-17 | Injury | STAAR SURGICAL COMPANY | 2024-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ,SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND INCREASED IOP. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED TO BE THE DEVICE "LENS WAS NOT FITTING FOR THE PATIENT". Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18567793&pc=MTA | 2023826-2024-00372 | 2023-11-17 | Injury | STAAR SURGICAL COMPANY | 2024-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -13.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ,SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND INCREASED IOP. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED TO BE THE DEVICE "LENS WAS NOT FITTING FOR THE PATIENT". Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18568014&pc=MTA | 2023826-2024-00401 | 2023-12-21 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -12.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULT, ELEVATED IOP (INTRAOCULAR PRESSURE) AND ANGLE CLOSURE. IT WAS REPORTED THAT MEDICAL OR SURGICAL INTERVENTION IS NOT NECESSARY. FOR ADDITIONAL INFORMATION THE REPORTER STATED " NO SYMPTOMS, WILL LIKELY NOT EXCHANGE FOR NOW". Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: ANGLE CLOSURE. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12276620&pc=MTA | 2023826-2021-03058 | 2021-07-06 | Injury | STAAR SURGICAL COMPANY | 2021-08-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6; -11.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; ELEVATED IOP; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; RETAINED VISCOELASTIC/OCUCOAT; AND ENLARGEMENT OF PI (SURGICAL) WAS PERFORMED ALONG WITH PARACENTESIS TO REDUCE IOP. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND LATER THE SAME DAY REPLACED WITH A SHORTER LENGTH LENS. THE PROBLEM WAS RESOLVED AND "PATIENT'S VAULT HAS SLOWLY DECREASED TO ACCEPTABLE LEVEL WITH DECREASING IOP". THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED "RETAINED VISCOELASTIC IN SMALL EYE & THICK BROWN IRIS". Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12271376&pc=MTA | 2023826-2021-03057 | 2021-07-06 | Injury | STAAR SURGICAL COMPANY | 2021-08-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6; -11.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; ELEVATED IOP; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; RETAINED VISCOATELASTIC/OCUCOAT; AND ENLARGEMENT OF PI (SURGICAL) WAS PERFORMED ALONG WITH PARACENTESIS TO REDUCE IOP. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND LATER THE SAME DAY REPLACED WITH A SHORTER LENGTH LENS. IT WAS NOTED POSTOPERATIVELY THE PATIENT HAD SHALLOW AC 1000MM VAULT BUT INDICATED THE PROBLEM WAS RESOLVED AND "PATIENT'S VAULT HAS SLOWLY DECREASED TO ACCEPTABLE LEVEL WITH DECREASING IOP". THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED "RETAINED VISELASTIC IN SMALL EYE & THICK BROWN IRIS". Manufacturer Narrative: CORRECTED DATA: B5 - (B)(6) 2021 SHOULD BE CORRECTED TO (B)(6) 2021. CLAIM#: (B)(4). Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: (B)(4) SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, RETAINED OCUCOAT, SHALLOW AC. HEALTH EFFECT IMPACT CODE: (B)(4) ENLARGEMENT OF PI (SURGICAL), BURPED PARACENTESIS, CLAIM#: (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082457&pc=MTA | 2023826-2021-04967 | 2021-11-17 | Injury | STAAR SURGICAL COMPANY | 2021-12-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS -18.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR: "EXCESSIVE VAULT, SHALLOW ANTERIOR CHAMBER, WTW IMPRECISE."Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS -18.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR: "EXCESSIVE VAULT, SHALLOW ANTERIOR CHAMBER, WTW IMPRECISE." Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS -18.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS AND THE PROBLEM RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR: "EXCESSIVE VAULT, SHALLOW ANTERIOR CHAMBER, WTW IMPRECISE." CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10537758&pc=MTA | 2023826-2020-02068 | 2020-08-10 | Injury | STAAR SURGICAL COMPANY | 2020-09-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO HIGH PRESSURE. THE LENS WAS EXCHANGED FOR A DIFFERENT SIZE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SPECIMEN CUP, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS RESIDUE ON LENS. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - ANGLE CLOSURE. H6: HEALTH EFFECT IMPACT CODE: 4625 - ADDITIONAL SURGERY; YAG - ADDITION OF NEW PI. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. ADDITIONAL SURGERY WAS PERFORMED, ADDITION OF NEW PI BY YAG AND ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS NOT RESOLVED. SEE MFR REPORT # 2023826-2020-02119 FOR EXCHANGED LENS. THE CAUSE OF THE EVENT WAS DUE TO PATIENT-RELATED FACTOR. CLAIM # (B)(4). Manufacturer Narrative: DEVICE BLA NUMBER: NA. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18341285&pc=MTA | 2023826-2023-05659 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-12-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: NARROW ANTERIOR CHAMBER, ELEVATED INTRAOCULAR PRESSURE WITH PIGMENT DISPERSION AFTER ICL IMPLANTATION WAS REPORTED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: NARROW ANTERIOR CHAMBER, PIGMENT DISPERSION (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14056490&pc=MTA | 2023826-2022-01147 | 2022-03-04 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -5.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED PIGMENT DISPERSION. ON (B)(6) 2022 ADDITIONAL SURGICAL INTERVENTION (YAG PI) AND MEDICAL INTERVENTION ( PRED FORTE, TIMO COMOD, LOTEMAX) WERE PERFORMED/PERSCRIBED PRIOR TO ICL EXPLANTATION. IT WAS REPORTED THAT ELEVATED IOP: 20MMHG WAS ASSESSED. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND THE PROBLEM RESOLVED. STATUS OF THE EYE (SIGNS/SYMPTOMS) AT LAST VISIT: BCVA: 20/20; UCVA: 20/20; IOP: 16MMHG; AC QUIET AND DEEP. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17846197&pc=MTA | 2023826-2023-04221 | 2023-06-13 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. THE PROBLEM WAS RESOLVED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-04220. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - ANGLE CLOSURE. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # 731544 | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17846540&pc=MTA | 2023826-2023-04220 | 2023-06-13 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO ELEVATED INTRAOCULAR PRESSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PIGMENT DISPERSION. THE LENS WAS REPLACED WITH A SHORTER LENS BUT THE PROBLEM WAS NOT RESOLVED. SEE MFR. REPORT # 2023826-2023-04221 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PIGMENT DISPERSION. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13552675&pc=MTA | 2023826-2022-00438 | 2021-12-02 | Injury | STAAR SURGICAL COMPANY | 2022-02-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED LATER THAT SAME DAY DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE (IOP) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15068918&pc=MTA | 2023826-2022-02299 | 2022-06-10 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE (B)(4). DEVICE CODE (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17725919&pc=MTA | 2023826-2023-04004 | 2023-06-28 | Injury | STAAR SURGICAL COMPANY | 2023-09-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -13.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. ON THIS SAME DATE THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IT WAS ADDITIONALLY NOTED "THE PATIENT UNDERWENT RIGHT EYE ICL IMPLANTATION SURGERY ON (B)(6) 2023. POSTOPERATIVELY ELEVATED INTRAOCULAR PRESSURE WAS OBSERVED EXCEEDING 50MMHG. THE ANTERIOR CHAMBER WAS SHALLOW WITH EXECESSIVE VAULTING. THE LENS WAS REMOVED ON THE SAME AFTERNOON CONSIDERING THE PATIENT'S UNIQUE OCULAR CONDITION, CHARACTERIZED BY A SMALL CORNEAL DIAMETER BUT A SMALL CORNEAL DIAMETER BUT A LARGE POSTERIOR CHAMBER SPACE AND SIGNS OF CILLARY BODY ATROPHERY LIKE CHANGES. AFTER INTERNAL DISCUSSION IT WAS SUGGESTED TO POSTPONE THE SURGERY. HOWEVER DUE TO THE STRONG WILLINGNESS OF THE PATIENT. IT IS RECOMMENDED THE PATIENT SEEK FURTHER TREATMENT AT A MORE EXPERIENCED HIGHER-LEVEL HOSPITAL AFTER A RECOVERY PERIOD OF THREE MONTHS". Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - ANGLE CLOSURE. H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, IN LIQUID. VISUAL INSPECTION FOUND NO DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17908804&pc=MTA | 2023826-2023-04480 | 2023-08-19 | Injury | STAAR SURGICAL COMPANY | 2023-10-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -8.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP. THIS EXPLANT RESOLVED THE PROBLEM CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (AGE<21) SHOULD HAVE BEEN INCLUDED. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, DRY, WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: ANGLE CLOSURE. H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17909001&pc=MTA | 2023826-2023-04481 | 2023-08-19 | Injury | STAAR SURGICAL COMPANY | 2023-10-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP. THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: ANGLE CLOSURE . H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18176646&pc=MTA | 2023826-2023-05183 | 2023-09-12 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -8.0 DIOPTER INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED WITH SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, ELEVATED IOP AND ANGLE CLOSURE AFTER SURGERY. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6- CLICICAL CODE4581- SIGNIFICAT REDUCTION OF IRIDOCORNEAL ANGLES; ANGLE CLOSURE. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14525966&pc=MTA | 2023826-2022-01543 | 2022-03-23 | Injury | STAAR SURGICAL COMPANY | 2022-05-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS OF -5.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. EXCESSIVE VAULT; ELEVATED IOP; AND SIGNIFICANT ENDOTHELIAL CELL LOSS WAS OBSERVED. IOP CONTROLLED DAY AFTER SURGERY. LENS REMAINS IMPLANTED. EYE IS STABLE AND UNDER OBSERVATION. CAUSE IS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE:UNK. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12140355&pc=MTA | 2023826-2021-02386 | 2021-06-08 | Injury | STAAR SURGICAL COMPANY | 2021-07-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE REPORTER INDICATED THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS FOR A SHORTER LENS. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS QUIET AND PATIENT PLEASED WITH OUTCOME. THE PIS WERE ENLARGED BY (B)(6) ON (B)(6) 2021. SEE MFR. REPORT#: 2023826-2021-04212 FOR REPLACEMENT LENS. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - ANGLE CLOSURE. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12817115&pc=MTA | 2023826-2021-04361 | 2021-07-23 | Injury | STAAR SURGICAL COMPANY | 2021-11-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -7.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14382243&pc=MTA | 2023826-2022-01508 | 2022-04-08 | Injury | STAAR SURGICAL COMPANY | 2022-05-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: REPORTEDLY A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS OF -3.5 DIOPTER WAS REMOVED DUE TO HIGH VAULT. IT WAS REPLACED WITH A SHORTER LENGTH LENS ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022. Manufacturer Narrative: PATIENT IDENTIFIER: UNKNOWN,AGE OR DATE OF BIRTH:UNKNOWN, SEX::UNKNOWN, WEIGHT: UNKNOWN, ETHNICITY:UNKNOWN, RACE:UNKNOWN. DATE OF EVENT: UNK. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS OF -3.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REPORTEDLY EXCESSIVE VAULT, ELEVATED IOP, AND NARROW ANGLES WERE FOUND. THE LENS WAS EXPLANTED ON (B)(6) 2022, AND REPLACED ON THE SAME DAY IN A SEPARATE SURGERY WITH A SHORTER LENGTH LENS OF A DIFFERENT MODEL. THIS RESOLVED THE PROBLEM. H6- HEALTH EFFECT- CLINICAL CODE 4581: NARROW ANGLES. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16330098&pc=MTA | 2023826-2023-00399 | 2022-12-10 | Injury | STAAR SURGICAL COMPANY | 2023-02-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT, ELEVATED IOP WITHOUT PUPIL BLOCK, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, ANGLE CLOSURE WITH ELEVATED IOP, AND PIGMENT DISPERSION WERE OBSERVED. CAUSE REPORTED AS UNKNOWN. REPORTEDLY, "HIGHER VAULT THAN EXPECTED PROBABLY DUE TO LONGER LENS." ON (B)(6) 2022, THE LENS WAS REPOSITIONED, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023, THE REPORTER SAID THAT THE DOCTOR GAVE UP EXPLANTING THE LENS BECAUSE IN THE LAST CONTROLS, IT COULD BE SEEN THAT THE VAULT HAD DECREASED, AND SHE DECIDED TO CONTINUE EVALUATING THE PATIENT. Manufacturer Narrative: CORRECTED DATA: H6: HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES; ANGLE CLOSURE; PIGMENT DISPERSION. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12052902&pc=MTA | 2023826-2021-02043 | 2021-05-18 | Injury | STAAR SURGICAL COMPANY | 2021-06-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 AND EXCHANGED FOR A SHORTER LENS. THE LENS WAS DISCARDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED DUE TO LARGE VAULT, SHALLOW ANTERIOR CHAMBER, ELEVATED INTRAOCULAR PRESSURE (IOP) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED WITH THE IMPLANT OF THE REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE (ICL TOO LONG). POST-OP THE PATIENT HAD MILD HALOS AND THE EYE WAS QUIET. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, SHALLOW ANTERIOR CHAMBER. CLAIM # (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6795690&pc=MTA | 2023826-2017-01321 | 2017-06-28 | Injury | STAAR SURGICAL COMPANY | 2017-08-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Intraocular Pressure Increased; Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, ELEVATED IOP. IN ONE WEEK POST-OP, THE SURGEON STATED THAT THE PATIENT HAS BEEN TREATED BY MEDICATIONS: COSOPT, SPERSADEX, AND DIAMOX. MULTIPLE ATTEMPS TO GET UPDATED INFORMATION HAS BEEN UNSUCCESSFUL. Manufacturer Narrative: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES): THE PATIENT'S ELEVATED IOP WAS TREATED WITH MEDICATIONS: COSOPT, SPERSADEX AND DIAMOX. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6795695&pc=MTA | 2023826-2017-01322 | 2017-06-28 | Injury | STAAR SURGICAL COMPANY | 2017-08-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Intraocular Pressure Increased; Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, ELEVATED IOP. IN ONE WEEK POST-OP, THE SURGEON STATED THAT THE PATIENT IOP IS 26.7MMHG, AND HAS BEEN TREATED BY TRUSOPT, AND TOBRADEX MEDICATIONS. THE SURGEON ALSO NOTED THAT THE PATIENT IOP IN TWO WEEKS POST-OP IS 20.6 MMHG. MULTIPLE ATTEMPS TO GET ADDITIONAL INFORMATION HAS BEEN UNSUCCESSFUL. Manufacturer Narrative: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES): THE SURGEON NOTED THAT THE PATIENT'S IOP IN TWO WEEKS POST-OP IS 20.6MM HG AFTER THE PATIENT WAS TREATED WITH MEDICATIONS: TRUSOPT AND TOBRADEX. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12761919&pc=MTA | 2023826-2021-04212 | 2021-06-24 | Injury | STAAR SURGICAL COMPANY | 2021-11-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021, THE SURGEON ENLARGED THE PIS BY YAG DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP) (28MMHG) AND SLIGHT VA BLURRINESS PER PATIENT. THE IOP RESOLVED AND WAS 25MMHG SAME DAY. BY (B)(6) 2021 THE IOP WAS 14MMHG. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2021-02386. Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK RACE: UNK EXPLANT DATE: NA (B)(4). CLAIM # (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17343177&pc=MTA | 2023826-2023-02983 | 2023-03-14 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ANGLE CLOSURE WITH ELEVATED IOP (INTRAOCULAR PRESSURE) AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. FOR CAUSE DETAILS THE REPORTER STATES " SULCUS TO SULCUS SMALLER THAN THE ESTIMATE". Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: ANGLE CLOSURE. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17343354&pc=MTA | 2023826-2023-02984 | 2023-03-14 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. FOR CAUSE DETAILS THE REPORTER STATES " SULCUS TO SULCUS SMALLER THAN THE ESTIMATE". Manufacturer Narrative: H6- HEALTH EFFECT - CLINICAL CODE: "1937" SHOULD BE REMOVED. CLAIM# (B)(4). Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135946&pc=MTA | 2023826-2021-00043 | 2020-12-07 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM6_13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -1.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, ANGLE CLOSURE, AND ELEVATED IOP. ON (B)(6) 2020 THE LENS WAS REPOSITIONED. REPORTEDLY THE IOP AND VAULT ARE BETTER AFTER REPOSITIONING. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-UPDATED INFORMATION: AS OF (B)(6) 2020, UCVA DISTANCE IS 0.4, READING DISTANCE IS GOOD, IOP IS 16/16MMHG AND VAULT IS STILL HIGH AT 1243. CLAIM# (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: AS OF (B)(6) 2021, THE VAULT IS (B)(4). TRANSIENT LOSS OF BCVA REPORTED. CLAIM# (B)(4). Manufacturer Narrative: B5: PER UPDATED INFORMATION, THE LENS WAS EXPLANTED ON (B)(6) 2021. REPORTEDLY, "THE PATIENT IS OK" POST EXCHANGE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135969&pc=MTA | 2023826-2021-00044 | 2020-12-07 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM6_13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -1.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, ANGLE CLOSURE, AND ELEVATED IOP. ON (B)(6) 2020 THE LENS WAS REPOSITIONED. REPORTEDLY THE IOP AND VAULT ARE BETTER AFTER REPOSITIONING. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-UPDATED INFORMATION: AS OF (B)(6) 2020, UCVA DISTANCE IS 0.4, READING DISTANCE IS GOOD, IOP 16/16MMHG AND VAULT IS STILL HIGH AT 1184. CLAIM# (B)(4). Manufacturer Narrative: B5-UPDATED INFORMATION: AS OF 03-FEB-2021 THE VAULT IS 1093. TRANSIENT LOSS OF BCVA REPORTED. CLAIM# (B)(4). Manufacturer Narrative: B5: PER UPDATED INFORMATION, THE LENS WAS EXPLANTED ON (B)(6) 2021. REPORTEDLY, "THE PATENT IS OK." POST EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15195867&pc=MTA | 2023826-2022-02598 | 2022-07-08 | Injury | STAAR SURGICAL COMPANY | 2022-08-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND BLURRED VISION. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. MEDICATION WAS ADMINISTERED. THE LENS WAS DISCARDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THERE WAS HIGH PRESSURE, ZONULAR INSTABILITY, IRIS PROLAPSE, IRIS ATROPHY AND SHALLOWING OF AC. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - ZONULAR INSTABILITY, IRIS PROLAPSE, IRIS ATOPHY, SHALLOWING OF A/C. H6: HEALTH EFFECT IMPACT CODE: 4644 - TOPICAL AND ORAL ANTI-OCULAR HYPERTENSIVE MEDICATION. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16750641&pc=MTA | 2023826-2023-01227 | 2023-02-19 | Injury | STAAR SURGICAL COMPANY | 2023-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (PD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP(60MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESED ON 19/FEB/2023), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, BLURRED VISION, AND EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT IMPACT CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGLES. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10544158&pc=MTA | 2023826-2020-02118 | 2020-08-17 | Injury | STAAR SURGICAL COMPANY | 2020-09-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER, INTO THE PATIENT'S EYE ON (B)(6) 2020. ON (B)(6) 2020 THE SURGEON EXPLANTED THE LENS DUE TO ELEVATED IOP. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND THAT THE LENS IS CURVED IN. CLAIM#: (B)(4). Manufacturer Narrative: B5: ADDITIONAL INFORMATION: LENS HAD BEEN IMPLANTED INTO THE OD EYE. ADDITIONAL SYMPTOMS REPORTED: EXCESSIVE VAULT, ANGLE CLOSURE, BLURRED VISION. ADDITION OF NEW PI-YAG, PHACO-EMULSIFICATION AND IOL IMPLANTATION. PERFORMED REFRACTIVE LENS EXCHANGE WITH PUPILLOPLASTY AND POSTERIOR SYNECHIAELYSIS ON (B)(6) 2020. STATUS: PSEUDOPHAKIA, ENCOURAGED TO USE +1.25 READERS TO HELP WITH COMPUTER AND PHONE. CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR, "EYE COULDN'T HANDLE DEVICE." H6: LINICAL CODE 4581 (ANGLE CLOSURE). CLAIM# (B)(4). Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10551350&pc=MTA | 2023826-2020-02119 | 2020-08-17 | Injury | STAAR SURGICAL COMPANY | 2020-09-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, INTO THE PATIENT'S EYE ON (B)(6) 2020. ON (B)(6) 2020 THE SURGEON EXPLANTED THE LENS DUE TO ELEVATED IOP. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: LENS HAD BEEN IMPLANTED IN THE OS EYE. ADDITIONAL SYMPTOMS REPORTED: EXCESSIVE VAULT, ANGLE CLOSURE, BLURRED VISION. ADDITION OF NEW PI-YAG, PHACO-EMULSIFICATION AND IOL IMPLANTED. PERFORMED REFRACTIVE LENS EXCHANGE WITH PUPILLOPLASTY ON (B)(6) 2020. STATUS: PSEUDOPHAKIA, "GAVE SCL'S +0.75 BASE CURVE 8.4 AND WILL SCHEDULE LASIK IN THE NEAR FUTURE." H6-CLINICAL CODE 4581 (ANGLE CLOSURE) CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SPECIMEN WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND RESIDU ON THE LENS AND THAT THE LENS HAS A CURVED SHAPE. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16309075&pc=MTA | 2023826-2023-00445 | NA | Injury | STAAR SURGICAL COMPANY | 2023-02-06 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -9.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING WITH ELEVATED INTRAOCULAR PRESSURE; GLARE/HALOES; BLURRED VISION. ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THIS SAME DATE IN A SEPARATE SURGERY A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED; IT WAS ADDITIONALLY INDICATED AN INCISION ENLARGEMENT AND ADDITION SUTURE WAS PERFORMED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "SHALLOW VAULT DUE TO LENGTH". Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14057555&pc=MTA | 2023826-2022-01148 | NA | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -4.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. ON (B)(6)2021 SECONDARY YAG WAS PERFORMED AND THE PROBLEM RESOLVED. IT WAS REPORTER THAT INTRAOCULAR PRESSURE WAS TAPER DOWN TO 17MMHG. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#:(B)(4). Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE: 4581- RAINBOW LIGHT. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -4.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE; BLURRED VISION AND HALOS/RAINBOW LIGHT. ON (B)(6) 2021 SECONDARY YAG WAS PERFORMED AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT MEDICATION WAS USED TO LOWER IOP. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. CLAIM# (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5089564&pc=MTA | 2023826-2015-01252 | 2015-08-19 | Injury | STAAR SURGICAL COMPANY | 2015-09-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN HIS RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT REPORTED THE SURGERY WENT WELL BUT AT ONE DAY POST-OP, HE COULDN'T SEE OUT OF HIS EYE, HE HAD TROUBLE FOCUSING AND IT WAS LIKE LOOKING THROUGH A WHITE SHEET OF PAPER. HE WENT IN FOR A FOLLOW-UP VISIT THAT SAME DAY AND IT WAS NOTED THE PERIPHERAL IRIDOTOMIES WERE NOT OPEN AND THE PRESSURE WAS UP. THE PERIPHERAL IRIDOTOMIES WERE OPENED AND THE PATIENT WAS SENT HOME. ON (B)(6) 2015 THE PATIENT HAD THE SAME PROBLEM AND HIS VISION WAS BLURRY. HE WENT BACK TO HIS DOCTOR AND THE DOCTOR DECIDED TO EXPLANT THE LENS THAT SAME DAY. HE WAS STILL HAVING PROBLEMS WITH HIS EYE AND HIS DOCTOR GAVE HIM DROPS. THE FACILITY REPORTED THE LENS WAS EXPLANTED DUE TO PERSISTENT ANGLE CLOSURE. Manufacturer Narrative: PER MEDICAL REVIEW - ANGLE CLOSURE FOLLOWING PIOL IMPLANTATION IS A RARE COMPLICATION. STAAR RECOMMENDS PREVENTION BY PERFORMING TWO LASER PERIPHERAL IRIDOTOMIES BEFORE SURGERY, AND IN THIS CASE, THE PIS HAD BEEN PERFORMED BUT THEY WERE NOT OPEN. EVEN AFTER THE PIS WERE OPENED UP WITH YAG LASER, IOP ROSE AGAIN. ONE POSSIBLE MECHANISM IS INITIAL PUPIL BLOCK DUE TO PUPIL-PIOL CONTACT THAT REMAINED BECAUSE SYNECHIAE DEVELOPED. ANOTHER POSSIBLE MECHANISM IS CILIOLENTICULAR BLOCK, BUT THERE ARE NO UBM OR OCT REPORTS TO DETERMINE THE CAUSE OF PERSISTENT ANGLE CLOSURE. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE INCORRECT DIMENSIONS OF THE IMPLANTED LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY ADVERSE EVENT OR CONDITION WOULD HAVE CAUSED INCORRECT SIZING OF THE LENS RESULTING IN ELEVATED IOP AND PERSISTENT ANGLE CLOSURE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6715656&pc=MTA | 2023826-2017-01142 | 2017-06-06 | Injury | STAAR SURGICAL COMPANY | 2017-07-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, -6.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO BLURRED VISION AND ELEVATED IOP. AN AHMED VALVE WAS IMPLANTED IN THE PATIENT. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: PATIENT IS FINE. BCVA IS 10/10 AND IOP IS CONTROLLED. PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7467360&pc=MTA | 2023826-2018-00604 | 2018-03-26 | Injury | STAAR SURGICAL COMPANY | 2018-04-26 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS LOSS OF BCVA AND ELEVATED IOP. THE LENS REMAINS IMPLANTED. REPORTEDLY, TIMOLOL WAS ADMINISTERED TO THE PATIENT IN ORDER TO LOWER IOP. THE SURGEON NOTES FLUCTUATING REFRACTION. NO UPDATED INFORMATION IS AVAILABLE AT THIS TIME. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -13.0 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS LOSS OF BCVA AND ELEVATED IOP. REPORTEDLY TIMOLOL WAS ADMINISTERED TO THE PATIENT IN ORDER TO LOWER IOP. THE SURGEON NOTES FLUCTUATING REFRACTION. THE REPOSITIONING WAS PERFORMED BUT DID NOT RESOLVE THE PROBLEM. AN EXCHANGE WAS PERFORMED "AROUND (B)(6)" FOR ANOTHER VICM5_13.2 LENS OF -13.0 DIOPTER. THE VA WAS STILL LOW AT 0.65 AND THE PROBLEM WAS NOT RESOLVED. THE REPORTER STATED THAT "THEY DON'T THINK THAT THE ICL IS CAUSING THE LOW VA. THEY ARE THINKING NOW THAT OPACITIES IN THE VITREOUS ARE CAUSING THE PROBLEMS (THE LENS AND MACULA LOOK GOOD)." (B)(4). METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: B5 - UPDATED TO: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -13.0 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS LOSS OF BCVA AND ELEVATED IOP. REPORTEDLY TIMOLOL WAS ADMINISTERED TO THE PATIENT IN ORDER TO LOWER IOP. THE SURGEON NOTES FLUCTUATING REFRACTION. REPOSITIONING WAS PERFORMED ON (B)(6) 2018 BUT DID NOT RESOLVE THE PROBLEM. AN EXCHANGE WAS PERFORMED (B)(6) 2018 FOR ANOTHER VICM5_13.2 LENS OF -13.0 DIOPTER. THE VA WAS STILL LOW AT 0.65 AND THE PROBLEM WAS NOT RESOLVED. THE REPORTER STATED THAT "THEY DON'T THINK THAT THE ICL IS CAUSING THE LOW VA. THEY ARE THINKING NOW THAT OPACITIES IN THE VITREOUS ARE CAUSING THE PROBLEMS (THE LENS AND MACULA LOOK GOOD)." D6B - (B)(6) 2018. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6661714&pc=MTA | 2023826-2017-01012 | 2017-02-27 | Injury | STAAR SURGICAL COMPANY | 2017-06-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, DIOPTER -12.50 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO THE PATIENT EXPERIENCING EXCESSIVE VAULT, LOSS OF BCVA, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. SECONDARY INTERVENTION INCLUDES ENLARGEMENT OF INCISION. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF CLEAR RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6802614&pc=MTA | 2023826-2017-01350 | 2017-07-07 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. AFTER IMPLANTATION, THE PATIENT EXPERIENCED EXCESSIVE VAULT, LOSS OF BCVA, ANGLE CLOSURE WITH ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND BLURRED VISION. AT A LATER DATE THE SURGEON EXPLANTED THE LENS, CUT IT IN ORDER TO SHORTEN IT, THEN RE-IMPLANTED THE LENS BACK INTO THE EYE. IT WAS REPORTED THAT THE VAULT AND EYE ARE GOOD NOW. A LENS EXCHANGE HAS BEEN PLANNED BUT NOT PERFORMED. Manufacturer Narrative: PMA 50(K): - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: OFF-LABEL, UNAPPROVED, OR CONTRAINDICATED USE (KERATOCONUS EYE). (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7441301&pc=MTA | 2023826-2018-00565 | 2017-10-17 | Injury | STAAR SURGICAL COMPANY | 2018-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICH12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON REPORTS EXCESSIVE VAULT, REFRACTIVE SURPRISE, LOSS OF BCVA, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID; IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, SIGNIFICANT REDUCTION OF ICA). (B)(4) OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.81MM)]. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7113305&pc=MTA | 2023826-2017-01883 | 2017-11-08 | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO VAULTING AND INCREASED INTRAOCULAR PRESSURE. A SHORTER LENGTH LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED , A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM -9.5 DIOPTER IMPLANTABLE COLLAMER LENS ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING, ELEVATED IOP, NARROWING OF ANGLE, SHALLOWING OF ACD. THE SURGEON ALSO PERFORMED A SECONDARY YAG ON (B)(6) 2017. THE PATIENT HAD BLURRY VISION DURING THE TIME OF POST-OP. IMPLANTED DATE IS CORRECTED TO (B)(6) 2017. (B)(4). CORRECTED DATA: METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND REPLACED WITH SHORTER LENS. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7267534&pc=MTA | 2023826-2018-00209 | 2017-12-22 | Injury | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), REFRACTIVE SURPRISE, PIGMENT DISPERSION, AND UNREACTIVE (FIXED) PUPIL. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTION: LENS WAS IMPLANTED (B)(6) 2017. THE LENS WAS EXCHANGED WITH A SMALLER LENS ON (B)(6) 2017 DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP. THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL OBSERVATION FOUND THE LENS HAPTIC MISSING A PIECE AND RESIDUE ON LENS. CONCLUSION: OFF-LABEL USE [ANTERIOR CHAMBER DEPTH < 3.00MM (2.84MM)]. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320724&pc=MTA | 2023826-2018-00319 | 2018-01-03 | Injury | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, ANGLE CLOSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), BLURRED VISION, AND IRIS ATROPHY. SURGEON ALSO STATES HE PERFORMED IRRIGATION/ASPIRATION, PHACO-EMULSIFICATION, AND IOL IMPLANTATION. PATIENT IS NOW PSEUDOPHAKIC. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED AND REPLACED WITH A NON-STAAR LENS. AFTER SURGERY DAY 1 THERE WAS NO ANGLE CLOSURE (ONLY ELEVATED IOP), "BUT THEN THE IRIS WAS GETTING CLOSER TO THE CORNEA." DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(6). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7504918&pc=MTA | 2023826-2018-00693 | 2018-04-19 | Injury | STAAR SURGICAL COMPANY | 2018-05-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM MICL13.7, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO OBSERVATION OF EXCESSIVE VAULT WITH BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTIONS: ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO OBSERVATION OF EXCESSIVE VAULT WITH BLURRED VISION AND ELEVATED IOP. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A PIECE OF A HAPTIC MISSING. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5037267&pc=MTA | 2023826-2015-01139 | 2015-07-22 | Injury | STAAR SURGICAL COMPANY | 2015-08-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND INCREASED ELEVATED INTRAOCULAR PRESSURE (UP TO 56MMHG). MEDICATIONS WERE ADMINISTERED BUT DID NOT HELP AND A YAG OF THE PERIPHERAL IRIDOTOMIES WAS PERFORMED BUT THE SURGEON DECIDED TO EXPLANT THE LENS. THE PATIENT EXPERIENCED NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT HAD ADDITIONAL SYMPTOMS OF BLURRY VISION AND NAUSEA. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. THE SURGEON INDICATED THE CAUSE OF THE EVENT WAS A MISCALCULATION/RECOMMENDED LENS SIZE. Manufacturer Narrative: THE PATIENT HAD SHALLOW ANTERIOR CHAMBER WITH NARROWED ANGLES OS, AND INCREASED IOP RESULTING FROM AN OVER-SIZED LENS. THE SURGEON DETERMINED SIZING USING THE FDA APPROVED LABELING. THE RISE IN IOP RESOLVED WITH LENS EXCHANGE, THEREFORE THE CHANGES IN ANTERIOR CHAMBER CONFIGURATION AND INCREASED IOP WERE RELATED TO THE PRESENCE OF THE OUTSIZE LENS IN THE EYE. BASED ON THE COMPLAINT HISTORY, MEDICAL REVIEW, EVALUATION OF THE RETURNED PRODUCT AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5731066&pc=MTA | 2023826-2016-00819 | 2016-04-22 | Injury | STAAR SURGICAL COMPANY | 2016-06-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -5.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT, ANGLE CLOSURE (WITH ELEVATED IOP), BLURRED VISION, AND PERMANENTLY DILATED PUPIL. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTIONAL DATA: (B)(4). CODE NOT APPLICABLE, NO OTHER ADVERSE EVENT REPORTED OUTSIDE OF SECONDARY SURGERY. (B)(4). SHOULD BE IN PATIENT CODE, NOT DEVICE CODE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5681611&pc=MTA | 2023826-2016-00759 | NA | Injury | STAAR SURGICAL COMPANY | 2016-05-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER . THE PATIENT EXPERIENCED BLURRY VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP CONDITION IS GOOD AND BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, THE DEVICES ASSOCIATED WITH THE WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, WAS MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL AND DEVICE HISTORY RECORD REVIEWS, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5923610&pc=MTA | 2023826-2016-01211 | 2016-08-01 | Injury | STAAR SURGICAL COMPANY | 2016-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -07.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, ELEVATED INTRAOCULAR PRESSURE AND REFRACTIVE SURPRISE. THE LENS WAS IMPLANTED UPSIDE DOWN. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: -DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC BROKEN AND BENT. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9417003&pc=MTA | 2023826-2019-02307 | 2018-07-05 | Injury | STAAR SURGICAL COMPANY | 2019-12-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. BLURRED VISION AND CONSTANT VERTICAL BLACK LINE WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY "BLURRING OF VISION SINCE SURGERY ((B)(6) 2018) AND A CONSTANT VERTICAL BLACK ONE WAS SEEN." PER REPORTER ON (B)(6) 2019, DOCTOR HAS REQUESTED A REPLACEMENT LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED TO: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. BLURRED VISION, ELEVATED IOP, AND CONSTANT VERTICAL BLACK LINE WERE OBSERVED. REPORTEDLY "BLURRING OF VISION SINCE SURGERY ((B)(6) 2018) AND A CONSTANT VERTICAL BLACK LINE WAS SEEN." THE ELEVATED IOP 'WAS BOUGHT UNDER CONTROL WITH DROPS' AND 'PBI WAS ALSO DONE.' THE LENS WAS REPOSITIONED ON (B)(6) 2019, BUT THE LINE REMAINED AND BECAME HORIZONTAL. THE LENS REMAINS IMPLANTED WITH PLANNED EXCHANGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. PATIENT CODE 3191 - SECONDARY INTERVENTION: MEDICAL INTERVENTION (DROPS), LENS REPOSITIONING, PBI. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8242379&pc=MTA | 2023826-2019-00102 | 2018-12-07 | Injury | STAAR SURGICAL COMPANY | 2019-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM, MICL13.7, -14.0 DIOPTER, IMPLANTABLE COLLMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS LATER REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, PUPIL BLOCK, ANGLE CLOSURE, INTRAOCULAR PRESSURE, AND BLURRED VISION. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM, PATIENT WAS STABLE. PATIENT IS CURRENTLY ON ALPHAGAN. IN THE REPORTERS OPINION THE CAUSE OF EVENT IS DUE TO THE RECOMMENDED DIAMETER WAS TOO LARGE. Manufacturer Narrative: ADDITONAL DATA: CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8242395&pc=MTA | 2023826-2019-00101 | 2018-12-07 | Injury | STAAR SURGICAL COMPANY | 2019-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM, MICL13.7, -15.0 DIOPTER, IMPLANTABLE COLLMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS LATER REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, PUPIL BLOCK, ANGLE CLOSURE, INTRAOCULAR PRESSURE, AND BLURRED VISION. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM, PATIENT WAS STABLE. PATIENT IS CURRENTLY ON ALPHAGAN. IN THE REPORTERS OPINION THE CAUSE OF EVENT IS DUE TO THE RECOMMENDED DIAMETER, WAS TOO BIG. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: CORRECTED TO "(B)(6) 2018" IN THE INITIAL MDR. PATIENT CODE: "2137" NOT INCLUDED IN THE INITIAL MDR. DEVICE CODE: "1401" NOT INCLUDED IN THE INITIAL MDR. ADDITIONAL DATA: CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8269193&pc=MTA | 2023826-2019-00142 | 2019-01-11 | Injury | STAAR SURGICAL COMPANY | 2019-01-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, LOSS OF BEST CORRRECTED VISUAL ACUITY (BCVA), ELEVATED INTRAOCULAR PRESSURE AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO RECOMMENDED LENS WAS TOO LARGE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10002877&pc=MTA | 2023826-2020-00983 | 2019-08-01 | Injury | STAAR SURGICAL COMPANY | 2020-04-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON THE SAME DATE IN A SECOND SURGERY, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND BLURRED VISION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8774163&pc=MTA | 2023826-2019-01188 | 2019-06-10 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS EXCESSIVE VAULT, LOSS OF BCVA, PUPIL BLOCK, ELEVATED IOP, "FLAT CAMERA," AND ANGLE CLOSURE. ELEVATED IOP WAS TREATED WITH "MEDICAL DROPS" AND PAIN PILLS. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9794960&pc=MTA | 2023826-2020-00430 | 2019-12-26 | Injury | STAAR SURGICAL COMPANY | 2020-03-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -12.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ELEVATED INTRAOCULAR PRESSURE (IOP) AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. H6- DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13998035&pc=MTA | 2023826-2022-00825 | NA | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS OF -13.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXESSIVE VAULT, ELEVATED IOP, BLURRED VISION, GLARE/HALOES, UNREACTIVE (FIXED) PUPIL OF 5MM, AND SHALLOW ANTERIOR CHAMBER WERE REPORTED. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON (B)(6) 2022 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS OTHER. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11416096&pc=MTA | 2023826-2021-00507 | 2021-02-09 | Injury | STAAR SURGICAL COMPANY | 2021-03-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS, BLURRED VISION, GLARE/HALOES, AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT-RELATED FACTOR, "UNUSUAL ANATOMY". REPORTEDLY, "ALL LOOKS GOOD, MUCH LESS GLARE IN NEW LENS PER PATIENT." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ELEVATED IOP, BLURRED VISION, GLARE/HALOES, AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT-RELATED FACTOR, "UNUSUAL ANATOMY". REPORTEDLY, "ALL LOOKS GOOD, MUCH LESS GLARE IN NEW LENS PER PATIENT." D9- CHANGED TO (B)(6) 2021. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6440682&pc=MTA | 2023826-2017-00502 | 2016-12-15 | Injury | STAAR SURGICAL COMPANY | 2017-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF ICA, REFRACTIVE SURPRISE, LOSS OF BCVA, PIGMENT DISPERSION, UNREACTIVE PUPIL (FIXED), GLARE/HALOES, BLURRED VISION, AND IRIS ATROPHY. THE LENS WAS EXCHANGED FOR A SHORTER LENS, SAME DIOPTER. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: METHOD, RESULTS CODES: UPDATED: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6802597&pc=MTA | 2023826-2017-01349 | 2017-07-07 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -6.00 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. AFTER IMPLANTATION, PATIENT EXPERIENCED EXCESSIVE VAULT, ANGLE CLOSURE WITH ELEVATED IOP, GLARE/HALOS, AND BLURRED VISION. A LENS EXCHANGE HAS BEEN PLANNED BUT NOT PERFORMED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6546105&pc=MTA | 2023826-2017-00768 | 2017-01-09 | Injury | STAAR SURGICAL COMPANY | 2017-05-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN EVO+ COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED A VISIBLE DISTORTION TO THE IRIS. NO FURTHER INFORMATION HAS BEEN PROVIDED. AT THE DATE OF MDR SUBMISSION, MULTIPLE ATTEMPTS TO CONTACT REPORTER HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VICM5_13.2, 13.2MM, -10.00 DIOPTER. PATIENT ALSO EXPERIENCED EXCESSIVE VAULT, LOSS OF BCVA, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOES. LENS REMAINS IMPLANTED. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9746415&pc=MTA | 2023826-2020-00379 | 2019-11-25 | Injury | STAAR SURGICAL COMPANY | 2020-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT, ELEVATED IOP, GLARE/HALOES AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. RETURNED PRODUCT EVALUATION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322382&pc=MTA | 2023826-2021-03131 | 2021-07-12 | Injury | STAAR SURGICAL COMPANY | 2021-08-15 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, UNREACTIVE (FIXED) PUPIL, ELEVATED IOP, AND DISCOMFORT. AFTER EXPLANT THE PUPIL AND IOP ARE BACK TO NORMAL. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4625074&pc=MTA | 2023826-2015-00385 | NA | Injury | STAAR SURGICAL COMPANY | 2015-03-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Discomfort; No Code Available | UNK | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) IN (B)(6), 2014. THE REPORTER INDICATED THE SIZE OF THE LENS WAS TOO BIG AND THE ANTERIOR CHAMBER WAS SHALLOW. THE INTRAOCULAR PRESSURE WAS A PROBLEM AND A 2ND PERIPHERAL IRIDECTOMY WAS PERFORMED. THE PATIENT EXPERIENCED SOME DISCOMFORT. THE LENS WILL BE EXPLANTED AND EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE FACILITY REPORTED THE LENS WAS IMPLANTED ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS.Event Description: ADDITIONAL INFORMATION RECEIVED - THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. THE EXCHANGE OF THE LENS TO A SHORTER LENS RESOLVED THE PROBLEM. Manufacturer Narrative: PER MEDICAL REVIEW - MOST LIKELY THE MICL LENS SIZING ERROR CAUSED THE EXCESSIVE LENS VAULTING AND CONSEQUENT IOP SPIKE ASSOCIATED WITH NARROW ANGLE. ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING WAS A CONSEQUENCE OF WRONG LENS USE WHICH INCLUDE THE FOLLOWING: IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENT BASED UPON DIFFERENT TECHNIQUES, POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN A INDIVIDUAL CASE, IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST AND IMPROPER LENS LABEL. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: PT AGE AND WEIGHT: UNK. EVENT DATE: UNK. EXPIRATION DATE - UNK. EXPLANT DATE: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE ROOT CAUSE OF THE COMPLAINT. THE LENS SIZING ERROR MOST LIKELY CAUSED THE EXCESSIVE LENS VAULTING. (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18288508&pc=MTA | 2023826-2023-05482 | 2023-11-21 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Intraocular Pressure Increased; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICM513.7 IMPLANTABLE COLLAMER LENS OF -4.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. PUPIL BLOCK WITH ELEVATED IOP AND GLARES/HALOES ARE OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4) | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10816077&pc=MTA | 2023826-2020-02662 | 2020-04-07 | Injury | STAAR SURGICAL COMPANY | 2020-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Optical Problem | Intraocular Pressure Increased; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON 06-APR-2020. AT ONE DAY POST-OP, THE VAULT AND VISION WAS IMPROVED BUT WITH MYDRIASIS AND HALOS. THE PATIENTS INTRAOCULAR PRESSURE WAS ELEVATED BUT WAS TREATED AGGRESSIVELY WITH DIAMOX AND TOPICAL DROPS. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2020-00945. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE PATIENT IS DOING WELL WITH MILD RESIDUAL ASTIGMATISM, THAT WE HOPE TO ADDRESS WITH POSSIBLY PRK. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10515451&pc=MTA | 2023826-2020-02022 | NA | Injury | STAAR SURGICAL COMPANY | 2020-09-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Optical Problem; No Apparent Adverse Event | Intraocular Pressure Increased; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. PRESSURE SPIKES WERE OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONA INFORMATION: B5- THE REPORTER INDICATED THAT A 13.7MM MICL 13.7 OF -7.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ELEVATED IOP; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS REPORTED. PATIENT STATED "SHE NOTICED HALOES AROUND ANY LIGHTS. THE LENS REMAINS IMPLANTED; HOWEVER THE SURGEON PERFORMED ENLARGEMENT OF PI; AND IT IS REPORTED THAT THE SURGEON HAS GONE BACK AND PERFORMED A SECOND IRIDOTOMY. PATIENT STATES NO PAIN OR VISION CHANGES. THE LEFT EYE PRESSURE IS OK. D6A: IMPLANT DATE: (B)(6) 2020. H1-SERIOUS . H6- CLINICAL CODE:4581-SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. HEALTH EFFECT:4625- ENLARGEMENT OF PI; IRIDOTOMY AND SECOND IRIDOTOMY. H6-INVESTIGATION CODE 4110: LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 IN PREVIOUS MDR THE LENS DIOPTER REPORTED -7.0 IS CORRECTED TO -11.5. CLAIM# (B)(4). Manufacturer Narrative: NO APPARENT HARM OCCURRED IN RELATION TO THE ADVERSE EVENT. UNABLE TO ENTER CODE. (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11795155&pc=MTA | 2023826-2021-01396 | NA | Injury | STAAR SURGICAL COMPANY | 2021-05-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Intraocular Pressure Increased; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: IT WAS REPORTED THAT A 12.6MM, VICMO12.6, -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT RIGHT EYE (OD) AS A REPLACEMENT LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF ICA BUT IT DID NOT RESOLVE THE PROBLEM. FOR STATUS OF THE EYE " PRESENCE OF HALOS; UNREACTIVE PUPIL" WAS REPORTED. " DOCTOR BELIEVES THAT HIGH IOP, MAY BE DUE TO CORTICOSTEROIDS OR MANIPULATION" WAS REPORTED FOR ADDITIONAL INFORMATION. ON (B)(6) 2021 THE REPORTER REPORTED "CURRENT CONDITION: PATIENT COMPLAINING OF GLARE; LITTLE REACTIVE PUPILS AND PIO 18. THE USE OF CORTICOSTEROIDS IS PART OF THE POSTOPERATIVE TREATMENT. A PROGRESSIVE REDUCTION WAS MADE AND STOPPED YESTERDAY, (B)(6) 2021. CURRENTLY USING ANTI-GLAUCOMA DORZOLAMIDE EYE DROPS AND BRIMONIDINE. THE DOCTOR WILL TRY, NEXT WEEK, TO ADAPT THE CONTACT LENSES TO REDUCE THE RESIDUAL DEGREE AND CHECK IF THE VISUAL COMPLAINT IMPROVES." ON (B)(6) 2021 THE REPORTER REPORTED " PATIENT IS A LITTLE MORE SATISFIED. LAST REVIEW ON (B)(6) 2021: DOCTOR HAD TO USE PILOCARPINE, TO CLOSE THE PUPIL A LITTLE, WHICH WAS 6MM. 0.73 VAULT. THE PATIENT REMAINS UNDER OBSERVATION AND EVALUATION WEEKLY". IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5- IT WAS LATER REPORTED THAT "AFTER A WHILE THE GLARE DISAPPEARED." CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM # (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5955095&pc=MTA | 2023826-2016-01266 | 2016-03-25 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE, GLARE AND HALOS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THAT THE HAPTIC WAS BROKEN. UPON DIMENSIONAL INSPECTION, THE LENS WAS FOUND TO BE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5955096&pc=MTA | 2023826-2016-01267 | 2016-03-25 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -14.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE, GLARE AND HALOS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. THIS PRODUCT WAS NOT MARKETED IN THE US. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6028643&pc=MTA | 2023826-2016-01460 | 2016-02-06 | Injury | STAAR SURGICAL COMPANY | 2016-10-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, AND GLARES/HALOES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. FOREIGN MATERIAL - HAIR ON THE LENS SURFACE PRESENT. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6028650&pc=MTA | 2023826-2016-01459 | 2016-02-06 | Injury | STAAR SURGICAL COMPANY | 2016-10-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, AND GLARES/HALOES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND FOREIGN MATERIAL (SPOT) ON SURFACE, BUT NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6200082&pc=MTA | 2023826-2016-01824 | 2016-10-13 | Injury | STAAR SURGICAL COMPANY | 2016-12-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOES, AND ANGLE CLOSURE WITH ELEVATED IOP. AS OF THE DATE OF MDR FILING, THE LENS HAS NOT YET BEEN EXPLANTED AND REMAINS IMPLANTED IN THE PATIENT'S EYE. Manufacturer Narrative: ADDITIONAL INFORMATION: ADDITIONAL EVENT DESCRIPTION: LENS DIOPTER: -17.0/+1.5/70 (SPHERE/CYLINDER/AXIS). THE SURGEON ADMINISTERED IOP LOWERING MEDICATION. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9988375&pc=MTA | 2023826-2020-00945 | 2020-03-26 | Injury | STAAR SURGICAL COMPANY | 2020-04-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE SURGEON HAD PERFORMED A LIMBAL RELAXING INCISION BUT IT WAS LEAKING SO THE SURGEON SUTURED THE INCISION CLOSED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THERE WAS IMPROVED VISION AND VAULT BUT WITH MYDRIASIS AT DAY ONE AND PATIENT COMPLAINED OF HALOS. INTRAOCULAR PRESSURE WAS ELEVATED BUT WAS TREATED AGGRESSIVELY WITH DIAMOX AND TOPICAL DROPS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A ZIP LOC BAG. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D9- "(B)(6) 2020" SHOULD BE CORRECTED TO "(B)(6) 2020" IN THE PREVIOUS MDR. H6- PATIENT CODE 1937 SHOULD BE REMOVED FROM INITIAL MDR. H6- PATIENT CODE 3191- (MYDRIASIS) SHOULD BE REMOVED FROM INITIAL MDR. H6- DEVICE CODE 3001 SHOULD BE REMOVED FROM INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5572295&pc=MTA | 2023826-2016-00437 | 2015-10-20 | Injury | STAAR SURGICAL COMPANY | 2016-04-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | Intraocular Pressure Increased; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -17.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS). THE REPORTER INDICATED THE LENS DISLOCATED/SUBLUXED, WITH THE PATIENT EXPERIENCING ELEVATED INTRAOCULAR PRESSURE AND GLARES/HALOS. THE PATIENT COULD SEE A CRESCENT SHAPE ON THE RIGHT SIDE OF THEIR VISUAL FIELD, WHICH, WITH REFLECTION OF LIGHT, WAS MAKING IT IMPOSSIBLE FOR THEM TO DRIVE AT NIGHT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7404316&pc=MTA | 2023826-2018-00506 | 2018-02-08 | Injury | STAAR SURGICAL COMPANY | 2018-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased; Iritis; No Code Available | P030016 | NA | 1 | Event Description: THE SURGEON INDICATED THAT HE IMPLANTED A 12.6MM MICL12.6, -8.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 IRITIS WAS OBSERVED. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PERSISTENT IRITIS AND HIGH IOP (INTRAOCULAR PRESSURE). THE SURGEON REPORTED TO USE COMBIGAN, DUREZOL, PRED FORTE, AND LOTEMAX OVER THE COURSE OF TREATMENT. IT WAS REPORTED THAT AN IRIDECTOMY WAS PERFORMED. THE SURGEON REPORTED THAT HE BELIEVED THE CAUSE OF THE IRITIS WAS PIGMENT DISPERSION FROM ICL CHAFING OF POSTERIOR IRIS. ON (B)(6) 2018 THE SURGEON REPORTED THAT HE WAS NOT CERTAIN OF WHY THE INFLAMMATION CONTINUED. HE STATED THAT THERE WERE SOME PIGMENTARY CELLS, BUT MOST WERE WHITE CELLS AND THAT THE LENS SIZING WAS CORRECT. HE STATED THAT THE PATIENT IS DOING WELL IN SOFT CONTACT LENSES, INFLAMMATION IS RESOLVING. HE ALSO REPORTED THAT THEY WILL BE PERFORMING A PROPER FITTING FOR THE PATIENT IN THE FUTURE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A CARTRIDGE CASE. LENS WAS STUCK TO TAPE. VISUAL INSPECTION FOUND TAPE RESIDUE ON LENS. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: PATIENT CODE:(B)(4) SECONDARY SURGICAL INTERVENTION,EXPLANT. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7631046&pc=MTA | 2023826-2018-00931 | 2018-02-28 | Injury | STAAR SURGICAL COMPANY | 2018-06-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Intraocular Pressure Increased; Iritis; Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM MICL13.2, -12.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFTT EYE (OS) ON (B)(6) 2018. ON (B)(6) /2018 THE LENS WAS REMOVED DUE TO OBSERVATION OF EXCESSIVE VAULT, ELEVATED IOP (INTRAOCULAR PRESSURE), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, GLARE/HALOES, AND IRIS ATROPHY. ADDITIONAL SURGICAL INTERVENTION (ENLARGEMENT, ADDITION OF NEW PI AND REPOSITIONING) WERE PERFORMED. IT WAS REPORTED THAT AT PATIENTS LAST FOLLOW UP ((B)(6) 2018) THE PATIENT WAS ON FOUR DIFFERENT GLAUCOMA DROPS, IOP 25MMHG. ADDITIONAL INFORMATION REPORTED, PI PATENT, PATIENT REPORTS DECREASED GLARE. EXPLANTED ICL AFTER ROTATION, IRIS CELL AND FLARE INCREASED THE IOP. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A THERMA SCIENTIFIC CONTAINER. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: ADDITION OF NEW PI WAS NOT REPORTED. . Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THAT THE PATIENT HAD IRITIS DUE TO THE RUBBING OF THE LENS. THE LENS WAS ROTATED AND THE IRITIS STILL CONTINUED. ANTIBIOTICS AND ANTI-INFLAMMATORIES WERE USED TO ADDRESS THE INFLAMMATION. THE PROBLEM OF THE CELL AND FLARE WAS RESOLVED AFTER THE REMOVAL OF THE ICL, THERE WERE NO CELL AND FLARE POST REMOVAL. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5458648&pc=MTA | 2023826-2016-00197 | NA | Injury | STAAR SURGICAL COMPANY | 2016-02-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; Loss of Vision; Discomfort; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED AN MICL IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE. AT THE SAME DAY OF SURGERY AND ONE DAY POST-OP THE PATIENT WAS DOING WELL AND VA WAS 20/15. THE PATIENT TRAVELLED OUT OF STATE THAT SAME WEEK OF SURGERY AND REPORTED A SLIGHT DECLINE IN VISION AND DISCOMFORT TO THE EYE. THE PATIENT WAS PRESCRIBED A STEROID AND PRESSURE DROP. THE PATIENT WAS ADVISED TO SEE A LOCAL DOCTOR AND PRESENTED WITH ANGLE CLOSURE. THE PATIENT RETURNED TO THE SURGEON THE FOLLOWING WEEK AND PRESENTED WITH A SHALLOW CHAMBER. THE SURGEON HAD DIFFICULTY OPENING AN ADDITIONAL PI, TEMPORIZING HER PRESSURE BUT THE PATIENT WANTED THE ICL REMOVED. THE LENS WAS EXPLANTED AND THE PATIENT HAS RECOVERED HER VISION WELL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: INFORMATION RECEIVED - THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE. BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7875111&pc=MTA | 2023826-2018-01395 | 2018-08-15 | Injury | STAAR SURGICAL COMPANY | 2018-09-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Missing Value Reason | MTA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, MICL13.2, -14.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXCHANGED ON (B)(6) 2018 WITH A SHORTER LENGTH LENS DUE TO PUPIL BLOCK WITH ELEVATED IOP (INTRAOCULAR PRESSURE), TOO LARGE LENS WITH AC SHALLOWING, HIGH VAULTING AND NARROWING OF ANGLE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). LENS WAS RETURNED DRY, IN A SPECIMEN CONTAINER. VISUAL INSPECTION FOUND A TORN HAPTIC AND RESIDUE ON THE LENS SURFACE. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12005327&pc=MTA | 2023826-2021-01927 | 2021-05-16 | Injury | STAAR SURGICAL COMPANY | 2021-06-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Nausea; Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ANGLE CLOSURE AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE SURGEON ENLARGED THE PIS. THE PATIENT COMPLAINED OF EYE PAIN AND NAUSEA. ANOTHER LENS WAS NOT IMPLANTED. PATIENT WAS TAKING KRATOM OR POSSIBLY OTHER MEDS/DRUGS, PT ADMITTED TO USING "EDIBLE" CBD GUMMY AND KRATOM. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A CONTACT LENS CASE. VISUAL INSPECTION FOUND A HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. (B)(4). CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6832459&pc=MTA | 2023826-2017-01382 | 2017-08-02 | Injury | STAAR SURGICAL COMPANY | 2017-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Intraocular Pressure Increased; Nausea; Pain; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT RETURNED THE NEXT DAY, (B)(6) 2017, IN PAIN, NAUSEOUS, WITH A PRESSURE SPIKE (60MMHG) AND STEAMY EYE. ANOTHER IRIDOTOMY WAS PERFORMED AND PRESSURE DROPPED DOWN TO 40MMHG. THE CHAMBER WAS SHALLOW, WITH SOME ANGLE CLOSURE AND THE SURGEON DECIDED TO EXPLANT THE LENS THE SAME DAY. A SHORTER LENS (13.2MM) WAS IMPLANTED BUT THE PRESSURE AGAIN SPIKED AND THE LENS WAS EXPLANTED SAME DAY ((B)(6) 2017) - SEE MDR #2023826-2017-01383. A THIRD LENS WAS NOT IMPLANTED. THE LENS WAS DISCARDED. THE PATIENT IS DOING FINE. Manufacturer Narrative: (B)(4). DEVICE MANUFACTURER DATE: IT IS CORRECTED TO 23-DEC-2016. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17777874&pc=MTA | 2023826-2023-04107 | 2023-07-15 | Injury | STAAR SURGICAL COMPANY | 2023-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS. THE REPORTER STATED " PREOPERATIVE BASAL IOP WAS HIGHER THAN NORMAL, AND NO FURTHER EXAMINATION WAS PERFORMED TO EXCLUDE GLAUCOMA. AFTER THE CRYSTAL IS REMOVED, THE INTRAOCULAR PRESSURE IS TEMPORARILY STABILIZED TO NORMAL VALUE AFTER DRUG TREATMENT. THE LENS WAS REMOVED AND TRANSFERRED TO THE GLAUCOMA DEPARTMENT FOR FURTHER CONSULTATION." Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO CORNEAL ANGLES. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11939033&pc=MTA | 2023826-2021-01672 | 2021-05-11 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2, -5.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY, OBJECTIVE SLE: TRACE DMF. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN SPECIMEN CUP. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. H6: CLINCAL CODE - 1937 CORRECTED TO 4582. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986305&pc=MTA | 2023826-2021-01864 | 2021-05-19 | Injury | STAAR SURGICAL COMPANY | 2021-06-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), ANGLE CLOSURE, AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986314&pc=MTA | 2023826-2021-01865 | 2021-05-19 | Injury | STAAR SURGICAL COMPANY | 2021-06-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), ANGLE CLOSURE, AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4, | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12669504&pc=MTA | 2023826-2021-04033 | 2021-09-28 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). AQUEOUS WAS RELEASED FROM THE CHAMBER. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B3- CORRECTED TO (B)(6) 2021. B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. AQUEOUS WAS RELEASED FROM THE CHAMBER. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. ELEVATED INTRAOCULAR PRESSURE (IOP) WAS REPORTED AS A POST-OPERATIVE CONDITION. THE CAUSE OF THE EVENT WAS UNKNOWN. H6- HEALTH EFFECT- CLINICAL CODE: REMOVED CODE 1937 AND REPLACED WITH CODE 4582. CLAIM# (B)(4). Manufacturer Narrative: B5: PER EMAIL ON 22-FEB-2022, IT WAS CLARIFIED THAT ELEVATED IOP WAS EXPERIENCED AFTER THE INITIAL LENS WAS IMPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14401934&pc=MTA | 2023826-2022-01487 | 2022-04-08 | Injury | STAAR SURGICAL COMPANY | 2022-05-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM MICL13.7 -5.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO HIGH VAULT AND REPLACED WITH A SHORTER LENGTH LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 13.7MM MICL13.7 -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO HIGH VAULT AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS LATER REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS "OTHER", FOR CAUSE DETAILS THE REPORTER REPORTED " HIGH VAULT AND NARROW ANGLES". CLAIM # (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15099731&pc=MTA | 2023826-2022-02398 | 2022-06-07 | Injury | STAAR SURGICAL COMPANY | 2022-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF -9.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT:UNKNOWN, ETHNICITY:UNKNOWN, RACE:UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: CLAIM #(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17295510&pc=MTA | 2023826-2023-02791 | 2023-05-30 | Injury | STAAR SURGICAL COMPANY | 2023-07-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5 - "THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT." SHOULD BE CORRECTED TO "THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP." IN THE INITIAL MDR. H6 - HEALTH EFFECT CLINICAL CODE 1937: INTRAOCULAR PRESSURE INCREASED. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15285150&pc=MTA | 2023826-2022-02891 | NA | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -8.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED DUE TO A HIGH VAULT. A REPLACEMENT LENS WAS IMPLANTED OF UNKNOWN MODEL AND SIZE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. H6-PRODUCT EVALUATION: LENS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: B5: ELEVATED INTRAOCULAR PRESSURE WITH MEDICAL INTERVENTION WAS ALSO REPORTED. H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17561598&pc=MTA | 2023826-2023-03510 | NA | Injury | STAAR SURGICAL COMPANY | 2023-08-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A REPLACEMENT 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF A -4.5 WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)6) 2023. THE PROBLEM WAS RESOLVED. REPORTEDLY, "A FEW STRIAE ON THE ENDOTHELIUM." Manufacturer Narrative: A4-A6:UNK . H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17684361&pc=MTA | 2023826-2023-03760 | 2023-03-11 | Injury | STAAR SURGICAL COMPANY | 2023-09-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT AND ELEVATED IOP. LATER ON (B)(6) 2023 A SHORTER LENGTH LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11065610&pc=MTA | 2023826-2020-03104 | 2020-11-24 | Injury | STAAR SURGICAL COMPANY | 2020-12-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement; No Apparent Adverse Event | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED UPSIDE DOWN AND REMOVED A 12.6MM VICMO12.6, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE ON (B)(6) 2020. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2020, "ANOTHER LENS WILL BE IMPLANTED BUT THE TIMING OF THE OPERATION IS UNCERTAIN." Manufacturer Narrative: ADDITIONAL INFORMATION. D10 - INJECTOR MODEL#MSI-PF,LOT#UNK, CARTRIDGE MODEL#SFC-45,LOT#UNK, FOAMTIPPLUNGER MODEL#FTP,LOT#UNK. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE ON (B)(6) 2020. THE LENS WAS IMPLANTED AND REMOVED INTRA-OP WITHIN THE SAME SURGERY DUE TO THE LENS IMPLANTED UPSIDE DOWN. AT POST-OP ELEVATED IOP WAS OBSERVED DUE TO POSTOPERATIVE MEDICATION. "AFTER ADJUSTING THE MEDICATION, THE PATIENT FELT NORMAL AND DID NOT RETURN TO THE HOSPITAL FOR EXAMINATION AGAIN." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9913311&pc=MTA | 2023826-2020-00744 | 2018-04-26 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PROGRESSIVE INCREASE IN INTRAOCULAR PRESSURE (IOP). THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - RETURNED IN A SPECIMEN CUP, DRY WITH CLEAR SURIGCAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7860761&pc=MTA | 2023826-2018-01247 | 2018-08-10 | Injury | STAAR SURGICAL COMPANY | 2018-09-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. UPON ASSING ON (B)(6) 2018, ELEVATED IOP OF 39 MMHG WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. LENS WAS EXPLANTED ON (B)(6) 2018 WITH NO PLANS TO EXCHANGE LENS. AFTER REMOVAL OF LENS, THE PATIENT'S INTRAOCULAR PRESSURE RETURNED TO NORMAL. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8337166&pc=MTA | 2023826-2019-00248 | 2018-10-03 | Injury | STAAR SURGICAL COMPANY | 2019-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE AND UNREACTIVE PUPIL (FIXED). THE SURGEON ATTEMPTED TO FIX THE PUPIL BLOCK SURGICALLY ON (B)(6) 2018 BUT WAS UNSUCCESSFUL. A SYNECHIA WAS LATER DISCOVERED AT 4 O'CLOCK POSITION. THE PUPIL IS STILL NOT REACTING TO LIGHT AFTER EXPLANT. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO FIX THE PUPIL BLOCK SURGICALLY - THE SURGEON WENT IN THROUGH THE TEMPORAL INCISION AND FLUSHED WITH ANAESTHESIA AND MIOCHOL AND LOOSENED A SYNECHIA. THERE WERE CHECKUPS TO SEE IF THERE WERE ANY CHANGES. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8287616&pc=MTA | 2023826-2019-00180 | 2018-12-20 | Injury | STAAR SURGICAL COMPANY | 2019-01-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO INCREASED POST-OP PRESSURE. ON (B)(6) 2019 A SHORTER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. NO PATIENT INJURY IS NOTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN GAUZE. THERE WAS CLEAR SURGICAL RESIDUE AND DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8441979&pc=MTA | 2023826-2019-00438 | 2019-01-16 | Injury | STAAR SURGICAL COMPANY | 2019-03-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8668952&pc=MTA | 2023826-2019-01045 | 2019-02-19 | Injury | STAAR SURGICAL COMPANY | 2019-06-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9152716&pc=MTA | 2023826-2019-01856 | 2019-05-20 | Injury | STAAR SURGICAL COMPANY | 2019-10-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT 13.2MM, MICL13.2, -14.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED ON (B)(6) 2019 AND EXPLANTED ON (B)(6) 2019 DUE TO ELEVATED PRESSURE THAT COULD NOT BE RESOLVED WITH ADDITIONAL PI'S AND MEDS.ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT 13.2MM, MICL13.2, -14.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO ANGLE CLOSURE WITH ELEVATED IOP THAT COULD NOT BE RESOLVED WITH ADDITIONAL PI AND MEDICATION. THE REPORTER STATED " INCREASED IOP D/T EITHER ANGLE CLOSURE FROM ANTERIORLY ROTATED CILIARY BODY VS ATYPICAL AQUEOUS MISDIRECTION." STATUS OF THE EYE: DOING WELL AFTER ICL REMOVAL. IT WAS REPORTED THAT THE ANGLES ARE OPEN. IOP IS NORMAL. NO LOSS OF BEST CORRECTED VISION BUT PATIENT IS BACK TO WEARING GLASSES. PATIENT CODE 3191: ANGLE CLOSURE. CLAIM # (B)(4). Manufacturer Narrative: DATE OF EVENT: "(B)(6) 2019" SHOULD BE IN THE PREVIOUS MDR. CONCOMITANT MEDICAL PRODUCTS: "INJECTOR MODEL: MSIPF - INJ LOT: 1449232. CARTRIDGE MODEL: SFC45 - CART LOT: 1446792. INJECTOR SYSTEM MODEL: FTP - INJ SYSTEM LOT: 1448653" SHOULD BE IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9153161&pc=MTA | 2023826-2019-01857 | 2019-05-20 | Injury | STAAR SURGICAL COMPANY | 2019-10-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT 13.2MM, MICL13.2, -15.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED ON (B)(6) 2019 AND EXPLANTED ON (B)(6) 2019 DUE TO ELEVATED PRESSURE THAT COULD NOT BE RESOLVED WITH ADDITIONAL PI'S AND MEDS.ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: EVENT: THE REPORTER INDICATED THAT 13.2MM, MICL13.2, -15.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON 05 MAR 2019. ON 21 MAY 2019 THE LENS WAS EXPLANTED DUE TO ANGLE CLOSURE WITH ELEVATED IOP THAT COULD NOT BE RESOLVED WITH ADDITIONAL PI AND MEDICATION. THE REPORTER STATED "INCREASED IOP D/T EITHER ANGLE CLOSURE FROM ANTERIORLY ROTATED CILIARY BODY VS ATYPICAL AQUEOUS MISDIRECTION." STATUS OF THE EYE: DOING WELL AFTER ICL REMOVAL. IT WAS REPORTED THAT THE ANGLES ARE OPEN. IOP IS NORMAL. NO LOSS OF BEST CORRECTED VISION BUT PATIENT IS BACK TO WEARING GLASSES. PATIENT CODE 3191: ANGLE CLOSURE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B3- "05/20/2019" SHOULD BE IN THE PREVIOUS MDR. D11- "INJECTOR MODEL: MSIPF - INJ LOT: 1449232; CARTRIDGE MODEL: SFC45 - CART LOT: 1446792; INJ SYSTEM MODEL: FTP - INJ SYSTEM LOT: 1448653" SHOULD BE IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8841538&pc=MTA | 2023826-2019-01378 | 2019-07-03 | Injury | STAAR SURGICAL COMPANY | 2019-07-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -7.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019, THREE WEEKS POST OP THE SURGEON NOTICED DILATED PUPIL. ELEVATED IOP WAS NOTICED ON (B)(6) 2019. SURGICAL INTERVENTION (YAG PIS) AND MEDICAL INTERVENTION (PREDNISOLONE STEROIDS, MIOTICS, MYDRIATICS, MANNITOL, AND ORAL MEDICATION) WERE PERFORMED/PRESCRIBED BUT DID NOT RESOLVE THE ISSUE, THE REPORTER STATED " FROM (B)(6) 2019 NOT RESPONDING, EVEN TO MYDRIATICS, FORM (B)(6) 2019 IOP NOT RESPONDING FOR ANY MEDICATION". FOR LAST STATUS OF PUPIL AND IOP THE REPORTER STATED "PRESSURE IS GETTING UNDER CONTROL AND CONTINUING TOPICAL ANTI GLAUCOMA DROPS AND PATIENT IS STILL UNDER OBSERVATION. PUPIL IS DILATED AND FIXED, NOT RESPONDING TO LIGHT". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K #: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8841649&pc=MTA | 2023826-2019-01379 | 2019-07-03 | Injury | STAAR SURGICAL COMPANY | 2019-07-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -6.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019, THREE WEEKS POST OP THE SURGEON NOTICED DILATED PUPIL. ELEVATED IOP WAS NOTICED ON (B)(6) 2019. SURGICAL INTERVENTION (YAG PIS) AND MEDICAL INTERVENTION (PREDNISOLONE STEROIDS, MIOTICS, MYDRIATICS, MANNITOL, AND ORAL MEDICATION) WERE PERFORMED/PRESCRIBED BUT DID NOT RESOLVE THE ISSUE, THE REPORTER STATED " FROM (B)(6) 2019 NOT RESPONDING, EVEN TO MYDRIATICS, FORM (B)(6) 2019 IOP NOT RESPONDING FOR ANY MEDICATION". FOR LAST STATUS OF PUPIL AND IOP THE REPORTER STATED "PRESSURE IS GETTING UNDER CONTROL AND CONTINUING TOPICAL ANTI GLAUCOMA DROPS AND PATIENT IS STILL UNDER OBSERVATION. PUPIL IS DILATED AND FIXED, NOT RESPONDING TO LIGHT". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K #: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10222176&pc=MTA | 2023826-2020-01405 | 2019-09-20 | Injury | STAAR SURGICAL COMPANY | 2020-07-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10222185&pc=MTA | 2023826-2020-01406 | 2019-09-20 | Injury | STAAR SURGICAL COMPANY | 2020-07-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9507349&pc=MTA | 2023826-2019-02475 | 2019-11-13 | Injury | STAAR SURGICAL COMPANY | 2019-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -8.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO HIGH PRESSURE. THERE WAS NO PATIENT INJURY. THE REPORTER STATED THAT THE SURGEON WILL BE REPLACING THE LENS WITH A SHORTER LENGTH LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5: THE LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: DEVICE WAS RETURNED IN A SPECIMEN CUP, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT WITH RESIDUE ON THE LENS. CLAIM#:(B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9492475&pc=MTA | 2023826-2019-02445 | 2019-11-19 | Injury | STAAR SURGICAL COMPANY | 2019-12-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -7.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019 AND THE PATIENT EXPERIENCED ELEVATED IOP. A VAULT VALUE WAS REPORTED AND PUPIL BLOCK WITH ELEVATED IOP (28 MMHG). ADDITIONAL SURGICAL INTERVENTION (YAG, IRIDOTOMIA, IRIS SUTURE) WERE PERFORMED. IT WAS REPORTED THAT THE IOP WAS DECREASED. CAUSE OF EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10222180&pc=MTA | 2023826-2020-01410 | 2020-01-06 | Injury | STAAR SURGICAL COMPANY | 2020-07-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10222182&pc=MTA | 2023826-2020-01409 | 2020-01-06 | Injury | STAAR SURGICAL COMPANY | 2020-07-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10484156&pc=MTA | 2023826-2020-02011 | 2020-05-08 | Injury | STAAR SURGICAL COMPANY | 2020-09-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -2.25 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO PUPIL BLOCK WITH ELEVATED IOP AND THE PROBLEM RESOLVED. THE REPORTER STATED THAT THE PATIENT FELT UNCOMFORTABLE AND WANTED THE LENS REMOVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6341607&pc=MTA | 2023826-2017-00241 | NA | Injury | STAAR SURGICAL COMPANY | 2017-02-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S EYE AND IT WAS REPORTED THE INTRAOCULAR PRESSURE WAS ELEVATED AND THE ANGLE WAS TOO NARROW. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9485143&pc=MTA | 2023826-2019-02444 | NA | Injury | STAAR SURGICAL COMPANY | 2019-12-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM,VICM5_13.7, -10.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD). THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE AND PI HAD TO BE PERFORMED TO BRING THE PRESSURE UNDER CONTROL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS WAS IMPLANTED INTO THE RIGHT EYE (OD) ON (B)(6) 2019. SHOULD BE ADDED IN THE INITIAL MDR. D6- (B)(6) 2019 SHOULD BE IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. IMPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9169871&pc=MTA | 2023826-2019-01893 | 2019-09-05 | Injury | STAAR SURGICAL COMPANY | 2019-10-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM MICL12.6, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE ON (B)(6) 2019. THE REPORTER REPORTED PRE-OP AND POST OP ENDOTHELIAL CELL COUNT, ELEVATED IOP THAT LASTED 4 HOURS AND INCISION ENLARGEMENT. THE LENS WAS REMOVED AND EXCHANGED WITHIN THE SAME SURGERY AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THAT A 12.6MM MICL12.6, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS TORE BROKE DURING INJECTION/DELIVERY INTO THE EYE ON (B)(6) 2019. THE REPORTER REPORTED PRE-OP AND POST OP ENDOTHELIAL CELL COUNT, ELEVATED IOP THAT LASTED 4 HOURS AND INCISION ENLARGEMENT. THE LENS WAS REMOVED AND EXCHANGED WITHIN THE SAME SURGERY AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. THE REPORTER STATED THAT THERE WAS NO SIGNIFICANT LOSS OF ENDOTHELIAL CELLS, PATIENTS CORNEA IS CRYSTAL CLEAR. (B)(4). Manufacturer Narrative: B5 - POST-OP THE REPORTER STATED PATIENT OFFERED NO COMPLAINT OTHER THEN SLIGHTLY BLURRED VISION, WHICH WAS RESOLVED BY RECEIVING A SMALL PRESCRIPTION FOR PRE-EXISTING CORNEAL ASTIGMATISM. H6 - DEVICE CODE: DEVICE CODE: 3189 SHOULD BE CORRECTED TO 2993. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND ONEOF THE HAPTICS TORN. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9350566&pc=MTA | 2023826-2019-02188 | 2019-04-03 | Injury | STAAR SURGICAL COMPANY | 2019-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Unintended Movement | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH/DIOPTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. CLAIM#: (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS. -10.00 DIOPTER IN THE PATIENTS RIGHT EYE (OD) ON 06-MAR-2019. ELEVATED IOP; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; AND PIGMENT DISPERSION WAS NOTICED. ON 03-APR-2019 THE LENS WAS EXPLANTED. THE CAUSE WAS REPORTED AS USER ERROR; LENS IMPLANTED UPSIDE DOWN. A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. H6-PATIENT CODE 3191: PIGMENT DISPERSION . CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-DEVICE CODE (B)(6) IN PREVIOUS MDR IS NOT APPLICABLE. ADDITIONAL INFORMATION: H6- DEVICE CODE (B)(6) FOR LENS FLIPPED UPSIDE DOWN. CLAIM# 714950. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5549349&pc=MTA | 2023826-2016-00389 | 2014-04-24 | Injury | STAAR SURGICAL COMPANY | 2016-04-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 -13.50 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2013. EXCESSIVE VAULT WITH ELEVATED INTRAOCULAR PRESSURE RISE WAS NOTED ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 ANS NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: ON (B)(6) 2016, PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -13.50 (B)(4). LENS WORK ORDER SEARCH. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4578351&pc=MTA | 2023826-2015-00257 | 2014-09-26 | Injury | STAAR SURGICAL COMPANY | 2015-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO ELEVATED IOP. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW & WORK ORDER SEARCH. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS WAS RETURNED DRY AND NO VISIBLE DAMAGE WAS OBSERVED. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE LENGTH AND WIDTH WERE RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. A LENS WORK ORDER SEARCH WAS PERFORMED AND REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND. MEDICAL REVIEW - ELEVATED IOP AFTER ICL IMPLANTATION MAY OCCUR DUE TO: NON PATIENT/FUNCTIONING IRIDECTOMIES (TOO SMALL, TOO PERIPHERAL AND/OR NOT FULLY PERMEABLE, BLOCKED BY VISCO), REMAINING VISCOELASTIC IN THE POSTERIOR CHAMBER, OVERSIZE ICL WITH ANGLE CLOSURE, TOO NARROW ANGLES/CRODED AC DUE TO SMALL EYE ANATOMY OR TISSUE ABNORMALITIES (PRESENCE OF IRIS/CILIARY BODY CYSTS), MALPOSITIONED FOOTPLATES, UPSIDE DOWN ICL, ECT. THERE IS NOT ENOUGH INFORMATION TO DETERMINE THE EXACT CAUSE. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, WE WERE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4578988&pc=MTA | 2023826-2015-00226 | 2014-09-26 | Injury | STAAR SURGICAL COMPANY | 2015-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS NOT RETURNED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ELEVATED IOP EARLY AFTER ICL IMPLANTATION MAY OCCUR DUE TO: NON PATENT/FUNCTIONING IRIDECTOMIES (TOO SMALL, TOO PERIPHERAL AND/OR NOT FULLY PERMEABLE, BLOCKED BY VISCOELASTIC), REMAINING VISCOELASTIC IN THE POSTERIOR CHAMBER, OVERSIZED ICL WITH ANGLE CLOSURE, TOO NARROW ANGLES/CROWDED AC DUE TO SMALL EYE ANATOMY PREOPERATIVELY, UNEXPECTED ABNORMAL ANATOMY OR TISSUE ABNORMALITIES (PRESENCE OF IRIS/CILIARY BODY CYSTS), MALPOSITIONED FOOTPLATES, UPSIDE-DOWN ICL, ETC. PROMPT MANAGEMENT OF IOP RISE AND ITS CAUSE IS CRITICAL TO REDUCE OR ELIMINATE RISKS OF DAMAGE TO THE EYE AND VISION. GENERAL MANAGEMENT INCLUDES IOP-LOWERING DRUGS, BURPING THE INCISION, EVACUATING REMAINING FLUIDS FROM THE AC, ENLARGING THE EXISTING PI OR MAKING A NEW ONE, REMOVING OR REPLACING INADEQUATE SIZE LENSES, ETC. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4917149&pc=MTA | 2023826-2015-00839 | 2015-03-19 | Injury | STAAR SURGICAL COMPANY | 2015-07-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND ELEVATED IOP WAS OBSERVED POSTOPERATIVELY. A TRABECULETOMY WAS PERFORMED AND THE LENS WAS EXPLANTED A WEEK LATER. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. Manufacturer Narrative: (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5146056&pc=MTA | 2023826-2015-01247 | 2015-06-27 | Injury | STAAR SURGICAL COMPANY | 2015-10-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | PO30016 | NA | 1 | Event Description: THE REPORTER INDICATED THE PRIMARY LENS WAS EXCHANGED FOR THIS LENS. THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER IN THE PATIENT'S LEFT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (44MMHG). THE LENS WAS EXPLANTED AND NO OTHER LENS WAS IMPLANTED. POST-OP VISIT ON (B)(6) 2015 - BCVA WAS 20/25. SEE MFR # 2023826-2015-01246 FOR PRIMARY ICL. Manufacturer Narrative: DIOPTER: -13.50, DEVICE EVALUATED BY MANUFACTURER? NO -LENS NOT RETURNED. (B)(4). METHOD CODE(S): EVALUATION METHOD: WORK ORDER SEARCH. EVALUATION CODE(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS CODE(S)- (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5146356&pc=MTA | 2023826-2015-01245 | 2015-06-27 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | PO30016 | NA | 1 | Event Description: THE REPORTER INDICATED THE PRIMARY LENS WAS EXCHANGED FOR THIS LENS. THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (44MMHG). THE LENS WAS EXPLANTED AND NO OTHER LENS WAS IMPLANTED. POST-OP VISIT ON (B)(6) 2015 - BCVA WAS 20/25. SEE MFR # 2023826-2015-01244 FOR PRIMARY ICL. Manufacturer Narrative: DIOPTER: -13.50. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5309171&pc=MTA | 2023826-2015-01693 | 2015-12-03 | Injury | STAAR SURGICAL COMPANY | 2015-12-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE FACILITY REPORTED THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE, ANGLE CLOSURE, SHALLOWING OF THE ANTERIOR CHAMBER DEPTH AND PUPILLARY BLOCK. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/80. Manufacturer Narrative: RESULT: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE EXCESSIVE VAULT OF THE IMPLANTED LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY ADVERSE EVENT OR CONDITION WOULD HAVE CAUSED DAMAGE TO THE LENS. THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. PER MEDICAL REVIEW - REPORTEDLY, MICL (13.2 MM) WAS EXPLANTED ON THE FIRST POST-OP DAY TO ADDRESS ELEVATED IOP. NO LENS WAS IMPLANTED AS A REPLACEMENT. ACCORDING TO COMPLETED CUSTOMER COMPLAINT QUESTIONNAIRE FOR SURGERY, DATED ON (B)(6) 2016, THE ADVERSE EVENT WAS ATTRIBUTED TO DEVICE (EXCESSIVE VAULTING) THAT HAS CAUSED/CONTRIBUTED TO NARROWING OF THE ANGLE AND SUBSEQUENT IOP ELEVATION. IT SHOULD BE NOTED THAT AT THE TIME OF THE SURGERY PATIENT WAS 19 YEARS OLD AND PER DFU: "INDICATIONS: THE VISIAN ICL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE" AND THEREFORE, THERE IS NO SUFFICIENT DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6036776&pc=MTA | 2023826-2016-01471 | 2016-02-03 | Injury | STAAR SURGICAL COMPANY | 2016-10-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S LEFT EYE ON(B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THAT THE OPTIC WAS TORN/SPLIT, THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL WAS OBSERVED TO BE ON THE LENS SURFACE (BLOOD). UPON DIMENSIONAL INSPECTION, THE LENS WAS FOUND TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5579354&pc=MTA | 2023826-2016-00467 | 2016-03-11 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE REPORTER STATED AT ONE-DAY POST-OP, THE PATIENT RETURNED WITH CLOSED ANGLES AND HIGH PRESSURE (40MMHG). THE PERIPHERAL IRIDECTOMIES (PIS) WERE PATENT, BUT THE SURGEON PERFORMED 2 ADDITIONAL PIS AND DILATED THE PUPILS. A COUPLE OF HOURS LATER, THE VAULT WAS GOOD AND THE PRESSURE WAS IN THE 20S. TO DATE THE VAULT IS STILL A LITTLE HIGH BUT THE PATIENT IS DOING WELL. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL BE MONITORED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT CODE(S): (NO FAILURE DETECTED) : BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW, AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5579395&pc=MTA | 2023826-2016-00466 | 2016-03-11 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER STATED AT ONE-DAY POST-OP, THE PATIENT RETURNED WITH CLOSED ANGLES AND HIGH PRESSURE (40MMHG). THE PERIPHERAL IRIDECTOMIES (PIS) WERE PATENT, BUT THE SURGEON PERFORMED 2 ADDITIONAL PIS AND DILATED THE PUPILS. A COUPLE OF HOURS LATER, THE VAULT WAS GOOD AND THE PRESSURE WAS IN THE 20S. TO DATE THE VAULT IS STILL A LITTLE HIGH BUT THE PATIENT IS DOING WELL. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL BE MONITORED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: (B)(4). RESULT CODE(S): (OPERATIONAL PROBLEM) : A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW RESULT CODE(S): (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW, AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5221692&pc=MTA | 2023826-2015-01523 | NA | Injury | STAAR SURGICAL COMPANY | 2015-11-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT WAS BROUGHT BACK TO SURGERY DUE TO ELEVATED INTRAOCULAR PRESSURE AND HAD A SECONDARY PROCEDURE TO CREATE A NEW PERIPHERAL IRIDOTOMY BUILT IT WAS NOT SUCCESSFUL. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO ELEVATED INTRAOCULAR PRESSURE AND INTERMITTENT ANGLE CLOSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS (12.6MM) BUT THE PROBLEM WAS NOT RESOLVED. THE PATIENT IS STILL HAVING ISSUES WITH ELEVATED INTRAOCULAR PRESSURE AND THE PLAN IS TO EXPLANT THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): PER MEDICAL REVIEW: REPORTEDLY, VISIAN ICL (13.2 MM) WAS EXPLANTED/EXCHANGED TWO MONTHS POSTOPERATIVELY FOR A SHORTER LENGTH ICL (12.6 MM) TO ADDRESS VAULTING AND SUBSEQUENT ELEVATED IOP AND INTERMITTENT ANGLE CLOSURE. TWO DAYS PRIOR TO LENS EXCHANGE THE SURGEON PERFORMED ANOTHER IRIDOTOMY. ACCORDING TO REPORT FROM THE FACILITY, DATED ON (B)(6) 2015, ELEVATED IOP PERSISTED DESPITE INTERVENTION AND THE SURGEON WAS PLANNING TO EXPLANT/EXCHANGE ICL (12.6 MM) FOR A SHORTER ICL (12.1 MM). NO FURTHER INFORMATION WAS RECEIVED TO DATE. BASED ON THE RESULTS OF THE DHR REVIEW, THE AVAILABLE INFORMATION GIVEN WITHIN THIS COMPLAINT, AND WORK ORDER SEARCH, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO ROOT CAUSE IS DETERMINED. ADDITIONAL INFORMATION HAS BEEN REQUESTED FROM THE CUSTOMER, HOWEVER, NONE HAS BEEN FORTHCOMING. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: RESULT: (OPERATIONAL PROBLEM) : VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN, WITH A PIECE TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. EVENT DATE: UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5596422&pc=MTA | 2023826-2016-00508 | NA | Injury | STAAR SURGICAL COMPANY | 2016-04-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A PRESSURE SPIKE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS IMPLANTED ON (B)(6) 2016 AND REMOVED FROM THE LEFT EYE (OS) ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED NARROWING OF THE ANGLE, SHALLOWING OF THE ANTERIOR CHAMBER DEPTH AND A LITTLE PAIN. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/15. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS NOT THE DEVICE BUT WAS DUE TO LENS SIZE SELECTION (TOO LARGE). Manufacturer Narrative: (B)(4). Manufacturer Narrative: IMPLANTED DATE - (B)(6) 2016. EXPLANTED DATE - (B)(6) 2016. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5686518&pc=MTA | 2023826-2016-00770 | NA | Injury | STAAR SURGICAL COMPANY | 2016-05-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED MICL IMPLANTABLE COLLAMER LENSES IN BOTH EYES AND THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. THERE WAS AN ANTERIOR CHAMBER REACTION IN THE LEFT EYE AND THE PUPIL IS PERMANENTLY DILATED. THE PATIENT WAS ADMINISTERED MEDICATION FOR THE PRESSURE. THE REPORTER INDICATED TO DATE, THE LEFT EYE IS STILL DILATED BUT THE PRESSURE HAS IMPROVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR REPORT # 2023826-2016-00771 FOR THE OTHER EYE. Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, WAS IMPLANTED IN THE RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT'S VISUAL ACUITY WAS 20/15 -2 AND HAS BEEN REFERRED TO A GLAUCOMA SPECIALIST. Manufacturer Narrative: EXPIRATION DATE - UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5686526&pc=MTA | 2023826-2016-00771 | NA | Injury | STAAR SURGICAL COMPANY | 2016-05-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED MICL IMPLANTABLE COLLAMER LENSES IN BOTH EYES AND THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. THERE WAS AN ANTERIOR CHAMBER REACTION IN THE LEFT EYE AND THE PUPIL IS PERMANENTLY DILATED. THE PATIENT WAS ADMINISTERED MEDICATION FOR THE PRESSURE. THE REPORTER INDICATED TO DATE, THE LEFT EYE IS STILL DILATED BUT THE PRESSURE HAS IMPROVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR REPORT # 2023826-2016-00770 FOR THE OTHER EYE.Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, WAS IMPLANTED IN THE LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT'S VISUAL ACUITY WAS 20/20 AND HAS BEEN REFERRED TO A GLAUCOMA SPECIALIST. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4) PUPIL DILATED. IMPLANTED. (B)(4). THE LENS WAS NOT RETURNED. Manufacturer Narrative: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6498299&pc=MTA | 2023826-2017-00548 | 2016-05-21 | Injury | STAAR SURGICAL COMPANY | 2017-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -16.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXPLANTED DUE TO ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). AFTER EXPLANT, IOP RETURNED TO NORMAL LEVEL. TO DATE, THE SURGEON DOES NOT PLAN TO EXCHANGE THE LENS. THE SURGEON STATED THAT THE PATIENT MAY NOT BE SUITABLE FOR LENS IMPLANTATION. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: CONCLUSION CODE: THE SURGEON STATED THAT THE PATIENT MIGHT NOT HAVE BEEN A SUITABLE CANDIDATE FOR THE LENS. PRE-OPERATIVE INTRAOCULAR PRESSURE WAS VERY HIGH (40 MMHG). (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PRODUCT EVALUATION-FOUND THAT THE LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC SPLIT AND THE HAPTIC TORN, BROKEN, BENT, AND DEFORMED. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6505320&pc=MTA | 2023826-2017-00549 | 2016-05-21 | Injury | STAAR SURGICAL COMPANY | 2017-04-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXPLANTED DUE TO ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). THE SURGEON STATED THAT EVEN THOUGH THE VAULT VALUE WAS NORMAL AFTER IMPLANT, THE IOP HAD A HIGH VALUE. AFTER EXPLANT, IOP RETURNED TO NORMAL LEVEL. TO DATE, THE SURGEON DOES NOT PLAN TO EXCHANGE THE LENS. THE SURGEON STATED THAT THE PATIENT MAY NOT BE SUITABLE FOR LENS IMPLANTATION. Manufacturer Narrative: CORRECTED DATA: CONCLUSION CODE:THE SURGEON STATED THAT THE PATIENT MIGHT NOT HAVE BEEN A SUITABLE CANDIDATE FOR THE LENS. PRE-OPERATIVE INTRAOCULAR PRESSURE WAS VERY HIGH (40 MMHG). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION-LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND DEFORMED AND THE OPTIC SPLIT. ONLY A PORTION OF THE LENS WAS RETURNED, A LARGE PIECE WAS MISSING. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6801085&pc=MTA | 2023826-2017-01285 | 2016-07-25 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTED INDICATED THAT A VICMO12.1 IMPLANTABLE COLLAMER LENS WITH A -9.50 DIOPTER WAS IMPLANTED ON (B)(6)2017 IN THE LEFT EYE (OS). THE PATIENT WAS (B)(6) YEARS OLD AT THE TIME (THIS MODEL IS CONTRAINDICATED FOR USE IN PATIENTS UNDER THE AGE OF 21) THE PATIENT'S IOP ELEVATED WITHOUT PUPIL BLOCK OR ANGLE CLOSURE, SO THE SURGEON EXPLANTED THE LENS ON (B)(6)2017. THIS IS THE SAME PATIENT AS MDR # 2023826-2017-01286. Manufacturer Narrative: LENS WAS RETURNED WITH SOME MOISTURE IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON THE LENS. (B)(4). Manufacturer Narrative: INITIAL LENS WAS IMPLANTED ON (B)(6) 2016. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: THIS DEVICE IS CONTRAINDICATED FOR USE IN PATIENTS UNDER 21 YEARS ((B)(6) YEAR OLD PATIENT). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6801118&pc=MTA | 2023826-2017-01286 | 2016-07-25 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTED INDICATED THAT A VICMO12.1 IMPLANTABLE COLLAMER LENS WITH A -9.50 DIOPTER WAS IMPLANTED ON (B)(6)-2017 IN THE RIGHT EYE (OD). THE PATIENT WAS (B)(6) YEARS OLD AT THE TIME (THIS MODEL IS CONTRAINDICATED FOR USE IN PATIENTS UNDER THE AGE OF 21) THE PATIENT'S IOP ELEVATED WITHOUT PUPIL BLOCK OR ANGLE CLOSURE, SO THE SURGEON EXPLANTED THE LENS ON (B)(6)2017. THIS IS THE SAME PATIENT AS MDR # 2023826-2017-01285. Manufacturer Narrative: PRODUCT EVALUATION FOUND THE LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: LENS EXPLANT RESOLVED THE PROBLEM. PATIENT'S POST-OP BCVA WAS 20/20 ON (B)(6) 2017. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). CONCLUSION: THIS DEVICE IS CONTRAINDICATED FOR USE IN PATIENTS UNDER 21 YEARS ((B)(6) YEAR OLD PATIENT). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6447102&pc=MTA | 2023826-2017-00514 | 2016-12-30 | Injury | STAAR SURGICAL COMPANY | 2017-03-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. AFTER EXPLANT, IOP RETURNED TO NORMAL LEVEL. TO DATE, THE SURGEON DOES NOT PLAN TO EXCHANGE THE LENS. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6430473&pc=MTA | 2023826-2017-00478 | 2017-02-23 | Injury | STAAR SURGICAL COMPANY | 2017-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THAT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, WAS EXPLANTED DUE TO EXTREME HIGH PRESSURE. THE LENS WAS REPLACED ON (B)(6) 2017 WITH A SHORTER LENGTH LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THAT THE PATIENT IS DOING VERY WELL AFTER ICL EXCHANGE. THE LENS EXCHANGE FOR A SHORTER LENS RESOLVED THE PROBLEM. PATIENT CURRENT BSCVA IS 20/20. (B)(4). Manufacturer Narrative: LENS WAS RETURNED DRY, IN A POUCH. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLEAR RESIDUE AND LENT WAS OBSERVED ON THE LENS SURFACE. (B)(4). SECONDARY SURGICAL INTERVENTION, EXCHANGE. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6524681&pc=MTA | 2023826-2017-00625 | 2017-03-31 | Injury | STAAR SURGICAL COMPANY | 2017-04-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THAT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, WAS EXPLANTED AND REPLACED WITH A SHORTER LENGTH LENS WITHIN THE SAME SURGERY BECAUSE IT WAS CAUSING PRESSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2017 DUE TO ELEVATED INTRAOCULAR PRESSURE, EXCESSIVE VAULT, NARROWING OF THE ANGLE, AND SHALLOWING OF THE ACD. ADDITIONAL SURGICAL INTERVENTION ( LASER PERIPHERAL IRIDECTOMY) AND MEDICAL INTERVENTION (DIAMOX) WERE PERFORMED/PRESCRIBED. THE PATIENTS POST-OP UNCORRECTED VISUAL ACUITY WAS 20/50 RESIDUAL ASTIGMATISM AT 1 DAY POST -OP OF ICL EXCHANGE. THE REPORTER STATED THAT THE EVENT WAS PRODUCT-RELATED. (B)(4). IMPLANTED DATE WAS CORRECTED TO (B)(6) 2017 , LENS EXCHANGE WAS A SECONDARY SURGERY. CLAIM# (B)(4). Manufacturer Narrative: PREVIOUS IMPLANT DATE, (B)(6) 2017, IS INCORRECT. THE CORRECT IMPLANT DATE IS (B)(6) 2017. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2017 DUE TO ELEVATED INTRAOCULAR PRESSURE, EXCESSIVE VAULT, NARROWING OF THE ANGLE, AND SHALLOWING OF THE ACD. ADDITIONAL SURGICAL INTERVENTION ( LASER PERIPHERAL IRIDECTOMY) AND MEDICAL INTERVENTION (DIAMOX) WERE PERFORMED/PRESCRIBED. THE PATIENTS POST-OP UNCORRECTED VISUAL ACUITY WAS 20/50 RESIDUAL ASTIGMATISM AT 1 DAY POST -OP OF ICL EXCHANGE. THE REPORTER STATED THAT THE EVENT WAS PRODUCT-RELATED. (B)(4). IMPLANTED DATE WAS CORRECTED TO (B)(6) 2017 , LENS EXCHANGE WAS A SECONDARY SURGERY. CLAIM# (B)(4). | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6812325&pc=MTA | 2023826-2017-01339 | 2017-07-04 | Injury | STAAR SURGICAL COMPANY | 2017-08-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2017 THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -7.0, INTO THE PATIENT'S LEFT (OS) EYE. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER AND SIMILAR DIOPTER LENS DUE TO ELEVATED IOP. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DATE RECEIVED BY MANUFACTUTRER IN THE INTIAL SHOULD BE (B)(4) 2017. (B)(4). Manufacturer Narrative: DATE OF EVENT: IT IS CORRECTED TO "UNK" FROM "(B)(6) 2017". THE REPORTER STATED THAT THE ADVERSE EVENT (ELEVATED IOP) WAS NOTICED 2-3 WEEKS BEFORE THE EXPLANT DATE; (B)(6) 2017. NO DEFINITE DATE WAS PROVIDED. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050888&pc=MTA | 2023826-2017-01706 | 2017-09-21 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.2, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR ANOTHER LENS DUE HIGH PRESSURE AND FIXED DILATED PUPIL. THE BAG WAS ALSO NEEDED TO BE REFORMED. THE PATIENT'S CURRENT CONDITION IS THAT THE EYE HAS A BIT OF EDEMA BUT IS OVERALL FINE WITH THE SURGERY. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A LIQUID INSIDE A LENS CASE/ VIAL. VISUAL INSPECTION FOUND A PIECE OF LENS HAPTIC MISSING. METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. (B)(4). Manufacturer Narrative: PATIENT CODE: (B)(4) FIXED DILATED PUPIL, SECONDARY SURGERY, LENS EXCHANGE. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7356133&pc=MTA | 2023826-2018-00438 | 2018-01-29 | Injury | STAAR SURGICAL COMPANY | 2018-03-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S RIGHT EYE (OD). REPORTEDLY, "THERE IS AN IOP PROBLEM" AND THE PATIENT "IS ON MEDICATION FOR IT." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: REPORTEDLY, THE ICL THAT THE SURGEON IMPLANTED WAS A 13.2MM VICM5_13.2 ICL, DIOPTER -13.5. THE PATIENT WAS GIVEN MEDICATION ("DROPS") AND THE IOP WENT FROM 47.2MMHG ON (B)(6) 2018 TO 20.8MMHG ON (B)(6) 2018 "WITH 3 DROPS," AND THEN TO 12.4MMHG ON (B)(6) 2018 "WITH NO DROPS." (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6224163&pc=MTA | 2023826-2017-00004 | 2016-11-10 | Injury | STAAR SURGICAL COMPANY | 2017-01-04 | MTA | INTRAOCULAR LENS | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). AS OF THE DATE OF MDR REPORTING, THE LENS HAS NOT BEEN RETURNED AND THEREFORE NO PRODUCT EVALUATION HAS OCCURRED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED. | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978364&pc=MTA | 2023826-2017-01627 | 2017-02-03 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO ELEVATED INTRAOCULAR PRESSURE. THE PATIENT WAS PRESCRIBED MEDICATION FOR THE PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. USER FACILITY IS INCORRECT, SHOULD BE DISTRIBUTOR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978385&pc=MTA | 2023826-2017-01628 | 2017-02-03 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO ELEVATED INTRAOCULAR PRESSURE. A SUTURE WAS REQUIRED. THE PATIENT WAS PRESCRIBED MEDICATION FOR THE PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND. (B)(4). Manufacturer Narrative: REPORT SOURCE: USER FACILITY IS INCORRECT; SHOULD BE DISTRIBUTOR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5975346&pc=MTA | 2023826-2016-01316 | 2016-07-19 | Injury | STAAR SURGICAL COMPANY | 2016-09-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: - ON (B)(6) 2016, THE PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5724414&pc=MTA | 2023826-2016-00793 | NA | Injury | STAAR SURGICAL COMPANY | 2016-06-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED IOP . THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THE HAPTIC TO BE BROKEN AND BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873701&pc=MTA | 2023826-2016-01061 | 2016-05-16 | Injury | STAAR SURGICAL COMPANY | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Insufficient Information | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS UNABLE TO IMPLANT A A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, INTO THE SULCUS OF THE PATIENT'S RIGHT EYE (OD) DUE TO POSITIVE VITREO PRESSURE. THE LENS WAS INSERTED BUT THE PRESSURE DISPLACED THE LENS OUT OF THE PROPER POSITION IN THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY AND THERE WAS NO APPARENT INJURY. Manufacturer Narrative: DEVICE EVALUATION: EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN AND FIBERS ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6197215&pc=MTA | 2023826-2016-01805 | 2016-01-28 | Injury | STAAR SURGICAL COMPANY | 2016-12-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD CODE(S): NO TESTING METHOD PERFORMED, CLAIM EVALUATED BASED ON REPORTER RESPONSE TO COMPLAINT QUESTIONNAIRE. CONCLUSION CODE: DFU REQUIRES PATIENT'S ANTERIOR CHAMBER DEPTH (ACD) TO BE EQUAL OR GREATER THAN 3.0 MM. PATIENT (ACD) MEASURED AT 2.94 MM. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6036767&pc=MTA | 2023826-2016-01474 | 2016-02-03 | Injury | STAAR SURGICAL COMPANY | 2016-10-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: AS OF (B)(6) 2016, PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CONCLUSION: AS PER DFU FOR THIS LENS MODEL, VICLS ARE CONTRAINDICATED FOR PATIENTS UNDER 21 YEARS OF AGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6795677&pc=MTA | 2023826-2017-01319 | 2016-04-28 | Injury | STAAR SURGICAL COMPANY AG | 2017-08-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -3.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT EXPERIENCED EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION AND NARROWING OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A SHORTER LENS, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). OFF LABEL USE (THIS LENS MODEL IS CONTRADICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM). (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6209910&pc=MTA | 2023826-2016-01829 | 2016-06-20 | Injury | STAAR SURGICAL COMPANY | 2016-12-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXCHANGED DUE TO EXCESSIVE VAULT. Manufacturer Narrative: THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE PATIENT UNDERWENT YAG AND THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THE CORRECT DIOPTER REPORTED FOR VTICMO12.6 IMPLANTABLE COLLAMER LENS WAS -10.00/+1.5/009 DIOPTER. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED. | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6410776&pc=MTA | 2023826-2016-01644 | 2016-08-05 | Injury | STAAR SURGICAL COMPANY | 2017-03-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE (WITH ELEVATED IOP) AND HIGH OCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5983954&pc=MTA | 2023826-2016-01325 | 2016-08-31 | Injury | STAAR SURGICAL COMPANY | 2016-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. AT A POST-OP VISIT ON (B)(6) 2016, THE SURGEON NOTED AN INCREASE IN THE OCULAR PRESSURE AND LENS WAS CLOUDY. THE SURGEON ATTEMPTED TO RE-POSITION THE LENS BUT WITH NO SUCCESS. THE LENS WAS EXPLANTED AND A SHORTER LENS WAS IMPLANTED. THERE WERE NO COMPLICATIONS. Manufacturer Narrative: CORRECTION - THE REPORTER INDICATED THERE WAS NO INCREASE IN INTRAOCULAR PRESSURE. ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2016 FROM THE PATIENT'S RIGHT EYE (OD). THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE PATIENT HAD A DISTORTED PUPIL DUE TO CROWDED ANGLE. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE PATIENT IS DOING EXCELLENT AND THE POST-OP VISUAL ACUITY WAS 20/25. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6460525&pc=MTA | 2023826-2017-00570 | 2016-10-01 | Injury | STAAR SURGICAL COMPANY | 2017-04-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION-LENS WAS RETURNED WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6460543&pc=MTA | 2023826-2017-00571 | 2016-10-01 | Injury | STAAR SURGICAL COMPANY | 2017-04-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP. ON (B)(6) 2016, THE LENS WAS EXCHANGED FOR SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6224164&pc=MTA | 2023826-2017-00005 | 2016-10-15 | Injury | STAAR SURGICAL COMPANY | 2017-01-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.50 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT REPORTED EXCESSIVE VAULTING AND ELEVATED IOP. ON (B)(6) 2016, THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION NO LENS EVALUATION HAS BEEN PERFORMED, AS THE PRODUCT HAS NOT BEEN RETURNED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6361169&pc=MTA | 2023826-2017-00266 | 2016-10-27 | Injury | STAAR SURGICAL COMPANY | 2017-02-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -8.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED IOP. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2016, THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: LENS RETURNED DRY IN A SMALL CENTRIFUGE VIAL WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND DRY CLEAR RESIDUE AND BLACK DEBRIS ON THE LENS AND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PATIENT: SECONDARY SURGICAL INTERVENTION, LENS WAS EXCHANGED FOR SHORTER LENS. CLAIM #(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6172497&pc=MTA | 2023826-2016-01758 | 2016-11-05 | Injury | STAAR SURGICAL COMPANY | 2016-12-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ANGLE CLOSURE. THE PERIPHERAL IRIDOTOMIES WERE ENLARGED. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: PATIENT INFO: FEMALE. EXPLANT DATE: (B)(6) 2016. EVENT DATE: UNK. ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE CAUSE OF THE EVENT WAS UNKNOWN. PATIENT'S CURRENT CONDITION IS BCVA 20/20 +2 AND VAULT IS 3/4 CT. RESULT (OPERATIONAL PROBLEM) : VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE MISSING FROM THE LENS OPTIC. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6172498&pc=MTA | 2023826-2016-01757 | 2016-11-05 | Injury | STAAR SURGICAL COMPANY | 2016-12-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -4.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ANGLE CLOSURE. THE PERIPHERAL IRIDOTOMIES WERE ENLARGED. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: PATIENT INFO: FEMALE. EXPLANT DATE: (B)(6) 2016. EVENT DATE: UNK. ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE CAUSE OF THE EVENT WAS UNKNOWN. PATIENT'S CURRENT CONDITION IS BCVA 20/20 +2 AND VAULT IS 3/4 CT. RESULT: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6191984&pc=MTA | 2023826-2016-01801 | 2016-11-07 | Injury | STAAR SURGICAL COMPANY | 2016-12-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -3.50 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. SECONDARY SURGICAL INTERVENTION, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, LENS EXPLANTED. ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6281933&pc=MTA | 2023826-2017-00174 | 2017-01-06 | Injury | STAAR SURGICAL COMPANY | 2017-01-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED EVO VISIAN ICL LENSES BILATERALLY IN A PATIENT AND BOTH EYES PRESENT HIGH VAULTING. THE LENSES REMAIN IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE INCISION WAS ENLARGED AND ADDITIONAL SUTURE WAS REQUIRED. THE LENS EXCHANGED RESOLVED THE PROBLEM. THE PATIENTS POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6407568&pc=MTA | 2023826-2017-00412 | 2017-01-16 | Injury | STAAR SURGICAL COMPANY | 2017-03-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED LOW VAULT, AND ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION-LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6312700&pc=MTA | 2023826-2017-00228 | 2017-01-17 | Injury | STAAR SURGICAL COMPANY | 2017-02-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A MICL12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE REPORTER STATED THE CAUSE WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING. WORK ORDER SEARCH: NO SIMILAR CLAIM WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6439068&pc=MTA | 2023826-2017-00506 | 2017-01-23 | Injury | STAAR SURGICAL COMPANY | 2017-03-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR/BLACK RESIDUE ON THE PRODUCT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6424135&pc=MTA | 2023826-2017-00454 | 2017-01-25 | Injury | STAAR SURGICAL COMPANY | 2017-03-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED LOW VAULT, ELEVATED IOP, AND HAD ADDITIONAL REFRACTIVE TREATMENT. THE LENS WAS EXPLANTED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM12.6 IMPLANTABLE COLLAMER LENS; -5.0 DIOPTER; IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED LOW VAULTING, ELEVATED IOP, AND HAD ADDITIONAL REFRACTIVE TREATMENT. THE LENS WAS EXPLANTED. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM12.6 IMPLANTABLE COLLAMER LENS; -5.0 DIOPTER; IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED LOW VAULTING, ELEVATED IOP, AND HAD ADDITIONAL REFRACTIVE TREATMENT. THE LENS WAS EXPLANTED. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM12.6 IMPLANTABLE COLLAMER LENS; -5.0 DIOPTER; IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED LOW VAULTING, ELEVATED IOP, AND HAD ADDITIONAL REFRACTIVE TREATMENT. THE LENS WAS EXPLANTED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6389631&pc=MTA | 2023826-2017-00348 | 2017-02-07 | Injury | STAAR SURGICAL COMPANY | 2017-03-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THAT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, WAS IMPLANTED ON (B)(6) 2017 AND EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE REPORTER INDICATED THAT THERE WAS NO PATIENT INJURY AND THE EVENT WAS NOT DEVICE RELATED BUT THE EXACT CAUSE WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED STUCK INSIDE CARTRIDGE. CARTRIDGE WAS RETURNED TAPED TO LAF CARD. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE DEVICE, WHITE/GREY DRY RESIDUE WAS OBSERVED ON CARTRIDGE, AND A STUCK LENS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6527418&pc=MTA | 2023826-2017-00701 | 2017-03-29 | Injury | STAAR SURGICAL COMPANY | 2017-04-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -14 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ACCORDING TO THE SURGEON THE PATIENT EXPERIENCED EXCESSIVE VAULTING BECAUSE THE LENS WAS IMPLANTED IN A "PSEUDO PHAKIC EYE". THE PATIENT ALSO EXPERIENCED ELEVATED IOP. THE SHORTER LENS WAS ALREADY ORDERED AND SCHEDULED TO EXCHANGE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS IMPLANTED ON PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION AND CORRECTED: THE LENS WAS IMPLANTED ON PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS, AND THE PROBLEM WAS RESOLVED. ADDITIONAL PATIENT CODE (B)(4) (NEW INCISION): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CONCLUSION: OFF LABEL USE (BELOW 21 YEARS OF AGE AT THE TIME OF IMPLANT, LENS IMPLANT IN A PSEUDO PHAKIC EYE). (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6579894&pc=MTA | 2023826-2017-00860 | 2017-04-11 | Injury | STAAR SURGICAL COMPANY | 2017-05-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, DIOPTER -18.00 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO THE PATIENT EXPERIENCING EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/30 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/30. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6609924&pc=MTA | 2023826-2017-00921 | 2017-04-11 | Injury | STAAR SURGICAL COMPANY | 2017-06-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER AND SIMILAR DIOPTER LENS, DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND PIGMENT DISPERSION. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6565217&pc=MTA | 2023826-2017-00812 | 2017-04-14 | Injury | STAAR SURGICAL COMPANY | 2017-05-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, CLOSED ANGLE AND SHALLOW ANTERIOR CHAMBER. THE PRESSURE WAS 34MMHG BUT WAS BROUGHT DOWN TO NORMAL WITH DIAMOX. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN THE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6565218&pc=MTA | 2023826-2017-00811 | 2017-04-14 | Injury | STAAR SURGICAL COMPANY | 2017-05-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, CLOSED ANGLE AND SHALLOW ANTERIOR CHAMBER. THE PRESSURE WAS 34MMHG BUT WAS BROUGHT DOWN TO NORMAL WITH DIAMOX. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN THE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6645493&pc=MTA | 2023826-2017-00983 | 2017-05-16 | Injury | STAAR SURGICAL COMPANY | 2017-06-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER INDICATED AT THE INITIAL SURGERY THE PATIENT HAD SLIGHT IRIS PROLAPSE, WHICH WAS MANAGED AT THE TIME. AT ONE DAY POST-OP, THE PATIENT'S INTRAOCULAR PRESSURE WAS HIGH, WITH SHALLOW CHAMBER AND ELEVATED VAULT. THE PATIENT HAD RECURRING IRIS PROLAPSE. THE PATIENT RETURNED TO SURGERY AND THE IRIS PROLAPSE WAS MANAGED. BUT THE SURGEON DECIDED TO EXPLANT THE LENS, WHICH WAS REMOVED ON (B)(6) 2017 AND EXCHANGED FOR A SHORTER LENS. THE PATIENT IS DOING WELL AND THE LENS VAULT IS NORMAL. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6712101&pc=MTA | 2023826-2017-01186 | 2017-06-06 | Injury | STAAR SURGICAL COMPANY | 2017-07-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. THERE ARE NO PLANS TO EXCHANGE THE LENS. Manufacturer Narrative: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6728124&pc=MTA | 2023826-2017-01213 | 2017-06-20 | Injury | STAAR SURGICAL COMPANY | 2017-07-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON EXPLANTED A 13.2 MM MICL13.2 IMPLANTABLE COLLAMER LENS ,-5.0 DIOPTER, THAT WAS IMPLANTED LAST YEAR AND EXCHANGED IT FOR A SHORTER LENGTH LENS. THE REPORTER STATED THAT THE LENS WAS BEING REPLACED DUE TO IT BEING BIG, PROGRESSIVE ANTERIOR SYNECHIAE FROM HIGH VAULT, NARROWING OF THE ANGLE, ANGLE CLOSURE, AND SHALLOWING OF THE ACD. ADDITIONAL SURGICAL INTERVENTION (LASER IRIDOTOMY) WAS PERFORMED. THE PATIENTS POST- OP VA WAS 20/20. THE REPORTER INDICATED THAT A SURGICAL SYNECHIOLYSIS WAS PERFORMED AT THE SAME TIME OF THE LENS EXCHANGE. Manufacturer Narrative: THE LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT ALSO EXPERIENCED HIGH IOP. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RECEIVED DRY, IN SPECIMEN CONTAINER. VISUAL INSPECTION FOUND A TEAR ON A HAPTIC, DEBRIS, RED,CLEAR RESIDUE ON LENS SURFACE, AND PORTION OF A HAPTIC MISSING. (B)(4). WORK ORDER SEARCH: N O SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6767795&pc=MTA | 2023826-2017-01250 | 2017-07-06 | Injury | STAAR SURGICAL COMPANY | 2017-08-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER. IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS HAD A HIGH VAULT AND DUE TO CONCERNS OVER ANGLE CLOSURE, THE LENS WAS EXPLANTED ON (B)(6) 2017. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE AND NARROWING OF THE ANGLE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S CURRENT CONDITION WAS 20/20 UCVA, WITH GREAT INTRAOCULAR PRESSURE. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. (B)(4). Manufacturer Narrative: THE LENS WAS EXCHANGED FOR A SHORTER LENS AND UPON IMPLANTATION THERE WAS A RESIDUAL VAULTING. THE SURGEON ENLARGED PI AND RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6890291&pc=MTA | 2023826-2017-01469 | 2017-07-13 | Injury | STAAR SURGICAL COMPANY | 2017-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, -15.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON THE MORNING OF (B)(6) 2017. EXCESSIVE VAULTING AND HIGH IOP WAS NOTICED AND WAS EXPLANTED LATER THAT EVENING. THE PATIENT'S POST-OP BCVA WAS 20/40. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND BENT. DIMENSIONAL EVALUATION FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6818689&pc=MTA | 2023826-2017-01323 | 2017-07-26 | Injury | STAAR SURGICAL COMPANY | 2017-08-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.7, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO INCORRECT SIZING. THE LENS WAS EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017 DUE TO EXCESSIVE VAULT, ELEVATED (INTRAOCULAR PRESSURE) IOP, NARROWING OF ANGLES, ANGLE CLOSURE AND SHALLOWING OF ANTERIOR CHAMBER DEPTH (ACD). AN IRIDECTOMY WAS PERFORMED. THE LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6818692&pc=MTA | 2023826-2017-01324 | 2017-07-26 | Injury | STAAR SURGICAL COMPANY | 2017-08-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.7, -15.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO INCORRECT SIZING. THE LENS WAS EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: FEMALE. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ELEVATED IOP (INTRAOCULAR PRESSURE), NARROWING OF ANGLES, ANGLE CLOSURE AND SHALLOWING OF ACD (ANTERIOR CHAMBER DEPTH). AN IRIDECTOMY WAS PERFORMED. CORRECTED DATA: DATE OF IMPLANTATION WAS CORRECTED TO "(B)(6) 2017". THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL 13.7, -15.5 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. DEVICE EVALUATION: THE LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DEVICE MANUFACTURE DATE WAS CORRECTED TO "10-DEC-2015". (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7069472&pc=MTA | 2023826-2017-01802 | 2017-08-17 | Injury | STAAR SURGICAL COMPANY | 2017-11-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/10. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6873954&pc=MTA | 2023826-2017-01448 | 2017-08-18 | Injury | STAAR SURGICAL COMPANY | 2017-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGERY IMPLANTED A 13.7 MM VICMO13.7, -17.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. EXCESSIVE VAULTING, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS NOTICED. THE SURGEON DECIDED TO OBSERVE THE PATIENT AS THE PATIENT IS ASYMPTOMATIC AND IS HAPPY WITH VISUAL ACUITY. ON (B)(6) 2017, THE PATIENT'S BCVA IS 20/30, UCVA IS 20/40, AND IOP IS 15 MM HG WITHOUT TREATMENT. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT WITH PREVIOUS CORNEAL PATHOLOGY. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6920137&pc=MTA | 2023826-2017-01540 | 2017-08-24 | Injury | STAAR SURGICAL COMPANY | 2017-10-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP), NARROWING OF THE ANGLE, SHALLOWING OF THE ANTERIOR CHAMBER DEPTH (ACD) AND ACUTE ANGLE CLOSURE. THE PERIPHERAL IRIDECTOMIES WERE ENHANCED. THE LENS WAS EXCHANGED WITH A SHORTER LENS. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20 -2 AND THE PATIENT IS HAPPY. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE REPORTER STATED THE PERIPHERAL IRIDECTOMIES WERE "TAPPED/ENHANCED" DUE TO THE ELEVATED INTRAOCULAR PRESSURE BUT WAS UNSUCCESSFUL. THERE WERE ADDITIONAL PATIENT SYMPTOMS; PATIENT'S VISION WENT "WHITE". THE LENS EXCHANGE WAS PERFORMED ON (B)(6) 2017. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS FOUND. LENS WAS RETURNED IN LIQUID IN A CENTRIFUGE VIAL INSIDE A BIOHAZARD BAG. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A CENTRIFUGE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050133&pc=MTA | 2023826-2017-01769 | 2017-10-18 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7113646&pc=MTA | 2023826-2017-01885 | 2017-11-08 | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTED INDICATED THAT THE SURGEON IMPLANTED A MICL13.2,-16.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED (INTRAOCULAR PRESSURE) IOP, NARROWING OF THE ANGLE AND SHALLOWING OF ANTERIOR CHAMBER DEPTH (ACD). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PATIENT POST-OP BCVA IS 20/40. Manufacturer Narrative: SERIAL NUMBER - PREVIOUS SERIAL NUMBER WAS INCORRECT, (B)(4) IS THE CORRECT SERIAL NUMBER. (B)(4) Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7197904&pc=MTA | 2023826-2018-00068 | 2017-12-08 | Injury | STAAR SURGICAL COMPANY | 2018-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGERY IMPLANTED A MICL12.6, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO THE LENS BEING TOO BIG. THE LENS WAS EXCHANGED FOR SHORTER LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL TAPED TO PURCHASE RETURNED FORM. VISUAL INSPECTION FOUND THE HAPTIC TORN AND THE LENS MISSING A PIECE OF HAPTIC AND OPTIC. (B)(4) Manufacturer Narrative: ADDITIONAL DATA: DATE OF BIRTH: (B)(6). GENDER: FEMALE. DATE OF EVENT: (B)(6) 2017. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL12.6 IMPLANTABLE COLLAMER LENS OF -11.0 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP ASSESSED ON (B)(6) 2017. (B)(4) Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7281453&pc=MTA | 2023826-2018-00262 | 2017-12-22 | Injury | STAAR SURGICAL COMPANY | 2018-02-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RED AND CLEAR RESIDUE ON THE LENS SURFACE. CONCLUSION: THIS LENS IS CONTRAINDICATED FOR USE IN PATIENTS WITH ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7356136&pc=MTA | 2023826-2018-00431 | 2018-01-04 | Injury | STAAR SURGICAL COMPANY | 2018-03-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA).THE PROBLEM WAS NOT RESOLVED DUE TO THE REPLACEMENT LENS DID NOT UNFOLD PROPERLY AFTER IMPLANT. SEE MFR REPORT# 2023826-2018-00429 FOR CASE RESOLUTION. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (RED AND CLEAR RESIDUE ON LENS) AND THE LENS HAS A CURVED SHAPE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))] (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7324462&pc=MTA | 2023826-2018-00359 | 2018-01-18 | Injury | STAAR SURGICAL COMPANY | 2018-03-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT HE IMPLANTED A 13.2 MM MICL13.2, -4.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT RIGHT EYE (OD) ON (B)(6) 2018. ON POST OP DAY 1 EXCESSIVE VAULT, ELEVATED IOP AND ANGLE CLOSURE WAS DISCOVERED. THE PATIENT WAS PLACED ON COMBIGAN TO LOWER IOP. AT DAY 3 AND 5 THE ELEVATED VAULT PERSISTED WITH NARROWING OF THE ANGLE. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTED STATED THAT THE PATIENT IS HAPPY WITH GOOD CURRENT VISION. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND HAPTIC TEARS. CLAIM # (B)(4). Manufacturer Narrative: (B)(4) | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7324571&pc=MTA | 2023826-2018-00360 | 2018-01-18 | Injury | STAAR SURGICAL COMPANY | 2018-03-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT HE IMPLANTED A 13.2MM MICL13.2, -4.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT LEFT EYE (OS) ON (B)(6) 2018. ON POST OP DAY 1 EXCESSIVE VAULT, ELEVATED IOP AND ANGLE CLOSURE WAS DISCOVERED. THE PATIENT WAS PLACED ON COMBIGAN TO LOWER IOP. AT DAY 3 AND 5 THE ELEVATED VAULT PERSISTED WITH NARROWING OF THE ANGLE. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTED STATED THAT THE PATIENT IS HAPPY WITH GOOD CURRENT VISION. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND HAPTIC TEARS. THIS IS A RE-SUBMISSION TO CORRECT MFR # FOR SUPPLEMENTAL #1 IN WHICH THE MFR NUMBER WAS REPORTED INCORRECT. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7356143&pc=MTA | 2023826-2018-00434 | 2018-01-24 | Injury | STAAR SURGICAL COMPANY | 2018-03-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, ELEVATED IOP, AND "PROTUBERANT" IRIS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A TORN HAPTIC AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7558973&pc=MTA | 2023826-2018-00834 | 2018-02-28 | Injury | STAAR SURGICAL COMPANY | 2018-05-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED IN A 2ND SURGERY THE SAME DAY DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8172167&pc=MTA | 2023826-2018-01850 | 2018-05-19 | Injury | STAAR SURGICAL COMPANY | 2018-12-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS EXCESSIVE VAULT, PUPIL BLOCK, ELEVATED IOP, AND ENLARGEMENT OF OR ADDITION OF PI. IT'S REPORTED THAT ON (B)(6) 2018 THE LENS WAS EXPLANTED. IT IS UNKNOWN AS TO WHETHER A REPLACEMENT LENS WAS IMPLANTED DUE TO CONFLICTING INFORMATION REPORTED. ATTEMPTS TO OBTAIN CLARIFICATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7647578&pc=MTA | 2023826-2018-00984 | 2018-06-06 | Injury | STAAR SURGICAL COMPANY | 2018-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO IMPROPER FIT. ANOTHER LENS WAS NOT IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE AND ELEVATED INTRAOCULAR PRESSURE. THE INCISION WAS ENLARGED AND SUTURE WAS REQUIRED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. THE LENS WAS RETURNED DRY, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: ONE SIMILAR COMPLAINT TYPE EVENT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7683601&pc=MTA | 2023826-2018-01019 | 2018-06-09 | Injury | STAAR SURGICAL COMPANY | 2018-07-12 | MTA | MICL | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE SURGEON INDICATED THAT A12.6MM MICL12.6, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: ANGLE CLOSURE. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: ON (B)(6) 2018 THE LENS WAS EXPLANTED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING AND ANGLE CLOSURE WITH ELEVATED IOP. ANOTHER LENS WAS IMPLANTED ON THE SAME DAY. ON THE PATIENT'S LAST VISIT ON (B)(6) 2018, THE STATUS OF THE EYE WAS COMFORTABLE WITH POOR PUPIL CONSTRICTION AND MINIMAL HALO AND GLARE. (LENS EXCHANGE) TO (LENS EXPLANT) - PREVIOUS CODE IS NOT APPLICABLE, LENS WAS EXPLANTED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE VAULT WAS IMPROVED, BUT IT WAS STILL HIGH. IOP (INTRAOCULAR PRESSURE) IS WNL (WITHIN NORMAL LIMITS). IT IS BELIEVED THAT THE HIGH IOP (INTRAOCULAR PRESSURE) FROM PUPILARY BLOCK CAUSED THE CONSTRICTION ISSUE. PUPIL CONSTRICTION AND GLARE IMPROVED BUT ARE STILL PRESENT. (B)(4). Manufacturer Narrative: LENS WAS RETURNED DRY IN CARTRIDGE TRAY. VISUAL INSPECTION FOUND A TORN HAPTIC AND FOREIGN MATERIAL ON THE SURFACE OF THE LENS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7706773&pc=MTA | 2023826-2018-01067 | 2018-06-20 | Injury | STAAR SURGICAL COMPANY | 2018-07-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 12.6MM MICL12.6, -16.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED BY A SURGEON THAT WAS ON CALL DUE TO OBSERVATION OF EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PUPIL BLOCK AND ANGLE CLOSURE WITH ELEVATED IOP (INTRAOCULAR PRESSURE). ADDITIONAL SURGICAL INTERVENTION (ENLARGEMENT AND ADDITION OF NEW PI, YAG) WERE PERFORMED. ON (B)(6) 2018 AN EXCHANGED FOR A SHORTER LENGTH LENS WAS PERFORMED. THIS EXCHANGE RESOLVE THE PROBLEM. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7903843&pc=MTA | 2023826-2018-01372 | 2018-06-25 | Injury | STAAR SURGICAL COMPANY | 2018-09-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. EXCESSIVE VAULT WAS OBSERVED AND ELEVATED IOP (45MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE ASSESSED ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7954052&pc=MTA | 2023826-2018-01504 | 2018-09-06 | Injury | STAAR SURGICAL COMPANY | 2018-10-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -15.5 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. AT ONE DAY POST-OP, THE LENS WAS TRYING TO PROTRUDE THROUGH THE PUPIL AND WAS KINKED (TOO BIG). THE PUPIL WOULD NOT CONSTRICT. THE LENS WAS EXPLANTED ON (B)(6) 2018 BUT THE PUPIL REMAINED DILATED. THE INCISION WAS ENLARGED SLIGHTLY TO REMOVE THE LENS AND SUTURES WERE REQUIRED. ANOTHER LENS WAS NOT IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED AT 1 DAY POST-OP, SHOWED HIGH CENTRAL VAULT. THE INTRAOCULAR PRESSURE (IOP) WAS HIGH FOR FIRST FEW DAYS, YET 1 DAY POST-OP, THE EYE WAS SOFT WITH FLAT ANTERIOR CHAMBER. 2 DAY POST-OP, IOP 28MMHG. THE SURGEON FELT THE ORBSCAN WTW MAY HAVE BEEN MEASURED INCORRECTLY. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, TAPED TO A PLASTIC CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF RESIDUE ON THE LENS. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8176802&pc=MTA | 2023826-2018-01854 | 2018-09-18 | Injury | STAAR SURGICAL COMPANY | 2018-12-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -17.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) (30 MMHG) AND PUPIL DILATION. THE PATIENT WAS TREATED WITH MEDICATION. DUE TO THE PATIENTS SYMPTOMS, THE SURGEON PERFORMED A YAG IRIDOTOMY, WITH ADDITIONAL MEDICAL THERAPY, BUT IT WAS UNSUCCESSFUL. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. AFTER ICL EXCHANGE, THE IOP WAS VERY HIGH (>40MMHG) AND THE SURGEON SUSPECTED SENSITIVITY TO TOPICAL STEROID THERAPY. AFTER TAPERING OFF THE STEROID AND IOP-LOWERING THERAPIES, THE IOP COMPLETELY NORMALIZED. Manufacturer Narrative: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8048978&pc=MTA | 2023826-2018-01612 | 2018-10-10 | Injury | STAAR SURGICAL COMPANY | 2018-11-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO OBSERVATION OF EXCESSIVE VAULT AND ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND FIBER ON THE LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8113339&pc=MTA | 2023826-2018-01752 | 2018-10-25 | Injury | STAAR SURGICAL COMPANY | 2018-11-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, WHICH CAUSED AN INCREASE IN PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: PREVIOUS MDR STATED: "EXCESSIVE VAULT WHICH CAUSED AN INCREASE IN PRESSURE". CORRECTED TO : "EXCESSIVE VAULT ONLY" IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8217921&pc=MTA | 2023826-2019-00009 | 2018-11-16 | Injury | STAAR SURGICAL COMPANY | 2019-01-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: "TTHE" REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -8.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP (INTRAOCULAR PRESSURE) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS #1- TWO LENSES WERE RETURNED UNIDENTIFIED IN THE SAME BAG. LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS; PRESENCE OF RESIDUE ON LENS SURFACE. MDR # 2023826-2019-00010 FOR CROSS REFERENCE. LENS #2- TWO LENSES WERE RETURNED UNIDENTIFIED IN THE SAME BAG. LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A TORN HAPTIC AND PRESENCE OF RESIDUE ON LENS SURFACE. MDR # 2023826-2019-00009. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8218075&pc=MTA | 2023826-2019-00010 | 2018-11-16 | Injury | STAAR SURGICAL COMPANY | 2019-01-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: TTHE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP (INTRAOCULAR PRESSURE) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS #1- TWO LENSES RETURNED UNIDENTIFIED IN SAME BAG. LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A TORN HAPTIC; PRESENCE OF RESIDUE ON LENS SURFACE. MDR # 2023826-2019-00009 FOR CROSS REFERENCE. LENS #2- TWO LENSES RETURNED UNIDENTIFIED IN SAME BAG. LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. MDR# 2023826-2019-00010 FOR CROSS REFERENCE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8768409&pc=MTA | 2023826-2019-01184 | 2018-12-03 | Injury | STAAR SURGICAL COMPANY | 2019-07-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13 MM, VICMO13.2, -17.00 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT, ELEVATED IOP AND PIGMENT DISPERSION. A REPLACEMENT LENS WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. THE REPORTER ALSO REPORTER THAT WHEN THE LENS WAS REMOVED THEY NOTICED THAT ONE IF THE HAPTICS HAD A FISSURE (INCOMPLETE BROKE). THE REPORTER STATED " THE LENS WAS REMOVED DURING A SECOND SURGERY, SURGEON TRIED TO MAINTAIN THE IOP USING MEDICATION. BECAUSE THIS TREATMENT WASN'T EFFECTIVE DURING 4-5 MONTHS, FINALLY DECIDED TO REMOVE THE LENS DUE TO POTENTIAL DAMAGE OF RETINA. AN OCT-SCAN WAS PERFORMED PRIOR TO SECOND SURGERY AND RAISED QUESTION ABOUT ICL STABILITY DUE TO A SLIGHTLY ASYMMETRY, INITIALLY INTERPRETED AS A NASAL POSTERIOR SYNECHIAE. THE MOMENT OF SECOND SURGERY WAS POSTPONED FOR A WHILE DUE TO ATTEMPT TO CONTROL IOP WITH MEDICATION, AS REFRACTIVE OUTCOME WAS ACCEPTABLE AND PATIENT LIVE AND WORK ABROAD". Manufacturer Narrative: VISUAL INSPECTION ALSO FOUND RESIDUE ON THE LENS. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, IOP TREATED WITH MEDICATION CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8219260&pc=MTA | 2023826-2019-00044 | 2018-12-05 | Injury | STAAR SURGICAL COMPANY | 2019-01-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM, MICL13.7, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE ON (B)(6) 2018. ELEVATED IOP (INTRAOCULAR PRESSURE) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS OBSERVED, DATE ASSESSED: (B)(6) 2018. LENS REMAINS IMPLANTED, SURGEON HAD PLANNED TO EXCHANGE THE LENS ON (B)(6) 2018 FOR A SHORTER LENGTH LENS BUT THE SURGERY WAS POSTPONED BY THE PATIENT AND HAS NOT RESCHEDULED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM, MICL13.7, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE ON (B)(6) 2018. ELEVATED IOP (INTRAOCULAR PRESSURE) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS OBSERVED, DATE ASSESSED: (B)(6) 2018. THE LENS WAS REMOVED ON (B)(6) 2019 AND REPLACED WITH A SHORTER LENGTH LENS. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM AND THE PATIENT IS FINE. THE EXPLANTED LENS WAS DISCARDED AND WILL NOT BE RETURNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SECONDARY SURGICAL INTERVENTION, EXPLANT. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8356222&pc=MTA | 2023826-2019-00297 | 2018-12-14 | Injury | STAAR SURGICAL COMPANY | 2019-02-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A13.2MM, VICMO13.2, -14.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP (DATE ASSESSED: (B)(6) 2018). THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. PRESENCE OF RESIDUE ON LENS SURFACE WAS ALSO OBSERVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8656475&pc=MTA | 2023826-2019-00970 | 2018-12-23 | Injury | STAAR SURGICAL COMPANY | 2019-05-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND A PIECE OF LENS OPTIC AND HAPTIC TORN/BROKEN OFF. UNABLE TO PERFORM DIMENSIONAL INSPECTION SINCE HAPTIC/OPTIC MISSING. (B)(4). DATE OF EVENT: (B)(6) 2018. THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED ON (B)(6) 2019 FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. IMPLANT DATE: (B)(6) 2018. EXPLANT DATE: (B)(6) 2018. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372544&pc=MTA | 2023826-2019-00330 | 2019-01-30 | Injury | STAAR SURGICAL COMPANY | 2019-02-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER STATES ADDITIONAL PATIENT SYMPTOMS OF ANGLE CLOSURE AND ELEVATED IOP AND ALSO NOTES "ANISOCORIA ON OS." REPORTEDLY, ANISOCORIA WAS NOTICED AFTER LENS EXPLANT. BEFORE SURGERY, "PUPIL SIZE WAS FINE." IT IS UNKNOWN AS TO WHETHER THE LENS CAUSED THIS OR NOT. DISTRIBUTOR WAS CHECKED IN ERROR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8633167&pc=MTA | 2023826-2019-00877 | 2019-02-06 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND UNREACTIVE (FIXED) PUPIL. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MFR: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10093969&pc=MTA | 2023826-2020-01120 | 2019-02-06 | Injury | STAAR SURGICAL COMPANY | 2020-05-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8471422&pc=MTA | 2023826-2019-00458 | 2019-03-01 | Injury | STAAR SURGICAL COMPANY | 2019-04-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6, -9.5 DIOPTER , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNELA ANGLES, AND ELEVATED IOP (24 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A BIOHAZARD ZIPLOC BAG, IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8488607&pc=MTA | 2023826-2019-00541 | 2019-03-06 | Injury | STAAR SURGICAL COMPANY | 2019-04-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -12.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE. ON (B)(6) 2019 A SHORTER LENGTH REPLACEMENT LENS WAS IMPLANTED AND THEY STATED THAT THE PROBLEM RESOLVED. HOWEVER, FOR STATUS OF THE EYE THEY REPORTED STATED H CELL FLARE. Manufacturer Narrative: IT WAS REPORTED THAT THE CELL FLARE/INFLAMMATION HAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: EXPLANTED: (B)(6) 2019. (B)(4). DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. THERE WAS DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON LENS SURFACE. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8633168&pc=MTA | 2023826-2019-00878 | 2019-03-12 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. CURRENTLY, NO REPLACEMENT LENS HAS BEEN IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MFR: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS DEFORMED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9279660&pc=MTA | 2023826-2019-02066 | 2019-03-12 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: H3- DEVICE EVALUATION: RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8586235&pc=MTA | 2023826-2019-00745 | 2019-03-18 | Injury | STAAR SURGICAL COMPANY | 2019-05-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A CASE/VIAL WITH CLEAR SURGICAL RESIDUE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8586259&pc=MTA | 2023826-2019-00746 | 2019-03-18 | Injury | STAAR SURGICAL COMPANY | 2019-05-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL IN LIQUID. VISUAL INSPECTION FOUND LENS HAPTIC TORN AND BROKEN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8588494&pc=MTA | 2023826-2019-00777 | 2019-04-11 | Injury | STAAR SURGICAL COMPANY | 2019-05-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. HIGH IOP (INTRAOCULAR PRESSURE) AND A VAULT VALUE WAS REPORTED. PATIENT BCVA IS 20/20, LENS REMAINS IN THE EYE AND SURGEON PLANS TO EXCHANGE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON 21 MAR 2019. HIGH IOP AND A VAULT VALUE WAS REPORTED. ON 25 APR 2019 THE LENS WAS REMOVED AND A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED. THE REPLACEMENT LENS RESOLVED THE PROBLEM, PATIENT IS HAPPY. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXPLANT .(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8591429&pc=MTA | 2023826-2019-00778 | 2019-04-11 | Injury | STAAR SURGICAL COMPANY | 2019-05-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. HIGH IOP (INTRAOCULAR PRESSURE) AND A VAULT VALUE WAS REPORTED. PATIENT BCVA IS 20/20, LENS REMAINS IN THE EYE AND SURGEON PLANS TO EXCHANGE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIA WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMGE TO THE LENS. CLAIM #: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON 21 MAR 2019. HIGH IOP AND A VAULT VALUE WAS REPORTED. ON 25 APR 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. THIS EXCHANGE RESOLVED THE PROBLEM, PATIENT IS HAPPY. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317909&pc=MTA | 2023826-2019-02125 | 2019-05-09 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -10.5, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH THE SAME SIZE LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8957671&pc=MTA | 2023826-2019-01535 | 2019-05-14 | Injury | STAAR SURGICAL COMPANY | 2019-09-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS, ELEVATED IOP (27 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, WITH EXCESSIVE VAULT WERE OBSERVED ON (B)(6) 2019. LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9913404&pc=MTA | 2023826-2020-00721 | 2019-06-15 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8939422&pc=MTA | 2023826-2019-01558 | 2019-07-06 | Injury | STAAR SURGICAL COMPANY | 2019-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5 12.6, -05.00, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. PUPIL BLOCK WITH ELEVATE IOP (50 MMHG) AND EXCESSIVE VAULT WERE OBSERVED ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS "IRIS PROLAPSE OCCURRED DURING SURGERY. OVD REMAINED IN THE EYE AFTER SURGERY. ON THE DAY OF SURGERY, POSTOPERATIVE PUPILLARY MYDRIASIS OCCURRED DUE TO INCREASED INTRAOCULAR PRESSURE." PER THE REPORTER ON (B)(6) 2019, "THE DOCTOR WILL PRESCRIBE AN IRIS IMAGE CONTACT LENS TO THE PATIENT. ON THE DAY OF THE SURGERY THE DOCTOR SUGGESTED TO THE PATIENT THAT RE-OPERATE TO REMOVE OVD, BUT THE PATIENT REFUSED BECAUSE OF SEVERE EYE PAIN." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 3191: NO CODE AVAILABLE (MEDICATION: GLYCEOL, DIAMOX). ON THE DAY OF THE SURGERY, MEDICATION WAS PRESCRIBED TO LOWER THE IOP (GLYCEOL 200ML AND DIAMOX) IN THE MORNING AND EVENING. THERE WAS NO POSTOPERATIVE TREATMENT TO MANAGE THE IOP DUE TO THE PATIENT'S IOP LOWERED THE NEXT DAY. THE SURGEON STATED THE CAUSE OF THE IRIS PROLAPSE WAS DUE TO "VITREOUS BODY PRESSURE INCREASED BECAUSE OF THE PATIENT WITH STRONG TENSION DURING THE OPERATION. THERE WAS A POSSIBILITY THAT THE INCISION IS LINER AND SHORT". CLAIM#: (B)(4). Manufacturer Narrative: NO CODE AVAILABLE (IRIS PROLAPSE, PUPILLARY MYDRIASIS, EYE PAIN). (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317874&pc=MTA | 2023826-2019-02139 | 2019-07-13 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND ANGLE CLOSURE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PREVIOUSLY REPORTED ANGLE CLOSURE. UPDATED/CORRECTED INFO: PATIENT DID NOT EXPERIENCE ANGLE CLOSURE. PLEASE DISREGARD. PATIENT CODE 3191: ANGLE CLOSURE WAS REPORTED. THIS SHOULD BE REMOVED. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8908374&pc=MTA | 2023826-2019-01488 | 2019-07-23 | Injury | STAAR SURGICAL COMPANY | 2019-08-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE (IOP). AT 4 HOUR IOP CHECK, THE IOP WAS 51 MMHG DUE TO ANGLE CLOSURE FROM EXCESSIVE VAULTING. IOP WAS CONTROLLED BY TAPPING PARACENTESIS AND COMBIGAN DROPS. THE PATIENT WAS TAKEN BACK TO THE OPERATING ROOM TO EXPLANT THE ICL. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: RETURNED IN LIQUID IN SPECIMEN CUP. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8910129&pc=MTA | 2023826-2019-01467 | 2019-07-25 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM513.2, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ELEVATED IOP (24 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, WITH EXCESSIVE VAULT WAS OBSERVED ON (B)(6) 2019. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS WAS REPOSITIONED ON (B)(6) 2019 TO 90 DEGREES VERTICAL AND THE PROBLEM RESOLVED. POSTOPERATIVE REPORT STATED "VAULTING IS BETTER, LENS STAYS IN THE EYE, NEXT EXAMINATION IN ONE YEAR". PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS REPOSITIONING) CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8910138&pc=MTA | 2023826-2019-01468 | 2019-07-25 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM513.2, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ELEVATED IOP (25 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, WITH EXCESSIVE VAULT WAS OBSERVED ON (B)(6) 2019. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - LENS WAS REPOSITIONED 90 DEGREES TO VERTICLE POSITION ON (B)(6) 2019 AND IT RESOLVED THE PROBLEM. POSTOPERATIVE REPORT STATED "VAULTING IS BETTER AFTER ROTATION, NEXT EXAM IN 1 YEAR". H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS REPOSITIONING). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9588875&pc=MTA | 2023826-2020-00121 | 2019-10-10 | Injury | STAAR SURGICAL COMPANY | 2020-01-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE REPORTER STATED THAT THE PATIENT EXPERIENCED ELEVATED IOP AND EXCESSIVE VAULT. IOP WAS CONTROLLED WITH ANTI IOP DROPS, BCVA 6/12. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 OF -11.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE REPORTER STATED THE PATIENT EXPERIENCED EXCESSIVE VAULT AND ELEVATED IOP. THE ELEVATED IOP WAS CONTROLLED WITH ANTI IOP DROPS. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H3-DEVICE EVALUATION: LENS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308814&pc=MTA | 2023826-2020-01525 | 2019-12-04 | Injury | STAAR SURGICAL COMPANY | 2020-07-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019 DUE TO ELEVATED IOP (40MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE AND EXCESSIVE VAULT. ON (B)(6) 2020 A SHORTER LENGHT LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS "LENS WAS LARGER THAN IT SHOULD HAVE BEEN." Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9891527&pc=MTA | 2023826-2020-00598 | 2020-03-04 | Injury | STAAR SURGICAL COMPANY | 2020-03-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON 26/FEB/2020. PUPIL BLOCK WITH ELEVATED IOP (48MMHG) AND ANGLE CLOSURE WITH ELEVATED IOP (48MMHG) WERE ASSESSED ON (B)(6) 2020, AND EXCESSIVE VAULT WAS OBSERVED. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE (EXCESSIVE VAULT). REPORTEDLY, "PATIENT HAPPY." Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910966&pc=MTA | 2023826-2020-00675 | 2020-03-06 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -4.5 DIOPTER, INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. ON THE SAME DATE, IN A SEPARATE SURGERY, THE SURGEON EXCHANGED THE LENS WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE "(SIZE)." Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN LIQUID IN LENS CASE/ VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10214025&pc=MTA | 2023826-2020-01383 | 2020-04-15 | Injury | STAAR SURGICAL COMPANY | 2020-06-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -8.50 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS EXPLANTED DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP (INTRAOCULAR PRESSURE) AND THE PROBLEM RESOLVED. FOR STATUS OF THE EYE THE REPORTER STATED " THE CORNEA IS TRANSPARENT, THE PUPIL IS 3.5X3.0MM, THE PUPIL IS INSENSITIVE TO LIGHT. THE LENS IS TRANSPARENT, AND THER IS A LITTLE PIGMENT ON THE SURFACE OF THE ANTERIOR CAPSULE". ON (B)(6) 2020 THE REPORTER STATED " THE PIGMENTATION IS NOT INCREASED. THIS OBSERVATION IS NOT BEYOND NORMAL". Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: B4 - CORRECTED TO (B)(6) 2020 IN INITIAL MDR G3 - CORRECTED TO (B)(6) 2020 IN INITIAL MDR CLAIM# (B)(4). Manufacturer Narrative: SUPPLEMENTAL MDR#2 IS NOT APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7527986&pc=MTA | 2023826-2018-00758 | NA | Injury | STAAR SURGICAL COMPANY | 2018-05-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE PATIENT REPORTED HAD AN ICL LENS IMPLANTED ON (B)(6) 2018 AND AT ONE DAY POST-OP HAD HIGH EYE PRESSURE. THE FACILITY REPORTED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE INTRAOCULAR PRESSURE WAS STABLE. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. CORRECTED INFORMATION: PATIENT IDENTIFIER SHOULD BE "(B)(6)", (B)(6) IS INCORRECT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LENS CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8500521&pc=MTA | 2023826-2019-00587 | NA | Injury | STAAR SURGICAL COMPANY | 2019-04-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -10.50 WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED IOP (INTRAOCULAR PRESSURE). ADDITIONAL SURGICAL INTERVENTION (ENLARGEMENT OF PI OR ADDITION OF NEW PI/ YAG) WAS REPORTED TO BE PERFORMED. ON (B)(6) 2019 THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM RESOLVED. Manufacturer Narrative: (B)(4). | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10013329&pc=MTA | 2023826-2020-00963 | NA | Injury | STAAR SURGICAL COMPANY | 2020-04-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ELEVATED IOP (25 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED ON (B)(6) 2020, EXCESSIVE VAULT, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WERE OBSERVED. THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: B5 - "THE CAUSE OF EVENT IS REPORTED AS OTHER - NEW WTW CALCULATION METHOD." SHOULD BE INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10013330&pc=MTA | 2023826-2020-00964 | NA | Injury | STAAR SURGICAL COMPANY | 2020-04-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ELEVATED IOP (22 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED ON (B)(6) 2020, EXCESSIVE VAULT, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WERE OBSERVED. THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. CLAIM#: (B)(4). Manufacturer Narrative: B5 - "THE CAUSE OF EVENT IS REPORTED AS OTHER - NEW WTW CALCULATION METHOD." SHOULD BE INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9689153&pc=MTA | 2023826-2020-00281 | 2019-12-17 | Injury | STAAR SURGICAL COMPANY | 2020-02-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019 "DUE TO EYE PRESSURE TOO HIGH." NO REPLACEMENT LENS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# : (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THE A12.6MM MICL 12.6 OF -16.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019 DUE TO PRESSURE SPIKE SECONDARY TO HIGH VAULT. NARROW ANGLES, PIGMENT AND BLOOD IN THE ANGLES. PATIENTS IS NOW BACK TO NORMAL. DOCTOR IS PLANNING ON PLACING A SMALLER ICL. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED AT THIS TIME. CORRECTED DATA: DEVICE AVAILABLE FOR EVALUATION DATE IN PREVIOUS MDR SHOULD BE CORRECTED TO (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: G3- DATE IN PREVIOUS MDR IS CORRECTED TO (B)(6) 2020. DEVICE CODE- 1583 SHOULD BE INCLUDED IN PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9648851&pc=MTA | 2023826-2020-00206 | NA | Injury | STAAR SURGICAL COMPANY | 2020-01-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem; Unintended Movement | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: A DAMAGED (TEAR ON HAPTIC) 13.2MM, MICL 13.2, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS RETURNED TO STAAR. " OS (TANYA MOORE ) REPLACED OLD WITH NEW ICL", WAS HANDWRITTEN ON THE LENS PACKAGING BOX. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2012. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR THE SAME MODEL LENS, LENGTH AND DIOPTER DUE TO LOW VAULT OVERTIME, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PARTIAL ANGLE CLOSURE (FOCAL ANGLE CLOSURE) AND PROGRESSIVE NARROWING OF ANGLE WITH ELEVATED IOP (THAT REQUIRED MEDICAL INTERVENTION). THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THIS EVENT WAS DUE TO THE LENS BEING IMPLANTED UPSIDE DOWN. THE REPORTER STATED " THE ORIGINAL ICL WAS INADVERTENTLY INVERTED ON INSERTION. THIS WAS NOT RECOGNIZED UNTIL SEVERAL YEARS AFTER SURGERY, SINCE THE PATIENT HAD GOOD VISION AND WAS TEMPORARILY LOST TO FOLLOW UP. SHE WAS REFERRED BACK TO ME INTERNALLY FROM A COLLEAGUE WHO NOTED FOCAL ANGLE CLOSURE IN ONE EYE ONLY AND INCREASING IOP. UBM REVEALED AN INVERTED ICL IN THAT EYE, WHICH I EXCHANGED SUCCESSFULLY. I BELIEVE THAT THE SMALLER INCISION BUTTERFLY STYLE INJECTOR I WAS USING AT THE TIME FOR SEVERAL MONTHS (ON ADVICE OF THE LOCAL REP AT THAT TIME) CONTRIBUTED TO THE INVERSION OF THE ICL." Manufacturer Narrative: H6- PATIENT CODE 3191: SIGNIFICANT REDUCTION OF ICA, PARTIAL ANGLE CLOSURE (FOCAL ANGLE CLOSURE), NARROWING OF ANGLE, SECONDARY MEDICAL (MEDS) AND SURGICAL (EXCHANGE) INTERVENTION. H6- WORK ORDER SEARCH: A WORK ORDER SEARCH WAS PERFORMED FOR THE CORRECT SERIAL NUMBER AND NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF BIRTH: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE VIAL. VISUAL INSPECTION FOUND A TEAR ON A HAPTIC. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4554740&pc=MTA | 2023826-2015-00134 | 2014-08-04 | Injury | STAAR SURGICAL COMPANY | 2015-02-27 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED THE NEXT DAY DUE TO ELEVATED IOP ASSOCIATED WITH EXCESSIVE VAULTING. SIGNIFICANT REDUCTION OF IRIDO-CORNEALS ANGLES WAS NOTED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA WAS 20/22. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEONS DETERMINES THA THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4599006&pc=MTA | 2023826-2015-00303 | 2014-08-16 | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO ELEVATED IOP ASSOCIATED EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UCVA WAS 20/20. Manufacturer Narrative: (B)(4). EVALUATION CODES RESULTS (OTHER). A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.) TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THA THE LENS IS EXPLANTED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5182255&pc=MTA | 2023826-2015-01326 | 2014-11-05 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO 13.7 IMPLANTABLE COLLAMER -13.50 DIOPTER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. EXCESSIVE VAULTING ASSOCIATED WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED ON (B)(6) 2014. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS WITH A SIMILAR DIOPTER SIZE ON (B)(6) 2014. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT IS NOT MARKETED IN THE U.S. DIOPTER: -13.50. (B)(4). CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4629139&pc=MTA | 2023826-2015-00375 | 2014-11-28 | Injury | STAAR SURGICAL COMPANY | 2015-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? LENS NOT RETURNED. EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4574078&pc=MTA | 2023826-2015-00240 | 2014-12-22 | Injury | STAAR SURGICAL COMPANY | 2015-03-05 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE AND ANGLE CLOSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. THE REPORTER STATED THE EVENT WAS CAUSED BY THE PATIENT TAKING ORAL ANTIHISTAMINE/DECONGESTANT (CAUSING PUPIL DILATION AND ANGLE CROWDING), THIS WAS REVERSED AFTER BEING PRESCRIBED PILOCARPINE DROPS AND EYE MASSAGE. Manufacturer Narrative: (B)(4). EVALUATION, METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD - (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS - (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - ACCORDING TO USER FAILURE MODE ANALYSIS, IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING WAS A CONSEQUENCE OF WRONG LENS USE WHICH INCLUDE THE FOLLOWING: IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENT BASED UPON DIFFERENT TECHNIQUES, POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN A INDIVIDUAL CASE, IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST AND IMPROPER LENS LABEL. EXCESSIVE VAULTING COULD CAUSE ANGLE NARROWING, CLOSURE AND HIGH IOP. ANTIHISTAMINE MEDICATION COULD HAVE THE MYDRASIS EFFECT THAT EXACERBATED THE NARROWING AND CLOSURE OF ANTERIOR ANGLE. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4574128&pc=MTA | 2023826-2015-00241 | 2014-12-22 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-05 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE AND ANGLE CLOSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. THE REPORTER STATED THE EVENT WAS CAUSED BY THE PATIENT TAKING ORAL ANTIHISTAMINE/DECONGESTANT (CAUSING PUPIL DILATION AND ANGLE CROWDING), THIS WAS REVERSED AFTER BEING PRESCRIBED PILOCARPINE DROPS AND EYE MASSAGE. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - ACCORDING TO USER FAILURE MODE ANALYSIS, IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING WAS A CONSEQUENCE OF WRONG LENS USE WHICH INCLUDE THE FOLLOWING: IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENT BASED UPON DIFFERENT TECHNIQUES, POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN A INDIVIDUAL CASE, IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST AND IMPROPER LENS LABEL. EXCESSIVE VAULTING COULD CAUSE ANGLE NARROWING, CLOSURE AND HIGH IOP. ANTIHISTAMINE MEDICATION COULD HAVE THE MYDRASIS EFFECT THAT EXACERBATED THE NARROWING AND CLOSURE OF ANTERIOR ANGLE. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4513699&pc=MTA | 2023826-2015-00151 | 2014-12-31 | Injury | STAAR SURGICAL COMPANY | 2015-02-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND THE LENS WAS EXPLANTED DUE TO LOW VAULTING AND LOW OPACITY. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THERE IS NOT ENOUGH CLINICAL DATA TO DETERMINE THE EXACT CAUSE OF THE EVENT, HOWEVER, CANNOT RULE OUT LOW VAULTING SECONDARY TO A SHORT LENS AS A CONTRIBUTING FACTOR. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). STAAR RECOMMENDS THE LENS TO BE LEFT IMPLANTED IF NO CATARACTS ARE NOTED OR NO PROGRESSION OF ANTERIOR SUB-CAPSULAR CATARACT. THE SERIAL NUMBER WAS NOT PROVIDED SO WE WERE UNABLE TO PERFORMED WORK ORDER SEARCH. NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5012294&pc=MTA | 2023826-2015-01027 | 2015-01-31 | Injury | STAAR SURGICAL COMPANY | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (UP TO 30MMHG). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5016393&pc=MTA | 2023826-2015-01032 | 2015-02-14 | Injury | STAAR SURGICAL COMPANY | 2015-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (UP TO 25MMHG). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5016398&pc=MTA | 2023826-2015-01033 | 2015-02-24 | Injury | STAAR SURGICAL COMPANY | 2015-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (UP TO 28MMHG). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5086251&pc=MTA | 2023826-2015-01208 | 2015-03-06 | Injury | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THE HAPTIC TO BE TORN/BROKEN/BENT/DEFORMED. (B)(4). CONCLUSION: PER DFU, THIS LENS MODEL IS CONTRAINDICATED IN THE PATIENT'S AGE BEING UNDER 21 YEARS. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119297&pc=MTA | 2023826-2015-01298 | 2015-03-20 | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119166&pc=MTA | 2023826-2015-01299 | 2015-04-01 | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND A SHALLOW CHAMBER. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE OTHER COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5076256&pc=MTA | 2023826-2015-01128 | 2015-04-03 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -10.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. SUBSEQUENTLY, THE PATIENT HAD ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT NOT MARKETED IN THE U.S. DIOPTER: -10.00. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4) NEW INCISION. SECONDARY SURGERY. LENS EXPLANTED. VAULTING. BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THIS EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119168&pc=MTA | 2023826-2015-01291 | 2015-04-20 | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5174763&pc=MTA | 2023826-2015-01390 | 2015-05-16 | Injury | STAAR SURGICAL COMPANY | 2015-10-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER ICL LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: (B)(4). NO LENS RETURNED, ONLY EMPTY BAG. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5124781&pc=MTA | 2023826-2015-01295 | 2015-05-21 | Injury | STAAR SURGICAL COMPANY | 2015-10-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5146041&pc=MTA | 2023826-2015-01246 | 2015-05-26 | Injury | STAAR SURGICAL COMPANY | 2015-10-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | PO30016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS WITH A SIMILAR DIOPTER SIZE. THIS DID NOT RESOLVE THE PROBLEM. POST-OP VISIT ON (B)(6) 2015 - UCVA WAS 20/26. THE SECONDARY ICL WAS EXPLANTED ON (B)(6)2015. SEE MFR # 2023826-2015-01247 SECONDARY ICL. Manufacturer Narrative: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DIOPTER: -13.50, DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD CODE(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODE(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S) : (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5146063&pc=MTA | 2023826-2015-01244 | 2015-05-26 | Injury | STAAR SURGICAL COMPANY | 2015-10-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | PO30016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS WITH A SIMILAR DIOPTER SIZE. THIS DID NOT RESOLVE THE PROBLEM. POST-OP VISIT ON (B)(6) 2015 - UCVA WAS 20/26. THE SECONDARY ICL WAS EXPLANTED ON (B)(6) 2015. SEE MFR # 2023826-2015-01245 SECONDARY ICL. Manufacturer Narrative: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. CORRECTED DATA: IT WAS INITIALLY REPORTED THAT THE PATIENT'S POST-OP VISIT WAS ON (B)(6) 2015, AND THIS IS INCORRECT. THE PATIENT'S POST-OP VISIT WAS ON (B)(6) 2015. (B)(4); NO OTHER ADVERSE EVENTS REPORTED OUTSIDE OF LENS EXCHANGE AND ELEVATED IOP. (B)(4); CODE NOT APPLICABLE; SHOULD BE IN PATIENT CODE, NOT DEVICE CODE. (B)(4). Manufacturer Narrative: LENS NOT RETURNED. (B)(4). EVALUATION METHOD: LENS WORK ORDER SEARCH. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119300&pc=MTA | 2023826-2015-01300 | 2015-06-12 | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5151143&pc=MTA | 2023826-2015-01249 | 2015-08-03 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING ASSOCIATED WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS WITH A SIMILAR DIOPTER SIZE. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015 - UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -07.00. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5287706&pc=MTA | 2023826-2015-01627 | 2015-08-22 | Injury | STAAR SURGICAL COMPANY | 2015-12-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5287692&pc=MTA | 2023826-2015-01638 | 2015-08-22 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -09.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) (NEW INCISION) IS INCORRECT. NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5409020&pc=MTA | 2023826-2016-00088 | 2015-08-29 | Injury | STAAR SURGICAL COMPANY | 2016-02-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULTING WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2015 FOR A SHORTER LENGTH LENS WITH A SIMILAR DIOPTER. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER:-10.00. PT WEIGHT.: UNK. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD CODE(S): (EVALUATION METHOD): WORK ORDER SEARCH. RESULTS CODE(S): (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5109247&pc=MTA | 2023826-2015-01265 | 2015-08-31 | Injury | STAAR SURGICAL COMPANY | 2015-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS WITH A 15.50 DIOPER IN THE PATIENT'S RIGHT EYE (OD) ON(B)(6) 2015. ON (B)(6) 2015, THE PATIENT SHOWED INADEQUATE VAULTING, ELEVATED IOP AND PIGMENT DISPERSION. SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED WHEN FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2015. LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -15.50 DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5248750&pc=MTA | 2023826-2015-01577 | 2015-10-20 | Injury | STAAR SURGICAL COMPANY | 2015-11-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, IRIS ISCHEMIA, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER STATED THE CAUSE OF THE EVENT WAS DUE TO SIZING ISSUE. THE PATIENT HAS PERSISTENT ANISOCORIA/ IRIS PARESIS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 AND THE PATIENT IS VERY HAPPY WITH THE IMPROVED VISION. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. IN ADDITION, THE REPORTER STATED, THE CAUSE OF THE EVENT WAS DUE TO SIZING ISSUE. SINCE THE EXCHANGE WITH THE SMALLER SIZE LENS RESOLVED THE ISSUE, HENCE, IT IS VERY LIKELY THE CAUSE OF THE EVENT IS DUE TO LENS SIZING. PER MEDICAL REVIEW - REPORTEDLY, ICL (13.7MM) WAS EXPLANTED/EXCHANGED FOR ANOTHER SHORTER ICL (13.2 MM) EIGHT DAYS POSTOPERATIVELY TO ADDRESS STEEP VAULT AND SUBSEQUENT ELEVATED IOP THAT WAS FIRST NOTED 12 HOURS AFTER IMPLANTATION. IRIDOTOMIES WERE PATENT. ACCORDING TO THE COMPLETED CUSTOMER COMPLAINT QUESTIONNAIRE FOR SURGERY, SIGNED BY SURGEON ON 11/23/15, THE MOST LIKELY CAUSE OF THE REPORTED EVENTS (VAULTING AND IOP SPIKE) WAS SURGERY RELATED SIZING ISSUE. THE LENS WAS NOT MISLABELED AND THERE WAS NO OTHER LENS PROBLEM. HOWEVER, THE POSTOPERATIVE COMPLICATIONS (ANISOCORIA AND IRIS PARESIS) WERE ATTRIBUTED TO THE DEVICE PER SURGEON`S NOTE. HE ALSO REPORTED THAT IRIS PARESIS WAS IMPROVING. HE ALSO STATED THAT DESPITE ADVERSE EVENT, AFTER LENS EXCHANGE, THE PATIENT WAS VERY HAPPY WITH IMPROVED VISION AND SCHEDULED SURGERY FOR OTHER EYE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND THE MEDICAL REVIEW, THE MOST LIKELY CAUSE OF THE REPORTED EVENTS (VAULTING AND IOP SPIKE) WAS SURGERY RELATED SIZING ISSUE. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5611474&pc=MTA | 2023826-2016-00525 | 2015-10-20 | Injury | STAAR SURGICAL COMPANY | 2016-04-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PUPIL DISTORTION AND PERIPHERAL ANTERIOR SYNECHIAE. A SECONDARY YAG WAS PERFORMED TO ENLARGE/ADDITION OF PIS. THE PATIENT IS TAKING MEDICATION FOR THE PRESSURE. Manufacturer Narrative: (B)(4). DEVICE EVALUATIONS: THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND PEN MARKINGS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EVENT PROBLEM AND EVALUATION CODES: PATIENT CODE: (B)(4) INCISION OPENED. SECONDARY SURGERY. ENLARGEMENT/ADDITION IRIDECTOMY. PERIPHERAL ANTERIOR SYNECHIAE. PUPIL DISTORTION. EXPLANTED. VAULTING. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5613773&pc=MTA | 2023826-2016-00526 | 2015-10-23 | Injury | STAAR SURGICAL COMPANY | 2016-04-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PERIPHERAL ANTERIOR SYNECHIAE. A SECONDARY YAG WAS PERFORMED TO ENLARGE/ADDITION OF PIS. THE PATIENT IS TAKING MEDICATION FOR THE PRESSURE. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS WITH PEN MARKINGS AND FIBERS ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5643957&pc=MTA | 2023826-2016-00587 | 2015-10-25 | Injury | STAAR SURGICAL COMPANY | 2016-05-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -8.0 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT AND ELEVATED IOP WERE REPORTED, THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAPTIC TORN AND BENT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5408378&pc=MTA | 2023826-2016-00086 | 2015-11-03 | Injury | STAAR SURGICAL COMPANY | 2016-02-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULTING WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2015 FOR A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER:-12.50 (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5439668&pc=MTA | 2023826-2016-00153 | 2015-11-10 | Injury | STAAR SURGICAL COMPANY | 2016-02-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, - 09.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND AN UNREACTIVE PUPIL (FIXED). THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT HAD ELEVATED INTRAOCULAR PRESSURE (30MMHG) POST-OP. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (PROCESS EVALUATION): WORK ORDER SEARCH. (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5399905&pc=MTA | 2023826-2016-00080 | 2015-11-22 | Injury | STAAR SURGICAL COMPANY | 2016-01-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5432675&pc=MTA | 2023826-2016-00149 | 2015-11-25 | Injury | STAAR SURGICAL COMPANY | 2016-02-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5579290&pc=MTA | 2023826-2016-00474 | 2015-12-21 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE AND THERE WAS A RISE IN INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -10.0. PT WEIGHT: UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5602214&pc=MTA | 2023826-2016-00513 | 2015-12-22 | Injury | STAAR SURGICAL COMPANY | 2016-04-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536876&pc=MTA | 2023826-2016-00313 | 2015-12-28 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4 -9.50 DIOPTER IMPLANTABLE COLLAMER LENS ON (B)(6) 2015 IN TO THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE AND IOP WAS ELEVATED. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015 AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DIOPTER: -9.50. (B)(4). RESULTS: A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM #(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5575699&pc=MTA | 2023826-2016-00461 | 2015-12-30 | Injury | STAAR SURGICAL COMPANY | 2016-04-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PATIENT'S POST-OP ((B)(6) 2016) UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION - THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL ON LENS SURFACE (UNKNOWN). (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5628400&pc=MTA | 2023826-2016-00552 | 2016-01-01 | Injury | STAAR SURGICAL COMPANY | 2016-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.0 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT AND ELEVATED IOP WERE REPORTED. THE LENS WAS EXPLANTED ON (B)(6) 2016. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: COMPLETE DIOPTER VALUE IS: -14.0/+3.0/X092. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536839&pc=MTA | 2023826-2016-00310 | 2016-01-08 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 -12.50 DIOPTER IMPLANTABLE COLLAMER LENS ON (B)(6) 2016 IN TO THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE AND IOP WAS ELEVATED. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2016 AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -12.50 (B)(4) SECONDARY SURGERY. NEW INCISION. EXPLANTED. VAULTING. METHOD: WORK ORDER SEARCH. RESULTS: A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5577811&pc=MTA | 2023826-2016-00439 | 2016-01-16 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -07.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE WAS NOTED. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS WITH A SIMILAR DIOPTER. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2016, UCVA WAS 20/16. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -07.00 (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5826262&pc=MTA | 2023826-2016-00252 | 2016-01-18 | Injury | STAAR SURGICAL COMPANY | 2016-07-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ELEVATED INTRAOCULAR PRESSURE, PIGMENT DISPERSION AND AN UNREACTIVE PUPIL (FIXED). THE LENS WAS REPOSITIONED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/6/9. Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2016 FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN AND THE LENS HAD PEN MARKINGS ON IT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5525073&pc=MTA | 2023826-2016-00328 | 2016-01-19 | Injury | STAAR SURGICAL COMPANY | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY AND BENT, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE FOREIGN MATERIAL ON LENS SURFACE (FIBERS). (B)(4). THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5575700&pc=MTA | 2023826-2016-00463 | 2016-01-20 | Injury | STAAR SURGICAL COMPANY | 2016-04-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/29. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) (NEW INCISION) IS INCORRECT. NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5575705&pc=MTA | 2023826-2016-00464 | 2016-01-20 | Injury | STAAR SURGICAL COMPANY | 2016-04-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/29. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5668556&pc=MTA | 2023826-2016-00734 | 2016-02-10 | Injury | STAAR SURGICAL COMPANY | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE AND ELEVATED INTRAOCULAR PRESSURE. ADDITIONAL LASER PERIPHERAL IRIDOTOMIES WERE PERFORMED. THE REPORTER INDICATED THEY FELT THERE WAS POSSIBLE AQUEOUS MISDIRECTION. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: RESULT: (OPERATIONAL PROBLEM): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5724417&pc=MTA | 2023826-2016-00812 | 2016-03-28 | Injury | STAAR SURGICAL COMPANY | 2016-06-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO AN EXCESSIVE VAULT, ELEVATED IOP, NARROWING OF THE ANGLE, AND SHALLOWING OF THE ACD. THE LENS WAS EXCHANGED FOR A 12.6MM ICL AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). RESULTS: (OPERATIONAL PROBLEM) - VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THERE WAS EVIDENCE OF DARK COLOR RESIDUE ON THE LENS HAPTIC. THE LENS WAS RETURNED IN LIQUID. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5796321&pc=MTA | 2023826-2016-00906 | 2016-04-28 | Injury | STAAR SURGICAL COMPANY | 2016-07-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN AND SCRATCHED. (B)(4). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5796325&pc=MTA | 2023826-2016-00907 | 2016-04-28 | Injury | STAAR SURGICAL COMPANY | 2016-07-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: (B)(4). PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5817535&pc=MTA | 2023826-2016-00930 | 2016-06-21 | Injury | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: AS PER THE DFU OF THIS LENS MODEL, IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRADICTED. THIS PATIENT HAD AN ACD OF 2.915 MM AT THE TIME OF IMPLANTATION. CORRECTED INFORMATION: (B)(4) NOT APPLICABLE, SHOULD BE UNDER PATIENT CODE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5817536&pc=MTA | 2023826-2016-00932 | 2016-06-21 | Injury | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CONCLUSION: OFF-LABEL USE: PATIENT ACD BELOW 3.0 MM AND HAD A PRE-EXISTING CONDITION (PREVIOUS CORNEAL OR REFRACTIVE SURGERY). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5983960&pc=MTA | 2023826-2016-01326 | 2016-07-26 | Injury | STAAR SURGICAL COMPANY | 2016-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ANGLE CLOSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EXPLANTED DATE - (B)(6) 2016. THE REPORTER INDICATED THE LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20 AND VISION IS BETTER. (B)(4). RESULT CODE(S): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS CODE(S): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5983970&pc=MTA | 2023826-2016-01327 | 2016-07-26 | Injury | STAAR SURGICAL COMPANY | 2016-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ANGLE CLOSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THE LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD). THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20 BUT IS IMPROVING. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5957285&pc=MTA | 2023826-2016-01286 | 2016-08-13 | Injury | STAAR SURGICAL COMPANY | 2016-09-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (AT 1 DAY POST-OP) AND ANGLE CLOSURE. THE SURGEON HAD DIFFICULTY "TUCKING" THE FEET OF THE ICL DURING THE INITIAL SURGERY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/400. Manufacturer Narrative: (B)(4). BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS CUT INTO TWO PIECES, WITH A PIECE OF HAPTIC MISSING. THE LENS WAS RETURNED IN LIQUID. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8356345&pc=MTA | 2023826-2019-00279 | 2019-01-09 | Injury | STAAR SURGICAL COMPANY | 2019-02-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO MECHANICAL FAILURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THERE WAS NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE CAUSE OF THE EVENT WAS DUE TO SIZING. THE LENS EXCHANGE RESOLVED THE PROBLEM. DATE OF EVENT: (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: EVENT- "(B)(6) 2018" IN THE ORIGINAL MDR SHOULD BE CORRECTED TO "(B)(6) 2018" IMPLANTED DATE SHOULD BE CORRECTED TO (B)(6) 2018 IN THE ORIGINAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, WRAPPED IN GAUZE IN A ZIPLOCK BAG. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. THERE WAS DEBRIS ON THE LENS. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8356327&pc=MTA | 2023826-2019-00280 | 2019-01-16 | Injury | STAAR SURGICAL COMPANY | 2019-02-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO MECHANICAL FAILURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THERE WAS NO PATIENT INJURY. THE LENS WAS DISCARDED BY THE FACILITY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE CAUSE OF THE EVENT WAS DUE TO SIZING. THE LENS EXCHANGE RESOLVED THE PROBLEM. CORRECTED DATA: DATE OF EVENT: (B)(6) 2019. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4822523&pc=MTA | 2023826-2015-00644 | NA | Injury | STAAR SURGICAL COMPANY | 2015-06-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS THE SAME DAY OF EXPLANT.Event Description: ADDITIONAL INFORMATION RECEIVED - THE PATIENT EXPERIENCED NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH BEFORE EXPLANT OF THE LENS. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/15. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE ROOT CAUSE OF THE COMPLAINT. THE LENS SIZING ERROR MOST LIKELY CAUSED THE EXCESSIVE LENS VAULTING. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. EVENT DATE: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. THE INJECTOR, CARTRIDGE AND FOAM TIP PLUNGER WERE NOT RETURNED. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: MEDICAL REVIEW - IN THIS CASE THE IOP ELEVATION, ANGLE NARROWING AND SHALLOW ANTERIOR CHAMBER WERE THE CONSEQUENCES OF EXCESSIVE VAULTING OF THE LENS. THE TIMELY MANAGEMENT AND LENS EXCHANGE WERE NECESSARY TO PREVENT ANY FURTHER DAMAGE CAUSED BY THE EVENT. ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING WAS A CONSEQUENCE OF WRONG LENS USE WHICH INCLUDE THE FOLLOWING: IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENT BASED UPON DIFFERENT TECHNIQUES, POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN A INDIVIDUAL CASE, IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST AND IMPROPER LENS LABEL. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5089547&pc=MTA | 2023826-2015-01217 | NA | Injury | STAAR SURGICAL COMPANY | 2015-09-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER STATED THE CAUSE OF THE EVENT WAS DUE TO VARIED WHITE-TO-WHITE MEASUREMENTS. THERE WERE NO ADDITIONAL SYMPTOMS RELATED TO THE EVENT. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - REPORTEDLY, MICL(13.2 MM) WAS EXPLANTED/EXCHANGED FOR A SHORTER MICL NINE DAYS POSTOPERATIVELY TO ADDRESS EXCESSIVE VAULT AND SUBSEQUENT ELEVATED IOP. ACCORDING TO THE COMPLETED CUSTOMER COMPLAINT QUESTIONNAIRE FOR SURGERY THE EVENT WAS NOT PRODUCT RELATED BUT IT WAS CAUSED BY PROCEDURE RELATED FACTOR (THE SURGEON`S CHOICE OF LONGER LENS BASED ON VARIED WHITE- TO- WHITE MEASUREMENTS). UPON LENS EXCHANGE ALL SYMPTOMS RESOLVED AND UCVA WAS 20/20 CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND THE MEDICAL REVIEW, THE CAUSE OF THE EVENT WAS DUE TO A PROCEDURE RELATED FACTOR (THE SURGEON`S CHOICE OF LONGER LENS BASED ON VARIED WHITE- TO- WHITE MEASUREMENTS). UPON LENS EXCHANGE ALL SYMPTOMS RESOLVED AND UCVA WAS 20/20. THE EVENT WAS NOT PRODUCT RELATED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5125138&pc=MTA | 2023826-2015-01321 | NA | Injury | STAAR SURGICAL COMPANY | 2015-10-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED DUE TO THE LENS WAS TOO LARGE AND ELEVATED INTRAOCULAR PRESSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE FACILITY REPORTED THE LENS WAS EXPLANTED ON (B)(6) 2015, DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER. ADDITIONAL PERIPHERAL IRIDECTOMIES WERE PERFORMED. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: BASED ON THE RESULTS OF THE DHR REVIEW, THE AVAILABLE INFORMATION GIVEN WITHIN THIS COMPLAINT, WORK ORDER SEARCH, AND THE PRODUCT EVALUATION, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. AS REPORTED ON THE MEDICAL REVIEW THE CAUSE OF THE EVENT IS UNKNOWN. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. PER MEDICAL REVIEW: REPORTEDLY, MICL (12.6 MM) WAS EXPLANTED/EXCHANGED, 24 DAYS POSTOPERATIVELY, FOR A SHORTER LENGTH ICL (12.1MM) TO ADDRESS EXCESSIVE VAULT AND SUBSEQUENT ANGLE CLOSURE AND ELEVATED IOP. SECONDARY SURGICALLY INTERVENTION (ADDITIONAL PERIPHERAL IRIDOTOMY) WAS PERFORMED PRIOR TO LENS EXCHANGE. NO REPORTED POSTOPERATIVE SEQUELAE. ACCORDING TO THE SAME REPORT, CAUSE OF THE EVENT WAS UNKNOWN. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, EVALUATION OF THE RETURNED PRODUCT AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER - (B)(6) (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5511748&pc=MTA | 2023826-2016-00283 | NA | Injury | STAAR SURGICAL COMPANY | 2016-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -09.5 DIOPTER, IN THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO BEING TOO LARGE. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITH NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT IS DOING GREATEvent Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT IS DOING GREAT Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY HAVE CREATED OR NOT DETECTED A NONCONFORMANCE OF THE LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY EXCEPTIONAL EVENT OR CONDITION WOULD HAVE CAUSED EXCESSIVE VAULT OF THE LENS. (B)(4). Manufacturer Narrative: METHOD: PROCESS EVALUATION: DEVICE HISTORY RECORD REVIEW. RESULT: (OPERATIONAL PROBLEM): BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY HAVE CREATED OR NOT DETECTED A NONCONFORMANCE OF THE LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY EXCEPTIONAL EVENT OR CONDITION WOULD HAVE CAUSED EXCESSIVE VAULT OF THE LENS. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5511760&pc=MTA | 2023826-2016-00284 | NA | Injury | STAAR SURGICAL COMPANY | 2016-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO BEING TOO LARGE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED FROM THE REPORTER, IT WAS INDICATED THAT THE 12.6MM MICL IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER WAS IMPLANTED ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE PATIENT EXPERIENCED ELEVATED IOP, NARROWING OF THE ANGLE, ANGLE CLOSURE, AND SHALLOWING OF THE ACD WITH PRE-OP IOP 20MM HG. AN IRIDECTOMY WAS PERFORMED AND THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND BOTH HAPTICS HAD PIECES TORN OFF AND MISSING. THE LENS WAS RETURNED IN A SMALL CLEAR BOTTLE IN LIQUID. PER MEDICAL REVIEW IT SHOULD BE NOTED THAT AT THE TIME OF IMPLANTATION THE PATIENT WAS (B)(6) AND PER DFU, INDICATIONS: "THE VISIAN ICL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE" AND THEREFORE, THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5668553&pc=MTA | 2023826-2016-00731 | NA | Injury | STAAR SURGICAL COMPANY | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO OVER VAULT AND HIGH PRESSURE SPIKE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE PATIENT COMPLAINED OF A HEADACHE. THE PUPIL CAME DOWN IN THE RIGHT EYE AND THE SURGEON WAS UNABLE TO REPLACE THE LENS DURING LENS REPLACEMENT SURGERY ON (B)(6) 2016. THE REPLACEMENT LENS SURGERY OCCURRED AT A LATER DATE. RESULTS: (OPERATIONAL PROBLEM): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE TORN OFF AND MISSING FROM ONE HAPTIC. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR/REDDISH RESIDUE. CONCLUSIONS: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(6). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5680337&pc=MTA | 2023826-2016-00757 | NA | Injury | STAAR SURGICAL COMPANY | 2016-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO INCORRECT SIZE. THE LENS WAS EXCHANGED FOR ANOTHER ICL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE AND SHALLOW ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE IRIS SPHINCTER WAS DAMAGED DUE TO THE PRESSURE. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO MISMEASUREMENT OF THE WHITE-TO-WHITE. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). RESULT: (OPERATIONAL PROBLEM) : VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5579677&pc=MTA | 2023826-2016-00444 | 2016-02-04 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -11.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ELEVATED INTRAOCULAR PRESSURE WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2016 FOR A SHORTER LENGTH LENS WITH A SIMILAR DIOPTER SIZE. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2016, UCVA WAS 20/25. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THERE WAS NO VISIBLE DAMAGE TO THE LENS. (B)(4). AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM IS CONTRAINDICATED. THIS PATIENT HAS AN ACD OF 2.89MM. (B)(4). Manufacturer Narrative: DIOPTER: -11.50. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9819152&pc=MTA | 2023826-2020-00452 | 2020-01-30 | Injury | STAAR SURGICAL COMPANY | 2020-03-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT WITH ANGLE CLOSURE AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9819155&pc=MTA | 2023826-2020-00453 | 2020-01-30 | Injury | STAAR SURGICAL COMPANY | 2020-03-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT WITH ANGLE CLOSURE AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4579225&pc=MTA | 2023826-2015-00223 | 2014-03-17 | Injury | STAAR SURGICAL COMPANY | 2015-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS DISLOCATED/SUBLUXED AND THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. LENS NOT RETURNED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ELEVATED IOP IN THE PRESENCE OF HIGH VAULT AFTER ICL IMPLANTATION MAY OCCUR DUE TO: NON PATENT/FUNCTIONING IRIDECTOMIES (TOO SMALL, TOO PERIPHERAL AND/OR NOT FULLY PERMEABLE, BLOCKED BY VISCOELASTIC), REMAINING VISCOELASTIC IN THE POSTERIOR CHAMBER, OVERSIZED ICL WITH ANGLE CLOSURE, TOO NARROW ANGLES/CROWDED AC DUE TO SMALL EYE ANATOMY PREOPERATIVELY, UNEXPECTED ABNORMAL ANATOMY OR TISSUE ABNORMALITIES (PRESENCE OF IRIS/CILIARY BODY CYSTS), MALPOSITIONED FOOTPLATES, ETC. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, IT WAS FOUND THERE WAS NOT ENOUGH CLINICAL DATA TO DETERMINE THE EXACT CAUSE OF THIS EVENT; HOWEVER, WE CANNOT RULE OUT EXCESSIVE VAULTING SECONDARY TO A LONG LENS AS A CONTRIBUTING FACTOR. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6836423&pc=MTA | 2023826-2017-01377 | 2017-08-01 | Injury | STAAR SURGICAL COMPANY | 2017-08-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER. THE LENS WAS EXPLANTED A COUPLE OF HOURS LATER DUE TO ELEVATED INTRAOCULAR PRESSURE. ANOTHER IRIDECTOMY WAS PERFORMED BUT THE PRESSURE WAS STILL ELEVATED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED AND REMOVED ON (B)(6) 2017 FROM THE PATIENT'S LEFT EYE (OS). THE INCISION WAS ENLARGED. THE LENS EXPLANT RESOLVED THE PROBLEM. AFTER THE LAST POST-OP VISIT, THE BCVA WAS 20/20 AND THERE WERE NO PATHOLOGICAL SYMPTOMS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A LIQUID INSIDE A CONTACT CASE WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF ITS HAPTIC AND PRESENCE OF DEBRIS ON THE PRODUCT. "ADDITIONAL IRIDECTOMY WAS PERFORMED BUT THE PRESSURE WAS STILL ELEVATED" IS CORRECTED TO " ADDITIONAL YAG IRIDOTOMY WAS PERFORMED BUT THE PRESSURE WAS STILL ELEVATED" IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9633715&pc=MTA | 2023826-2020-00177 | 2020-01-07 | Injury | STAAR SURGICAL COMPANY | 2020-01-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. "DOCTOR SAW THAT THE LENS WAS CRACKED." THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2020. NO INJURY TO PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A ZIP LOCK BAG WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "THE DOCTOR NOTICED LOW VAULT, REFRACTIVE SURPRISE AND ELEVATED IOP. THE EXCHANGE RESOLVED THE PROBLEM." SHOULD BE ADDED TO THE INITIAL MDR. H6- MEDICAL DEVICE PROBLEM CODE 1401 - REFRACTIVE SURPRISE. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: A2- (B)(6) 1987, SHOULD BE IN THE PREVIOUS MDR. B3- (B)(6) 2020, SHOULD BE IN THE PREVIOUS MDR. H6- CLINICAL CODE 1937 SHOULD BE IN PREVIOUS MDR. H6- DEVICE CODE 1583 SHOULD BE IN PREVIOUS MDR . CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8846146&pc=MTA | 2023826-2019-01278 | 2019-06-19 | Injury | STAAR SURGICAL COMPANY | 2019-07-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO 12.6, -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ELEVATED IOP (28 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE ASSESSED ON (B)(6) 2019. REPORTEDLY PATIENT IS "UNDER DORZOLAMIDE AND TIMOLOL EYE DROPS, TWICE DAILY. VAULT IS 350." LENS REMAINS IMPLANTED. PER PATIENTS LAST VISIT, (B)(6) 2019, IOP 20 MMHG, "CONTINUING ANTI GLAUCOMA MEDICATION AND STOPPED STEROIDS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8846151&pc=MTA | 2023826-2019-01279 | 2019-06-19 | Injury | STAAR SURGICAL COMPANY | 2019-07-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO 12.6, -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ELEVATED IOP (28 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE ASSESSED ON (B)(6) 2019. REPORTEDLY PATIENT IS "UNDER DORZOLAMIDE AND TIMOLOL EYE DROPS, TWICE DAILY. VAULT IS 350." LENS REMAINS IMPLANTED. PER PATIENTS LAST VISIT, (B)(6) 2019, IOP 20 MMHG, "CONTINUING ANTI GLAUCOMA MEDICATION AND STOPPED STEROIDS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE CODE 1494: OFF LABEL USE - ACD<3.0MM. PATIENT CODE 3191: NO CODE AVAILABLE - MEDICAL INTERVENTION, USE OF DORZOLAMIDE AND TIMOLOL EYE DROPS TO TREAT ELEVATED IOP. ADVERSE EVENT OR PRODUCT PROBLEM: CORRECTED TO ADVERSE EVENT FROM PRODUCT PROBLEM IN INITIAL MDR. DEVICE CODE 2993 ADDED; DEVICE CODE 3189 NO LONGER APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966497&pc=MTA | 2023826-2019-01640 | 2018-07-27 | Injury | STAAR SURGICAL COMPANY | 2019-09-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -11.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED, THE SURGEON REPORTING PUPIL BLOCK, ELEVATED IOP, PIGMENT DISPERSION, AND UNREACTIVE (FIXED) PUPIL. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: CORRECTED DATA: "DIOPTER -11.5" SHOULD BE CORRECTED TO "DIOPTER -12.0" IN INITIAL MEDWATCH REPORT. DEVICE CODE:1583 SHOULD BE REMOVED FROM INITIAL MEDWATCH REPORT, AS IT IS NOT APPLICABLE. IT WAS REPORTED THAT THE CAUSE WAS DUE TO ACUTE UVEAL REACTION, THE EXUDATION OF ANTERIOR CHAMBER, ANTERIOR CHAMBER PIGMENT LEAD TO PUPILLARY CLOSURE, THE PATIENT WAS ALERGIC. REPORTER INDICATED THERE WAS NO VAULTING AND PROBLEM HAS RESOLVED AFTER THE LENS EXPLANT. CLAIM#: (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8432797&pc=MTA | 2023826-2019-00442 | 2019-01-31 | Injury | STAAR SURGICAL COMPANY | 2019-03-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM MICL13.7, -10.5 DIOPTER , IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENTS EYE ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019 "DUE TO MECHANICAL VAULTING, TOO STEEP." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: OTHER RELEVANT HISTORY: ADDED INFORMATION (DENSE NS = NUCLEAR SCLENOTIC CATARACT). CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP WITH MOISTURE ON LENS, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8471518&pc=MTA | 2023826-2019-00457 | 2019-03-01 | Injury | STAAR SURGICAL COMPANY | 2019-04-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP (65 MMHG). THIS RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WAS RETURNED IN BIOHAZARD ZIP LOC BAG, IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9279656&pc=MTA | 2023826-2019-02067 | 2019-03-12 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT,ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8957673&pc=MTA | 2023826-2019-01536 | 2019-05-14 | Injury | STAAR SURGICAL COMPANY | 2019-09-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS, ELEVATED IOP (29 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, WITH EXCESSIVE VAULT WERE OBSERVED ON 31/MAY/2019. LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 OF -11.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT; ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS OBSERVED. LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PATIENT CODE 3191 FOR SIGNIFICANT REDUCTIO IN IRIDO-CORNEAL ANGLES IS NOT APPLICABLE. DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8895562&pc=MTA | 2023826-2019-01454 | 2019-07-09 | Injury | STAAR SURGICAL COMPANY | 2019-08-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -8.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PUPIL BLOCK WITH ELEVATED IOP. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DATE RECEIVED BY MANUFACTURER SHOULD HAVE BEEN REPORTED AS 16-JUL-2019 IN INTIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. DEVICE CODE 1494: OFF LABEL USE (AGE OF PATIENT>45 YEARS OLD). CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10445821&pc=MTA | 2023826-2020-01819 | 2020-05-19 | Injury | STAAR SURGICAL COMPANY | 2020-08-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10033107&pc=MTA | 2023826-2020-01027 | 2019-09-02 | Injury | STAAR SURGICAL COMPANY | 2020-05-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; No Code Available; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOSS OF BCVA (PRE-OP BCVA = 20/20 POST-OP BCVA = 20/17-18) AND PUPIL BLOCK WITH ELEVATED IOP WAS OBSERVED. LENS WAS REPOSITIONED AND ANTERIOR CHAMBER IRRIGATION/ EVACUATION OF VISCO/ FLUIDS DID NOT RESOLVE THE PROBLEM. REPORTEDLY, PATIENTS LAST VISIT ON (B)(6) 2019 MIDRIASIS (DILATED PUPILS) WAS OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B3 CORRECTED FROM (B)(6) 2019 TO (B)(6) 2019 IN INITIAL MDR. B5 - CORRECTED TO: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -13.50 D IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTED SAME DAY LOSS OF BCVA; PUPIL BLOCK WITH AN UNSPECIFIED ELEVATED IOP; UNREACTIVE (FIXED) PUPIL AND IRIS ATROPHY. ANTERIOR CHAMBER IRRIGATION/ EVACUATION OF VISCO FLUIDS AND A LENS REPOSITIONING WERE PERFORMED. IT WAS REPORTED THAT THE "ROTATION OF THE LENS" DID NOT RESOLVE THE PROBLEM AND CURRENTLY THE PATIENT HAS PUPIL BLOCK AND UNREACTIVE (FIXED) PUPIL. THE LENS WAS EXPLANTED ON (B)(6) 2019. POST-EXPLANT UCVA AND BCVA WERE REPORTED TO BE 20/18 WITH AND IOP OF 12MM HG AND "MIDRASIS" WAS NOTED. H6 - HEALTH EFFECTS CLINICAL CODE 4581: IRIS ATROPHY, UNREACTIVE PUPIL. H6 - HEALTH EFFECTS IMPACT CODE 4627. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476061&pc=MTA | 2023826-2020-01922 | 2020-07-23 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS). THE LENS WAS EXPLANTED AND EXCHANGED DUE TO HIGH VAULTING AND PUPILLARY BLOCK. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -15.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE AND IRIS ATROPHY. ADDITIONAL SURGERY WAS PERFORMED, ENLARGEMENT OF PI BY YAG AND ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. THE LENS WAS REPLACED FOR A SHORTER LENS (DIFFERENT MODEL - TMICL) ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. SOME MILD IRIS ATROPHY BUT PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - IRIS ATROPHY. H6: HEALTH EFFECT IMPACT CODE: 4625 - ENLARGEMENT OF PI - YAG; ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7648241&pc=MTA | 2023826-2018-00973 | 2018-04-26 | Injury | STAAR SURGICAL COMPANY | 2018-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Consequences Or Impact To Patient; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THAT A 12.1MM, MICL12.1, -12.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) UPSIDE DOWN, THE LENS WAS REMOVED AND THE BACKUP LENS WAS USED. THE REPORTER STATED THAT THERE WAS NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED INFORMATION: DATE OF EVENT WAS CHANGED TO (B)(6) 2018 PREVIOUS REPORTED DATE WAS INCORRECT. (B)(4). THE LENS WAS NOT IMPLANTED UPSIDE DOWN AS INITIALLY REPORTED. THE LENS WAS IMPLANTED ON (B)(6)) 2018 AND EXCHANGED ON (B)(6) 2018 DUE TO OBSERVATION OF LOW VAULT, ELEVATED IOP (INTRAOCULAR PRESSURE) WITHOUT PUPIL BLOCK AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE REPORTER STATED THAT THE CAUSE OF EVENT WAS DUE TO PATIENT RELATED FACTOR. (B)(4). Manufacturer Narrative: -IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD). (B)(4) Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS FOUND. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6447958&pc=MTA | 2023826-2017-00541 | 2017-01-12 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Known Impact Or Consequence To Patient; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12 DIOPTER, IN THE PATIENT'S LEFT EYE (OS). ON (B)(6) 2017 THE LENS WAS TORN BEFORE INJECTION INTO THE EYE, AND IT WAS NOT IMPLANTED. ON (B)(6) 2017 THE LENS WAS EXCHANGE, AND THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BE RETURNED YET. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS TORE/BROKE BEFORE INJECTION INTO THE EYE AND WAS NOT IMPLANTED. THE LENS WAS EXCHANGED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: THE DEVICE WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND THAT THE LENS WAS MISSING A PIECE OF THE HAPTIC. CORRECTED DATA: (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER STATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS TORE/BROKE BEFORE INJECTION INTO THE EYE AND WAS NOT IMPLANTED. THE PATIENT EXPERIENCED POST-OPERATIVE ELEVATED INTRAOCULAR PRESSURE. A DIFFERENT LENS OF THE SAME MODEL AND DIOPTER WAS IMPLANTED ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. RESULT: WORK ORDER SEARCH PERFORMED, NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7647516&pc=MTA | 2023826-2018-00983 | 2018-05-31 | Injury | STAAR SURGICAL COMPANY | 2018-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Known Impact Or Consequence To Patient; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE REPORTER INDICATED THE LENS WILL BE EXPLANTED AND EXCHANGED DUE TO BEING TOO HIGH. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE SURGEON ENLARGED THE PERIPHERAL IRIDECTOMIES USING YAG. THE INCISION WAS ENLARGED AND SUTURE WAS REQUIRED. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS EXCHANGE RESOLVED THE PROBLEM. (B)(4) Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8344402&pc=MTA | 2023826-2019-00266 | 2018-12-20 | Injury | STAAR SURGICAL COMPANY | 2019-02-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Known Impact Or Consequence To Patient; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. A LENS VAULT VALUE OF 1150 UM WAS REPORTED AND THE LENS REMAINS IMPLANTED WITH PLANS TO REMOVE AND EXCHANGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -6.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND HIGH IOP. THE PATIENT CONTROLLED THE HIGH IOP WITH ANTI-IOP DROPS UNTIL THE LENS WAS EXPLANTED. THE SURGEON PLANS TO IMPLANT A SHORTER LENGTH LENS. DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE ON LENS IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, LENS EXPLANT. PATIENT CODE 2692: CODE IN ORIGINAL MDR NO LONGER APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8172146&pc=MTA | 2023826-2018-01871 | 2018-05-19 | Injury | STAAR SURGICAL COMPANY | 2018-12-17 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; No Apparent Adverse Event | Intraocular Pressure Increased; No Known Impact Or Consequence To Patient; No Code Available | NA | NA | 1 | Event Description: A VICMO12.6, 12.6MM IMPLANTABLE COLLAMER LENS, DIOPTER -14.0 WAS RETURNED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL.Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -14.0 INTO THE PATIENT'S RIGHT (OD) EYE. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, LENS DISLOCATION/SUBLUXATION, PUPIL BLOCK, AND ELEVATED IOP. THE SURGEON ALSO REPORTS ENLARGEMENT OF PI. REPORTEDLY, THE PROBLEM IS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: TYPE OF REPORTABLE EVENT-CORRECTED. DEVICE EVALUATED BY MFR THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. METHOD CODE ADDED. RESULT CODE UPDATED. CONCLUSION CODE UPDATED. CLAIM # (B)(4). Manufacturer Narrative: PLEASE DISREGARD PREVIOUS STATEMENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6430437&pc=MTA | 2023826-2017-00461 | NA | Injury | STAAR SURGICAL COMPANY | 2017-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE PATIENT EXPERIENCED BLURRED VISION AND PAIN. A SHORTER LENS WAS IMPLANTED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS RETURNED WITH STICKY RESIDUE ON THE LENS AND UNABLE TO EVALUATE DUE TO RESIDUE. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6468284&pc=MTA | 2023826-2017-00598 | NA | Injury | STAAR SURGICAL COMPANY | 2017-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED BLURRED VISION AND PAIN. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (THE EVENTS WERE NOT REPORTED ON THE INITIAL MDR). THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND ANGLE CLOSURE. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6574198&pc=MTA | 2023826-2017-00833 | NA | Injury | STAAR SURGICAL COMPANY | 2017-05-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER; IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER STATED THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING; ELEVATED INTRAOCULAR PRESSURE; NARROWING OF THE ANGLE; ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT EXPERIENCED BLURRY VISION AND PAIN. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF LENS OPTIC AND HAPTIC TORN OFF AND MISSING. A PIECE OF THE OTHER HAPTIC WAS TORN OFF AND MISSING. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED A SHORTER LENS (13.2MM) WAS IMPLANTED ON (B)(6) 2017 AND THE LENS EXCHANGE RESOLVED THE PROBLEM. THE LENS VAULT WAS GOOD AND THE INTRAOCULAR PRESSURE WAS NORMALIZED. THE PATIENT IS DOING GOOD, WITH A VISUAL ACUITY OF 20/20. (B)(4). Manufacturer Narrative: (B)(4). ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6574253&pc=MTA | 2023826-2017-00834 | NA | Injury | STAAR SURGICAL COMPANY | 2017-05-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER; IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER STATED THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING; ELEVATED INTRAOCULAR PRESSURE; NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT EXPERIENCED BLURRY VISION AND PAIN. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN, WITH A PIECE TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED A SHORTER LENS (13.2MM) WAS IMPLANTED ON (B)(6) 2017 AND THE LENS EXCHANGE RESOLVED THE PROBLEM. THE LENS VAULT WAS GOOD AND THE INTRAOCULAR PRESSURE WAS NORMALIZED. THE PATIENT IS DOING GOOD, WITH A VISUAL ACUITY OF 20/20. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5661682&pc=MTA | 2023826-2016-00729 | 2016-04-19 | Injury | STAAR SURGICAL COMPANY | 2016-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO HIGH VAULT. THE LENS WAS NOT EXCHANGED FOR ANOTHER ICL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD). THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE PATIENT EXPERIENCED PAIN AND BLURRY VISION. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT CODE(S): (OPERATIONAL PROBLEM) : VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE. THERE WAS EVIDENCE OF DRY RESIDUE. (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, THE DEVICES ASSOCIATED WITH THE WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, WAS MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5661683&pc=MTA | 2023826-2016-00730 | 2016-04-19 | Injury | STAAR SURGICAL COMPANY | 2016-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO HIGH VAULT. THE LENS WAS NOT EXCHANGED FOR ANOTHER ICL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE PATIENT EXPERIENCED PAIN AND BLURRY VISION. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW RESULT CODE(S): (OPERATIONAL PROBLEM) : THE PRODUCT FOR THIS EVENT WAS RETURNED BUT WAS UNABLE TO EVALUATE DUE TO THE LENS WAS STUCK TO THE LID OF THE CONTAINER THE LENS WAS RETURNED IN. THE LENS WAS RETURNED DRY. (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, THE DEVICES ASSOCIATED WITH THE WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, WAS MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14138530&pc=MTA | 2023826-2022-01231 | 2022-02-11 | Injury | STAAR SURGICAL COMPANY | 2022-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -13.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND LATER ON THIS SAME DATE A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "BIGGER LENS". Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM #: (B)(4). Manufacturer Narrative: B5 - PATIENT EXPERIENCED PAIN AND ELEVATED INTRAOCULAR PRESSURE - SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6832461&pc=MTA | 2023826-2017-01383 | 2017-08-02 | Injury | STAAR SURGICAL COMPANY | 2017-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER STATED THE PATIENT WAS IN PAIN AND HAD ELEVATED INTRAOCULAR PRESSURE (PRESSURE SPIKE). THE PERIPHERAL IRIDOTOMIES WERE CLOSED, WITH A SHALLOW CHAMBER AND ANGLE CLOSURE. THE SURGEON DECIDED TO EXPLANT THE LENS THE SAME DAY. A THIRD LENS WAS NOT IMPLANTED. THE LENS WAS DISCARDED. THE PATIENT IS DOING FINE. THIS LENS WAS THE EXCHANGE LENS FOR A LONGER LENS IMPLANTED - SEE MDR #2023826-2017-01382. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611920&pc=MTA | 2023826-2019-00780 | 2019-04-16 | Injury | STAAR SURGICAL COMPANY | 2019-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ANGLE CLOSURE WITH ELEVATED IOP (35 MMHG) AND EXCESSIVE VAULT WAS OBSERVED. "AFTER AN ADDITIONAL LPI, WITH A ROUND OF DIAMOX, COMBIGAN, AND SIMBRINZA THE IOP DID RESPOND." ENLARGEMENT OF PI, YAG, AND THE ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, AND ADDITIONAL SUTURES WERE DONE ON (B)(6) 2018, PER REPORTER AS OF (B)(6) 2019 THE PATIENT HAS PAIN AND PHOTOPHOBIA. PER REPORTER AS OF (B)(6) 2019 THE PATIENT'S ACUITY IS CORRECTABLE TO 20/20 OU. REPORTEDLY "PATIENT IS NOT INTERESTED IN REPLACEMENT AT THIS TIME." Manufacturer Narrative: ENLARGEMENT OF THE PERIPHERAL IRIDOTOMIES WAS PERFORMED BY YAG LASER. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS WAS PERFORMED. AN ADDITIONAL SUTURE WAS REQUIRED. THE ICL WAS EXPLANTED ON (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: 1 SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611924&pc=MTA | 2023826-2019-00781 | 2019-04-16 | Injury | STAAR SURGICAL COMPANY | 2019-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ANGLE CLOSURE WITH ELEVATED IOP (35 MMHG) AND EXCESSIVE VAULT WAS OBSERVED. "AFTER AN ADDITIONAL LPI, WITH A ROUND OF DIAMOX, COMBIGAN, AND SIMBRINZA THE IOP DID RESPOND." ENLARGEMENT OF PI, YAG, AND THE ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, AND ADDITIONAL SUTURES WERE DONE ON (B)(6) 2018, PER REPORTER AS OF (B)(6) 2019 THE PATIENT HAS PAIN AND PHOTOPHOBIA. PER REPORTER AS OF (B)(6) 2019 THE PATIENT'S ACUITY IS CORRECTABLE TO 20/20 OU. REPORTEDLY "PATIENT IS NOT INTERESTED IN REPLACEMENT AT THIS TIME." Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM MICL 13.2 OF -9.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE(OS) ON (B)(6) 2019. ANGLE CLOSURE WITH ELEVATED IOP (35MMHG) AND EXCESSIVE VAULT WAS OBSERVED. "AFTER AN ADDITIONAL LPI WITH A ROUND OF DIAMOX; COMBIGAN AND SIMBRINZA THE IOP DID RESPOND." ENLARGEMENT OF PERIPHERAL IRIDOTOMIES; YAG AND THE ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS AND ADDITIONAL SUTURES WERE DONE AND LENS WAS EXPLANTED ON (B)(6) 2019. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9782966&pc=MTA | 2023826-2020-00405 | NA | Injury | STAAR SURGICAL COMPANY | 2020-03-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, - 8.50 DIOPTER, IN THE PATIENTS EYE ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS WAS TOO LARGE, HIGH PRESSURE SPIKE AND EYE PAIN. A SHORTER LENS WILL BE IMPLANTED WITHIN THE NEXT MONTH. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9782969&pc=MTA | 2023826-2020-00406 | NA | Injury | STAAR SURGICAL COMPANY | 2020-03-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, - 8.50 DIOPTER, IN THE PATIENTS EYE ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS WAS TOO LARGE, HIGH PRESSURE SPIKE AND EYE PAIN. A SHORTER LENS WILL BE IMPLANTED WITHIN THE NEXT MONTH. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5596421&pc=MTA | 2023826-2016-00507 | NA | Injury | STAAR SURGICAL COMPANY | 2016-04-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pain; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A PRESSURE SPIKE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS IMPLANTED AND REMOVED FROM THE RIGHT EYE (OD) DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED NARROWING OF THE ANGLE, SHALLOWING OF THE ANTERIOR CHAMBER DEPTH AND A LITTLE PAIN. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS NOT THE DEVICE BUT WAS DUE TO LENS SIZE SELECTION (TOO LARGE). Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). RESULT: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6762431&pc=MTA | 2023826-2017-01260 | 2017-07-07 | Injury | STAAR SURGICAL COMPANY | 2017-08-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2017 THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, DIOPTER -11.0, INTO THE PATIENT'S EYE. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING, ELEVATED IOP, AND ANGLE CLOSURE. ATTEMPTS TO OBTAIN FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: ADDITIONAL SYMPTOMS INCLUDE: IRIDECTOMY, BLURRED VISION, PAIN, PUPIL BLOCK. LENS EXCHANGE WAS PERFORMED ON THE RIGHT (OD) EYE. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN THE LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: PROBLEM WAS RESOLVED WITH LENS EXCHANGE. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. PREVIOUS IN MDR FOLLOW-UP#1: ADDITIONAL SYMPTOM: IRIDECTOMY. _CORRECTED: IRIDECTOMY IS NOT A SYMPTOM, IT WAS PERFORMED PRIOR ICL IMPLANTATION SURGERY. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6762437&pc=MTA | 2023826-2017-01262 | 2017-07-07 | Injury | STAAR SURGICAL COMPANY | 2017-08-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2017 THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -14.5, INTO THE PATIENT'S EYE. IT IS PLANNED TO REPLACE THE LENS WITH A SHORTER LENS DUE TO HIGH VAULTING AND PRESSURE SPIKE . ATTEMPTS TO OBTAIN FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. PREVIOUS IN MDR FOLLOW-UP#1: ADDITIONAL SYMPTOMS: PI REVISION, NARROWING AND CLOSURE OF ANGLES, IRIDECTOMY, PAIN, BLURRED VISION, PUPILLARY BLOCK. THE LENS WAS EXCHANGED FOR A SHORTER LENS, SURGERY OCCURRED IN THE RIGHT (OD) EYE. ADDITIONAL SYMPTOMS: NARROWING AND CLOSURE OF ANGLES, PAIN, BLURRED VISION, PUPILLARY BLOCK. IRIDECTOMY WAS PERFORMED PRIOR LENS IMPLANT SURGERY AND PI WAS REVISED PRIOR TO EXPLANTATION. THE LENS WAS EXCHANGED FOR A SHORTER LENS, SURGERY OCCURRED IN THE LEFT EYE (OS). (B)(4). Manufacturer Narrative: ADDITIONAL SYMPTOMS: PI REVISION, NARROWING AND CLOSURE OF ANGLES, IRIDECTOMY, PAIN, BLURRED VISION, PUPILLARY BLOCK. THE LENS WAS EXCHANGED FOR A SHORTER LENS, SURGERY OCCURRED IN THE RIGHT (OD) EYE. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6762432&pc=MTA | 2023826-2017-01261 | 2017-07-08 | Injury | STAAR SURGICAL COMPANY | 2017-08-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2017 THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -14.5, INTO THE PATIENT'S EYE. IT IS PLANNED TO REPLACE THE LENS WITH A SHORTER LENS DUE TO "A LOT" OF VAULTING AND PRESSURE SPIKE . ATTEMPTS TO OBTAIN FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: UPDATED INFORMATION: THE LENS WAS EXCHANGED IN THE RIGHT (OD) EYE ON (B)(6) 2017 FOR A SHORTER LENS. AFTER THE LENS EXCHANGE THE IOP DECREASED TO 13 (PREVIOUSLY SPIKED AT 68). ADDITIONAL SYMPTOMS INCLUDE NARROWING AND ANGLE CLOSURE, IRIDECTOMY, PAIN , BLURRED VISION, YAG REVISION, AND PUPIL BLOCK. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: PREVIOUS IN MDR FOLLOW-UP#1: THE LENS WAS EXCHANGED IN THE RIGHT (OD) EYE ON (B)(6) 2017 FOR A SHORTER LENS. AFTER THE LENS EXCHANGE THE IOP DECREASED TO 13 (PREVIOUSLY SPIKED AT 68). ADDITIONAL SYMPTOMS INCLUDE NARROWING AND ANGLE CLOSURE, IRIDECTOMY, PAIN , BLURRED VISION, YAG REVISION, AND PUPIL BLOCK. CORRECTED: THE LENS WAS EXCHANGED IN THE RIGHT (OD) EYE ON (B)(6) 2017 FOR A SHORTER LENS. AFTER THE LENS EXCHANGE THE IOP DECREASED TO 13 (PREVIOUSLY SPIKED AT 68). ADDITIONAL SYMPTOMS INCLUDE NARROWING AND ANGLE CLOSURE, PAIN, BLURRED VISION, AND PUPIL BLOCK. IRIDECTOMY WAS PERFORMED PRIOR LENS IMPLANT SURGERY AND THE SURGEON ALSO PERFORMED AC DECOMPRESSION AND YAG REVISION PRIOR TO EXPLANTATION. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4743979&pc=MTA | 2023826-2015-00474 | 2015-04-09 | Injury | STAAR SURGICAL COMPANY | 2015-05-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE - 50MM HG (PUPILLARY BLOCK), PAIN AND BLURRY VISION. ANOTHER PERIPHERAL IRIDECTOMY WAS PERFORMED. A SHORTER LENS WAS IMPLANTED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/25, THE CHAMBER WAS SHALLOW AND THE EYE IRRITATED. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - GIVEN THE REPORTED INFORMATION, THE EVENT WAS CAUSED BY EXCESSIVE VAULTING OF THE ICL LENS. THE HIGH IOP WAS THE CONSEQUENCE OF PUPILLARY BLOCK CAUSED BY EXCESSIVE VAULTING OF THE LENS. THE LENS EXPLANTATION, EXCHANGE AND ADDITIONAL IRIDOCTOMY WERE NECESSARY TO PREVENT FURTHER DAMAGE. ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. (B)(4). Manufacturer Narrative: METHOD (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. AS REPORTED ON THE MEDICAL REVIEW, EXCESSIVE VAULTING IS A CONSEQUENCE OF WRONG LENS USE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF WHITE TO WHITE MEASUREMENTS BASED ON TECHNIQUES, IMPROPER SULCUS TO SULCUS MEASUREMENT, AND PATIENT CONDITION, POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS, IRREGULAR SULCUS OR CILIARY SULCUS CYST). (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5804108&pc=MTA | 2023826-2016-00958 | 2016-06-02 | Injury | STAAR SURGICAL COMPANY | 2016-07-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pain; Pupillary Block; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO AN EXCESSIVE VAULT. THE PATIENT EXPERIENCED AN ELEVATED IOP, PUPILLARY BLOCK, POST-OP PAIN, AND CLOUDY VISUAL ACUITY (VA). THE REPORTER INDICATED THE EVENT WAS PRODUCT-RELATED AND THE DEVICE CONTRIBUTED TO THE EVENT. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE REPORTER STATED THE CAUSE OF THE EVENT WAS WHITE-TO-WHITE MISMEASUREMENT. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/25. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN VIAL IN LIQUID. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. DEVICE EVALUATED BY MANUFACTURER? NO - 81 (OTHER): LENS NOT RETURNED. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7119246&pc=MTA | 2023826-2017-01886 | 2017-11-16 | Injury | STAAR SURGICAL COMPANY | 2017-12-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pain; Retinal Detachment; Vitreous Hemorrhage; Halo; Vitrectomy; No Code Available | P030016 | NA | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2017 AN ICL WAS REMOVED DUE TO RETINAL DETACHMENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 OF -12.0 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO RETINAL TEAR LEADING TO ITS SUBSEQUENT DETACHMENT. THE PATIENT ALSO HAD ELEVATED IOP, NARROWING OF ANGLE AND EXPERIENCED "SOME HALOS AND FLOATING LIGHTS". ADDITIONAL PERIPHERAL IRIDECTOMY BROUGHT THE ELEVATED IOP BACK TO PRE-OP LEVEL AND THE ENLARGEMENT OF PERIPHERAL IRIDOTOMY PERFORMED INCREASED DEPTH. VITRECTOMY WAS ALSO PERFORMED. THE SURGEON REPORTED THAT THE PATIENT IS "CURRENTLY APHAKIC WILL NEED LENS IMPLANT". (B)(4). METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 OF -12.0 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2017. ON (B)(6) 2017 THE LENS (AND NATURAL LENS) WAS EXPLANTED DUE TO RETINAL TEAR LEADING TO ITS SUBSEQUENT DETACHMENT. LASER RETINOPEXY WAS PERFORMED FOLLOWED BY PNEUMATIC CRYOPEXY ONLY TO HAVE THE CURTAIN IN PATIENT'S VISION RETURN. VITRECTOMY PERFORMED LED TO VITREOUS HEMORRHAGE. A SECOND RETINAL TEAR WAS NOTICED LEADING TO A SECOND VITRECTOMY, SCLERAL BUCKLE AND "LENSECTOMY" OR EXPLANT OF THE MICL AND NATURAL LENS.THE PATIENT ALSO HAD ELEVATED IOP, NARROWING OF ANGLE, PAIN AND EXPERIENCED "SOME HALOS AND FLOATING LIGHTS". ADDITIONAL PERIPHERAL IRIDECTOMY BOUGHT THE ELEVATED IOP BACK TO PRE-OP LEVEL AND THE ENLARGEMENT OF PERIPHERAL IRIDOTOMY PERFORMED INCREASED DEPTH. THE SURGEON REPORTED THAT THE PATIENT IS "CURRENTLY APHAKIC WILL NEED LENS IMPLANT". (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7583196&pc=MTA | 2023826-2018-00869 | NA | Injury | STAAR SURGICAL COMPANY | 2018-06-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Patient Problem/Medical Problem; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2, -10.5 DIOPTER , IMPLANTABLE COLLAMER LENS AND RESULTED IN A VERY HIGH VAULT , ANGLE CLOSURE WITH ELEVATED IOP (INTRAOCULAR PRESSURE) WAS ALSO OBSERVED. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: CORRECTED DATA: CORRECT PATIENT IDENTIFIER IS (B)(6) NOT (B)(6). COMPLAINT IS AN ADVERSE EVENT; SECONDARY SURGERY WAS PERFORMED TO EXCHANGE LENS. THEREFORE, PRODUCT PROBLEM IS INCORRECT. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. PATIENT CODE:(B)(4) SECONDARY SURGICAL INTERVENTION, EXCHANGED, ANGLE CLOSURE. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11325783&pc=MTA | 2023826-2021-00423 | 2021-01-08 | Injury | STAAR SURGICAL COMPANY | 2021-02-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -15.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS EXPLANTED DUE TO EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED OP. THIS EXPLANT RESOLVED THE PROBLEM. THE REPORTER STATED " THE DOCTOR WILL IMPLANT ANOTHER LENS BUT HE DIDNT DECIDE THE DATE YET." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10537711&pc=MTA | 2023826-2020-02069 | 2020-08-10 | Injury | STAAR SURGICAL COMPANY | 2020-09-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pupillary Block | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO HIGH PRESSURE. THE LENS WAS EXCHANGED FOR A DIFFERENT SIZE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SPECIMEN CUP, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN AND THERE WAS RESIDUE ON LENS. H6: HEALTH EFFECT IMPACT CODE: 4625 - ADDITIONAL SURGERY; YAG - ADDITION OF NEW PI. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE. ADDITIONAL SURGERY WAS PERFORMED, ADDITION OF NEW PI BY YAG AND ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS NOT RESOLVED. SEE MFR REPORT # 2023826-2020-02118 FOR EXCHANGED LENS. THE CAUSE OF THE EVENT WAS DUE TO PATIENT-RELATED FACTOR. CLAIM # (B)(4). Manufacturer Narrative: DEVICE BLA NUMBER: NA. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4773364&pc=MTA | 2023826-2015-00555 | 2015-04-13 | Injury | STAAR SURGICAL COMPANY | 2015-05-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Malposition of Device | Intraocular Pressure Increased; Pupillary Block | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/80. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. PUPIL BLOCK WITH ELEVATED IOP IN THE PRESENCE OF A VICMO/VTICMO MODEL MAY OCCUR DUE TO: REMAINING VISCOELASTIC IN THE POSTERIOR CHAMBER, OVERSIZED ICL WITH ANGLE CLOSURE, TOO NARROW ANGLES/CROWDED AC DUE TO SMALL EYE ANATOMY PREOPERATIVELY, UNEXPECTED ABNORMAL ANATOMY OR TISSUE ABNORMALITIES (PRESENCE OF IRIS/CILIARY BODY CYSTS), BLOCKAGE OF THE PORT BY REMAINING VISCOELASTIC, OR BY AN INFLAMMATORY MEMBRANE, ETC. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12321641&pc=MTA | 2023826-2021-03143 | 2021-07-12 | Injury | STAAR SURGICAL COMPANY | 2021-08-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THAT A 13.2MM,MICL13.2, -13.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS EXPLANTED. PATIENT HAD A PUPILLARY BLOCK. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5-THE REPORTER INDICATE THAT A 13.2MM, MICL13.2, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGE FOR A SHORTER LENGTH LENS DUE TO PUPIL BLOCK WITH ELEVATED IOP. THIS EXCHANGE RESOLVED THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - PATIENT ALSO EXPERIENCED UNREACTIVE (FIXED) PUPIL - SHOULD HAVE BEEN INCLUDED. H6 - CODE 4628 SHOULD BE CORRECTED TO 4629. H6 - HEALTH EFFECT CLINICAL CODE: 4581- UNREACTIVE FIXED PUPIL SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15423694&pc=MTA | 2023826-2022-03076 | 2022-05-15 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL 13.2 IMPLANTABLE COLLAMER LENS OF -12.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE(OS) ON (B)(6) 2022. PUPIL BLOCK WITH ELEVATED IOP; ANGLE CLOSURE WITH ELEVATED IOP WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND AT A LATER DATE A REPLACEMENT LENS OF A DIFFERENT MODEL BUT SAME SIZE AND DIOPTER WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS DEVICE. Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10787190&pc=MTA | 2023826-2020-02583 | 2020-10-07 | Injury | STAAR SUGICAL COMPANY | 2020-11-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT 12.6MM, MICL12.6, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED ON (B)(6) 2020 AND THE PATIENT EXPERIENCED NARROW ANGLE DUE TO EXCESSIVE VAULT, LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT 12.6MM, MICL12.6,-9.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK, ANGLE CLOSURE AND IOP. ADDITIONAL SURGICAL INTERVENTION (ENLARGEMENT OF PI, YAG) AND MEDICAL INTERVENTION (DROP THERAPY) WAS ALSO PERFORMED/PRESCRIBED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS THE DEVICE, REPORTER STATED " TOO LONG LENS WITH EXCESSIVE VAULT PERIPHERAL IRIDOTOMY OCCLUDED WITH IRIS VAULT." FOR STATUS OF THE EYE THE REPORTER STATED " PATIENT FEELS VISION HAS RETURNED BACK TO NORMAL. PATIENTS IOP HAVE BEEN STABLE SINCE BEING OFF DROP THERAPY SINCE (B)(6) 2020. WILL LIKE TO SEE IOPS STABLE FOR A COUPLE MONTHS BEFORE PROCEEDING WITH ANOTHER ICL PROCEDURE." CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - 4625 - ENLARGEMENT OF PI, SHOULD HAVE BEEN SPECIFIED IN PREVIOUS MEDWATCH REPORTS FOR CODE 4625. B5 - "YAG": SHOULD BE REMOVED AS IT WAS NOT SPECIFIED WHAT TYPE OF SURGERY WAS PERFORMED. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10787942&pc=MTA | 2023826-2020-02584 | 2020-10-07 | Injury | STAAR SUGICAL COMPANY | 2020-11-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT 12.6MM, MICL12.6, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP (INTRAOCUAR PRESSURE), PUPIL BLOCK AND ANGLE CLOSURE. ADDITIONAL SURGICAL INTERVENTION (ENLARGEMENT OF PI, YAG) AND MEDICAL INTERVENTION WAS ALSO PERFORMED/ PRESCRIBED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS THE DEVICE " TOO LONG LENS". FOR STATUS OF THE EYE THE REPORTER STATED " EXPLANTED ICL'S CONT. IOP, LARGELY 20 + PIGMENT FROM ENLARGING PI'S." THE FOLLOWING ADDITIONAL INFORMATION WAS REPORTED " PT. REQUIRED MULTIPLE WOUND BURRS OVER MANY DAYS ALONG WITH MAX MED. THERAPY: LUMIGAN, ALPHAGAN, TIMOLOL, TRUSOPT DIAMOX 500 BID. NOW UNDER CONTROL ONLY LUMIGAN." Manufacturer Narrative: B5: THE REPORTER INDICATED THAT 12.6MM, MICL12.6,-9.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK, ANGLE CLOSURE AND IOP. ADDITIONAL SURGICAL INTERVENTION (ENLARGEMENT OF PI, YAG) AND MEDICAL INTERVENTION (DROP THERAPY) WAS ALSO PERFORMED/PRESCRIBED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS THE DEVICE, REPORTER STATED " TOO LONG LENS WITH EXCESSIVE VAULT PERIPHERAL IRIDOTOMY OCCLUDED WITH IRIS VAULT." FOR STATUS OF THE EYE THE REPORTER STATED " PATIENT FEELS VISION HAS RETURNED BACK TO NORMAL. PATIENTS IOP HAVE BEEN STABLE SINCE BEING OFF DROP THERAPY SINCE 23 OCT 2020. WILL LIKE TO SEE IOPS STABLE FOR A COUPLE MONTHS BEFORE PROCEEDING WITH ANOTHER ICL PROCEDURE." CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - 4625 - ENLARGEMENT OF PI, SHOULD HAVE BEEN SPECIFIED IN PREVIOUS MEDWATCH REPORTS FOR CODE 4625. B5 - "YAG": SHOULD BE REMOVED AS IT WAS NOT SPECIFIED WHAT TYPE OF SURGERY WAS PERFORMED. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13430229&pc=MTA | 2023826-2022-00283 | 2021-11-23 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE (IOP) AND ANGLE CLOSURE. THERE WAS ADDITION OF NEW PI AND ENLARGEMENT OF PI BY YAG ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO LENS CALCULATOR ERROR. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. (B)(4). CLAIM # (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822369&pc=MTA | 2023826-2023-04278 | 2023-08-22 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -10.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE AND UNREACTIVE (FIXED) PUPIL. CAUSE OF EVENT WAS REPORTED AS "DEVICE". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: UNREACTIVE/FIXED PUPIL. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: B5 - REPORTEDLY, PROBLEM IS RESOLVED. "CURRENTLY THE PATIENT IS PRESENTING A LIGHT CORNEAL EDEMA. A SPECULAR MICROSCOPY IS SCHEDULED IN 30-45 DAYS. THEN IF THE NUMBER AND SHAPE OF ENDOTHELIAL CELLS IS OK A SECOND IMPLANTATION OF AN ICL OF 12.1MM WIDE WILL BE PLANNED." CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15423696&pc=MTA | 2023826-2022-03075 | 2022-05-15 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM MICL 13.7 IMPLANTABLE COLLAMER LENS OF -7.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 13-MAY-2022. ON 06-JUN-2022 THE LENS WAS EXPLANTED DUE TO PUPIL BLOCK WITH ELEVATED IOP; ANGLE CLOSURE WITH ELEVATED IOP; ADDITION OF NEW PI. REPORTEDLY THE PROBLEM WAS RESOLVED. "PATIENT REFUSED TO HAVE A REPLACEMENT ICL IMPLANTED. WILL CONSIDER LASIK OR PRK IN FUTURE." CAUSE REPORTED AS DEVICE. Manufacturer Narrative: (B)(4).CLAIM#: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17672736&pc=MTA | 2023826-2023-03646 | 2023-05-19 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023, DUE TO EXCESSIVE VAULTING, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE (IOP) AND BLURRED VISION. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THERE WAS 1+ PIGMENT ANTERIOR CHAMBER, THE IRIS WAS RESOLVING. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17673554&pc=MTA | 2023826-2023-03647 | 2023-05-19 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023, DUE TO EXCESSIVE VAULTING, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE (IOP) AND BLURRED VISION. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THERE WAS 1+ PIGMENT ANTERIOR CHAMBER, THE IRIS WAS RESOLVING. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558798&pc=MTA | 2023826-2021-00703 | 2021-03-08 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, INTO THE PATIENT'S LEFT (OS) EYE. THE SURGEON REPORTS THE LENS "VAULTED TOO MUCH" AND HIGH PRESSURE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: B1-CORRECTED B5-ADDITIONAL INFORMATION: ON (B)(6) 2021 THE LENS WAS EXPLANTED. ADDITIONAL SYMPTOMS REPORTED: TRANSIENT LOSS OF BCVA, PUPIL BLOCK, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE AND UNKNOWN: "POSSIBLE FAILURE OF PI'S." H6-CLINICAL CODE 4581 [PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL]. HEALTH IMPACT CODE CORRECTED. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION-UNK. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14493426&pc=MTA | 2023826-2022-01674 | 2022-04-23 | Injury | STAAR SURGICAL COMPANY | 2022-05-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL13.2 -12.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT, PUPIL BLOCK, ANGLE CLOSURE, ELEVATED INTRAOCULAR PRESSURE, UNREACTIVE (FIXED) PUPIL, AND BLURRED VISION. ADDITIONAL MEDICAL INTERVENTION (AQUEOUS SUPPRESSANTS AND DIOMOX) WERE PRESCRIBED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH, SAME MODEL AND POWER LENS. IT WAS REPORTER THAT THE EXCHANGE RESOLVED THE PROBLEM. HOWEVER, IT WAS ALSO REPORTED THAT THE VAULT IS STILL HIGH. FOR STATUS OF THE EYE THE REPORTER STATED " QUIET, COMFORTABLE, STABLE, HAPPY PATIENT." IN THE REPORTERS OPINION THE DEVICE AND PATIENT-RELATED FACTOR WAS THE CAUSE OF THIS EVENT. Manufacturer Narrative: HEALTH EFFECT- CLINICAL CODE (B)(4). TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14493429&pc=MTA | 2023826-2022-01673 | 2022-04-23 | Injury | STAAR SURGICAL COMPANY | 2022-05-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL13.2 -11.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT, PUPIL BLOCK, ANGLE CLOSURE, ELEVATED INTRAOCULAR PRESSURE, UNREACTIVE (FIXED) PUPIL, AND BLURRED VISION. ADDITIONAL MEDICAL INTERVENTION (AQUEOUS SUPPRESSANTS AND DIOMOX) WERE PRESCRIBED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH, SAME MODEL AND POWER LENS. IT WAS REPORTER THAT THE EXCHANGE RESOLVED THE PROBLEM. HOWEVER, IT WAS ALSO REPORTED THAT THE VAULT IS STILL HIGH. FOR STATUS OF THE EYE THE REPORTER STATED " QUIET, COMFORTABLE, GOOD VISION, HAPPY." IN THE REPORTERS OPINION THE DEVICE AND PATIENT-RELATED FACTOR IS THE CAUSE OF THIS EVENT. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- "THE REPORTER INDICATED THAT A 13.2MM MICL13.2 -11.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT, PUPIL BLOCK, ANGLE CLOSURE, ANGLE CLOSURE, ELEVATED INTRAOCULAR PRESSURE, UNREACTIVE (FIXED) PUPIL AND BLURRED VISION. ADDITIONAL MEDICAL INTERVENTION (AQUEOUS SUPPRESSANTS AND DIOMOX) WAS PRESCRIBED. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND PROBLEM RESOLVED. STATUS OF THE EYE WAS REPORTED AS "QUIET, COMFORTABLE, GOOD VISION, HAPPY". HOWEVER IT WAS ALSO REPORTED THE VAULT IS STILL HIGH. IN THE REPORTER'S OPINION THE DEICE AND PATIENT RELATED FACTOR IS THE CAUSE OF THIS EVENT." H6- HEALTH EFFECT: IMPACT CODE- 4627 SHOULD BE CORRECTED TO 4629. D6B- (B)(6) 2022 CORRECTED TO (B)(6) 2022. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- "THE REPORTER INDICATED THAT A 13.2MM MICL13.2 -11.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT, PUPIL BLOCK, ANGLE CLOSURE, ANGLE CLOSURE, ELEVATED INTRAOCULAR PRESSURE, UNREACTIVE (FIXED) PUPIL AND BLURRED VISION. ADDITIONAL MEDICAL INTERVENTION (AQUEOUS SUPPRESSANTS AND DIOMOX) WAS PRESCRIBED. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON (B)(6) 2022 THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS AND PROBLEM RESOLVED. STATUS OF THE EYE WAS REPORTED AS "QUIET, COMFORTABLE, GOOD VISION, HAPPY". HOWEVER IT WAS ALSO REPORTED THE VAULT IS STIL HIGH. IN THE REPORTER'S OPINION THE DEICE AND PATIENT RELATED FACTOR IS THE CAUSE OF THIS EVENT." H6- HEALTH EFFECT: IMPACT CODE- 4627. D6B- (B)(6) 2022. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18509197&pc=MTA | 2023826-2024-00227 | 2023-10-16 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -15.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT, TRANSIENT LOSS OF BCVA, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, PUPIL BLOCK,ANGLE CLOSURE, UNREACTIVE (FIXED) PUPIL AND IRIS ATROPHY WAS REPORTED. ON (B)(6) 2023 THE LENS WAS EXPLANTED. MILD CORTICAL LENS OPACITY WAS OBSERVED ON (B)(6) 2023. IT WAS REPORTED THAT A MONOFOCAL LENS WAS IMPLANTED. IT WAS REPORTED THAT THE PROBLEM RESOLVED HOWEVER, IT WAS ALSO REPORTED THAT IT DID NOT RESOLVE. ADDITIONAL INFORMATION AND CLARIFICATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581:SIGNIFICANT REDUCTION OF ICA, ANGLE CLOSURE,UNREACTIVE/FIXED PUPIL, IRIS ATROPHY. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4727938&pc=MTA | 2023826-2015-00471 | 2015-03-25 | Injury | STAAR SURGICAL COMPANY | 2015-04-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE - 50MM HG (PAPILLARY BLOCK), NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT'S VISION WAS BLURRY. ANOTHER PERIPHERAL IRIDECTOMY WAS PERFORMED BUT WITH NO IMPROVEMENT. A SHORTER LENS WAS IMPLANTED ON (B)(6) 2015. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/15, THE PATIENT IS STABLE AND DOING WELL. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - ANTERIOR CHAMBER ANGLE CLOSURE, PUPILLARY BLOCKAGE AND INCREASED IOP ARE CONSEQUENCES OF EXCESSIVE VAULTING. ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: BASED ON THE RESULTS OF THE DHR REVIEW, THE AVAILABLE INFORMATION GIVEN WITHIN THIS COMPLAINT, PRODUCT EVALUATION AND WORK ORDER SEARCH, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. AS REPORTED ON THE MEDICAL REVIEW, IT IS LIKELY THE EXCESSIVE VAULTING WAS A CONSEQUENCE OF WRONG LENS USE. HOWEVER, NO DEFINITE ROOT CAUSE DETERMINED. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5305169&pc=MTA | 2023826-2015-01689 | 2015-11-01 | Injury | STAAR SURGICAL COMPANY | 2015-12-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, PUPILLARY BLOCK, NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT EXPERIENCED BLURRY VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE REPORTER STATED THE CAUSE OF THE EVENT WAS PROBABLY DUE TO SIZING ISSUE. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THIS LENS THAT WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. THE REPORTER STATED, THE CAUSE OF THE EVENT WAS DUE TO SIZING ISSUE. SINCE THE EXCHANGE WITH THE SMALLER SIZE LENS RESOLVED THE ISSUE, ALTHOUGH NOT CONFIRMED, IT IS LIKELY THE CAUSE OF THE EVENT IS DUE TO LENS SIZING. THE PRODUCT WAS NOT RETURNED TO STAAR FOR AN EVALUATION. MEDICAL REVIEW - REPORTEDLY, MICL( 13.2MM) WAS EXPLANTED/EXCHANGED APPROXIMATELY ONE AND A HALF MONTHS POSTOPERATIVELY FOR A SHORTER SIZE MICL (12.6 MM) TO ADDRESS HIGH VAULT AND SUBSEQUENT COMPLICATIONS (PUPILLARY BLOCK, ELEVATED IOP, BLURRY VISION). ACCORDING TO THE COMPLETED CUSTOMER COMPLAINT QUESTIONNAIRE FOR SURGERY, DATED ON (B)(6) 2015, THE CAUSE OF THE EVENT WAS UNKNOWN BUT WAS PROBABLY DUE TO SIZING ISSUE. PATIENT WAS UNDER MONITOR AND THE MOST CURRENT UCVA WAS 20/25. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7187682&pc=MTA | 2023826-2018-00043 | 2017-12-12 | Injury | STAAR SURGICAL COMPANY | 2018-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM MICL13.2, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO SPIKED IOP. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: PATIENT'S NAME IS CORRECTED TO (B)(6). THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS OF -9.5 DIOPTER IN THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, ELEVATED IOP (PUPILLARY BLOCK), NARROWING OF ANGLES, SHALLOWING OF ACD AND ANGLE CLOSURE. THERE WAS NO PATIENT INJURY. THE PATIENT CURRENTLY HAS BLURRED VISION, AN ENLARGED/ UNRESPONSIVE PUPIL (URRETS-ZAVALIA SYNDROME) AND PHOTOPHOBIA. THE REPORTER STATED THAT THE DOCTOR MIGHT SUGGEST SOFT CONTACTS TO BE USED BY PATIENT IN ORDER TO IMPROVE THE PATIENT CONDITION. (B)(4). CONCLUSION CODE: AS PER THE DFU, THIS LENS IS CONTRAINDICATED FOR IMPLANTATION IN PATIENTS ABOVE THE AGE OF 45 YEARS AND IN THOSE WITH ACD (ANTERIOR CHAMBER DEPTH) LESS THAN 3.0MM CORRECTED DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT, IN A SMALL POUCH TAPED TO A PAPER WITH FACILITY INFORMATION. VISUAL INSPECTION FOUND THE LENS RETURNED IN PIECES AND PRESENCE OF RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS IN A SMALL POUCH TAPED TO PAPER WITH FACILITY INFORMATION. VISUAL INSPECTION FOUND THE LENS RETURNED IN PIECES BEING UNABLE TO EVALUATE. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5710842&pc=MTA | 2023826-2016-00807 | NA | Injury | STAAR SURGICAL COMPANY | 2016-06-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Prolapse; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (PUPILLARY BLOCK), NARROWING OF THE ANGLE, ANGLE CLOSURE, SHALLOWING OF THE ANTERIOR CHAMBER AND IRIS PROLAPSE. THE PATIENT EXPERIENCED BLURRED VISION. A SECONDARY SURGERY WAS PERFORMED TO LASER THE PERIPHERAL IRIDOTOMIES AND THE LENS WAS REPOSITIONED BUT THE PATIENT STILL HAD PRESSURE ISSUES. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S PUPIL HAS REMAINED LARGE WITH SOME SHALLOWING OF THE ANTERIOR CHAMBER BUT THE ISSUE HAS IMPROVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). RESULTS CODE(S): (OPERATIONAL PROBLEM): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6365697&pc=MTA | 2023826-2017-00273 | 2017-01-04 | Injury | STAAR SURGICAL COMPANY | 2017-02-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Discomfort; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF ICA, PUPIL BLOCK. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS WAS PERFORMED. THERE IS NO PLAN TO EXCHANGE THE LENS. TO DATE, THE PATIENT IS A LITTLE UNCOMFORTABLE BECAUSE THE PUPIL DOES NOT CONSTRICT FULLY. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION-LENS WAS RETURNED DRY IN A CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS, AND DRY CLEAR RESIDUE AND DEBRIS ON LENS. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL INFOMATION: DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12200495&pc=MTA | 2023826-2021-02768 | 2021-06-17 | Injury | STAAR SURGICAL COMPANY | 2021-07-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pupillary Block; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AND REMOVED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, DUE TO THE LENS DID NOT SIT RIGHT IN THE EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, 12.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE (IOP) AND THE LENS DID NOT SIT RIGHT IN THE EYE. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. H6: HEALTH EFFECT - IMPACT CODE: 4627. THIS CODE WAS INCORRECT, THE LENS WAS INSERTED AND REMOVED DURING THE SAME SURGERY. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12708008&pc=MTA | 2023826-2021-04151 | 2021-08-23 | Injury | STAAR SURGICAL COMPANY | 2021-10-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -9.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING; PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5686676&pc=MTA | 2023826-2016-00776 | 2016-04-21 | Injury | STAAR SURGICAL COMPANY | 2016-05-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE REPORTER INDICATED THE PATIENT EXPERIENCED PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE AND PIGMENT DISPERSION. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5680351&pc=MTA | 2023826-2016-00766 | NA | Injury | STAAR SURGICAL COMPANY | 2016-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ANGLE CLOSURE, ELEVATED INTRAOCULAR PRESSURE, PRESSURE SPIKE AND PUPILLARY BLOCK. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE REPORTER STATED THE SURGEON DID NOT BELIEVE THE EVENT WAS LENS RELATED AND WAS NOT A VAULTING ISSUE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: RESULT: (OPERATIONAL PROBLEM) : VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5686683&pc=MTA | 2023826-2016-00777 | 2016-04-21 | Injury | STAAR SURGICAL COMPANY | 2016-05-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER INDICATED THE PATIENT EXPERIENCED PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE AND PIGMENT DISPERSION. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE IMPLANTED 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS HAS DIOPTER OF -7.00. WORK ORDER SEARCH-NO ADDITIONAL SIMILAR COMPLAINT EVENT WITHIN ASSOCIATE LOTS WAS FOUND. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6036794&pc=MTA | 2023826-2016-01468 | 2016-02-25 | Injury | STAAR SURGICAL COMPANY | 2016-10-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -04.50 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND PUPIL BLOCK WITH ELEVATED IOP. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT'S POST-OP ((B)(6) 2016) UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION - THE LENS WAS RETURNED DRY AND BENT IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED. AS PER DFU FOR THIS LENS MODEL, TICLS ARE CONTRAINDICATED FOR PATIENTS UNDER 21 YEARS OLD. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6085632&pc=MTA | 2023826-2016-01612 | 2016-02-29 | Injury | STAAR SURGICAL COMPANY | 2016-11-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON THE SAME DAY DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND PUPIL BLOCK. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6085636&pc=MTA | 2023826-2016-01607 | 2016-02-29 | Injury | STAAR SURGICAL COMPANY | 2016-11-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON THE SAME DAY DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND PUPIL BLOCK. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6350028&pc=MTA | 2023826-2017-00204 | 2016-11-08 | Injury | STAAR SURGICAL COMPANY | 2017-02-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING (VAULT VALUE OF 1200 MICROMETERS). THE PATIENT EXPERIENCED A PUPIL BLOCK (WITH ELEVATED IOP). THE PATIENT ALSO EXPERIENCED AN ENLARGEMENT OF PI OR ADDITION OF NEW PI (SURGICAL). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN, AND FIBERS ON LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6551447&pc=MTA | 2023826-2017-00775 | 2017-03-17 | Injury | STAAR SURGICAL COMPANY | 2017-05-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PUPIL BLOCK WITH ELEVATED IOP, PIGMENT DISPERSION. THE PATIENT ALSO UNDERWENT THE ENLARGEMENT OF PI OR ADDITION OF NEW PI, ENLARGEMENT OF INCISION, AND ADDITIONAL SUTURE SURGERY. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE DR. SAID THE PATIENT HAS BEEN FOLLOWED UP WITH ANOTHER LOCAL SURGEON AND REPORTED A NORMAL VAULT AFTER THE LENS EXCHANGED. AS OF THE DAY OF MDR SUBMISSION, THE LENS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION - PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CONTAINER WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6551455&pc=MTA | 2023826-2017-00776 | 2017-03-17 | Injury | STAAR SURGICAL COMPANY | 2017-05-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PUPIL BLOCK WITH ELEVATED IOP, PIGMENT DISPERSION. THE PATIENT ALSO UNDERWENT THE ENLARGEMENT OF PI OR ADDITION OF NEW PI, ANDENLARGEMENT OF INCISION SURGERY. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS NOT BEEN RETURNED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CONTAINER WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN AND BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, ENLARGEMENT OF PI OR ADDITION OF NEW PI, ENLARGEMENT OF INCISION, SIGNIFICANT REDUCTION OF ICA, PIGMENT DISPERSION). (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6595994&pc=MTA | 2023826-2017-00839 | 2017-04-06 | Injury | STAAR SURGICAL COMPANY | 2017-05-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (PUPILLARY BLOCK). THE PATIENT EXPERIENCED GLARE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6582973&pc=MTA | 2023826-2017-00863 | 2017-04-10 | Injury | STAAR SURGICAL COMPANY | 2017-05-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2, DIOPTER -6.00 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO THE PATIENT EXPERIENCING EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED IOP. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND CLEAR RESIDUE). (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND CLEAR RESIDUE). (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6900415&pc=MTA | 2023826-2017-01488 | 2017-06-30 | Injury | STAAR SURGICAL COMPANY | 2017-09-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE. THE PATIENT EXPERIENCED AN UNREACTIVE PUPIL (FIXED) AND IRIS ATROPHY. THE SURGEON PERFORMED PUPILLOPLASTY. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (FIBERS). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6872831&pc=MTA | 2023826-2017-01439 | 2017-08-18 | Injury | STAAR SURGICAL COMPANY | 2017-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.2, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, (PUPILLARY BLOCK) ELEVATED IOP, NARROWING OF THE ANGLE, ANGLE CLOSURE, AND SHALLOWING OF THE ACD. THE LENS WAS THEN EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. THE REPORTER STATED THAT THE SOFTWARE DOES NOT ALWAYS PREDICT CORRECT SIZE IMPLANT. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.2, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE EXCESSIVE VAULTING, ANGLE CLOSURE, (PUPILLARY BLOCK) ELEVATED IOP, NARROWING OF THE ANGLE, ANGLE CLOSURE, AND SHALLOWING OF THE ACD. (B)(4). Manufacturer Narrative: METHOD CODE: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.2, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, PUPILLARY BLOCK, ELEVATED IOP, NARROWING OF THE ANGLE, ANGLE CLOSURE, SHALLOWING OF THE ACD AND DILATED PUPIL. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND CLEAR RESIDUE ON THE LENS SURFACE. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978078&pc=MTA | 2023826-2017-01615 | 2017-08-27 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -4.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK, ELEVATED IOP, UNREACTIVE (FIXED) PUPIL, AND ANGLE CLOSURE. THE PROBLEM WAS RESOLVED. THERE IS NO PLAN TO EXCHANGE THE LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE ON IT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7504965&pc=MTA | 2023826-2018-00687 | 2017-09-28 | Injury | STAAR SURGICAL COMPANY | 2018-05-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -8.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS REMOVED ON (B)(6) 2017 DUE TO OBSERVATION OF EXCESSIVE VAULT., SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PUPIL BLOCK WITH ELEVATED IOP (INTRAOCULAR PRESSURE). THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: LENS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND A PIECE OF A HAPTIC MISSING. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7504990&pc=MTA | 2023826-2018-00688 | 2017-09-28 | Injury | STAAR SURGICAL COMPANY | 2018-05-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -8.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS REMOVED ON (B)(6) 2017 DUE TO OBSERVATION OF EXCESSIVE VAULT.,SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND PUPIL BLOCK WITH ELEVATED IOP (INTRAOCULAR PRESSURE). THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS SURFACE. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7246094&pc=MTA | 2023826-2018-00151 | 2017-12-21 | Injury | STAAR SURGICAL COMPANY | 2018-02-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM MICL12.6, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, (PUPILIARY BLOCK) ELEVATED IOP, NARROWING OF THE ANGLE, AND ANGLE CLOSURE. ONE SUTURE WAS USED ON REMOVAL. PHACO- EMULSIFICATION WAS PERFORMED ON THE PATIENT AND AN IOL LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. ITS IS NOTED THAT THE PATIENT WEIGHED HER OPTIONS AND DECIDED TO RECEIVE AN IOL OVER ANOTHER ICL. Manufacturer Narrative: CORRECTED DATA: COMBIGAN, TRAVATAN, AND DIAMOX WERE ADMINISTERED FOR TREATMENT OF THE ICL REMOVAL. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. PATIENT CODE: (B)(4) NARROWING OF THE ANGLE, ANGLE CLOSURE, ADDITION OF SUTURE, SECONDARY SURGERY, LENS EXPLANT, PHACO-EMULSIFICATION, IOL IMPLANT. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT UNDER THE AGE OF 21 YEARS OLD. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7246096&pc=MTA | 2023826-2018-00152 | 2017-12-21 | Injury | STAAR SURGICAL COMPANY | 2018-02-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM MICL12.6, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, (PUPILIARY BLOCK) ELEVATED IOP, NARROWING OF THE ANGLE, AND ANGLE CLOSURE. PHACO- EMULSIFICATION WAS PERFORMED ON THE PATIENT AND AN IOL LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. ITS IS NOTED THAT THE PATIENT WEIGHED HER OPTIONS AND DECIDED TO RECEIVE IOL OVER ANOTHER ICL. Manufacturer Narrative: THE PATIENT WAS GIVEN MEDICATION FOR TREATMENT (COMBIGAN, TRAVATAN AND DIAMOX). (B)(4) Manufacturer Narrative: IN FOLLOW-UP #1, CORRECTED TO 20-MAR-2018. (B)(4) Manufacturer Narrative: PATIENT CODE: (B)(4) NARROWING OF THE ANGLE, ANGLE CLOSURE, SECONDARY SURGERY, LENS EXPLANT, PHACO-EMULSIFICATION, IOL IMPLANT. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT UNDER THE AGE OF 21 YEARS OLD. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7602293&pc=MTA | 2023826-2018-00912 | 2018-05-15 | Injury | STAAR SURGICAL COMPANY | 2018-06-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM MICL13.2, -15.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(4) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED WITH A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PUPIL BLOCK AND ANGLE CLOSURE WITH ELEVATED IOP (INTRAOCULAR PRESSURE).ADDITIONAL SURGICAL INTERVENTION (YAG PI) WAS PERFORMED ON (B)(6) 2018. IT WAS REPORTED THAT AT PATIENTS LAST FOLLOW UP, VAULT AND IOP (INTRAOCULAR PRESSURE) WERE MUCH IMPROVED. Manufacturer Narrative: ADDITIONAL SURGICAL INTERVENTION OF ENLARGEMENT OF PI AND ADDITIONAL OF NEW PI (YAG) PERFORMED ON (B)(6) 2018. (B)(4) Manufacturer Narrative: (B)(4). WORK ORDER SEARCH- NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6786422&pc=MTA | 2023826-2017-01304 | NA | Injury | STAAR SURGICAL COMPANY | 2017-08-10 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE REPORTER STATED THAT THE LENS WAS EXCHANGED. Manufacturer Narrative: THE LENS WAS EXPLANTED DUE TO ELEVATED IOP, NARROWING OF THE ANGLE, ANGLE CLOSURE, PUPILLARY BLOCK, AND EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE RESULTS WERE EXCELLENT. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5429567&pc=MTA | 2023826-2016-00125 | 2015-11-21 | Injury | STAAR SURGICAL COMPANY | 2016-02-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, PUPIL BLOCK AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: CORRECTED INFORMATION: (B)(4). ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: ADDITIONAL INFORMATION ABOUT RETURNED PRODUCT EVALUATION: DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # 708578 H3 OTHER TEXT : LENS NOT RETURNED. | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5505906&pc=MTA | 2023826-2016-00269 | 2015-12-23 | Injury | STAAR SURGICAL COMPANY | 2016-03-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PUPIL BLOCK AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/18. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT AND DEFORMED. DRIED, DISCOLORED MATERIAL WAS FOUND ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND LENS MEASUREMENTS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5668563&pc=MTA | 2023826-2016-00735 | 2016-02-10 | Injury | STAAR SURGICAL COMPANY | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, PUPILLARY BLOCK AND ELEVATED INTRAOCULAR PRESSURE. ADDITIONAL LASER PERIPHERAL IRIDOTOMIES WERE PERFORMED. THE REPORTER INDICATED THEY FELT THERE WAS POSSIBLE AQUEOUS MISDIRECTION. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: RESULT: (OPERATIONAL PROBLEM): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. (B)(4). Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, PRODUCT EVALUATION, MEDICAL REVIEW AND DEVICE HISTORY RECORD REVIEW, WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5780974&pc=MTA | 2023826-2016-00886 | 2016-04-27 | Injury | STAAR SURGICAL COMPANY | 2016-07-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.0MM ICM130V4 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT, ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), PUPIL BLOCK (WITH ELEVATED IOP), AND UNREACTIVE PUPIL (FIXED). THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) IS 20/25. Manufacturer Narrative: CONCLUSION: IT SHOULD BE NOTED THAT AT THE TIME OF IMPLANTATION THE PATIENT WAS (B)(6) AND PER DFU INDICATIONS: "THE VISIAN ICL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE". THEREFORE, THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS. CORRECTED DATA: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REMEASURED AND WAS FOUND TO BE IN SPECIFICATION. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5072607&pc=MTA | 2023826-2015-01214 | 2015-08-12 | Injury | STAAR SURGICAL COMPANY | 2015-09-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Torn Material; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THE LENS TORE WHILE BEING LOADED DUE TO INCORRECT CARTRIDGE (FELT CARTRIDGE WAS TOO SMALL FOR THIS SIZE LENS). THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE LENS WAS TOO LARGE (EXCESSIVE VAULTING) AND ELEVATED INTRAOCULAR PRESSURE/PUPILLARY BLOCK. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/15 AND THE EXCHANGED ICL WAS IN A GOOD POSITION. THE REPORTER INDICATED THE SURGEON FELT THE EVENT WAS DUE A SIZING/LENS CALCULATION ISSUE. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW: REPORTEDLY, ICL (13.7 MM) WAS EXPLANTED/EXCHANGED FOR A SHORTER LENGTH ICL (13.2 MM) ONE WEEK POSTOPERATIVELY TO ADDRESS EXCESSIVE VAULTING AND SUBSEQUENT PUPILLARY BLOCK AND ELEVATED IOP. ACCORDING TO COMPLETED CUSTOMER COMPLAINT QUESTIONNAIRE FOR SURGERY, RECEIVED ON 8/24/15, THE EVENT WAS PRODUCT RELATED AND CAUSED BY LARGE DIAMETER. THE SAME REPORT STATED THAT UCVA AFTER REPLACEMENT WAS 20/15 AND THAT ICL WAS IN A GOOD POSITION. IT SHOULD BE NOTED THAT EXPLANTED ICL WAS USED AS A BACKUP LENS SINCE THE FIRST ONE WAS DAMAGED DURING LOADING BEFORE INSERTION. THE FIRST LENS DIDN'T HAVE PATIENT CONTACT. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403874&pc=MTA | 2023826-2020-01726 | 2020-06-23 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 | MTA | IMPLANTABLE COLLAMER LENS | Off-Label Use; Patient-Device Incompatibility; No Apparent Adverse Event | Intraocular Pressure Increased; Pupillary Block; No Known Impact Or Consequence To Patient; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT AN IMPLANTABLE COLLAMER LENS MAY BE REMOVED AND REPLACED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL.Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL13.2, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO PUPIL BLOCK, ELEVATED IOP, ANGLE CLOSURE, AND UNREACTIVE (FIXED) PUPIL. AN ADDITIONAL SURGICAL PI WAS ALSO ADDED. THE PROBLEM WAS RESOLVED. THE CAUSE IS REPORTED AS A PATIENT-RELATED FACTOR AND UNKNOWN: "POSSIBLY ICL TOO LARGE FOR EYE DESPITE CONFORMITY TO SIZING TABLE/GUIDELINES." Manufacturer Narrative: B5-DISREGARD PREVIOUS B5 STATEMENT. D2-UNABLE TO ENTER PRODUCT CODE MTA. G2-REPORT SOURCE ADDED. H6-HEALTH EFFECT CLINICAL CODE 4581 (UNREACTIVE (FIXED) PUPIL, ANGLE CLOSURE). H6-HEALTH EFFECT IMPACT CODE UNABLE TO ENTER: 4629 (SECONDARY SURGERY, LENS EXCHANGE); 4625 (ADDITION OF NEW PI). H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). Manufacturer Narrative: PRODUCT CODE UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW UNK ENTRY. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15032443&pc=MTA | 2023826-2022-02271 | 2022-06-09 | Injury | STAAR SURGICAL COMPANY | 2022-07-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; Toxic Anterior Segment Syndrome (TASS); Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -7.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON THE SAME DAY OF (B)(6) 2022 IN A SEPARATE SURGERY THE LENS WAS REMOVED DUE TO TOXIC ANTERIOR SEGMENT SYNDROME (TASS)AND ELEVATED IOP. DIAGNOSTIC PERFORMED: NON CONTACT TONOMETRY FLARE DIAGNOSIS. TREATMENT PERFORMED: DROPPING MEDICINE IN THE EYES. ELEVATED IOP AND UNREACTIVE (FIXED) PUPIL WAS ALSO OBSERVED. CAUSE REPORTED AS DEVICE. Manufacturer Narrative: (B)(4). DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18377615&pc=MTA | 2023826-2023-05714 | 2023-10-11 | Injury | STAAR SURGICAL COMPANY | 2023-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Uveitis; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1, -7.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE, PIGMENT DISPERSION, BLURRED VISOIN AND PROLONGED POSTOPERATIVE ANTERIOR UVEITIS. IT WAS REPORTED THAT THE IOP WAS CONTROLLED WITH MEDICATION. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - PIGMENT DISPERSION H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10786734&pc=MTA | 2023826-2020-02499 | 2020-03-11 | Injury | STAAR SURGICAL COMPANY | 2020-11-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -06.50 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON 11/MAR/2020. ELEVATED IOP (58.0 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS ASSESSED ON (B)(6) 2020. THE LENS WAS REMOVED AND EXCHANGED WITH A STRONGER POWER, SAME LENGTH LENS AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. REPORTEDLY, SURGEON "INADVERTENTLY IMPLANTED LENS PREPARED FOR OD." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: "THE LENS WAS REMOVED AND EXCHANGED" SHOULD BE CORRECTED TO "THE LENS WAS EXPLANTED AND IN A SEPARATE SURGERY WAS REPLACED". D9: DATE ENTERED IN SUPPLEMENTAL #1 SHOULD BE REMOVED AND DATE ENTERED IN D9 IN INITIAL MEDWATCH REPORT SHOULD BE KEPT. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS ALSO REPORTED THAT THE PATIENT EXPERIENCED GLARE/HALOES AND AC IRRIGATION. THIS SHOULD HAVE BEEN INCLUDED. H6 - 4625 - AC IRRIGATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308925&pc=MTA | 2023826-2021-00320 | 2021-01-05 | Injury | STAAR SURGICAL COMPANY | 2021-02-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2, -12.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO ELEVATED IOP (28MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESED ON (B)(6) 2020), PIGMENT DISPERSION, UNREACTIVE(FIXED) PUPIL, GLARE/HALOES, IRIS ATROPHY, AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS "EXCESSIVE VAULT CAUSING PIGMENTARY DISPERSION AND INCREASED EYE PRESSURE." REPORTEDLY, "VAULT IS NOW BETTER, ALONG WITH IOP. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- REPLACEMENT LENS (B)(6) DID NOT RESOLVE THE PROBLEM AND WAS EXPLANTED ON (B)(6)2022. PER EMAIL ON (B)(6) 2022, THE "REASON FOR EXPLANT- PATIENT COMPLAINED OF DISCOMFORT". CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK. RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# : (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5688983&pc=MTA | 2023826-2016-00253 | 2015-12-15 | Injury | STAAR SURGICAL COMPANY | 2016-05-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -11.5 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ELEVATED IOP, REDUCTION OF ICA, UNREACTIVE PUPIL, SHALLOW CHAMBER, AND GLARES/HALOES WERE REPORTED. A LENS EXCHANGE IS PLANNED BUT HAS NOT YET BEEN PERFORMED. PATIENT IS CURRENTLY ON ANTIGLAUCOMA AND ANTI-INFLAMMATORY EYE DROPS. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON 2015/12/15. ELEVATED INTRAOCULAR PRESSURE (IOP), REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), UNREACTIVE PUPIL, SHALLOW CHAMBER, AND GLARES/HALOES WERE REPORTED. THE SURGEON PERFORMED ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS AND INTRAOCULAR INJECTION (MEDICATION). PATIENT WAS ON ANTIGLAUCOMA AND ANTI-INFLAMMATORY EYE DROPS. THE LENS WAS EXPLANTED. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. AS PER DFU FOR THIS LENS MODEL, VICLS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7480272&pc=MTA | 2023826-2018-00657 | 2018-04-02 | Injury | STAAR SURGICAL COMPANY | 2018-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL13.2, -12.50, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO OBSERVATION OF EXCESSIVE VAULT, ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (DATE ASSESSED: (B)(6) 2018), UNREACTIVE (FIXED) PUPIL, GLARE/HALOES, AND BLURRED VISION. ADDITIONAL SURGICAL INTERVENTION ( ENLARGEMENT OF PI; YAG) AND MEDICAL INTERVENTION (INTRAOCULAR INJECTION) WERE PERFORMED/PRESCRIBED. THE REPORTED STATED THAT THE CAUSE OF THIS EVENT WAS UNKNOWN. "FIXED PUPIL" WAS REPORTED FOR STATUS OF THE EYE AT PATIENT'S LAST FOLLOW UP. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: AS OF (B)(6) 2018 BEST CORRECTED ACUITY 20/30, IOP IS ELEVATED TO 27UM. PATIENT IS BEING MONITORED FOR OCULAR HYPERTENSION. CORRECTIONS: PATIENT DID NOT HAVE INTRAOCULAR INJECTIONS. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CORRECTED DATA: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) . WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7480527&pc=MTA | 2023826-2018-00662 | 2018-04-05 | Injury | STAAR SURGICAL COMPANY | 2018-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 MM MICL13.2, -9.5, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO OBSERVATION OF EXCESSIVE VAULT, ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, GLARE/HALOES, AND BLURRED VISION. ADDITIONAL SURGICAL INTERVENTION (YAG) AND MEDICAL INTERVENTION (INTRAOCULAR INJECTION) WERE PERFORMED/PRESCRIBED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE STATUS OF THE EYE WAS REPORTED TO BE WITHIN NORMAL LIMITS. PATIENT IS BEING MONITORED FOR OCULAR HYPERTENSION. CORRECTIONS: PATIENT DID NOT HAVE INTRAOCULAR INJECTIONS. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WITHIN SPECIFICATION. CORRECTED DATA: CORRECT DATE IS (B)(6) 2018, (B)(6) 2018 IS INCORRECT. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6546111&pc=MTA | 2023826-2017-00770 | 2017-01-09 | Injury | STAAR SURGICAL COMPANY | 2017-05-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN EVO+ COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED A VISIBLE DISTORTION TO THE IRIS. NO FURTHER INFORMATION HAS BEEN PROVIDED. AT THE DATE OF MDR SUBMISSION, MULTIPLE ATTEMPTS TO CONTACT REPORTER HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THE SURGEON IMPLANTED A VICM5_13.2, 13.2MM, -8 DIOPTER LENS. PATIENT ALSO EXPERIENCED EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOES. LENS REMAINS IMPLANTED. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8307839&pc=MTA | 2023826-2019-00222 | 2019-01-10 | Injury | STAAR SURGICAL COMPANY | 2019-02-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS EXCESSIVE VAULT, ELEVATED IOP AND GLARE/HALOS. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: ADVERSE EVENT OR PRODUCT PROBLEM: CORRECTED FOR "ADVERSE EVENT" TO "PRODUCT PROBLEM" FOR INITIAL MDR. TYPE OF REPORTABLE EVENT: CORRECTED FROM "SERIOUS INJURY" TO "MALFUNCTION". CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL PATIENT CODE ADDED, 3191 - NO CODE AVAILABLE (SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES) PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9745394&pc=MTA | 2023826-2020-00378 | 2019-11-25 | Injury | STAAR SURGICAL COMPANY | 2020-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT, ELEVATED IOP AND GLARE/HALOES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: DEVICE CODE 3001: GLARE/HALOES. METHOD CODE 4110: LENS WORK ORDER SEARCH: ONE ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10105253&pc=MTA | 2023826-2020-01077 | NA | Injury | STAAR SURGICAL COMPANY | 2020-06-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -17.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS, PUPIL BLOCK WITH ELEVATED IOP, GLARE/HALOES, AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7835113&pc=MTA | 2023826-2018-01317 | 2018-02-19 | Injury | STAAR SURGICAL COMPANY | 2018-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL13.2 , -125 DIOPTER WAS IMPLANTED AND REMOVED ON (B)(6) 2018 DUE TO HIGH IOP AND UNHAPPY VISION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: TWO SIMILAR COMPLAINT TYPE EVENT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED ON (B)(6) 2018 DUE TO EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED IOP (DATE ASSESSED: (B)(6) 2018 AND (B)(6) 2018), AND UNREACTIVE PUPIL. ENLARGEMENT OF PI (YAG AND SURGICAL) WAS REPORTED. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER AND SHORTER LENGTH LENS. IT WAS REPORTED THAT THIS EXCHANGE RESOLVED THE PROBLEM. ADDITIONAL INFORMATION REPORTED: PATIENT HAD ICL OD WITHOUT COMPLICATIONS (B)(6) 2018. PATIENT HAD PRK OU WITHOUT COMPLICATION (B)(6) 2018. REPORTER STATED THAT POSSIBLE CAUSE OF EVENT WAS UGH SYNDROME DESPITE PATENT PI. (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: PATIENT CODE: SECONDARY SURGICAL INTERVENTION,EXPLANT. WORK ORDER SEARCH- NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7631451&pc=MTA | 2023826-2018-00930 | 2018-02-28 | Injury | STAAR SURGICAL COMPANY | 2018-06-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Visual Disturbances; No Consequences Or Impact To Patient; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM MICL13.2, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO OBSERVATION OF EXCESSIVE VAULT, ELEVATED IOP (INTRAOCULAR PRESSURE), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION, AND GLARE/HALOES. ADDITIONAL SURGICAL INTERVENTION (ENLARGEMENT AND ADDITION OF NEW PI) WERE PERFORMED. ADDITIONAL INFORMATION REPORTED; DECREASED CELL AND FLARE + REMOVAL. PATIENT REPORTS DECREASE IN GLARE. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A THERMO SCIENTIFIC CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN. CORRECTED DATA: PRE-EXISTING CONDITION: AXIAL LENGTH GREATER THAN 29. CONDITION SHOULD HAVE BEEN MENTIONED IN INITIAL MDR. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. PATIENT CODE: (B)(4) SECONDARY SURGICAL PROCEDURE, ENLARGEMENT OF PI, ADDITION OF NEW PI, EXPLANT. PATIENT CODE: (B)(4) SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6308851&pc=MTA | 2023826-2017-00226 | NA | Injury | STAAR SURGICAL COMPANY | 2017-02-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure, Delayed, Uncontrolled | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE RIGHT EYE (OD). THE PATIENT REPORTED EPISODE OF ELEVATED EYE PRESSURE OF UNDETERMINED CAUSE, NOT ONGOING. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6309027&pc=MTA | 2023826-2017-00227 | NA | Injury | STAAR SURGICAL COMPANY | 2017-02-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure, Delayed, Uncontrolled | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS,-9.5 DIOPTER, IN THE LEFT EYE (OS). THE PATIENT REPORTED EPISODE OF ELEVATED EYE PRESSURE OF UNDETERMINED CAUSE, NOT ONGOING. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5989110&pc=MTA | 2023826-2016-01351 | 2016-07-12 | Injury | STAAR SURGICAL COMPANY | 2016-09-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure, Delayed, Uncontrolled; Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION - THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6001930&pc=MTA | 2023826-2016-01398 | 2015-08-31 | Injury | STAAR SURGICAL COMPANY AG | 2016-10-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure, Delayed, Uncontrolled; Intraocular Pressure Increased; Pupillary Block; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM13.7 IMPLANTABLE COLLAMER LENS, -16.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, PUPILLARY BLOCK, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ANGLE CLOSURE. A SECONDARY YAG WAS PERFORMED TO ENLARGE OR ADDITION OF THE PERIPHERAL IRIDECTOMIES. Manufacturer Narrative: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN/BENT/DEFORMED. FOREIGN MATERIAL (UNIDENTIFIED SPOTS) WAS PRESENT ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5016061&pc=MTA | 2023826-2015-01061 | 2015-05-26 | Injury | STAAR SURGICAL COMPANY | 2015-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure, Delayed, Uncontrolled; Missing Value Reason | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -15.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON 06/20/2015 DUE TO INACCURATE WHITE TO WHITE MEASUREMENT. THE LENS WAS EXCHANGED FOR A SMALLER LENS LENGTH WITH A SIMILAR DIOPTER AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -15.00 (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6024713&pc=MTA | 2023826-2016-01447 | 2016-09-08 | Injury | STAAR SURGICAL COMPANY | 2016-10-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure, Delayed, Uncontrolled; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE REPORTER INDICATED A POSSIBLE CAUSE OF THE EVENT MAY BE ANATOMY AND LENS SIZE. THE PATIENT IS DOING WELL AND UCVA WAS 20/20. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5989262&pc=MTA | 2023826-2016-01355 | 2016-09-14 | Injury | STAAR SURGICAL COMPANY | 2016-09-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure, Delayed, Uncontrolled; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5989275&pc=MTA | 2023826-2016-01356 | 2016-09-14 | Injury | STAAR SURGICAL COMPANY | 2016-09-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure, Delayed, Uncontrolled; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. A PIECE OF THE OTHER HAPTIC WAS TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6265479&pc=MTA | 2023826-2017-00095 | 2016-12-01 | Injury | STAAR SURGICAL COMPANY | 2017-01-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure, Delayed, Uncontrolled; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXTREME VAULTING, WHICH CAUSED PRESSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: PATIENT COMPLAINED OF PAIN AND AFTER EXAMINATION, EXTREME VAULTING WAS FOUND CAUSING HIGH PRESSURE. LENS WAS EXPLANTED AND A LENS OF SHORTER LENGTH WAS IMPLANTED. THE PATIENT IS DOING WELL AND BCVA WAS 20/20 POST SECOND ICL IMPLANT. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7405164&pc=MTA | 2023826-2018-00522 | 2018-02-09 | Injury | STAAR SURGICAL COMPANY | 2018-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure, Delayed, Uncontrolled; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED.. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND THE TORN HAPTIC PIECE MISSING. THERE WAS PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED SHORTER LENS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5989101&pc=MTA | 2023826-2016-01363 | NA | Injury | STAAR SURGICAL COMPANY | 2016-09-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure, Delayed, Uncontrolled; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6140269&pc=MTA | 2023826-2016-01708 | NA | Injury | STAAR SURGICAL COMPANY | 2016-12-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure, Delayed, Uncontrolled; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5680335&pc=MTA | 2023826-2016-00758 | NA | Injury | STAAR SURGICAL COMPANY | 2016-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure, Delayed, Uncontrolled; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO INCORRECT SIZE. THE LENS WAS EXCHANGED FOR ANOTHER ICL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: RESULT: (OPERATIONAL PROBLEM): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF ONE HAPTIC WERE TORN OFF AND MISSING. THE OTHER HAPTIC WAS BENT. THE LENS WAS RETURNED IN LIQUID AND THERE WAS EVIDENCE OF SURGICAL RESIDUE. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE AND SHALLOW ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE IRIS SPHINCTER WAS DAMAGED DUE TO THE PRESSURE. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO MISMEASUREMENT OF THE WHITE-TO-WHITE. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. EVENT PROBLEM AND EVALUATION CODES: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6251929&pc=MTA | 2023826-2017-00058 | NA | Injury | STAAR SURGICAL COMPANY | 2017-01-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure, Delayed, Uncontrolled; Pain; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE AND NARROWING OF THE ANGLE. THE PATIENT EXPERIENCED PAIN AND WAS TREATED WITH MEDICATION. A REFRACTIVE LENS EXCHANGE WAS PERFORMED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10496543&pc=MTA | 2023826-2020-01974 | NA | Injury | STAAR SURGICAL COMPANY | 2020-09-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Pain | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VISIAN ICL LENS, -17.0 DIOPTER INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. THE SURGEON REPORTS PAIN IN THE CILIARY ARCH. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. REPORTEDLY, THE DOCTOR THINKS POSSIBLY THIS PAIN IS NOT DUE TO THE ICL. Manufacturer Narrative: B5-UPDATED INFORMATION: REPORTEDLY, THE PATIENT IS "PERFECT" NOW AND WITHOUT SYMPTOMS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11547918&pc=MTA | 2023826-2021-00739 | 2021-02-12 | Injury | STAAR SURGICAL COMPANY | 2021-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Defective Device; Unintended Movement | Pain; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE PATIENT WAS REPORTED WITH HYPEROPIA AND ASTIGMATISM, WHICH CAUSED BAD AND BLURRED VISION. ALSO REPORTED PERMANENT LOSS OF BCVA. THE LENS WAS ALSO REPORTED AS HAVING FLIPPED. THE PATIENT HAD A SMALL EYE, THERE WAS PUPILAR CONSTRICTION AND THE PUPIL BEGAN TO CLOSE DURING THE SURGERY. THE IRIS GOT EXTRUDED OUT THE WOUND DURING THE SURGERY. THE LENS WAS EXPLANTED AND RE-IMPLANTED ON (B)(6) 2021. PERIBULBAR BLOCK WAS ADMINISTERED DURING 2ND SURGERY. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT IS HAPPY WITH THE RESULTS OF THE SURGERY. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: IT WAS ALSO INDICATED THE PATIENT EXPERIENCED PAIN; NON CONFORMITY THAT WASN'T SPECIFIED WHAT TYPE; LENS WAS IMPLANTED UPSIDE DOWN/FLIPPED, LENS WAS EXPLANTED ON 03-MARCH-2021 AND THE DAY AFTER WAS RE-IMPLANTED. H6: 4581- LOSS OF BCVA, IRIS PROLAPSE. CLAIM#: (B)(4). Manufacturer Narrative: IT WAS LATER CLARIFIED THAT THE LENS WAS REMOVED ON (B)(6) 2021 AND PROPERLY REIMPLANTED, INTRAOPERATIVELY AND NOT IN A SEPARATE SURGERY. THIS HAD RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN. ETHNICITY: UNKNOWN. RACE: UNKNOWN. DATE REC¿D BY MFR: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11557439&pc=MTA | 2023826-2021-00740 | 2021-02-12 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Defective Device; Unintended Movement | Pain; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE PATIENT WAS REPORTED WITH HYPEROPIA AND ASTIGMATISM, WHICH CAUSED BAD AND BLURRED VISION. ALSO REPORTED PERMANENET LOSS OF BCVA. THE LENS WAS ALSO REPORTED AS HAVING FLIPPED. THE PATIENT HAD A SMALL EYE, THERE WAS PUPILAR CONSTRICTION AND THE PUPIL BEGAN TO CLOSE DURING THE SURGERY. THE LENS WAS EXPLANTED AND RE-IMPLANTED ON (B)(6) 2021. PERIBULBAR BLOCK WAS ADMINISTERED DURING THE 1ST SURGERY (INITIAL) AND 2ND SURGERY (RE-IMPLANT). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. TWO HOURS AFTER THE RE-IMPLANT SURGERY, ANTERIOR CHAMBER A LITTLE BIT NARROW AND PUPIL DILATED. THE PATIENT IS HAPPY WITH THE RESULTS OF THE SURGERY. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS ALSO INDICATED THAT THE PATIENT EXPERIENCED PAIN. REPORTER INDICATED THERE WAS PRODUCT NONCONFORMITY BUT THIS WAS NOT SPECIFIED ON. LENS WAS IMPLANTED UPSIDE DOWN/FLIPPED. IT WAS LATER CLARIFIED THAT THE LENS WAS EXPLANTED ON (B)(6) 2021, AND THE NEXT DAY RE-IMPLANTED. H6 - 4581 - LOSS OF BCVA. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5/H10- IT WAS LATER CLARIFIED THAT THE LENS WAS REMOVED ON (B)(6) 2021 AND PROPERLY REIMPLANTED, INTRAOPERATIVELY AND NOT IN A SEPARATE SURGERY. THIS HAD RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4917237&pc=MTA | 2023826-2015-00799 | 2015-06-17 | Injury | STAAR SURGICAL COMPANY | 2015-07-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO NO LENS VAULT. THERE WAS NO LENS OPACITY. THE PATIENT EXPERIENCED BLURRY VISION AND EYE PAIN. THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/15.Event Description: ADDITIONAL INFORMATION RECEIVED - THE FACILITY INDICATED THE DATE OF INITIAL SURGERY WAS (B)(6) 2015. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE ROOT CAUSE OF THE COMPLAINT. USE OF THE WRONG LENS WAS MOST LIKELY THE CAUSE OF LACK OF VAULT. (B)(4). Manufacturer Narrative: (B)(4). EVALUATION CODES: METHOD: WORK ORDER SEARCH RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: CORRECTION - IMPLANTED DATE - (B)(6) 2015. METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - THE CONSEQUENCE OF INADEQUATE VAULTING COULD BE CATARACT THAT WARRANT LENS EXCHANGE TO PREVENT FURTHER DAMAGE. ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING WAS A CONSEQUENCE OF WRONG LENS USE WHICH INCLUDE THE FOLLOWING: IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENT BASED UPON DIFFERENT TECHNIQUES, POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN A INDIVIDUAL CASE, IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST AND IMPROPER LENS LABEL. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6074717&pc=MTA | 2023826-2016-01575 | 2016-10-18 | Injury | STAAR SURGICAL COMPANY | 2016-11-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS REPOSITIONED ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE LENS WAS EXCHANGED FOR A SHORTER LENS BUT THAT LENS ALSO IS VAULTED. THE PATIENT'S POST-OP BSCVA WAS 20/30 +1, WITH BLURRY VISION AND PAIN. Manufacturer Narrative: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14107495&pc=MTA | 2023826-2022-01095 | 2022-01-11 | Injury | STAAR SURGICAL COMPANY | 2022-04-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Pain; Discomfort | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 FOR RELIEF OF PAIN AND DISCOMFORT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12226996&pc=MTA | 2023826-2021-02915 | 2021-07-14 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Pain; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -9 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTED PIGMENT DISPERSION AND PATIENT EXPRESSED A BIT OF PAIN/DISCOMFORT. AT A FOLLOW UP DATE AN OCT SCAN WAS PERFORMED AND IT WAS INDICATED THAT IOP WAS NORMAL AND PATIENT IS OKAY, THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED "THIS IRIS SHOULD BE FRAGILE BECAUSE OD IS OK, 13.2MM ICL WAS IMPLANTED AS WELL". Manufacturer Narrative: (B)(4). PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8628687&pc=MTA | 2023826-2019-00889 | 2018-06-15 | Injury | STAAR SURGICAL COMPANY | 2019-05-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Pain; Discomfort; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM, VICM5_12.1, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT EXPERIENCED DISCOMFORT AND ACHE, ON (B)(6) 2018 THE LENS WAS EXPLANTED. IT WAS REPORTED THAT THE PROBLEM RESOLVED AFTER EXPLANTING THE LENS. THERE IS NO PROBLEM WITH THE PATIENT. AT PATIENTS LAST VISIT, BCVA WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY WITH RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND; NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: EVENT DESCRIPTION (EXPLANT DATE) "(B)(6) 2018" SHOULD BE CORRECTED TO "(B)(6) 2019" IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9186507&pc=MTA | 2023826-2019-01860 | 2019-08-01 | Injury | STAAR SURGICAL COMPANY | 2019-10-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Pain; Discomfort; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM5 12.6, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY "THE PATIENT HAD A PAIN OF THE EYE AND THERE WAS NO TREATMENT FOR THE PAIN." ON (B)(6) 2019 THE LENS WAS REMOVED. PER THE REPORTER ON (B)(6) 2019, "AFTER REMOVING THE LENS", THE PATIENT IS IN GOOD CONDITION. LENS REMOVAL RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: PATIENT CODE 3191- SECONDARY SURGICAL INTERVENTION; LENS EXPLANT. RESULT CODE 213 CORRECTED TO 114. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE ON LENS, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9186511&pc=MTA | 2023826-2019-01861 | 2019-08-01 | Injury | STAAR SURGICAL COMPANY | 2019-10-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Pain; Discomfort; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM5 12.6, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON 20/FEB/2018. REPORTEDLY "THE PATIENT HAD A PAIN OF THE EYE AND THERE WAS NO TREATMENT FOR THE PAIN." ON 10/AUG/2019 THE LENS WAS REMOVED. PER THE REPORTER ON 22/SEP/2019, "AFTER REMOVING THE LENS", THE PATIENT IS IN GOOD CONDITION. LENS REMOVAL RESOLVED THE PROBLEM. Manufacturer Narrative: RESULT CODE 213 IS NOT APPLICABLE AND SHOULD BE CORRECTED TO RESULT CODE 114 IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE ON LENS, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT CODE 3191: 'EXCHANGE" SHOULD BE CORRECTED TO "EXPLANT" IN THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5922831&pc=MTA | 2023826-2016-01210 | 2015-11-04 | Injury | STAAR SURGICAL COMPANY | 2016-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Pain; Discomfort; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO 13.2 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO PATIENT PAIN AND DISCOMFORT. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (NOT POSSIBLE TO SPECIFY) ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS IN A PATIENT UNDER THE AGE OF 21 YEARS IS CONTRAINDICATED. THIS PATIENT WAS (B)(6) AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5917994&pc=MTA | 2023826-2016-01169 | 2015-11-04 | Injury | STAAR SURGICAL COMPANY | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Pain; Discomfort; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO 13.2 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO PATIENT DISCOMFORT AND PAIN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL WAS OBSERVED TO BE ON THE LENS (NOT POSSIBLE TO SPECIFY). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WEIGHT: UNK. (B)(4) DEVICE NOT RETURNED. | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6094953&pc=MTA | 2023826-2016-01636 | 2016-10-18 | Injury | STAAR SURGICAL COMPANY | 2016-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Pain; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED LATER THAT SAME DAY DUE TO PATIENT WAS IN EXTREME PAIN. THE LENS WAS NOT REPLACED WITH ANOTHER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9721911&pc=MTA | 2023826-2020-00312 | 2020-01-13 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-18 | MTA | IMPLANTABLE COLLAMER LENS | Contamination /Decontamination Problem | Pain; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. APPROXIMATELY 1-MONTH AFTER SURGERY THE SURGEON REPORTS WHITE DEPOSITS FOUND ON THE LENS AND THE PATIENT HAS PAIN AND CONGESTION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6296898&pc=MTA | 2023826-2017-00165 | 2017-01-02 | Injury | STAAR SURGICAL COMPANY | 2017-02-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Pain; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -10 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE SURGEON PLANS TO EXCHANGE LENS FOR A LARGER SIZE DUE TO LOW VAULTING. LENS REMAINS IMPLANTED. THE PATIENT HAS PAIN IN THEIR RIGHT EYE (OD). Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED ON 01/25/2017 AND EXCHANGED WITH A LONGER LENS. THE PROBLEM WAS RESOLVED. SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. (B)(4). Manufacturer Narrative: FOLLOWUP #1 WAS INADVERTENTLY SKIPPED AND FOLLOWUP #2 AND #3 WERE SUBMITTED INSTEAD. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). LENS NOT RETURNED. | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7261496&pc=MTA | 2023826-2018-00174 | 2018-01-10 | Injury | STAAR SURGICAL COMPANY | 2018-02-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Pain; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -10.50, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE PATIENT BEGAN TO COMPLAIN OF OCULAR PAIN. REPORTEDLY, ALSO NOTED IS EXCESSIVE VAULT AND THAT THE LENS IS "COMPROMISING THE ANGLES." THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: LENS WAS EXCHANGED WITH A DIFFERENT POWER, VAULT WAS FINE (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6775540&pc=MTA | 2023826-2017-01298 | 2017-07-07 | Injury | STAAR SURGICAL COMPANY | 2017-08-08 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Pain; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, PUPILLARY BLOCK (ELEVATED IOP), BLURRED VISION, AND PAIN. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: THE EXCHANGE RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A CASE/VIAL. THE LENS VIAL WAS RETURNED TAPED TO LAF CARD. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7146624&pc=MTA | 2023826-2017-01945 | 2016-11-16 | Injury | STAAR SURGICAL COMPANY | 2017-12-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Pain; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6, -7 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE, GLARES, HALOES, INDUCED ASTIGMATISM, AND CILIARY GROOVE PAIN. THE LENS WAS EXCHANGED FOR A DIFFERENT LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10285658&pc=MTA | 2023826-2020-01486 | 2020-06-22 | Injury | STAAR SURGICAL COMPANY | 2020-07-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Pain; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 LENS, -17.0 DIOPTER INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED, THE SURGEON REPORTING LENS ROTATION, LENS DISLOCATION/SUBLUXATION, CILIARY PAIN, LENS DECENTRATION, TYNDALL 2+, AND DOUBLE IMAGE. TOPICAL STERIODS WERE ALSO ADMINISTERED. REPOSITIONING DID NOT RESOLVE THE PROBLEM. REPORTEDLY, ON (B)(6) 2020 THE LENS WAS EXPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, HOWEVER 1 BROKEN HAPTIC WAS DISCOVERED UPON EXPLANT WHICH COULD BE A CONTRIBUTING CAUSE. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: REPORTEDLY, A REPLACEMENT LENS WAS IMPLANTED AT A LATER DATE. PAIN PERSISTS. REFERENCE FILE# (B)(4) FOR REPLACEMENT LENS INFORMATION. H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. THERE WAS DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9385347&pc=MTA | 2023826-2019-02271 | 2019-10-30 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Pain; Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICM5_12.1, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. PATIENT REPORTED, "ISSUES WITH BINOCULAR VISION, NASAL PAIN, AND A CONSTANTLY CONTRACTION OF THE CILIARY BODY." LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - PATIENT CODE: 3191 - NO CODE AVAILABLE (CONTRACTION OF THE CILIARY BODY) WAS REPORTED IN MANUFACTURER NARRATIVE, BUT CODE SHOULD HAVE ALSO BEEN SELECTED IN H6. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9385361&pc=MTA | 2023826-2019-02272 | 2019-10-30 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Pain; Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICM5_12.1, -08-50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. PATIENT REPORTED, "ISSUES WITH BINOCULAR VISION, NASAL PAIN, AND A CONSTANTLY CONTRACTION OF THE CILIARY BODY." LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - PATIENT CODE 3191: NO CODE AVAILABLE (CONTRACTION OF THE CILIARY BODY) WAS REPORTED IN THE MANUFACTURER NARRATIVE, BUT SHOULD HAVE ALSO BEEN SELECTED IN H6 SECTION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6487527&pc=MTA | 2023826-2017-00618 | NA | Injury | STAAR SURGICAL COMPANY | 2017-04-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Pupillary Block; Blurred Vision; No Known Impact Or Consequence To Patient; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, AND THE LENS WAS NOTED AS HAVING AN EXCESSIVE VAULT. THE PLAN IS TO EXPLANT AND EXCHANGE THE LENS FOR A SHORTER LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS IMPLANTED ON (B)(6) 2017 IN THE PATIENT'S RIGHT EYE (OD), -11.0 DIOPTER. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO HYPEROPIC SURPRISE, EXCESSIVE VAULTING, NARROWING OF THE ANGLE AND ANGLE CLOSURE. A YAG PERIPHERAL IRIDOTOMY WAS PERFORMED TO RELIEVE PUPIL BLOCK. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENTS POST-OP BSCVA WAS 20/25. (B)(4). CORRECTED DATA: THE IMPLANTED DATE IS (B)(6) 2017 NOT (B)(6) 2017. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL, IN A BIOHAZARD BAG. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE. CLAIM # (B)(4). Manufacturer Narrative: CONCLUSION CODE: (VISIAN® IMPLANTABLE COLLAMER® LENSES ARE INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE. THIS PATIENT WAS (B)(6) YEARS OF AGE). (B)(4) Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6629958&pc=MTA | 2023826-2017-00949 | 2017-05-10 | Injury | STAAR SURGICAL COMPANY | 2017-06-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Pupillary Block; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE SURGEON COMMENTED THAT THE PATIENT ALSO EXPERIENCED PUPIL BLOCK, AND SUSPICION OF PIGMENT DISPERSION. THE DR. PLANS TO EXCHANGE THE LENS FOR THE PATIENT. AS OF THE DAY OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. MULTIPLE ATTEMPTS TO CONTACT TO CUSTOMER HAVE BEEN UNSUSCCESSFUL. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7403689&pc=MTA | 2023826-2018-00510 | 2018-03-06 | Injury | STAAR SURGICAL COMPANY | 2018-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED. THE SURGEON REPORTS THE FOLLOWING: EXCESSIVE VAULT WITH "NEAR CORNEAL TOUCH IN PERIPHERY," "INTRAOPERATIVE PUPILLARY BLOCK WITH EXTRUSION OF IRIS THROUGH WOUNDS," LENS REPOSITIONING AND ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS. REPORTEDLY, THE PROBLEM WAS RESOLVED WITH EXPLANT. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL PLASTIC BAG AND HAD CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND FOUND THE PRESENCE OF FIBERS AND BLACK SPOTS ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT, LENS REPOSITIONING, NEAR CORNEAL TOUCH IN PERIPHERY, EXTRUSION OF IRIS THROUGH WOUNDS, ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6540585&pc=MTA | 2023826-2017-00741 | 2016-11-21 | Injury | STAAR SURGICAL COMPANY | 2017-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Vitreous Floaters; Headache; Blurred Vision; Visual Disturbances; Halo; Discomfort; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -2.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED IRIS DISTORTION, DISCOMFORT, HEADACHES, GLARE/HALOES, BLURRED VISION, DOUBLE VISION OR "GHOSTING", "VITREOUS FLOATERS", PUPIL REMAINS DILATED AND NONRESPONSIVE, AND OTHER MISCELLANEOUS VISUAL DISTURBANCES. THE PATIENT IS NOT HAPPY WITH THE RESULT OF THE IMPLANT. ON (B)(6) 2017 THE PATIENT BEGAN PHYLOCARPINE DROPS TREATMENT AND THE DR. IS MONITORING THE OCCURENCE. AS OF THE DATE OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PATIENT ALSO EXPERIENCED EXCESSIVE VAULT, LOSS OF BCVA, AND IRIS ATROPHY. (B)(6). (B)(4). Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW - BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. "PHYLOCARPINE" PILOCARPINE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17882069&pc=QCB | 2023826-2023-04269 | 2023-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF-4.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE ON (B)(6) 2023. BLURRED VISION; SEVERE DISTENDING SENSATION;PAIN; DIZZINESS WAS REPORTED AFTER SURGERY. ON (B)(6) 2023 THE LENS WAS REPOSITIONED THEN ON SAME DAY REMOVED AND THE PROBLEM WAS RESOLVED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: H6-CLINICAL CODE 4581- SEVERE DISTENDING SENSATION; DIZZINESS. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17882376&pc=QCB | 2023826-2023-04270 | 2023-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF-4.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE(OS) ON 07-JUL-2023. BLURRED VISION; SEVERE DISTENDING SENSATION;PAIN; DIZZINESS WAS REPORTED AFTER SURGERY. ON (B)(6) 2023 THE LENS WAS REPOSITIONED THEN ON SAME DAY REMOVED AND THE PROBLEM WAS RESOLVED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLINICAL CODE 4581- SEVERE DISTENDING SENSATION; DIZZINESS. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13362153&pc=QCB | 2023826-2022-00348 | 2021-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7;-11.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE; AND ON (B)(6) 2021 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16465175&pc=MTA | 2023826-2023-00767 | 2023-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE REPORTER NOTES A PLANNED LENS EXPLANT AND REPLACEMENT DUE TO EXCESSIVE VAULT. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP (INTRAOCULAR PRESSURE). A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6:1494-OFF LABEL: PROGRESSIVE MYOPIA. D9: RETURN DATE- 05-APR-2023 CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16595289&pc=MTA | 2023826-2023-00985 | 2023-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE REPORTER NOTES A PLANNED LENS EXPLANT AND REPLACEMENT DUE TO EXCESSIVE VAULT. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP (INTRAOCULAR PRESSURE). A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: 1494-OFF LABEL- PROGRESSIVE MYOPIA D9: RETURN DATE: 05-APR-2023 (B)(4) | Lens Exchange | Resolved